## Slide 1: <FPY improvement on FTN testing in Venture>
Owner | Luo HuaZhen/Victor Montero | ROI |  |  | Needle seal force test failure drop from 11% to 0%; Cost saving 376.6 USD/year | 
Date | 05 Apr 2021 |  |  |  |  | 
Purpose |  |  |  |  |  | 
This project is to improve FTN sample manager test First Pass Yield by reducing the Needle Seal Force test failure rate from 11% to 0% through tightening spring rate spec. |  |  |  |  |  | 
Progress |  | Project Status |  |  |  | 
Q4 2019, Needle seal force is identified as area to improve. Q1- Q2 2020, an evaluation with high side of spring rate was conducted in Venture. The result shows positive. Q3-Q4 2020, Milford sustaining engineering team conducted evaluation in Milford after getting positive evaluation result from Venture. Q1 2021, evaluation result shows positive in Milford, but team wants optimize the spec for spring rate as observed some batch differences although results shows almost 0 fallout. |  | Q1 | Q2 | Q3 | Q3 | Q4
 |  |  |  | Target to complete | Target to complete | 
Next Steps |  | Risks & Issues |  |  |  | 
Q2 2021, Milford will review extended evaluation result and release updated spring rate drawing to finalize this improvement. |  | Evaluation progress was affected by Covid-19 |  |  |  | 
Notes: Look at having different colour schemes for different pillars

## Slide 2
Needle Seal Force Test – Force Stack up Analysis
Spring occupy 60.7% of Test Acceptance specification (< 20 lb/in).
Raw springs are measured from 3rd party, the compression force is at low side (62 lb/in, spec : 59 ~ 72 lb/in).
Proposed : To use spring at high side tolerance – 65.5 ~ 72 lb/in, to raise up seal force 5~10 lb/in.
Man
Method
Material
Machine
Operator
No new trainee
Seating Port Locking Nut on spring (1.17lb/in)
Position Wash Cup (2.6lb/in)
Spring Tolerance (12.14lb/in)
Control PCBA ( X lb/in)
Load Cell (X lb/in)
Fixture (Wash Cup misalignment)(X lb/in)
Test Software (Needle travelling distance )(X lb/in)
Discipline
As per WI
Needle Seal Force
< 20 lb/in
Parts | Force Ib/in | % of 20 lb/in
Spring Tolerance | 12.14 | 60.70
Port Seating | 1.17 | 5.85
Wash cup position | 2.60 | 13.0
Fixture | X | 20.45
Test Software | X | 
Load Cell | X | 
Control PCBA | X | 
Total |  | 100
Note : X – To be studied / defined
Area to improve

## Slide 3: Evaluation result shows positive reading with high side of spring rate
Needle seal force test failure rate dropped from 11% to 0 since started evaluation in Q1 2020.
Spring force reading variance from different batches was observed during evaluation. Although all test data shows it’s within the spring rate spec (< USL 68 lb/in and > LSL 48 lb/in).
An extended evaluation is conducting in Milford to optimize the spring rate spec before releasing new spring drawing with tighten spec.