## Slide 1: <FPY improvement on FTN testing in Venture>
Owner | Luo HuaZhen | ROI |  |  | Needle seal force test failure drop from 11% to 0%; Cost saving 376.6 USD/year | 
Date | 05 May 2020 |  |  |  |  | 
Purpose |  |  |  |  |  | 
This project is to improve FTN sample manager test First Pass Yield by reducing the Needle Seal Force test failure from 11% to 0% by tightening spring rate spec from 65.7lb/in (+/- 10%) to 63.5lb/inch+10%/-0%. |  |  |  |  |  | 
Progress |  | Project Status |  |  |  | 
Q4 2019, Needle seal force failure was identified as area to improve Q1 2020, provided positive evaluation data to Milford Sustaining Engineering Q2 2020, Milford is working on Spring rate spec update |  | Q1 | Q2 | Q3 | Q3 | Q4
 |  |  |  | Target to complete | Target to complete | 
Next Steps |  | Risks & Issues |  |  |  | 
Q3 2020, implementation of new spring rate spec |  | To update parts spec, have to go through a cycle of evaluation and across function to get it done Evaluation progress was affected by Covid-19 |  |  |  | 
Notes: Look at having different colour schemes for different pillars

## Slide 2
Needle Seal Force Test – Force Stack up Analysis
Spring occupy 60.7% of Test Acceptance specification (< 20 lb/in).
Raw springs are measured from 3rd party, the compression force is at low side (62 lb/in, spec : 59 ~ 72 lb/in).
Proposed : To use spring at high side tolerance – 65.5 ~ 72 lb/in, to raise up seal force 5~10 lb/in.
Man
Method
Material
Machine
Operator
No new trainee
Seating Port Locking Nut on spring (1.17lb/in)
Position Wash Cup (2.6lb/in)
Spring Tolerance (12.14lb/in)
Control PCBA ( X lb/in)
Load Cell (X lb/in)
Fixture (Wash Cup misalignment)(X lb/in)
Test Software (Needle travelling distance )(X lb/in)
Discipline
As per WI
Needle Seal Force
< 20 lb/in
Parts | Force Ib/in | % of 20 lb/in
Spring Tolerance | 12.14 | 60.70
Port Seating | 1.17 | 5.85
Wash cup position | 2.60 | 13.0
Fixture | X | 20.45
Test Software | X | 
Load Cell | X | 
Control PCBA | X | 
Total |  | 100
Note : X – To be studied / defined

## Slide 3
Needle Seal Force Test – Spring 63.5lb/inch (+10% and -0)
19 Dec
Based on shifted test data,
Max Reading: 65.79 lb/in
Min Reading: 51.22 lb/in
USL: 68lb/in
LSL: 48lb/in

## Slide 4: <FPY improvement on BSM/QSM testing in Venture>
Owner | Luo HuaZhen | ROI |  |  | Reduce failure rate from 12.5% to 6% | 
Date | 29 May 2020 |  |  |  |  | 
Purpose |  |  |  |  |  | 
This project is to improve the FPY on BSM/QSM testing by empty beaker |  |  |  |  |  | 
Progress |  | Project Status |  |  |  | 
 |  | Q1 | Q2 | Q3 | Q3 | Q4
 |  |  |  |  | Target to complete | Target to complete
Next Steps |  | Risks & Issues |  |  |  | 
Notes: Look at having different colour schemes for different pillars

## Slide 5: <FPY improvement on FTN testing in Venture>
Owner | Luo HuaZhen | ROI |  |  | Reducing failure rate from H – 2.79%, I – 3.52% to 1% | 
Date | 29 May 2020 |  |  |  |  | 
Purpose |  |  |  |  |  | 
Improve the FPY on FTN testing by reducing Preblank test failure rate. Implemented Preblank preparation procedure. |  |  |  |  |  | 
Progress |  | Project Status |  |  |  | 
 |  | Q1 | Q2 | Q3 | Q3 | Q4
 |  |  |  |  | Target to complete | Target to complete
Next Steps |  | Risks & Issues |  |  |  | 
Notes: Look at having different colour schemes for different pillars