## FORM
Section 1: Introduction
Requester | CHAN MYAE SOE | Date |  | 2020-11-03
Evaluation# |  | Version #00
Category | Others | TMD/ NCR # |  | NA
Affected P/N(s) | 186017000, 186017016 | P/N Description(s) |  | ACQUITY ARC QSM Family
Section 2: Evaluation Assessment
Problem Statement (Describe briefly of the issue. Any impacts to the production capacity, output, resources, or product designs. Include serial numbers and document numbers with revision where applicable. List down potential causes.)
0.3ml/min flow rate accuracy failure was observed at TE82943-1 is improved from 20% to 12% failure rate from Aug'20 to Oct'20 after re-routing the waste tubing. The evaluation is to monitor if failure rate can be further down to 8% which is comparable to the other 5 testers. 20 instruments data to be collected at TE82943-1 to understand the failure rate and symptoms For this activity, production output and resources are not affected The evaluation shall be completed if the overall failure rate from Aug'20 to end of this evaluation is <=8%
Evaluation Plan (Actions must be listed in a sequential manner.)
Action Owner | Action Description |  |  |  | Expected Completion Date
Chan | 1) To collect 0.3ml/min flowrate accuracy data of 20 instruments from TE82943-1 2) After item1, compile the data to compare with other stations with equate test numbers 3) Evaluation shall be completed if item 1 failure rate from Aug'20 to end of evalution achieve less than or equal 8% , 4)If failure rate >=8%, swapping weighing balance with another station TE82943-24 to see if failure follows the balance or not. VAF should perform for both TE8294-1 and TE82943-24 if swapping weighing balance. |  |  |  | 2021-01-31
End of Evaluation
Section 3: Evaluation Approvals
| Name / Title |  | Signature |  | Date
CMO Mfg Lead
Waters Mfg Engineer
Waters Quality Engineer

## Sheet1
TMD
First Article
NCR
Failure Analysis
Others