# INSTALLATION QUALIFICATION REPORT

For the Alliance Final Test System #6

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Chan Yee Hwa | Manufacturing Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |
| Michael Cotter | Principal Test Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Senior Operations Quality Engineer |  |  |

##### The “Prepared By” signature denotes that this document accurately describes the qualification activities.

##### The “Reviewed By” signature[s] denotes agreement with the description of the qualification activities.

##### The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of Change |
| --- | --- | --- |
| 00 | Chan Yee Hwa | First document for approval |

# Purpose

##### The Alliance Final Test System #6  is located in the WPPL Flextronics Contract Manufacturer Site in Singapore.  The station is used to perform the final functional test of the following part numbers: 186269506 and 186269505.  In the final test, the UUT is connected to the test station and electrical, mechanical, and fluidic performance of the unit is tested.

##### The test station is equipped with a Prime/Purge Module, two (2) Temperature Control Modules, four (4) Column Heater Modules, four (4) 2489 Detectors and an Empower Work Station.

##### This Installation Qualification is being performed as required by VAL001270-06JUN2018-VAF, which can be found in the VAL001270 directory of SAP.

# Scope of Report

##### This Installation Qualification Report details the qualification activities undertaken for the Alliance Final Test System #6  based on the requirements of the Installation Qualification Procedure 730004135, Ver. 00.

##### This Installation Qualification Report defines the documents used to qualify the Alliance Final Test System #6 .

# Summary of Results - Statement of Fitness for Intended Use

##### The Alliance Final Test System #6  was tested in accordance with the Installation Qualification Procedure 730004135, Ver. 00.  All Test Scripts were performed and all objective evidence requirements have been met.

##### Waters now has the necessary documentation in place to ensure that the Alliance Final Test System #6  complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

##### Following the approval of this Qualification Report, the Alliance Final Test System #6  can be considered fully qualified and fit for intended use.

# Glossary

##### The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| CTQ | Critical to Quality |
| EQID | Equipment Identification |
| Hz | Hertz |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| N/A | Not Applicable |
| P/N | Part Number |
| SAP | Enterprise Resource Planning Software |
| S/N | Serial Number |
| TSP | Test Specification and Procedure |
| UUT | Unit Under Test |
| VAC | Volts Alternating Current |
| VAF | Validation Assessment Form |

# Training

##### The operator of the Alliance Final Test System #6  was trained according to the Alliance Separations Module / Altus A10 Solvent and Sample Module, TSP WAT271129, REV. 33, VER. 35.  The executor of this qualification was trained according to the Installation Qualification Procedure 730004135, Ver. 00.  Operator and executor training records are provided as Attachment 7.1 and 7.2, respectively.

# Qualification Results

## Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.3.

#### The Alliance Separations Module / Altus A10 Solvent and Sample Module, TSP WAT271129, REV. 33, VER. 35 was confirmed to be in SAP with a status of Release General, as evidenced below:

## Test Script IQ-02: Site Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.4.

#### The ambient temperature and humidity were measured using an Hisamatsu TN-2500 Thermohygrograph, EQID-0418 and were determined to be within specification as defined by the Alliance 2695 Separations Module Operator’s Guide 721008073 VER. 00.  The calibration record of the Thermohygrograph is provided as Attachment 7.5.

#### The electrical supply was measured using a Fluke 179 Multimeter, EQID-0246, and was determined to be within specification as defined by the Alliance 2695 Separations Module Operator’s Guide 721008073 VER. 00.  The calibration record is provided as Attachment 7.6.

#### The Test Station was confirmed to have more than the thirteen (13) dedicated electrical outlets required.

#### The space and test bench of the Alliance Final Test System #6  were confirmed to have adequate space for testing and service.  The bench was confirmed to be large and strong enough to support all test equipment.  A picture of the Test Station work area is provided as Attachment 7.7.

## Test Script IQ-03: Instrumentation Calibration Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed script is provided as Attachment 7.8.

#### Each component of The Alliance Final Test System #6 was confirmed to be in either the Calibration Manager Database or the Preventive Maintenance Database with a twelve (12) month maintenance cycle.  Screen captured from both the databases are provided as Attachment 7.9 and 7.10.

#### Note that since this is a newly installed test system, calibration or preventive maintenance has not yet been performed. The calibration dates listed in the Calibration Database reflects date of production.

## Test Script IQ-04: Software and Instrument Firmware Version was not performed as part of this qualification.

## The Alliance Final Test System #6   does not have custom software and does not download instrument firmware.

## Test Script IQ-05: General Safety and Alarms was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.11.

#### It was confirmed that there were no mechanical safety pinch points or sharp edges.

#### An electrical safety test was performed using a Chroma 19032 Electrical Safety Analyzer, EQID-1492.  Each component of The Alliance Final Test System #6  was confirmed to pass an Electrical Safety Hi-Pot/Ground Continuity test.  The calibration record is provided as Attachment 7.12.

#### It was confirmed that the glass bottles in use at the test station were secured.

#### It was confirmed that the waste tanks in use at the test station were secured.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.13, Attachment 7.15 and Attachment 7.17 for test port #1, #2 and #3 respectively.

#### The e2695 with Heater/Cooler (P/N: 186269506, S/N:  F18SM4615A) was chosen as the UUT for this Installation Qualification because it verifies the full functionality of the Alliance Final Test System #6 .

#### One (1) UUT (P/N: 186269506, S/N:  F18SM4615A) was tested on each available test port of the Alliance Final Test System #6 with all tests passing.  The passing test record is provided as Attachment 7.14, Attachment 7.16 and Attachment 7.18 for test port #1, #2 and #3 respectively.

## This table provides a summary of each test script as performed according to 730004135, Ver. 00 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachment 7.3 | None |
| IQ-02 | Pass | Attachments 7.4, 7.5, 7.6, and 7.7 | None |
| IQ-03 | Pass | Attachments 7.8, 7.9, and 7.10 | None |
| IQ-05 | Pass | Attachments 7.11and 7.12 | None |
| IQ-06 | Pass | Attachments 7.13, 7.14, 7.15,7.16, 7.17 and 7.18 | None |

# Attachments

## Completed Operator Training Record TSP WAT271129, REV. 33, VER. 35

## Completed Executor Training Record 730004135, Ver. 00

## Completed Test Script IQ-01: Associated Document Verification

## Completed Test Script IQ-02: Site Verification

## Objective Evidence IQ-02: Hisamatsu TN-2500 Thermohygrograph (EQID-0418) Calibration Record

## Objective Evidence IQ-02: Fluke 179 Multimeter (EQID-0246) Calibration Record

## Objective Evidence IQ-02: Picture of Test Station Work Area

## Completed Test Script IQ-03: Instrument Calibration Verification

## Objective Evidence IQ-03: Screen Capture of Alliance Final Test System #6  in Calibration Master Database

## Objective Evidence IQ-03: Screen Capture of Alliance Final Test System #6  in Preventive Maintenance Master Database

## Completed Test Script IQ-05: General Safety and Alarms

## Objective Evidence IQ-05: Chroma 19032 Electrical Safety Analyzer (EQID-1492) Calibration Record

## Completed Test Script IQ-06: Functional Verification of UUT Part #186269506 (S/N:  F18SM4615A) at Port #1

## Objective Evidence IQ-06: Passing Test Record for UUT Part #186269506 (S/N:  F18SM4615A) at Port #1

## Completed Test Script IQ-06: Functional Verification of UUT Part #186269506 (S/N:  F18SM4615A) at Port #2

## Objective Evidence IQ-06: Passing Test Record for UUT Part #186269506 (S/N:  F18SM4615A) at Port #2

## Completed Test Script IQ-06: Functional Verification of UUT Part #186269506 (S/N:  F18SM4615A) at Port #3

## Objective Evidence IQ-06: Passing Test Record for UUT Part #186269506 (S/N:  F18SM4615A) at Port #3

# Deviation from Test Equipment Qualification

##### There were no deviations generated during the execution of the Installation Qualification Procedure 730004135, Ver. 00.

# References

##### The current approved versions of the following documents are referenced within this report:

| Document Number | Document Title |
| --- | --- |
| 730004135, Ver. 00 | Installation Qualification Procedure |
| TSP WAT271129, REV. 33, VER. 35 | Alliance Separations Module / Altus A10 Solvent and Sample Module |
| VAL001270-06JUN2018-VAF | Validation Assessment Form for Alliance Test System |
| WAT046538, Ver. 09 | Generic Electrical Safety Test Procedure |
| 721008073, VER. 00 | Alliance 2695 Separations Module Operator’s Guide |