# INSTALLATION QUALIFICATION REPORT

For the Alliance Final Test System #6

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Chan Yee Hwa | Manufacturing Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |
| Michael Cotter | Principal Test Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Senior Operations Quality Engineer |  |  |

##### The “Prepared By” signature denotes that this document accurately describes the qualification activities.

##### The “Reviewed By” signatures denote agreement with the description of the qualification activities.

##### The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of Change |
| --- | --- | --- |
| 00 | Chan Yee Hwa | First document for approval |
| 01 | Chan Yee Hwa | Add fourth ball valve to Alliance Final Test System #6 |

# Purpose

The Alliance Final Test System #6  is located in the WPPL Flextronics Contract Manufacturer Site in Singapore.  The station is used to perform the final functional test of the following part numbers: 186269506 and 186269505.  In the final test, the UUT is connected to the test station and electrical, mechanical, and fluidic performance of the unit is tested.

##### The test station is equipped with a Prime/Purge Module, two (2) Temperature Control Modules, four (4) Column Heater Modules, four (4) 2489 Detectors and an Empower Work Station.  A fourth ball valve is being installed as part of the Prime/ Purge Module of the Alliance Final Test System #6 at the Flextronics Contract Manufacturer Site in Singapore; this Installation Qualification is being performed as required by VAL001270-05JUL18-VAF, which can be found in the VAL001270 directory of SAP.

# Scope of Report

##### This Installation Qualification Report details the qualification activities undertaken for the Alliance Final Test System #6 based on the requirements of the Installation Qualification Procedure 730004135, Ver. 00.

##### This Installation Qualification Report defines the documents used to qualify the Alliance Final Test System #6.

# Summary of Results - Statement of Fitness for Intended Use

##### The fourth port of Alliance Final Test System #6 was tested in accordance with the Installation Qualification Procedure 730004135, Ver. 00.  All required Test Scripts were performed and all objective evidence requirements have been met.

##### Waters now has the necessary documentation in place to ensure that the fourth port of Alliance Final Test System #6 complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

##### Following the approval of this Qualification Report, the fourth port of Alliance Final Test System #6 can be considered fully qualified and fit for intended use.

# Glossary

##### The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| CTQ | Critical to Quality |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| N/A | Not Applicable |
| P/N | Part Number |
| SAP | Enterprise Resource Planning Software |
| S/N | Serial Number |
| TSP | Test Specification and Procedure |
| UUT | Unit Under Test |
| VAF | Validation Assessment Form |
| WPPL | Waters Pacific Private Limited |

# Training

##### The operator of the Alliance Final Test System #6 was trained according to the Alliance Separations Module / Altus A10 Solvent and Sample Module, TSP WAT271129, REV. 33, VER. 35.  The executor of this qualification was trained according to the Installation Qualification Procedure 730004135, Ver. 00.  Operator and executor training records are provided as Attachment 7.1 and 7.2, respectively.

# Qualification Results

## Test Script IQ-01: Associated Documentation Verification, Test Script IQ-02: Site Verification, Test Script IQ-03: Instrumentation Calibration Verification, Test Script IQ-04: Software and Instrument Firmware Version, and Test Script IQ-05: General Safety and Alarms were not performed as part of this qualification according to VAL001270-05JUL18-VAF.

## All result can be found in the previously executed VAL001270-IQR_Iteration 00, which can be found in VAL001270 directory of SAP. This remains a valid qualification for Alliance Final Test System #6.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.3.

#### The e2695 with Heater/Cooler (P/N: 186269506, S/N:  G18SM4074A) was chosen as the UUT for this Installation Qualification because it verifies the full functionality of the Alliance Final Test System #6.

#### One (1) UUT (P/N: 186269506, S/N:  G18SM4074A) was tested on fourth test port of the Alliance Final Test System #6with all tests passing.  The passing test record is provided as Attachment 7.4.

## This table provides a summary of each test script as performed according to 730004135, Ver. 00 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-06 | Pass | Attachments 7.3 and 7.4 | None |

# Attachments

## Completed Operator Training Record TSP WAT271129, REV. 33, VER. 35

## Completed Executor Training Record 730004135, Ver. 00

## Completed Test Script IQ-06: Functional Verification of UUT Part #186269506 (S/N:  G18SM4074A) at Port #4

## Objective Evidence IQ-06: Passing Test Record for UUT Part #186269506 (S/N:  G18SM4074A) at Port #4

# Deviation from Test Equipment Qualification

##### There were no deviations generated during the execution of the Installation Qualification Procedure 730004135, Ver. 00.

# References

##### The current approved versions of the following documents are referenced within this report:

| Document Number | Document Title |
| --- | --- |
| 730004135, Ver. 00 | Installation Qualification Procedure |
| TSP WAT271129, REV. 33, VER. 35 | Alliance Separations Module / Altus A10 Solvent and Sample Module |
| VAL001270-05JUL18-VAF | Validation Assessment Form for Alliance Test System |
| VAL001270-IQR_Iteration 00 | Installation Qualification Report for VAL001270 |