# Validation Assessment

For the Waters Global Test System

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Rick Davis | System Owner |  |  |
| Emma Pierce | Test Validation Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Quality Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the planned Process Validation activities.

The “Reviewed By” signatures denote agreement with the details of the planned testing and/or deliverables to be performed in order to qualify the Waters Global Test System (GTS).

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

# Purpose

The purpose of the Validation Assessment Form (VAF) is to document the assessment of an engineering change to the released Waters Global Test System (TE0082988). This document shall also define the deliverable(s) required to maintain the equipment qualification, based on the engineering change. Qualification of the equipment shall ensure that it functions within its specifications.

# Background

The Waters Global Test System (GTS) is a general purpose test platform to be used together with a group of test fixtures for the different PCBA functionality test. The system is equipped with Power Distribution Units, Computer, Ethernet Switch, PXI Chassis, Oscilloscope, Multimeter, Function Generator, Load and Power Supply Units. The UUT will be connected on test fixture, the test fixture together with UUT will be tested using automated test software. The test report will be automatically generated upon completion of test.

# Description of Change

## A new duplicate test system, GTS TE82988-12, is being installed at the Venture Contract Manufacturer Site in Singapore to increase production capacity.  This does not affect the existing Waters Global Test Systems or the system controls and test methods.

# Risk assessment

## High Risk Assessment

## Please check the appropriate boxes below to determine the high risk assessed qualification or requalification requirements:

| ### Specification | ## No | ## Yes |
| --- | --- | --- |
| Is the change a test specification value being modified and/or added? |  |  |

| ### Software/Firmware | ## No | ## Yes |
| --- | --- | --- |
| Is the change a software modification that affects the way the system controls, and/or measures a test specification and/or PSC result? |  |  |
| Is the change due to firmware downloaded to the UUT by a test system? |  |  |
| Is the change a flash file update? |  |  |
| Is it a change to an existing, or the addition of a new, PCB Test Sequence? |  |  |

| ### Hardware (Mechanical, Electrical, Pneumatic, Hydraulic, etc.) | ## No | ## Yes |
| --- | --- | --- |
| Is the change a mechanical, electrical, pneumatic, and/or hydraulic modification that alters the way the system controls, and/or measures a test specification and/or PSC result? |  |  |

| ### Equipment/Process Moves | ## No | ## Yes |
| --- | --- | --- |
| Is the change a move of the equipment, test method or process, to a different facility? |  |  |

## Medium Risk Assessment

## Please check the appropriate boxes below to determine the medium risk assessed qualification or requalification requirements:

| ### Duplicate System or Process | ## No | ## Yes |
| --- | --- | --- |
| Is the change due to a duplicate system, test method, or process? |  |  |
| ### Replacement of Non-Identical | ## No | ## Yes |
| Is the change due to a replacement of a system component with a non-identical equivalent? |  |  |

| ### Computer Related or Replacement of Custom PCBs | ## No | ## Yes |
| --- | --- | --- |
| Is the change an update to the PC operating system? |  |  |
| Is the change a computer or printed circuit board replacement? |  |  |

| ### Software/Firmware | ## No | ## Yes |
| --- | --- | --- |
| Is the change a minor software bug fix or a minor improvement (i.e. a part number add)? |  |  |
| Is the change a test system instrument firmware update (i.e. update to a stack component)? |  |  |

| ### Equipment/Process Moves | ## No | ## Yes |
| --- | --- | --- |
| Is the change a move of the equipment, test method or process, within the Milford site? |  |  |

## This change results in a High, Medium, or Low Risk Assessment and the associated qualification or requalification requirements, as detailed in Section 6.0.

# Hazard Assessment:

No hazards are potentially introduced to the environment, operator, assembler, service technician or end user.

# verification, Qualification and/or validation:

| ## Equipment Design Verification (EDV) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does an EDV Protocol need to be created or revised due to this change? |  |  |
| Based on the Risk Assessment, does an EDV Protocol need to be executed due to this change? (Justify partial execution in the Comments/Justification section below.) |  |  |
| Comments/Justification The system is classified as fit for use. This change does not affect the test method or test system, therefore an EDV is not required. |  |  |

| ## Installation Qualification (IQ) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does an IQ Protocol need to be executed due to this change? (Justify partial execution in the Comments/Justification section below.) |  |  |
| Comments/Justification The Waters Global Test System will be qualified using the pre-approved Qualification Protocol 735004235. The other duplicate systems are classified as fit for use and are not affected by this change. |  |  |

| ## Test Method Validation (TMV) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does a TMV need to be executed due to this change? (Justify partial execution in the Comments/Justification section below.) |  |  |
| Comments/Justification This test equipment does not utilize an external measurement device and is not subject to operator induced variation. This change does not require a TMV as defined in the Process Validation Procedure 730001835. |  |  |
| ## Operational Qualification (OQ) | ## No | ## Yes |
| Based on the Risk Assessment, does an OQ Protocol need to be created or revised due to this change? |  |  |
| Based on the Risk Assessment, does an OQ need to be executed due to this change? (Justify partial execution in the Comments/Justification section below.) |  |  |
| Comments/Justification OQ is not required because 100% of critical to quality (CTQ) attributes are verified. |  |  |

| ## Performance Qualification (PQ) | ## No | ## Yes |
| --- | --- | --- |
| Based on the Risk Assessment, does a PQ Protocol need to be created or revised due to this change? |  |  |
| Based on the Risk Assessment, does a PQ need to be executed due to this change? (Justify partial execution in the Comments/Justification section below.) |  |  |
| Comments/Justification PQ is not required because 100% of critical to quality (CTQ) attributes are verified. |  |  |

## Electromagnetic Interference (EMI) Assessment:

This change only affects test equipment; there is no impact to product EMI.

# IMPLEMENTATION ASSESSMENT:

- The new equipment will entered into the onsite Venture Calibration Database. There is no impact to documentation or facilities/utilities.

- The qualification of this equipment has a high impact on production schedules, as the increased capacity will enable production to better meet production goals.

- The equipment will be installed by Venture personnel. The IQ will be executed by Venture personnel according to the pre-approved installation Qualification 735004135.

- No additional training will be required.

# Validation Assessment
Objective Evidence Completion

For the Waters Global Test System

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

REVIEWED BY (minimum of one reviewer, additional reviewers may be added as applicable)

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Rick Davis | System Owner |  |  |
| Emma Pierce | Test Validation Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Quality Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the planned Process Validation activities.

The “Reviewed By” signatures denote agreement with the details of the planned testing and/or deliverables to be performed in order qualify the Waters Global Test System.

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.