# INSTALLATION QUALIFICATION REPORT

For the Waters Global Test System (GTS)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Edwin Koh | Senior Operations Manager |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Kah Chin Tan | Senior Quality Assurance Manager |  |  |

##### The “Prepared By” signature denotes that this document accurately describes the qualification activities.

##### The “Reviewed By” signature denotes agreement with the description of the qualification activities.

##### The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Luo HuaZhen | First document for approval |

# Purpose

##### The Waters Global Test System (GTS) (TE82988-12) is located in the Venture Contract Manufacturer Site in Singapore.  The system is a general purpose test platform to be used together with a group of test fixtures for the different PCBA functionality test. The system is equipped with Power Distribution Units, Computer, Ethernet Switch, PXI Chassis, Oscilloscope, Multimeter, Function Generator, Load and Power Supply Units.

##### Once the UUT is connected together with test fixture, the operator selects the appropriate part number from the test software and begins the test.  When the test is complete, a test record is automatically generated and test data is uploaded to SigmaSure server.

##### This Installation Qualification is being performed as required by the VMP Submission Form TE0082988-12 VMP Submission Form which can be found in the TE0082988QX directory of SAP.

# Scope of Report

##### This Installation Qualification Report details the qualification activities undertaken for the Waters Global Test System (GTS) (TE82988-12) based on the requirements of the Installation Qualification Procedure 730004135, Ver. 02 and TE0082988-12 VMP Submission Form.

##### This Installation Qualification Report defines the documents used to qualify the Waters Global Test System (GTS) (TE82988-12).

# Summary of Results - Statement of Fitness for Intended Use

##### The Waters Global Test System (GTS) was tested in accordance with the Installation Qualification Procedure 730004135, Ver. 02 and TE0082988-12 VMP Submission Form.  All required Test Scripts were performed and all objective evidence requirements have been met.

##### Waters now has the necessary documentation in place to ensure that the Waters Global Test System (GTS) complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

##### Following the approval of this Qualification Report, the Waters Global Test System (GTS) (TE82988-12) can be considered qualified and fit for intended use.

# Glossary

##### The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| GTS | Global Test System |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| N/A | Not Applicable |
| P/N | Part Number |
| S/N | Serial Number |
| PCBA | Printed Circuit Board Assembly |
| SAP | Enterprise Resource Planning Software |
| UUT | Unit Under Test |
| WPPL | Waters Pacific Private Limited |
| VAC | Volts Alternating Current |
| VMP | Validation Master Plan |

# Training

##### The operator of the Waters Global Test System (GTS) (TE82988-12) was trained according to the Waters UK PCB 210001115 test procedure 50593, Ver. 03.  The executor of this qualification was trained according to the Installation Qualification Procedure 730004135, Ver. 02.  Both training records are provided as Attachments 7.1 and 7.2, respectively.

# Qualification Results

## Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Procedure 730004135, Ver. 03. The completed Test Script Data Sheet is provided as Attachment 7.3.

## 6.1.1 The Waters UK test procedure 50593, Ver.03 was confirmed to be in SAP with a status of Ready, as evidenced below:

## Test Script IQ-02: Site Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02. The completed Test Script Data is provided as attachment 7.4.

#### The ambient temperature and humidity were measured using a themohygraph and were determined to be within specification as defined. The calibration record is provided as attachment 7.5.

#### The electrical supply was measured using a Multimeter and was determined to be within specification as defined. The calibration record is provided as attachment 7.6

#### The Test Station was confirmed to have a dedicated power supply.

#### The space and test station were confirmed to have adequate space for testing and service. A picture of the Test Station work area is provided as attachment 7.7.

## Test Script IQ-03: Instrumentation Calibration Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02.  The completed script is provided as Attachment 7.8

#### Each component of test station was confirmed to be in either the Calibration Manager Database or the Preventive Maintenance Database with a twelve month maintenance cycle. Screen captured from the databases are provided as Attachment 7.9 and 7.10.

#### Note that since this is a newly installed test system, calibration or preventive maintenance has not yet been performed. The calibration dates listed in the Calibration Database reflects date of production.

## Test Script IQ-04: Software and Instrument Firmware Version was not performed according to TE0082988-12 VMP Submission Form.

## Test Script IQ-05: General Safety and Alarms was performed according to Installation Qualification Procedure 730004135, Ver. 03.  The completed script is provided as Attachment 7.11.

#### It was confirmed that there were no mechanical safety pinch points or sharp edges.

#### An electrical safety test was performed. Each component of test station was confirmed to pass.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00 and TE0082988-12 VMP Submission Form.  The completed Test Script Data Sheets are provided as Attachments 7.4.

#### One (1) GTS Self-Test Fixture TE83004 was tested using the Waters Global Test System (GTS) (TE82988-12) with all tests passing.  The passing test record is provided as Attachment 7.5.

#### One (1) UUT (P/N: 210001115, S/N: Gxxxxxxxxx) was tested using the Waters Global Test System (GTS) (TE82988-12) with all tests passing.  The passing test record is provided as Attachment 7.6.

## This table provides a summary of each test script as performed according to 730004135, Ver. 02 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 |  |  |  |
| IQ-02 |  |  |  |
| IQ-03 |  |  |  |
| IQ-05 |  |  |  |
| IQ-06 |  |  |  |

# Attachments

## Completed Operator Training Record 50593, Ver. 03

## Completed Executor Training Record 730004135, Ver. 02

## Completed Test Script IQ-01: Associated Document Verification

## Completed Test Script IQ-02: Site Verification

## Objective Evidence IQ-02

## Completed Test Script IQ-03: Instrument Calibration Verification

## Objective Evidence IQ-03:

## Completed Test Script IQ-04: General Safety and Alarms

## Objective Evidence IQ-04

## Completed Test Script IQ-05: General Safety and Alarms

## Objective Evidence IQ-05

## Completed Test Script IQ-06: Functional Verification

## Objective Evidence IQ-06: Passing Test Record for Self-Test Fixture TE83004

## Objective Evidence IQ-06: Passing Test Record for UUT Part # 210001115 (S/N: Gxxxxxxxx)

# Deviation from Test Equipment Qualification

##### There were no deviations generated during the execution of the Installation Qualification Procedure 730004135, Ver. 02.

# References

##### The current approved versions of the following documents are referenced within this report:

| Document Number | Document Title |
| --- | --- |
| 730004135, Ver. 02 | Installation Qualification Procedure |
| 50593, Ver. 03 | Waters UK PCB 210001115 test procedure |
| ###### TE0082988QX | Validation Assessment Form TE0082988-12 VMP Submission Form |