# INSTALLATION QUALIFICATION REPORT

For the Waters Global Test System (GTS)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Edwin Koh | Senior Operations Manager |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Kah Chin Tan | Senior Quality Assurance Manager |  |  |

##### The “Prepared By” signature denotes that this document accurately describes the qualification activities.

##### The “Reviewed By” signature denotes agreement with the description of the qualification activities.

##### The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Luo HuaZhen | First document for approval |

# Purpose

##### The Waters Global Test System (GTS) (TE82988-12) is located in the Venture Contract Manufacturer Site in Singapore.  The system is a general purpose test platform to be used together with a group of test fixtures for the different PCB functionality test. The system is equipped with Power Distribution Units, Computer, Ethernet Switch, PXI Chassis, Oscilloscope, Multimeter, Function Generator, Load and Power Supply Units.

##### Once the UUT is connected together with test fixture, the operator selects the appropriate part number from the test software and begins the test.  When the test is complete, a test record is automatically generated and test data is uploaded to SigmaSure server.

##### This Installation Qualification is being performed as required by the VMP Submission Form TE0082988-12 VMP Submission Form which can be found in the TE008298812 directory of SAP.

# Scope of Report

##### This Installation Qualification Report details the qualification activities undertaken for the Waters Global Test System (GTS) (TE82988-12) based on the requirements of the Installation Qualification Procedure 730004135, Ver. 02 and TE0082988-12 VMP Submission Form.

##### This Installation Qualification Report defines the documents used to qualify the Waters Global Test System (GTS) (TE82988-12).

# Summary of Results - Statement of Fitness for Intended Use

##### The Waters Global Test System (GTS) was tested in accordance with the Installation Qualification Procedure 730004135, Ver. 02 and TE0082988-12 VMP Submission Form.  All required Test Scripts were performed and all objective evidence requirements have been met.

##### Waters now has the necessary documentation in place to ensure that the Waters Global Test System (GTS) complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

##### Following the approval of this Qualification Report, the Waters Global Test System (GTS) (TE82988-12) can be considered qualified and fit for intended use.

# Glossary

##### The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| % | Percent |
| # | Number |
| ºC | Degrees Celsius |
| BOM | Bill of Materials |
| Cal | Calibration |
| cGMP | Current Good Manufacturing Practice |
| CTQ | Critical to Quality |
| Doc. | Document |
| EH&S | Environmental Health and Safety |
| EQID | Equipment Identification |
| GDP | Good Documentation Practices |
| GTS | Global Test System |
| Hz | Hertz |
| ID | Identification |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| N/A | Not Applicable |
| NI | National Instrument |
| PCB | Printed Circuit Board |
| P/N | Part Number |
| PXI | PCI eXtensions for Instrumentation |
| PCI | Peripheral Component Interconnect |
| S/N | Serial Number |
| PM | Preventive Maintenance |
| PCBA | Printed Circuit Board Assembly |
| SAP | Enterprise Resource Planning Software |
| TSP | Test and Specification Procedure |
| UUT | Unit Under Test |
| WPPL | Waters Pacific Private Limited |
| VAC | Volts Alternating Current |
| VMP | Validation Master Plan |
| VAF | Validation Assessment Form |

# Training

##### The operator of the Waters Global Test System (GTS) (TE82988-12) was trained according to the Waters UK PCB 210001115 test procedure 50593.  The executor of this qualification was trained according to the Installation Qualification Procedure 730004135, Ver. 02.  Both training records are provided as Attachments 7.1 and 7.2, respectively.

# Qualification Results

## Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02. The completed Test Script Data Sheet is provided as Attachment 7.3.

## 6.1.1 The 210001115 KESTEREL DCC(Kanye) Programming and Test Instructions, Rev. 03, Ver. 03 was confirmed to be in SAP with a status of Release General, as evidenced below:

## 6.1.2 The Waters GTS Preventative Maintenance and Calibration Procedure 735004006, Rev. A, Ver. 00 was confirmed to be in SAP with a status of Release General, as evidenced below:

## Test Script IQ-02: Site Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02. The completed Test Script Data is provided as Attachment 7.4.

## 6.2.1 The ambient temperature and humidity were measured using a Hygrometer and were determined to be within specification as defined. The measurement record and the calibration record are provided as Attachment 7.5.

## 6.2.2 The electrical supply was measured using a Multimeter and was determined to be within specification as defined. The measurement record and the calibration record are provided as Attachment 7.6.

## 6.2.3 The Test Station was confirmed to have a dedicated power supply.

## 6.2.4 The space and test station were confirmed to have adequate space for testing and service. A picture of the Test Station work area is provided as Attachment 7.7.

## 6.2.5 The vacuum is used to make mechanical and electrical connections between the UUT and fixture/fixture probes is confirmed has minimum vacuum of 23 inches Hg(78kPa). A picture of vacuum gauge with readings is provided as Attachment 7.8.

## Test Script IQ-03: Instrumentation Calibration Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02.  The completed script is provided as Attachment 7.9

## 6.3.1 All calibration required equipment have the certificate of the calibration as Attachment 7.10

## Test Script IQ-04: Software and Instrument Firmware Version was performed according to Installation Qualification Procedure 730004135, Ver. 02. The completed script is provided as Attachment 7.11

## 6.4.1 The Test Equipment Software version listed in 210001115 KESTEREL DCC(Kanye) Programming and Test Instructions, Rev. 03, Ver. 03 is Test Software Rev C and was verified to match the version listed on the test equipment program display.  A screen capture of the program display is provided as Attachment 7.12.

## 6.4.2 The Test Equipment self-test Software version listed in Waters GTS Preventative Maintenance and Calibration Procedure 735004006, Rev. A, Ver. 00 is NI TestStand 4.2.1 was verified to match the version listed on the test equipment program display.  A screen capture of the program display is provided as Attachment 7.13.

## Test Script IQ-05: General Safety and Alarms was performed according to Installation Qualification Procedure 730004135, Ver. 02.  The completed script is provided as Attachment 7.14.

## 6.5.1 It was confirmed that there were no mechanical safety pinch points or sharp edges.

## 6.5.2 It was confirmed that there were Safety Labels, Emergency Stop and Guarding equipped on this system.  A screen capture of the program display is provided as Attachment 7.15.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02 and TE0082988-12 VMP Submission Form.  The completed Test Script Data Sheets are provided as Attachments 7.16.

## 6.6.1 One (1) GTS Self-Test Fixture TE83004 was tested using the Waters Global Test System (GTS) (TE82988-12) with all tests passing.  A screen capture of test passing is provided as Attachment 7.17.

## 6.6.2 One (1) UUT (P/N: 210001115, S/N: 21/066) was tested using the Waters Global Test System (GTS) (TE82988-12) with all tests passing.  A screen capture of test passing is provided as Attachment 7.18. The passing test record was uploaded to Waters SigmaSure server and it’s provided as Attachment 7.19.

## This table provides a summary of each test script as performed according to 730004135, Ver. 02 along with the Attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachments 7.3 | None |
| IQ-02 | Pass | Attachments 7.4, 7.5, 7.6, 7.7 and 7.8 | None |
| IQ-03 | Pass | Attachments 7.9 and 7.10 | None |
| IQ-04 | Pass | Attachments 7.11, 7.12 and 7.13 | None |
| IQ-05 | Pass | Attachments 7.14 and 7.15 | None |
| IQ-06 | Pass | Attachments 7.16, 7.17, 7.18 and 7.19 | None |

# Attachments

## Completed Operator Training Record 210001115 KESTEREL DCC(Kanye) Programming and Test Instructions, Rev. 03, Ver. 03.

## Completed Executor Training Record 730004135, Ver. 02

## Completed Test Script IQ-01: Associated Document Verification

## Completed Test Script IQ-02: Site Verification

## Objective Evidence IQ-02: The measurement record and the calibration record of Hygrometer

## Objective Evidence IQ-02: The measurement record and the calibration record of FLUKE Multimeter

## Objective Evidence IQ-02: The pictures of the work space and test station bench

## Objective Evidence IQ-02: The pictures of the Vacuum Gauge with reading

## Completed Test Script IQ-03: Instrument Calibration Verification

## Objective Evidence IQ-03: The certificate of calibration for all calibration required equipment

## Completed Test Script IQ-04: Software and Instrument Firmware Version verification

## Objective Evidence IQ-04: Screen capture of the test program version

## Objective Evidence IQ-04: Screen capture of the GTS self-test program version

## Completed Test Script IQ-05: General Safety and Alarms

## Objective Evidence IQ-05: The pictures of System Safety Labels, Emergency Stop and Guarding

## Completed Test Script IQ-06: Functional Verification

## Objective Evidence IQ-06: Screen capture of passing test for Self-Test Fixture TE83004

## Objective Evidence IQ-06: Screen capture of passing test for UUT Part # 210001115 (S/N: 21/066)

## Objective Evidence IQ-06: The passing test record from SigmaSure Server for UUT Part # 210001115 (S/N: 21/066)

# Deviation from Test Equipment Qualification

##### There were no deviations generated during the execution of the Installation Qualification Procedure 730004135, Ver. 02.

# References

##### The current approved versions of the following documents are referenced within this report:

| Document Number | Document Title |
| --- | --- |
| 730004135 | Installation Qualification Procedure |
| 50593 | Waters UK PCB 210001115 test procedure |
| ###### TE008298812VMP | Validation Assessment Form TE0082988-12 VMP Submission Form |
| ###### 735004006 | Waters GTS Preventative Maintenance and Calibration Procedure |