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Purpose
The Waters Global Test System (GTS) for UK Waters PCB Testing (TE82988-12) is located in the Venture Contract Manufacturer Site in Singapore.  The system is a general purpose test platform to be used together with a group of test fixtures for the different PCBA functionality test. The system is equipped with Power Distribution Units, Computer, Ethernet Switch, PXI Chassis, Oscilloscope, Multimeter, Function Generator, Load and Power Supply Units. 
Once the UUT is connected together with test fixture, the operator selects the appropriate part number from the test software and begins the test.  When the test is complete, a test record is automatically generated and test data is uploaded to SigmaSure server. 
This Installation Qualification is being performed as required by the Validation Assessment Form TE0082988-07AUG2018-VAF which can be found in the TE0082988QX directory of SAP.
[bookmark: _Toc442432503]Scope of Report
This Installation Qualification Report details the qualification activities undertaken for the Waters Global Test System (GTS) for UK Waters PCB Testing (TE82988-12) based on the requirements of the Installation Qualification Procedure 730004135, Ver. 02 and TE0082988-07AUG2018-VAF.
This Installation Qualification Report defines the documents used to qualify the Waters Global Test System (GTS) for UK Waters PCB Testing (TE82988-12).
[bookmark: _Toc442432509]Summary of Results - Statement of Fitness for Intended Use
The Waters Global Test System (GTS) for UK Waters PCB Testing was tested in accordance with the Installation Qualification Procedure 730004135, Ver. 02 and TE0082988-07AUG2018-VAF.  All required Test Scripts were performed and all objective evidence requirements have been met.
Waters now has the necessary documentation in place to ensure that the Waters Global Test System (GTS) for UK Waters PCB Testing complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.
Following the approval of this Qualification Report, the Waters Global Test System (GTS) for UK Waters PCB Testing (TE82988-12) can be considered qualified and fit for intended use.
Glossary
The following definitions are used in this document:
	Term
	Explanation

	cGMP
	Current Good Manufacturing Practice

	GTS
	Global Test System

	IQ
	Installation Qualification

	IQR
	Installation Qualification Report

	N/A
	Not Applicable

	P/N
	Part Number

	S/N
	Serial Number

	PCBA
	Printed Circuit Board Assembly

	SAP
	Enterprise Resource Planning Software

	UUT
	Unit Under Test

	VAF
	Validation Assessment Form

	WPPL
	Waters Pacific Private Limited

	VAC
	Volts Alternating Current


[bookmark: _Toc442432507]Training
[bookmark: _Toc442432506][bookmark: _GoBack]The operator of the Waters Global Test System (GTS) for UK Waters PCB Testing (TE82988-12) was trained according to the Venture Functional Test Procedure for W1210000478 MI-WAT-WAT-099, Ver. 00 and 210001115 Kestrel DCC(Kanye) Programming and Test Instructions, MI-WAT-QDA-047, Ver. 00.  The executor of this qualification was trained according to the Installation Qualification Procedure 730004135, Ver. 02.  Both training records are provided as Attachments 7.1 and 7.2, respectively.
Qualification Results
Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02. The completed Test Script Data Sheet is provided as Attachment 7.3.  
Test Script IQ-02: Site Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02. The completed Test Script Data is provided as attachment 7.4.
The ambient temperature and humidity were measured using a themohygraph and were determined to be within specification as defined. The calibration record is provided as attachment 7.5.
The electrical supply was measured using a Multimeter and was determined to be within specification as defined. The calibration record is provided as attachment 7.6
The Test Station was confirmed to have a dedicated power supply.
The space and test station were confirmed to have adequate space for testing and service. A picture of the Test Station work area is provided as attachment 7.7.
Test Script IQ-03: Instrumentation Calibration Verification was performed according to Installation Qualification Procedure 730004135, Ver. 02.  The completed script is provided as Attachment 7.8
Each component of test station was confirmed to be in either the Calibration Manager Database or the Preventive Maintenance Database with a twelve month maintenance cycle. Screen captured from the databases are provided as Attachment 7.9 and 7.10.
Note that since this is a newly installed test system, calibration or preventive maintenance has not yet been performed. The calibration dates listed in the Calibration Database reflects date of production.
Test Script IQ-04: Software and Instrument Firmware Version was not performed according to TE0082988-07AUG2018-VAF. 
Test Script IQ-05: General Safety and Alarms was performed according to Installation Qualification Procedure 730004135, Ver. 02.  The completed script is provided as Attachment 7.11.
It was confirmed that there were no mechanical safety pinch points or sharp edges.
An electrical safety test was performed using a Chroma 19032 Electrical Safety Analyzer, EQID-1492. Each component of test station was confirmed to pass.

Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00 and TE0082988-07AUG2018-VAF.  The completed Test Script Data Sheets are provided as Attachments 7.4.
One (1) UUT (P/N: 210000478, S/N: Gxxxxxxxxx) was tested using the Waters Global Test System (GTS) for UK Waters PCB Testing (TE82988-12) with all tests passing.  The passing test record is provided as Attachment 7.5.
One (1) UUT (P/N: 210001115, S/N: Gxxxxxxxxx) was tested using the Waters Global Test System (GTS) for UK Waters PCB Testing (TE82988-12) with all tests passing.  The passing test record is provided as Attachment 7.6.
This table provides a summary of each test script as performed according to 730004135, Ver. 02 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.
	Script or Section
	Pass/Fail
	Objective Evidence
	Deviations

	IQ-01
	
	
	

	IQ-02
	
	
	

	IQ-03
	
	
	

	IQ-05
	
	
	

	IQ-06
	
	
	


[bookmark: _Toc442432505][bookmark: _Toc309978937][bookmark: _Toc442432504]Attachments
Completed Operator Training Record MI-WAT-WAT-099, Ver. 00 and 210001115 Kestrel DCC(Kanye) Programming and Test Instructions, MI-WAT-QDA-047, Ver.00
Completed Executor Training Record 730004135, Ver. 02
Completed Test Script IQ-01: Associated Document Verification 
Completed Test Script IQ-02: Site Verification 
Objective Evidence IQ-02 
Completed Test Script IQ-03: Instrument Calibration Verification 
Objective Evidence IQ-03: 
Completed Test Script IQ-05: General Safety and Alarms
 Objective Evidence IQ-05
 Completed Test Script IQ-06: Functional Verification
Objective Evidence IQ-06: Passing Test Record for UUT Part # 210000478 (S/N: Gxxxxxxxx)
Objective Evidence IQ-06: Passing Test Record for UUT Part # 210001115 (S/N: Gxxxxxxxx)
Deviation from Test Equipment Qualification
There were no deviations generated during the execution of the Installation Qualification Procedure 730004135, Ver. 02.
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