## flextronics_reviewed
| Correct as of: | 2020-06-05
No. | Product Family | Part Number | Serial No. | FW Version | Date Received
1 | 2489/TUV | 186015022 | K05UPT479M | 1.10.1174 | 2007 | In use for spare bench p/n:
2 | 2489/TUV | 186015028 | A08UPT943M | 1.60.2273 |  | Wavelength Verifcation Failure (Not In Use)(To Scrapped)
3 | 2489/TUV | 186015031 | C10TUV025N | 1.60.2274 | 2012 | In use for spare bench p/n:
4 | 2489/TUV | 186002487 | L0787E095M | 1.40.1966 | 2009 | In use for spare bench p/n:
5 | 2489/TUV | WAT081110 | L03487649M | 1.01 (Spare bench testing only) | 2005
6 | 2489/TUV | WAT081110 | H05487269M | 1.01 (Spare bench testing only) | 2007 | No Display (To Scrapped)
7 | 2489/TUV | WAT081110 | H05487273M | 1.01 (Spare bench testing only) | 2007 | Lamp Power Suppy (To Scrapped)
8 | 2489/TUV | WAT081110 | H05487251M | 1.01 (Spare bench testing only) | 2007 | No Display (To Scrapped)
9 | 2489/TUV | WAT081110 | H05487257M | Non good unit, (test bed – no Optic Bench) |  | (To Scrapped)
10 | 2489/TUV | WAT081110 | G98487535M |  |  | (To Scrapped)
11 | 2998/PDA/eLambda | 186002998 | D08998406N | 1.41.6164 | 2010 | In use for spare bench p/n:
12 | 2998/PDA/eLambda | 186299800 | B13998443A | 1.41.6164 | 2015 | In use for spare bench p/n:
13 | 2998/PDA/eLambda | 186015026 | B08UPD563M | Not Tested in Flex and no test report | 2010 | Not Tested in Flex and no test report
14 | 2998/PDA/eLambda | 186015030 | A09UPL372M | Not Tested in Flex and no test report | 2011 | Not Tested in Flex and no test report
15 | 2998/PDA/eLambda | 186015032 | BETUPD008N | 1.65.6229 | 2012 | In use for spare bench p/n:
16 | 2998/PDA/eLambda | 186015034 | G13C2P293N | Not Tested in Flex and no test report | 2015 | In use for spare bench p/n:
17 | 2489 UV/Visible Detector | 186248900 | K1787E723A | 1.69.2682 | 2017
18 | 2490 UV/Visible Detector | 186248900 | K1787E724A | 1.69.2682 | 2017
19 | 2491 UV/Visible Detector | 186248900 | K1787E725A | 1.69.2682 | 2017
20 | 2492 UV/Visible Detector | 186248900 | K1787E727A | 1.69.2682 | 2017
21 | 2493 UV/Visible Detector | 186248900 | K1787E728A | 1.69.2682 | 2017
22 | 2494 UV/Visible Detector | 186248900 | K1787E729A | 1.69.2682 | 2017
23 | 2495 UV/Visible Detector | 186248900 | K1787E730A | 1.69.2682 | 2017
24 | 2496 UV/Visible Detector | 186248900 | K1787E731A | 1.69.2682 | 2017
25 | 2497 UV/Visible Detector | 186248900 | K1787E732A | 1.69.2682 | 2017
26 | 2498 UV/Visible Detector | 186248900 | K1787E745A | 1.69.2682 | 2017

## flextronics (2)
|  | Updated : | 2020-08-03
No. | Category | Product Family | Part Number | Serial No. | FW Version | Date Received | Status | Remarks
1 | Known Good unit | 2489/TUV | 186015022 | K05UPT479M | 1.10.1174 | 2007 | Online | In use for spare bench p/n:700002554
2 | Known Good unit | 2489/TUV | 186015028 | A08UPT943M | 1.60.2273 | NA | To Scrap | 1. Not In Use 2. Wavelength Verifcation Failure
3 | Known Good unit | 2489/TUV | 186015031 | C10TUV025N | 1.60.2274 | 2012 | Online | In use for spare bench p/n: 700005245
4 | Known Good unit | 2489/TUV | 186002487 | L0787E095M | 1.40.1966 | 2009 | Online | In use for spare bench p/n: 279001150
5 | Known Good unit | 2489/TUV | WAT081110 | L03487649M | 1.01 (Spare bench testing only) | 2005 | To Scrap | 1. Defective unit 2. No Display
6 | Known Good unit | 2489/TUV | WAT081110 | H05487269M | 1.01 (Spare bench testing only) | 2007 | Online | Working condition p/n:WAS081147
7 | Known Good unit | 2489/TUV | WAT081110 | H05487273M | 1.01 (Spare bench testing only) | 2007 | Online | Working condition p/n:WAS081147
8 | Known Good unit | 2489/TUV | WAT081110 | H05487251M | 1.01 (Spare bench testing only) | 2007 | To Scrap | 1. Defective unit 2. Auto Restart
9 | Bad unit | 2489/TUV | WAT081110 | H05487257M | 1.01 (Spare bench testing only) | TBD | To Scrap | 1. Defective unit 2. Faulty PCB (Noise failure) 3. Lamp lighting failure
10 | Bad unit | 2489/TUV | WAT081110 | G98487535M | 1.01 (Spare bench testing only) | TBD | To Scrap | 1. Defective unit 2. Unable to power ON
11 | Known Good unit | 2998/PDA/eLambda | 186002998 | D08998406N | 1.41.6164 | 2010 | Online | In use for spare bench p/n:700011867 and 700003765
12 | Known Good unit | 2998/PDA/eLambda | 186299800 | B13998443A | 1.41.6164 | 2015 | Online | In use for spare bench p/n: 700003765 and 700003765
13 | Known Good unit | 2998/PDA/eLambda | 186015026 | B08UPD563M | Not Tested in Flex and no test report | 2010 | To Scrap | 1. Not in use. 2. Unable to test spare bench as Known good unit version. 3. Not compatible with current software of TE82882 4. Will encounter loss comm when running test
14 | Known Good unit | 2998/PDA/eLambda | 186015030 | A09UPL372M | Not Tested in Flex and no test report | 2011 | To Scrap | 1. Not in use. 2. Unable to test spare bench as Known good unit version. 3. Not compatible with current software of TE82882 4. Will encounter loss comm when running test
15 | Known Good unit | 2998/PDA/eLambda | 186015032 | BETUPD008N | 1.65.6229 | 2012 | Online | In use for spare bench p/n: 700002787 and 700011652
16 | Known Good unit | 2998/PDA/eLambda | 186015034 | G13C2P293N | Not Tested in Flex and no test report | 2015 | Yes | In use for spare bench p/n:700009532
17 | Reliability unit | 2489 UV/Visible Detector | 186248900 | K1787E723A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
18 | Reliability unit | 2490 UV/Visible Detector | 186248900 | K1787E724A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
19 | Reliability unit | 2491 UV/Visible Detector | 186248900 | K1787E725A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
20 | Reliability unit | 2492 UV/Visible Detector | 186248900 | K1787E727A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
21 | Reliability unit | 2493 UV/Visible Detector | 186248900 | K1787E728A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
22 | Reliability unit | 2494 UV/Visible Detector | 186248900 | K1787E729A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
23 | Reliability unit | 2495 UV/Visible Detector | 186248900 | K1787E730A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
24 | Reliability unit | 2496 UV/Visible Detector | 186248900 | K1787E731A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
25 | Reliability unit | 2497 UV/Visible Detector | 186248900 | K1787E732A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
26 | Reliability unit | 2498 UV/Visible Detector | 186248900 | K1787E745A | 1.69.2682 | 2017 | Reliability unit | Reliability unit need to 1. Replaced Ellipsoidal DT Mirror P/n: 402000170 2. 402000451 Deutium Lamp 3. Upgrade version software)
27 | Known Good unit | 1515 HplC | 186001515 | NA | NA | NA | Request to return to WPPL | 1. Not in use
28 | Known Good unit | 1525 Hplc | 186001525 | NA | NA | NA | Request to return to WPPL | 1. Not in use
28 | Known Good unit | 515 Hplc | WAT207000 | NA | NA | NA | Request to return to WPPL | 1. Not in use