## Investigation
Venture | [5 empty] | CEI | [5 empty] | Status
No. | PCBA Serial no. | Failure mode in Instrument | Failed at Instrument | Observation before sending to CEI | Action Update | Visual Inspection | U12 Measurement | Boot Up Verification | ICT Testing Result | Investigation | Corrective Action
1 | 190916240C | Fail function. Remaks firmware download fail | 186015000 - iBSM | 1.Send to CEI for repair Initial FA report (24 July 2017) Result: NDF 2. Unit was returned and tested in Venture Result: FAIL (TEST REBOOT FIRMWARE FOR PUMP B DOWNLOAD FAIL) Board was returned to CEI on Feb 2018. 3. No any Visual reject observed | Mar 23, 2018 Follow Up action plan from cei March 28, 2018 Follow up email requested from CEI on the status. Apr 09, 2018 Follow up with CEI on the plan for this RMA. Part to replacement provided on WK16 | No visual defect found | Good | 1. Boot up verification test slip - Done Result: Able to boot-up 2. Boot loader available? 3. MAC Address - As attached verification slip 4. IP Address - As attached verification slip 5. Application firmware - As attached verification slip | ICT Test Result: PASS No rework conducted | 1st FA: Change U12 (Dec 12, 2016 2nd FA: NDF 3rd FA: Ongoing Board passed all the test verification conducted at CEI, no failure observed | Board is multiple return with result NDF. Further verification & analysis will conducted by CA tan. Team decided to replace the board with FG sn 040418243C. Please advise approval to ship board sn 040418243C as a replacement board to close this RMA | Closed
2 | 210917768C | Firmware download Remarks: Beep Sound when power on unit | 186017000 - QSM | 1. This board was tried in FTN 5017 to check the detection of PCBA and to manual download firmware. 2. For both Without & With adding a jumper at pin 7 to force boot loader, below conditions are observed. - No beep sound when power on the unit. - PCBA was not detected by hyperterminal. - no command is able to type - Loader cannot scan CPU 2000 but can scan other instruments 3. No any Visual reject observed Conclusion: CPU 2000 board cannot be detected to manual download firmware. | Mar 23, 2018 Venture request RMA No. to CEI March 28, 2018 CEI send some queries on boards, pending RMA issuance. Apr 12, 2018 CEI Provided the RMA NO. RMA#25797 and to be collected on WK16. (Tentative: Tuesday) Apr 17, 2018 Board collected by CEI. Apr 20, 2018 CEI received the board May 14, 2018 Report/Discussion on the details of the report from CEI, board to be shipped back to GES on May 22, 2018. June 19, 2018 Unit Tested as PASS - 186015046 | No visual defect found | Good | 1. Boot up verification test slip result: Not able to boot-up. 2. Boot loader available - None 3. MAC address – None 4. IP Address - None 5. Application firmware - None *No response at all from the board -Proceed to retest ICT then retest to boot up. 2. Board able to boot up after PASSED ICT TEST. Boot up test as per attached below, no IP address & program loaded (reset during ICT). | Retested ICT 3X - PASSED No defect found. Confirmed U12, R168, R169, R175 & R177 are GOOD. | Board was initially passed test on 26 Sep 2017 prior shipped to GES (pls refer attached CEI ICT test record). As received board unable to boot up, no response at all from the board. Retest ICT, no fault found - PASSED. Then board able to boot up after ICT reset the program loaded on the board. | Request GES to investigate on the program loaded to the board that was caused board unable to boot up. | Closed
3 | 281017888C | Test reboot firmware for Pump B download fail | 186015000 | 1. This board was tried in Column Manager 5043 to check the detection of PCBA and to manual download firmware. 2. For both Without & With adding a jumper at pin 7 to force boot loader, below conditions are observed. - PCBA was not detected by hyperterminal. - no command is able to type - Loader cannot scan CPU 2000. 3. When tried to download firmware through tester, it failed at "power up/RS232 port" step. 4. Oberserve that the U18 Pin Bent. Conclusion: CPU 2000 board cannot be detected to manual download firmware and failed at when power up/ RS232 port checking. | Mar 23, 2018 Venture request RMA No. to CEI Apr 12, 2018 CEI Provided the RMA NO. RMA#25797 and to be collected on WK16. (Tentative: Tuesday) Apr 17, 2018 Board collected by CEI. Apr 20, 2018 CEI received the board May 14, 2018 Report/Discussion on the details of the report from CEI, board to be shipped back to GES on May 22, 2018. June 18, 2018 Unit Tested as PASS - 186015046 | Lead Bent loc.U18 | Good | 1. Boot up verification test slip result: Not able to boot-up. 2. Boot loader available - None 3. MAC address – None 4. IP Address - None 5. Application firmware - None *No response at all from the board -Proceed to retest ICT FAIL due to IC U18 lead bent 2. After replace U18 board passed test then Board able to boot up. Boot up test as per attached below, no IP address & program loaded (reset during ICT). | Retested ICT 3X - PASSED No defect found. Confirmed U12, R168, R169, R175 & R177 are GOOD. | Board was initially passed test on 4 Nov 2017 prior shipped to GES. As received observed U18 lead bent, as per test verification result the defect is captured at ICT test. Suspect U18 was 'knocked off' during installation to the instrument. As received board unable to boot up, no response at all from the board. Replaced U18 WA-332000377 board passed ICT test. Then board able to boot up after ICT reset the program loaded on the board. | Request GES to investigate on the program loaded to the board that was caused board unable to boot up & possible handling issue that may causing U18 lead bent. | Closed
4 | 101017321C | Firmware download Remarks:Beep Sound when power on unit | 186017000 | 1. This board was tried in FTN 5017 to check the detection of PCBA and to manual download firmware. 2. For both Without & With adding a jumper at pin 7 to force boot loader, below conditions are observed: - No beep sound when power on the unit. - PCBA was not detected by hyperterminal. - no command is able to type - Loader cannot scan CPU 2000 but can scan other instruments 3. No any Visual reject observed Conclusion: CPU 2000 board cannot be detected to manual download firmware. | Mar 23, 2018 Venture request RMA No. to CEI March 28, 2018 CEI send some queries on boards, pending RMA issuance. Apr 12, 2018 CEI Provided the RMA NO. RMA#25797 and to be collected on WK16. (Tentative: Tuesday) Apr 17, 2018 Board collected by CEI. Apr 20, 2018 CEI received the board May 14, 2018 Report/Discussion on the details of the report from CEI, board to be shipped back to GES on May 22, 2018. June 19, 2018 Unit Tested as PASS - 186015046 | Connector crack loc.J2 | Good | 1. Boot up verification test slip result: Not able to boot-up. 2. Boot loader available - None 3. MAC address – None 4. IP Address - None 5. Application firmware - None *No response at all from the board -Proceed to retest ICT then retest to boot up. 2. Board able to boot up after PASSED ICT TEST. Boot up test as per attached below, no IP address & program loaded (reset during ICT). | Retested ICT 3X - PASSED No defect found. Confirmed U12, R168, R169, R175 & R177 are GOOD. | As received, observed J2 crack. No requirement to use those connector at CEI during production assembly process. Most likely conncetor being crack during installation at instrument. Board was initially passed test on 13 Oct 2017 prior shipped to GES (pls refer attached CEI ICT test record). As received board unable to boot up, no response at all from the board. Retest ICT, no fault found - PASSED. Then board able to boot up after ICT reset the program loaded on the board. | Request GES to investigate on the program loaded to the board that was caused board unable to boot up. | Closed
5 | 101017309C | 1. Firmware download No detection PCBA sound 2. Power on initialization | 1. 186002441 2. 186015046 | 1. This board was re-tried on 5017 unit. Below conditions observed - PCBA was not detected by hyperterminal. - no command is able to type - Loader cannot scan CPU 2000 but can scan other instruments. - when tried to download firmware through Diagnostic mode, also cannot download firmware. It shows as "serial port is not responding" '- When rejected PCBA is removed, and replaced by a new PCBA, firmware could be downloaded. 3. This board was re-tried on 5043 unit. Below conditions observed: - PCBA was not detected by hyperterminal. - no command is able to type - Loader cannot scan CPU 2000. - When tried to download firmware through tester, it failed at "power up/RS232 port" step. - When a new PCBA was changed, firmware is able to downloaded. 4. Observed that J4 with crack. Conclusion: CPU 2000 board cannot be detected to manual download firmware and failed at when power up/ RS232 port checking. | Mar 23, 2018 Venture request RMA No. to CEI March 28, 2018 CEI send some queries on boards, pending RMA issuance. Apr 12, 2018 CEI Provided the RMA NO. RMA#25797 and to be collected on WK16. (Tentative: Tuesday) Apr 17, 2018 Board collected by CEI. Apr 20, 2018 CEI received the board May 14, 2018 Report/Discussion on the details of the report from CEI, board to be shipped back to GES on May 22, 2018. June 14, 2018 Unit Tested as PASS - 186015046 | Conn crack loc.J1, J2 & J4 | Good | 1. Boot up verification test slip result: Not able to boot-up. 2. Boot loader available - None 3. MAC address – None 4. IP Address - None 5. Application firmware - None *No response at all from the board -Proceed to retest ICT then retest to boot up. 2. Board able to boot up after PASSED ICT TEST. Boot up test as per attached below, no IP address & program loaded (reset during ICT). | Retested ICT 3X - PASSED No defect found. Confirmed U12, R168, R169, R175 & R177 are GOOD. | As received, observed J1, J2 & J4 crack. No requirement to use those connector at CEI during production assembly process. Most likely conncetor being crack during installation at instrument. Board was initially passed test on 13 Oct 2017 prior shipped to GES (pls refer attached CEI ICT test record). As received board unable to boot up, no response at all from the board. Retest ICT, no fault found - PASSED. Then board able to boot up after ICT reset the program loaded on the board. | Request GES to investigate on the program loaded to the board that was caused board unable to boot up. | Closed
6 | 130218583C | Intelligent Intake Valve HW not supported | 186015000 | 1. Failed during Alarm check of 186015000 with error message "Intelligent Intake Valve HW not supported". 2. Verified PCBA in BSM two units and same failure duplicated. 3. Used it in 186002441, Firmware was able to downloaded. But during the testing, contrast adjusted LCD became dim & also failed at temperature variation test . After changing to another 210000520, can resolve the problem. Performed verification with two units of 2441. Same failure and condition encountered. 4. Used it in 186015043, Firmware was able to be downloaded. But during the testing, there is no LED light either GREEN or RED for "RUN" LED.After to another 210000520, can resolve the problem. | Apr 27, 2018 Venture request RMA No. to CEI May 14, 2018 Follow up to CEI for update June 07, 2018 RMA No. 25965, board collected on June 13 July 17, 2018 board to ship back on July 24, 2018 July 30, 2018 Unit Tested as PASS - 186015085 | missing insulator loc.B1 WA-336000118 | Refer attached: | Boot up verification passed test passed test | ICT Test PASS | No Fault Found | N/A | Closed
7 | 090518861C | When power on UUT, there is no light at DS1 to DS4 Leds. When initialize UUT, CPU bd cannot detect | 186015046 | 1. During UUT initialize, bootloader cannot detect CPU 210000520. 2. Connection between 210000520 & 210000367 are checked. No issue found. 3. Cable connection between 210000520 & 210000367 are checked. No issue found. 4. Cable connection between 210000520 & 279003610 are checked. No issue found. 5. Checked 210000520 and found there is no light foe led DS1 to DS4. 6. PCBA 210000520 was verified in another UUT. Same failure can be duplicated. 7. Replaced with good 210000520 - PASSED, Confirmed reject 210000520. | Jul 10, 2018 Venture request RMA No. to CEI Jul 13, 2018 Received RMA NO. 26213 for return Aug 08, 2018 CEI collected boards | [6 empty] | Open