Release Notes PDA Device Replacement Release        May 19, 2017
Data System Requirements:
Firmware:    

          ACQ-PDA II (thermal):   v1.69.6282    checksum: 0xE29BE7DA

          2998:      v1.69.6278         checksum: 0xEDC5D76C
IMPORTANT NOTES:
The only material change in this release compared to the previous is with regards to the saturation latch circuitry initialization.  This is purely an electrical engineering construct.  To this end, the saturation latch circuitry register is now initialized with the value 0x8.
System Constraints and Limitations:

None known
Tracker Items that have been implemented
	SCR
	Title
	Description

	35511
	PDA Device Project – Photodiode Array supply chain


	This SCR address the primary purpose of this project.  The supply chain for photodiode array electronic devices has been disrupted and must be restored.  Project Plan 721007853 contains the information to drive the project.  Many primary products affected.




TIPS:
How do I check if my instrument has a Hammatsu or an OnSemi chip?

Once the instrument has initialized, type in command line:   dump dataman

At the bottom of the print out look for information shown below
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In this example, the instrument was a 2998 with a Hamamatsu chip installed.

