DOCUMENT TITLE: PM / Setup Electrical Safety Testers

DOCUMENT TYPE: Calibration Preventive Maintenance

EQUIPMENT NAME: Electrical Safety Analyzer QuadTech Guardian 6000 QuadTech Guardian 6000 Plus Chroma 19032

EQUIPMENT NUMBER: 735003029

| AUTHOR(S): Matt Brule | DATE: August 14, 2013 |
| --- | --- |
| REVISED BY: Matt Brule | DATE: July 01, 2016 |

Table of Contents:

Reference Documentation	3

Scope and notes about this procedure	3

1.	Calibration	4

2.	Setup (All tester types)	4

3.	QUADTECH Guardian 6000 Setup	5

4.	Quadtech Guardian 6000 Installation	10

5.	Quadtech Guardian 6000Plus or Chroma 19032 setup	11

6.	Quadtech Guardian 6000 PLUS / CHroma 19032 Installation	17

Programming PARAMETERS QUADTECH Guardian 6000	19

Programming parameters QuadTech Guardian 6000 Plus / Chroma 19032	20

Electrical safety Analyzer Equipment List	21

QuadTech Guardian 6000 Setup Record	22

QuadTech Guardian 6000 PLUS / Chroma 19032 Setup Record	23

# Equipment Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| N/A | N/A | N/A |

# Disposables Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| N/A | N/A | N/A |

# Reference Documentation

| Document Number | Description |
| --- | --- |
| WAT046538 | GENERIC HI-POT & GND CONT Test PROC. |
| 730001799 | Operating Procedure, Calibration and PM |

# Scope and notes about this procedure

This procedure details the programming of electrical safety test equipment for use in testing Waters’ product in manufacturing and at contract manufacturers worldwide.

There are three electrical safety analyzers that can be used to perform this testing:

- QuadTech Guardian 6000

- QuadTech Guardian 6000 Plus

- Chroma 19032

The QuadTech Guardian 6000, QuadTech Guardian 6000 Plus, and the Chroma 19032 are set up with a single program in memory location 1 named 12345678 to test all Waters’ product.

The programming parameters for memory location 1 in all testers are equal, there are some differences in nomenclature, and the more advanced units have additional capability that is not used, but the tests executed are equivalent.

All figures in this procedure are for reference only and are intended to be representative of actual testers and may not be exact.

- Calibration

  - Referring to 730001799, outsource the calibration of the electrical safety analyzer.

- Setup (All tester types)

Upon return from calibration (or anytime setup of a new or existing tester is required) follow the appropriate process below depending on the instrument model being setup.

  - Inspect all UUT interface cables and replace any that are worn.

  - Connect power cord

  - Power up instrument

    - Complete sections 3 and 4 for the QuadTech Guardian 6000.

    - Complete sections 5 and 6 for QuadTech Guardian 6000 Plus or the Chroma 19032.

- QUADTECH Guardian 6000 Setup

  - Key Lock

    - Look at the analyzers front panel to see if Key Lock is locked by observing the LED in Figure 1.  If Key Lock is locked go to step 3.1.2 to disable it, if not continue to step 3.2.

    - Type the 4 digit password from the calibration database then press the Off key.  The Key Lock LED will turn off, indicating that the Key Lock is unlocked.

Figure 1: QuadTech Guardian 6000 front panel with Key Lock and keypad

  - Remove all installed programs.

    - With the analyzer in standby status, (no lights blinking).

    - Press the Red Stop button.

    - Press Clear, 6, 0, 0, 0, Clear, the display will appear as in Figure 2.

Figure 2: QuadTech Guardian 6000 Main screen with Memory location 1 loaded at step 1

    - Press the Enter key.  The beeper will sound and all front panel LEDs will flash once, the display will return to STEP-01 in the upper left (Figure 3).  The other text on the display can be ignored.

Figure 3: QuadTech Guardian 6000 at Step 1

  - Use left / right arrow keys to adjust the display contrast if necessary.

  - Program Memory Location 1.

Use the keypad and display (Figure 4) to program.

Figure 4: QuadTech Guardian 6000 display and keypad

    - Press the Prog key, the screen will appear as in Figure 5.

Figure 5: QuadTech Guardian 6000 Select Step

    - Press Enter to begin programming.

    - Enter all parameter settings to match the table below.

Use the Enter key to scroll through each parameter.

Use the up and down arrow keys to scroll through the settings for each parameter and the numeric keys to enter numeric values.

From the Select Step screen (Figure 5), use the up and down arrow keys to toggle between test steps.

    - Record all of the parameter settings on the QuadTech Guardian 6000 Setup Record on page 22.

| PARAMETER | SETTING | DESCRIPTION |
| --- | --- | --- |
| Step 1 |  |  |
| Mode | GR | Ground Continuity / Bond test |
| Current | 20.00A | Test Current |
| High Limit | 100.0m | Test Upper Limit |
| Low Limit | Disable | Test Lower Limit |
| Test Time | 3.0s | Total test time |
| Step 2 |  |  |
| Mode | WDC | Dielectric Withstand (Hi-Pot) |
| Voltage | 2.121 kV | Test Voltage |
| High Limit | 5.00mA | Test Upper Limit |
| Low Limit | Disable | Test Lower Limit (Not used) |
| Arc Limit | Disable | Arc Limit (Not used) |
| Hold / Test Time | 3.0s | Time voltage is applied |
| Ramp Time | 2.0s | Time voltage takes to rise from 0 V to Test Voltage |

    - When programming is complete and all parameter settings are correct, press the Prog key to exit programming mode.

  - Store the program in Memory location 1.

    - Press the Sto key, use the arrow keys to select location memory location 1 and Press Enter

    - Enter ‘12345678’ and Press the Enter key, the screen will appear as in Figure 6.

Figure 6: QuadTech Guardian 6000 Store memory location 1

    - Press Enter to store the program the display will look like Figure 7.

Figure 7 QuadTech Guardian 6000 programmed

- Quadtech Guardian 6000 Installation

  - Apply preventative maintenance sticker per 730001799.

  - Install at point of use per test procedure WAT046538.

  - Demonstrate correct parameter settings to Test Engineering and Manufacturing.

    - From the initial screen, press the RcL key.

    - Press Enter twice to recall memory location 1, 12345678.

    - From the initial screen, press the Prog key.

    - If necessary, scroll to step 1 using the arrow keys

    - Use the Enter key to scroll through each parameter for step 1, repeat as necessary to review all parameters.  Test settings are summarized on page 19.

    - From the Select step screen use the arrow keys to select step 2.

    - Use the Enter key to scroll through each parameter for step 2, repeat as necessary to review all parameters.  Test settings are summarized on page 19.

    - Press the Prog key to exit.

    - Load the correct program: Memory location 1.

      - Press the Rcl key

      - Use the up and down arrow keys to scroll to the correct program.

      - Press Enter twice to recall the program.

      - Press the Clear key.

  - Lock the Key Lock.

    - From the initial screen, type the 4 digit password from the calibration database then press the Off key.  The Key Lock LED will illuminate.  The display will appear as in Figure 8

Figure 8: QuadTech Guardian 6000 Final setup screen

  - Sign and date the QuadTech Guardian 6000 Setup Record (Page 22).

  - Obtain signature on the QuadTech Guardian 6000 Setup Record of Test Engineering and the Manufacturing Technician that setup is correct, the correct program is loaded, and the keypad is locked.

  - Store the Setup Record in the calibration folder per procedure 730001799.

The Electrical Safety Analyzer is now ready for use.  Sections 5 and 6 only need to be completed for the QuadTech Guardian 6000 Plus or Chroma 19032.

- Quadtech Guardian 6000Plus or Chroma 19032 setup

The only difference on the front control panel between a QuadTech Guardian 6000 Plus and a Chroma 19032 is the name of the instrument, see Figure 9.  This section will refer to both testers, pictures will reference the QuadTech Guardian 6000 Plus.

Figure 9: Difference between QuadTech Guardian 6000 Plus and the Chroma 19032

The primary interface for the QuadTech Guardian 6000Plus and the Chroma 19032 is the blue LCD, it’s ‘soft keys’, and the numeric keypad. Figure 10 shows the interface at the MAIN MENU.

Figure 10: QuadTech Guardian 6000 Plus / Chroma 19032 Key soft keys and keypad

  - Key Lock

    - Look at the analyzers front panel to see if Key Lock is locked by observing ‘Lock’ on the bottom of the display is in reverse text (see Figure 11 left). If Key Lock is locked go to step 5.1.2 to unlock it, if not continue to step 5.2.

Figure 11: QuadTech Guardian 6000 Plus / Chroma 19032 Key Lock indicator and menu choice

    - From the Main menu (Figure 11), press the 4 key.

    - Type the 4 digit password from the calibration database then press Enter, the screen will return to the main menu with Key Lock unlocked.

  - Remove all installed programs.

    - From the MAIN MENU, press the MEMORY soft key.

    - Using the PAGE UP and PAGE DOWN soft keys, scroll through all 100 memory locations and verify that there are none programmed.  Use the DELETE soft key to delete any found.

Program locations 2- 10 in Figure 12 are not programmed, note there is no program name and 0 program steps. While program location 1 shows 2 steps and the name .12345678’

Figure 12: QuadTech Guardian 6000 Plus / Chroma 19032 Identifying a program

Program M emory Location 1.

  - Use the display and keypad (Figure 13) to program.

Figure 13: QuadTech Guardian 6000 Plus / Chroma 19032 display / keypad

    - From the MAIN MENU, press the PROGRAM soft key.

    - Enter all parameter settings to match the table below (also shown in Figure 14 for reference).

Use the up and down arrow keys to scroll through for each parameter, the soft keys to select setting, and the numeric keys to enter numeric values.

Use the UP / DOWN soft keys to navigate between test steps.

    - Record all parameter settings on the QuadTech Guardian 6000 Plus / Chroma 19032 Setup Record on page 23.

| PARAMETER | SETTING | DESCRIPTION |
| --- | --- | --- |
| Step 1 |  |  |
| Mode | GB | Ground Continuity / Bond test |
| Current | 20.00A | Test Current |
| High Limit | 100.0m | Test Upper Limit |
| Low Limit | OFF | Test Lower Limit |
| Test Time | 3.0s | Total Test Time |
| Twin Port | OFF | Not Used |
| Step 2 |  |  |
| Mode | DC | Dielectric Withstand (Hi-Pot) |
| Voltage | 2.121 kV | Test Voltage |
| High Limit | 5.00mA | Test Upper Limit |
| Low Limit | OFF | Not Used |
| Dwell Time | OFF | Not Used |
| Arc Limit | OFF | Not Used |
| Arc Filter | 3-230 kHz | Not Used |
| Test Time | 3.0s | Time voltage is applied |
| Ramp Time | 2.0s | Time voltage takes to rise from 0 V to Test Voltage |
| *Fall time | OFF | Not used |
| CHNL (H_L) | OFF | Not Used |

*Fall time is only present on the Chroma 19032 instruments, N/A Guardian 6000 Plus

Figure 14: QuadTech Guardian 6000 Plus / Chroma 19032 step 1 settings

    - From the MAIN MENU, press the MEMORY soft key.  The MEMORY SETUP screen will be displayed (Figure 15)

Figure 15: QuadTech Guardian 6000 Plus / Chroma 19032 memory setup

    - Press the STORE soft key.

    - Enter the program name 12345678, then press the ENTER key.

- Quadtech Guardian 6000 PLUS / CHroma 19032 Installation

  - Apply preventative maintenance sticker per 730001799.

  - Install at point of use, setup per test procedure WAT046538.

  - Demonstrate correct parameter settings to Test Engineering and Manufacturing Technician.

    - From the MAIN MENU, Press the MEMORY soft key.

    - Press the RECALL soft key with the program in memory location 1, 12345678 highlighted to recall it.

    - Press the ENTER key to confirm the recall.

    - Press the MENU key to display the MAIN MENU.

    - Press the PROGRAM soft key.

    - Verify test step 1 settings.

    - Press the UP soft key.

    - Verify test step 2 settings.

    - Press the MENU key to display the MAIN MENU

    - Press the MEMORY soft key.  The MEMORY SETUP screen will be displayed.

    - Press the RECALL soft key with the program in memory location 1, 12345678 highlighted to recall it.

    - Press the ENTER key to confirm the recall.

    - Press the MENU key to display the MAIN MENU.

  - Lock the Key Lock.

    - Press the 4 key to recall the Key Lock

    - Type the 4 digit password from the calibration database then press Enter.

    - Verify that the MESSAGE is RECALL LOCK : YES,  press the EXIT soft key.  The MAIN MENU will be displayed.

    - Press the TEST soft key.  The display will appear as in Figure 16, with the test program displayed (TEST-M001-12345678), and the LOCK text in reverse test.

Figure 16: QuadTech Guardian 6000 Plus / Chroma 19032 Final Setup Screen

  - Sign and date the QuadTech Guardian 6000 Plus / Chroma 19032 Setup Record.

  - Obtain signatures on the QuadTech Guardian 6000 Plus / Chroma 19032 Setup Record of Test Engineering and Manufacturing that setup is correct, the correct program is loaded, and the keypad is locked.

  - Store the Setup Record in the calibration folder

Electrical Safety Analyzer is now ready for use.

# Programming PARAMETERS QUADTECH Guardian 6000

Following is a summary of programming parameters for the QuadTech Guardian 6000.  This procedure details how to program and verify these parameters into memory location 1 named 12345678.  As well as provides record / sign off / and approval of these settings and the locking of the analyzer keypad before production use.

| Step 1 |  |  |
| --- | --- | --- |
| PARAMETER | Memory Location _1 12345678 | DESCRIPTION |
| Mode | GR | Ground Continuity / Bond test |
| Current | 20.00A | Test Current |
| High Limit | 100.0m | Test Upper Limit |
| Low Limit | Disable | Test Lower Limit |
| Test Time | 3.0s | Total test time |
| Twin Port | OFF | Channel selection (Not used) |
| Step 2 |  |  |
| Mode | WDC | Dielectric Withstand (Hi-Pot) |
| Voltage | 2.121 kV | Test Voltage |
| High Limit | 5.00mA | Test Upper Limit |
| Low Limit | Disable | (Not used) |
| Arc Limit | Disable | (Not used) |
| Hold Time | 3.0s | Time voltage is applied |
| Ramp Time | 2.0s | Time voltage takes to rise from 0 V to Test Voltage |

The critical parameters are in reverse text.

Waters instruments require a ground continuity test, an instantaneous measurement, on a routine production basis.  The test equipment used performs a ground bond test, which is the same as a ground continuity test, but, with current applied, and for a specified time.  This Ground Bond test is more stringent than the Ground Continuity requirement, due to the increased current and longer duration.

.

# Programming parameters QuadTech Guardian 6000 Plus / Chroma 19032

Following is summary of programming parameters for the QuadTech Guardian 6000 Plus and the Chroma 19032.  This procedure details how to program and verify these parameters into Memory location 1 named 12345678.  As well as provides record / sign off / and approval of these settings and the locking of the keypad before production use.

| Step 1 |  |  |
| --- | --- | --- |
| PARAMETER | Setting | DESCRIPTION |
| Mode | GR | Ground Continuity / Bond test |
| Current | 20.00A | Test Current |
| High Limit | 100.0m | Test Upper Limit |
| Low Limit | OFF | Test Lower Limit |
| Test Time | 3.0s | Total test time |
| Twin Port | OFF | (Not used) |
| Step 2 |  |  |
| Mode | DC | Dielectric Withstand (Hi-Pot) |
| Voltage | 2.121 kV | Test Voltage |
| High Limit | 5.00mA | Test Upper Limit |
| Low Limit | OFF | (Not used) |
| Dwell Time | OFF | (Not Used) |
| Arc Limit | OFF | (Not used) |
| Arc Filter | 3-230 kHz | Not Used |
| Test Time | 3.0s | Time voltage is applied |
| Ramp Time | 2.0s | Time voltage takes to rise from 0 V to Test Voltage |
| *Fall time | OFF | (Not used) |
| CHNL (H-L) | OFF | (Not used) |

*Fall time is only present on the Chroma 19032 instruments, N/A Guardian 6000 Plus

The critical parameters are in reverse text.

Waters instruments require a ground continuity test, an instantaneous measurement, on a routine production basis.  The test equipment used performs a ground bond test, which is the same as a ground continuity test, but, with current applied, and for a specified time.  This Ground Bond test is more stringent than the Ground Continuity requirement, due to the increased current and longer duration.

# Electrical safety Analyzer Equipment List

| Instrument Serial Number | Identifier | Manufacturer | Model | Location |
| --- | --- | --- | --- | --- |
| 8280865 | 86041 | Quadtech | Guardian 6000 | Flextronics |
| 1140656 | 90259 | Quadtech | Guardian 6000 | Milford |
| 1140659 | 90261 | Quadtech | Guardian 6000 | Milford |
| 1140657 | 90265 | Quadtech | Guardian 6000 | Milford |
| 1140663 | 90266 | Quadtech | Guardian 6000 | Milford |
| 1140662 | 90270 | Quadtech | Guardian 6000 | Milford |
| 1140660 | 90282 | Quadtech | Guardian 6000 | Flextronics |
| 5037654 | 91354 | Quadtech | Guardian 6000 | Milford |
| 7312754 | 91356 | Quadtech | Guardian 6000 | Venture |
| 5278621 | 91683 | Quadtech | Guardian 6000 | Venture |
| 1140658 | 91862 | Quadtech | Guardian 6000 | Milford (Regulatory) |
| 11173729 | 91890 | Quadtech | Guardian 6000 | Venture |
| 1140661 | 92599 | Quadtech | Guardian 6000 | Milford |
| 19032G400201 | 93237 | Chroma | 19032 | Milford |
| 19032G400202 | 93238 | Chroma | 19032 | Milford |
| 19032G400225 | 93260 | Chroma | 19032 | CEI |
| 19032G400264 | 93467 | Chroma | 19032 | Venture |
| 5037655 | 95037 | Quadtech | Guardian 6000 | Flextronics |
| G60000000174 | 95050 | Quadtech | Guardian 6000 Plus | Venture |
| 9111936 | 95051 | Quadtech | Guardian 6000 Plus | Milford |
| 6069766 | 98206 | Quadtech | Guardian 6000 | Flextronics |
| 8060121 | 98210 | Quadtech | Guardian 6000 Plus | Milford |
| 3505384 | 99099 | Quadtech | Guardian 6000 | Milford |
| 3505383 | 99102 | Quadtech | Guardian 6000 | Milford |
| 3505382 | 99103 | Quadtech | Guardian 6000 | Milford |
| 3505381 | 99104 | Quadtech | Guardian 6000 | Milford |

# QuadTech Guardian 6000 Setup Record

ID Number:

Performed By:						Date:

|  | Memory 1: 12345678 |  |  |
| --- | --- | --- | --- |
|  | PARAMETER | VALUE | SETTING |
| Step 1 | Mode | GR |  |
|  | Current | 20.00A |  |
|  | High Limit | 100.0m |  |
|  | Low Limit | Disable |  |
|  | Test Time | 3.0s |  |
| Step 2 | Mode | WDC |  |
|  | Voltage | 2.121 kV |  |
|  | High Limit | 5.00mA |  |
|  | Low Limit | Disable |  |
|  | Arc Limit | Disable |  |
|  | Test Time | 3.0s |  |
|  | Ramp Time | 2.0s |  |

Key Lock locked   (Key Lock LED illuminated)

Program displayed:						(M01)

Approved by:

Performer:								Date:

Verified By:  Test Engineering

Name:									Title

Signature:								Date:

End User Approval:  Manufacturing Technician

Name:									Title

Signature:								Date:

# QuadTech Guardian 6000 PLUS / Chroma 19032 Setup Record

ID Number:

Performed By:						Date:

|  | Memory 1: 12345678 |  |  |
| --- | --- | --- | --- |
|  | PARAMETER | VALUE | SETTING |
| Step 1 | Mode | GB |  |
|  | Current | 20.00A |  |
|  | High Limit | 100.0m |  |
|  | Low Limit | OFF |  |
|  | Test Time | 3.0s |  |
|  | Twin Port | OFF |  |
| Step 2 | Mode | DC |  |
|  | Voltage | 2.121 kV |  |
|  | High Limit | 5.00mA |  |
|  | Low Limit | OFF |  |
|  | Dwell Time | OFF |  |
|  | Arc Limit | OFF |  |
|  | Arc Filter | 3-230 kHz |  |
|  | Test Time | 3.0s |  |
|  | Ramp Time | 2.0s |  |
|  | *Fall Time | OFF |  |
|  | CHNL (H-L) | OFF |  |

*Fall time is only present on the Chroma 19032 instruments, N/A Guardian 6000 Plus

Key Lock locked   (‘Lock’ text in reverse)

Program displayed:						(TEST-M001-12345678)

Approved by:

Performer:								Date:

Verified By:  Test Engineering

Name:									Title

Signature:								Date:

End User Approval:  Manufacturing Technician

Name:									Title

Signature:								Date: