## Summary
Arc FTN - Phase 2A Training summary
Training Status |  |  |  | Improvement action status :
Planned Assy | Completed Assy | Pending Assy | Completed Training Hrs | Action Items | Closed | Open | To be discussed
9 | 6 | 3 | 57.5 | 15 | 5 | 9 | 1
Training build summary |  |  |  | Testing summary
Assy's build | Completed | Process issue | Re-work done | Assy's Tested | Completed | Test Failure | Re-work done | PASSED after Re-work
289010749 | 14 | 0 | 0 | 279001991 | 12 | 2 | 2 | 2
289010183 | 14 | 0 | 0 | 279003321 | 12 | 0 | 0 | 0
279001950 | 14 | 0 | 0 | 279003341 | 0 | 0 | 0 | 0
279001991 | 14 | 2 | 2 | 289010908 | 0 | 0 | 0 | 0
279003341 | 0 | 0 | 0
289004640 | 14 | 0 | 0
279003321 | 14 | 1 | 1
289010908 | 0 | 0 | 0
289010714 | 0 | 0 | 0

## Training plan
FTN Remote Training Plan :
Seq # | Phase 2A Training |  |  | Phase 2A Shipment to Milford | Testing Required | Tester Name/No. | Tester IQ Done | Material Availability for Qty 14 | 2021-03-08 | 2021-03-09 | 2021-03-10 | 2021-03-11 | 2021-03-12 | 2021-03-15 | 2021-03-16 | 2021-03-17 | 2021-03-18 | 2021-03-19 | 2021-03-22 | 2021-03-23 | 2021-03-24 | 2021-03-25 | 2021-03-26 | 2021-03-29 | 2021-03-30 | 2021-03-31 | 2021-04-01 | 2021-04-02 | 2021-04-05 | 2021-04-06 | 2021-04-07 | 2021-04-08 | 2021-04-09 | 2021-04-12 | 2021-04-13 | 2021-04-14 | 2021-04-15 | 2021-04-16 | 2021-04-19 | 2021-04-20 | 2021-04-21 | 2021-04-22 | 2021-04-23 | 2021-04-26 | 2021-04-27 | 2021-04-28 | 2021-04-29 | 2021-04-30
1 | SYRINGE SUPPORT | 289010749 | ASSY, SUPPORT SLEEVE W/ CUP | Assy 1 | No | NA | NA | Yes | bar: 2021-03-09
2 | SYRINGE HOUSING | 289010183 | ASSY, HOUSING, NEEDLE SEAL, FTN-R |  |  |  |  | Yes | bar: 2021-03-10
3 | THETA DRIVE | 279001950 | ASSY, THETA DRIVE, CPSM | Assy 2 | Yes | ACQUITY CPSM RTZ Assembly Tester - TE82995 | IQ in Progress | Yes | bar: 2021-03-08
4 | DATUM BASE W/ | 279001991 | ASSY, DATUM BASE W/ THETA DRIVE |  |  |  |  | Yes | bar: 2021-03-09
5 | SYRINGE DRIVE | 279003341 | ASSY, DRIVE, SYRINGE 1/4-20 | Assy 3 | Yes | ARC Metering Syringe Drive Tester - TE0084021 | Yes | Yes
|  |  | * Build with decoupled 289010714 |  |  |  |  | Yes
6 | NEEDLE CARRIAGE | 289004640 | ASSY, NEEDLE CARRIAGE ACQ FTN | Assy 4 | Yes | ACQUITY CPSM RTZ Assembly Tester - TE82995 | IQ in Progress | Yes | bar: 2021-03-10 to 2021-03-16
7 | RZZ | 279003321 | ASSY, RZZ MECH, CPSM 2, FTN-R | Assy 5 | Yes | ACQUITY CPSM RTZ Assembly Tester - TE82995 |  | Yes | bar: 2021-03-16 to 2021-03-17
|  |  | * Build with decoupled 289004640 |  |  |  |  | NA
7A | RTZ | 735002271 | ACQUITY CPSM RTZ ASSEMBLY TESTING | Testing | NA |  |  | NA | bar: 2021-03-17 to 2021-03-19
8 | MSV SHAFT ASSY | 289010908 | ASSY, SHAFT W/SPRING, MSV | Assy 6 | Yes | Metering Syringe Valve Leak Tester - TE0084040 | Yet to receive | Yes
9 | MSV | 289010714 | ASSY, SYRINGE VALVE, METERING |  |  |  |  | Yes
[11 empty] | C - Closed |  |  |  |  | P - Planned |  |  |  |  | O - On hold
Note on E12: 289004640 no shipped by milford. Used Training build assy to pair with 3321

## 2A Training status
FTN Remote Training Plan : | [6 empty] | Phase 2A : Sub-Assy Training Build Status
[7 empty] | Training completed and Built on | [14 empty] | 6.5Hrs / day
# | Phase 2A Training Assemblies |  |  | Sequence followed | Start Date | Trainer | Assy 1 | Assy 2 | Assy 3 | Assy 4 | Assy 5 | Assy 6 | Assy 7 | Assy 8 | Assy 9 | Assy 10 | Assy 11 | Assy 12 | Assy 13 | Assy 14 | End Date | Training Hrs/Assy | Comments
1 | SYRINGE SUPPORT | 289010749 | ASSY, SUPPORT SLEEVE W/ CUP | 2 | 2021-03-09 | Christine | C | C | C | C | C | C | C | C | C | C | C | C | C | C | 2021-03-09 | 5.5 | NA
2 | SYRINGE HOUSING | 289010183 | ASSY, HOUSING, NEEDLE SEAL, FTN-R |  | 2021-03-10 | Christine | C | C | C | C | C | C | C | C | C | C | C | C | C | C | 2021-03-10 |  | NA
3 | THETA DRIVE | 279001950 | ASSY, THETA DRIVE, CPSM | 1 | 2021-03-08 | Vy Levan | C | C | C | C | C | C | C | C | C | C | C | C | C | C | 2021-03-09 | 10.5 | NA
4 | DATUM BASE W/ | 279001991 | ASSY, DATUM BASE W/ THETA DRIVE |  | 2021-03-08 | Vy Levan | C | C | C | C | C | C | C | C | C | C | C | C | C | C | 2021-03-09 |  | 12 - Testing completed 2 - Testing pending
5 | SYRINGE DRIVE | 279003341 | ASSY, DRIVE, SYRINGE 1/4-20 * Build with decoupled 289010714
6 | NEEDLE CARRIAGE | 289004640 | ASSY, NEEDLE CARRIAGE ACQ FTN | 3 | 2021-03-10 | Vy Levan | C | C | C | C | C | C | C | C | C | C | C | C | C | C | 2021-03-16 | 23.5 | Build with 3321
7 | RZZ | 279003321 | ASSY, RZZ MECH, CPSM 2, FTN-R * Build with decoupled 289004640 | 4 | 2021-03-16 | Vy Levan | C | C | C | C | C | C | C | C | C | C | C | C | C | C | 2021-03-17 | 7.5 | 12 - Testing completed 2 - Testing pending
7A |  | 735002271 | ACQUITY CPSM RTZ ASSEMBLY TESTING ( TE82995-5 ) | 5 | 2021-03-17 | Vy Levan | C | C | C | C | C | C | C | C | C | C | C | C |  |  |  | 10.5
8 | MSV SHAFT ASSY | 289010908 | ASSY, SHAFT W/SPRING, MSV
9 | MSV | 289010714 | ASSY, SYRINGE VALVE, METERING
[6 empty] | C - Completed |  |  | IP - In Progress | [7 empty] | Completed training Hrs |  |  |  |  | 57.5
Note on V11: 400003719 - Due to shortage of 2 Pcs , Waters SG consigned from alternate vendor.
Note on E12: 289004640 not shipped by milford. Used Training build assy to pair with 3321

## Traineee Vs Assy
Trainee Vs Assy's trained
Assy's | Gary | David | Sylvester | Jernail | Total
279001950 | 4 | 2 | 4 | 4 | 14
279001991 | 4 | 2 | 4 | 4 | 14
289010749 | 4 | 2 | 4 | 4 | 14
289010183 | 4 | 2 | 4 | 4 | 14
289004640 ( step 1.1 to Step 4.3 ) | 1 | 1 | 1 | 1 | 4
289004640 ( step 4.4 to Step 5.13) | 1 | 1 | 1 | 1 | 4
289004640 ( step 1.1 to Step 5.13) | 2 | 0 | 1 | 1 | 4
289004640 ( step 5.14 to Step 8.8) | 3 | 1 | 2 | 2 | 8
289004640 ( Full steps ) | 1 | 1 | 2 | 2 | 6
279003321 | 4 | 2 | 4 | 4 | 14
279003321 + 279001991 RTZ Testing | 3 | 4 | 4 | 4 | 15

## Action items
CUSTOMER : WATERS |  |  |  | Open action |  | 9
PRODUCT : Arc FTN |  |  |  | Closed action |  | 5
Phase 2A - Remote Training Actions |  |  |  | To be discussed |  | 1
# | ASP | Task / Issue | Action Required | Owner | Due Date | Status | Comments
1 | 279001991 : Step 5.4 | Motor Rope and Pulley must be in parallel when assemble. Fixture to be modified. refer picture. | Fixture to be modified | Kwan |  | Closed
2 | 279001991 : Step 1.1 to 4.1 | Supporting fixture required for Datam base to support assembly from step 1.1 to 4.1 | New Fixture required To be discussed with Kwan. | Kwan |  | Open
3 | 279001991 : Step 7.1 | Test assembly per TSP - IS there any testing required ? How this was tested and paired with other assy to use in chassis assy 3378 ? | Discussed with Milford. 279001991 has to be tested with 279003321 Assy. | Michael |  | Closed | Clarified by Trainer Vy Levan : - 279001991 manufactured assy to be tested with 279001993 master assy. - 279003321 manufactured assy to be tested with 279001991 master assy. - 279001991 manufactured assy and 279003321 manufactured unit can be tested together. - Once tested, both PASSED assy can be used in any chassis assy. - Not an requirement to use both as pair in same chassis assy.
4 | 289010749 : Step 1.4a, 1.4b | Fixture 999041838 and instruction use is not in ASP. To be added | Proposed to ASP update for clarity | Michael |  | Open
5 | 289010749 : Step 1.2 | Gasket material is very thin and transparent. Not easy to handle due to electrostatic . Waters recommend to use DI Water or Keep the gasket in more visible surface and easy to handle. | Flex to work on | Gary |  | Open
6 | 289010749 : Step 1.3 (DWG of Tool : 28901749) | Using of Omega load cell to verify the spring force 60 to 100 LBF of the fixture not mentioned in ASP. How often the loadcell to be verified with Fixture ? PM requirement for Tool verify with Load cell ? | Freq - every 3 month - to be clarified | Kwan / Gary |  | Closed | As the runaway nature of the die spring is expected to be insignificant, we will do ramped up verification of a Quarterly check via PM for the first year (before the expiry of the load cell’s calibration). By doing so, we will affirm Milford’s (Steve) statement that this load cell is primarily used for set up of the tooling, which we have already done and not required then after.
7 | 289010183 : Critical Clean | Lamina Flow hood requirement for 289010183 assy | To be discussed with Milford team on this requirement | Michael | NA | Closed | Michael commented - Lamina flow hood not required. Critical clean gloves, handling and station cleaning is must.
8 | 289004640 : Page3 Note | Need to observe difficulty in using manual torque. If long cycle time or consistency required, recommend to use "Atlas Copco Micor torque system " | Flex to observe | Gary |  | Open
9 | 2890004640 - Step 3.1 | Hard to assemble the four(4) parts on the table with torque driver. It slips easily and hard to align and tight. Flex recommended to add fixture base to hold sample needle mounting base(415001487) to align needle retainer with curved spring and shoulder screw. | To be observed during manufacturing asssembly for the fixture requirement. | Gary |  | Open
10 | 289004640 - Step 6.1 | WAT410320 - M3 Washer - Shiny side must be down . Not in ASP. Update may required | Proposed to ASP update for clarity | Michael |  | Open
11 | 2890004640 - Step 6.22 | Fixture 999038335 - Fixture to be modifed to route the Flex cable based on Milford picture ( refer photo ). | Fixture to be modified and drawings to be updated. | Kwan/Gary |  | Closed | Modification done and Verified with Assy
12 | 279003321 - Step 1.1 | Item P/N & Picture Identification are not matching. ASP Update required | Proposed to ASP update for clarity | Michael |  | Open
13 | TE82995-5 - RTZ Tester | ACQUITY CPSM RTZ ASSEMBLY TESTER - Final dialogue pass showing " Select PASS or FAIL below" to complete the test. If Fail clicked, it showing PASSED | Tester Software to be updated. | Michael |  | Open
14 | TE82995-5 - RTZ Tester | Needle seal calibration failure seen due to Needle adjusted too low or high with Teaching pin during calibration.Hard to see and judge needle height. | Proposed to use "gap tool" for verification of needle adjustment gap with Teach pin. |  |  | To be discussed
15 | TE82995-5 - RTZ Tester | During well plate Piercing test, Needle not piercing aluminium foil but Milford well plate aluminium foils are pierced with needle . Is it needle piercing force issue ? | To be verified with Milford Test Engineering | Michael |  | Open
Milford Well plate
Flex Well plate

## Failure Observations
CUSTOMER : WATERS
PRODUCT : Arc SM-FTN |  |  |  | Trainer : | Vy Levan
Test Assy | Tester Used | # | Failure Observed | Failure root cause | Re-work performed | After Re-work | Corrective action plan
279003321 + 279001991 | TE82995-5 ( RTZ ASSEMBLY TESTER) | 1 | 279001991 : Assy 11 - Belt Tension failure detected | Due to high Belt tension < 1.0 | Belt tension adjusted using fixture. | Test Passed | TBD
|  | 2 | 279001991 : Assy 1 - Datum base unable to rotate | Flywheel not located properly. Misalignment of fly wheel may load the datumbase. | Flywheel located properly to run the test | Test Passed | TBD
|  | 3 | Test Failed at Needle seal calibration | Teach Pin calibration not done properly. Needle not center aligned to Teach pin hole center. | Re-calibration done. | Test Passed | TBD
|  | 4 | 279001991 : Assy 5 - Belt tension read 1.72 ( marginal pass ). Not advised to continue. It may fail in Final unit test. | Due to loose belt tension | Belt tension adjusted using fixture. | Test Passed | TBD
|  | 5 | Test Failed at Needle seal calibration | Needle pin adjusted too low during Teach pin calibration | Re-calibration done. | Test Passed | TBD

## 1991
| CUSTOMER : WATERS | [6 empty] | Trainer : | Vy Levan
PRODUCT : Arc SM-FTN
| Assy | Date | 279001950 - ASSY, THETA DRIVE, CPSM |  |  |  | 279001991 - ASSY, DATUM BASE W/ THETA DRIVE |  |  |  | 1991 Identifier
|  |  | Gary | David | Jernail | Sylvester | Gary | David | Jernail | Sylvester
x | Assy 1 | Start | 2021-03-08 |  |  |  | 2021-03-08 |  |  |  | 10MK1003
|  | End | 2021-03-08 |  |  |  | 2021-03-08
x | Assy 2 | Start |  |  |  | 2021-03-08 |  |  |  | 2021-03-08 | 10MK1009
|  | End |  |  |  | 2021-03-08 |  |  |  | 2021-03-08
x | Assy 3 | Start |  |  | 2021-03-08 |  |  |  | 2021-03-08 |  | 10MK1012
|  | End |  |  | 2021-03-08 |  |  |  | 2021-03-08
x | Assy 4 | Start |  | 2021-03-08 |  |  |  | 2021-03-08 |  |  | 10MK1010
|  | End |  | 2021-03-08 |  |  |  | 2021-03-08
x | Assy 5 | Start | 2021-03-08 | [5 empty] | 2021-03-08 |  | 10MK1005
|  | End | 2021-03-08 | [5 empty] | 2021-03-08
x | Assy 6 | Start |  |  |  | 2021-03-08 |  |  |  | 2021-03-08 | 10MK1002
|  | End |  |  |  | 2021-03-08 |  |  |  | 2021-03-08
x | Assy 7 | Start |  |  | 2021-03-08 |  |  |  | 2021-03-08 |  | 10MK1016
|  | End |  |  | 2021-03-08 |  |  |  | 2021-03-08
x | Assy 8 | Start |  | 2021-03-08 | [5 empty] | 2021-03-08 | 10MK1006
|  | End |  | 2021-03-08 | [5 empty] | 2021-03-08
x | Assy 9 | Start | 2021-03-08 |  |  |  | 2021-03-08 | 2021-03-09 |  |  | 10MK1011
|  | End | 2021-03-08 |  |  |  | 2021-03-08 | 2021-03-09
x | Assy 10 | Start |  |  | 2021-03-09 |  |  |  | 2021-03-09 |  | 10MK1013
|  | End |  |  | 2021-03-09 |  |  |  | 2021-03-09
x | Assy 11 | Start |  |  |  | 2021-03-09 |  |  |  | 2021-03-09 | 10MK1004
|  | End |  |  |  | 2021-03-09 |  |  |  | 2021-03-09
x | Assy 12 | Start | 2021-03-09 |  |  |  | 2021-03-09 |  |  |  | 10MK1008
|  | End | 2021-03-09 |  |  |  | 2021-03-09
| Assy 13 | Start |  |  | 2021-03-09 |  |  |  | 2021-03-09 |  | 10MK1007
|  | End |  |  | 2021-03-09 |  |  |  | 2021-03-09
| Assy 14 | Start |  |  |  | 2021-03-09 |  |  |  | 2021-03-09 | 10MK1014
|  | End |  |  |  | 2021-03-09 |  |  |  | 2021-03-09
Assy 1 to Assy 9 - Done as individual assy.
Assy 10 to Assy 14 - Done in sequence
Trainee Vs Assy trained
| Assy's |  | Gary | David | Sylvester | Jernail
| Assy : 279001950 |  | 4 | 2 | 4 | 4
| Assy : 279001991 |  | 4 | 2 | 4 | 4
Training Observations :
279001991 : Step 3.1
As per customer , Pulley can be tighten anywhere at motor shaft , but Flex suggest to tighten with Pulley motor shaft flat surface to avoid pulley slipping
279001991 : Step 4.1
Add " Hand tight the screws "
279001991 : Step 5.4
Motor Rope and Pulley must be in parallel.
Fixture to be modified. refer picture.
279001991 : Step 1.1 to 4.1
Supporting fixture required for Datam base to support assemmbly from step 1.1 to 4.1
279001991 : Step 7.1
Test assembly per TSP - IS there any testing required ? Clarified by Milford, it has to be tested with 4640 & 3321.
Points to remember :
279001991 : Step 1.1
Datam base , Flatness measurement is critical. Uneven surface may affect the sample selection. Lead to functional failure
279001991 : step 1.2
Tolerance for the Datam base insulator is not critical and it wont affect any function. Bu the insulator shall be in line and not to project out side of the surface.
Re-Work done :
279001991 :
- 2 assemblies had process issue. Motor assy not followed as per assy step 5.2. Noticed , Motor cable routing was done in wrong direction. Re-work done
Note on B16: Loose Belt tension detected ( 1.72 deg ). Re-work done, belt tension detect 1.521
Note on B28: Failed at Theta Belt tension test <1.0 . Rework done.belt tension detect 1.539

## 10183
CUSTOMER : WATERS
PRODUCT : Arc SM-FTN | [6 empty] | Trainer : | Christine hemmingsen
Assy | Date | 289010749 - ASSY, SUPPORT SLEEVE W/ CUP |  |  |  | 289010183 - ASSY, HOUSING, NEEDLE SEAL, FTN-R |  |  |  | 10183 Identifier
|  | Gary | David | Jernail | Sylvester | Gary | David | Jernail | Sylvester
Assy 1 | Start | 2021-03-09 | [6 empty] | 2021-03-10 | 10ML1002
| End | 2021-03-09 | [6 empty] | 2021-03-10
Assy 2 | Start | 2021-03-09 | [5 empty] | 2021-03-10 |  | 10ML1014
| End | 2021-03-09 | [5 empty] | 2021-03-10
Assy 3 | Start |  |  |  | 2021-03-09 | 2021-03-10 |  |  |  | 10ML1004
| End |  |  |  | 2021-03-09 | 2021-03-10
Assy 4 | Start |  |  |  | 2021-03-09 |  | 2021-03-10 |  |  | 10ML1013
| End |  |  |  | 2021-03-09 |  | 2021-03-10
Assy 5 | Start |  |  | 2021-03-09 |  |  |  |  | 2021-03-10 | 10ML1012
| End |  |  | 2021-03-09 |  |  |  |  | 2021-03-10
Assy 6 | Start |  |  | 2021-03-09 |  |  |  | 2021-03-10 |  | 10ML1003
| End |  |  | 2021-03-09 |  |  |  | 2021-03-10
Assy 7 | Start |  | 2021-03-09 |  |  | 2021-03-10 |  |  |  | 10ML1006
| End |  | 2021-03-09 |  |  | 2021-03-10
Assy 8 | Start |  | 2021-03-09 |  |  |  | 2021-03-10 |  |  | 10ML1005
| End |  | 2021-03-09 |  |  |  | 2021-03-10
Assy 9 | Start | 2021-03-09 | [6 empty] | 2021-03-10 | 10ML1001
| End | 2021-03-09 | [6 empty] | 2021-03-10
Assy 10 | Start | 2021-03-09 | [5 empty] | 2021-03-10 |  | 10ML1010
| End | 2021-03-09 | [5 empty] | 2021-03-10
Assy 11 | Start |  |  |  | 2021-03-09 | 2021-03-10 |  |  |  | 10ML1007
| End |  |  |  | 2021-03-09 | 2021-03-10
Assy 12 | Start |  |  |  | 2021-03-09 |  |  |  | 2021-03-10 | 10ML1011
| End |  |  |  | 2021-03-09 |  |  |  | 2021-03-10
Assy 13 | Start |  |  | 2021-03-09 |  |  |  | 2021-03-10 |  | 10ML1009
| End |  |  | 2021-03-09 |  |  |  | 2021-03-10
Assy 14 | Start |  |  | 2021-03-09 |  | 2021-03-10 |  |  |  | 10ML1008
| End |  |  | 2021-03-09 |  | 2021-03-10
10749 - All 14 completed first
Then followed 10183 to complete all
Trainee Vs Assy trained
Assy's |  | Gary | David | Sylvester | Jernail
Assy : 289010749 |  | 4 | 2 | 4 | 4
Assy : 289010183 |  | 4 | 2 | 4 | 4
Training Observations :
289010183 : Critical Clean
Lamina Flow hood requirements - Waters to confirm..
289010749 : Step 1.4a, 1.4b
Trainer explain the usage of Fixture 999041838 but not in ASP. Need to Add.
289010749 : Step 1.2
Gasket material is very thin and transparent. Not easy to handle due to electrostatic . Waters recommend to use DI Water or Keep the gasket in more visible surface and easy to handle.
289010749 : Step 1.3
Using of Omega load cell to verify the spring force not mentioned in the drawing. Drawing to be updated for Load cell verification.
Points to remember :
289010749 : Step 1.3
Support sleeve shall be secured/locked in to the fixture base using Front locking mechanism. If the locking mechanism doesn't work or support sleeve easily come out may affect the pressing operation or impact fluid function.
289010183 : step 1.11
90'deg bend approx. Must be inline with support sleeve.
289010183 : step 1.13
Sleeve support must fit nicely with locking nut through spring. Refer Do's & Dont's
289010183 : Step 1.14
Sleeve support must easy compress with housing when inserted. Hard pressing may impact fluid flow and fail functional test.

## 4640
CUSTOMER : WATERS
PRODUCT : Arc SM-FTN |  |  | Trainer : | Vy Levan |  |  | Trainer : | Vy Levan |  |  | Trainer : | Vy Levan
Assy | Date | 289004640 - ASSY, NEEDLE CARRIAGE ACQ FTN |  |  |  | 289004640 - ASSY, NEEDLE CARRIAGE ACQ FTN |  |  |  | 289004640 - ASSY, NEEDLE CARRIAGE ACQ FTN |  |  |  | 4640 Identifier
|  | Gary | David | Jernail | Sylvester | Gary | David | Jernail | Sylvester | Gary | David | Jernail | Sylvester
Stage 1 ( 3 split ) |  | From Step 1.1 to step 4.3 |  |  |  | From Step 4.4 to Step 5.13 |  |  |  | From Step 5.14 to Step 8.8
Assy 1 | Start | 2021-03-10 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12 |  |  |  | 11MM1002
| End | 2021-03-10 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12
Assy 2 | Start |  |  |  | 2021-03-10 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12 | 11MM1004
| End |  |  |  | 2021-03-10 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12
Assy 3 | Start |  | 2021-03-10 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12 |  |  | 11MM1006
| End |  | 2021-03-10 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12
Assy 4 | Start |  |  | 2021-03-11 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12 |  | 11MM1003
| End |  |  | 2021-03-11 |  |  |  | 2021-03-11 |  |  |  | 2021-03-12
Stage 2 ( 2 Split ) |  | 289004640 - ASSY, NEEDLE CARRIAGE ACQ FTN | [7 empty] | 289004640 - ASSY, NEEDLE CARRIAGE ACQ FTN
|  | Gary |  | David |  | Jernail |  | Sylvester |  | Gary | David | Jernail | Sylvester
|  | From Step 1.1 to step 5.13 | [7 empty] | From Step 5.14 to Step 8.8
Assy 5 | Start | 2021-03-11 | [7 empty] | 2021-03-12 |  |  |  | 11MM1005
| End | 2021-03-11 | [7 empty] | 2021-03-12
Assy 6 | Start | [6 empty] | 2021-03-11 |  |  |  |  | 2021-03-12 | 11MM1009
| End | [6 empty] | 2021-03-11 |  |  |  |  | 2021-03-12
Assy 7 | Start |  |  |  |  | 2021-03-11 | [5 empty] | 2021-03-12 |  | 11MM1007
| End |  |  |  |  | 2021-03-11 | [5 empty] | 2021-03-12
Assy 8 | Start | 2021-03-11 | [7 empty] | 2021-03-15 |  |  |  | 11MM1008
| End | 2021-03-11 | [7 empty] | 2021-03-15
Stage 3 ( Full ) |  | 289004640 - ASSY, NEEDLE CARRIAGE ACQ FTN
|  | Gary |  |  | David |  |  | Jernail |  |  | Sylvester
From Step 1.1 to step 8.8
Assy 9 | Start | [9 empty] | 2021-03-15
| End | [9 empty] | 2021-03-15
Assy 10 | Start | [6 empty] | 2021-03-15 | [5 empty] | 11MM1011
| End | [6 empty] | 2021-03-15
Assy 11 | Start | 2021-03-15 | [11 empty] | 11MM1010
| End | 2021-03-15
Assy 12 | Start | [9 empty] | 2021-03-15 |  |  | 11MN1012
| End | [9 empty] | 2021-03-15
Assy 13 | Start | [6 empty] | 2021-03-15 | [5 empty] | 11MM1013
| End | [6 empty] | 2021-03-15
Assy 14 | Start |  |  |  | 2021-03-16
| End |  |  |  | 2021-03-16
Trainee Vs Assy trained
Assy's |  | Split | Gary | David | Sylvester | Jernail
289004640 ( step 1.1 to Step 4.3 ) |  | Split 3 | 1 | 1 | 1 | 1
289004640 ( step 4.4 to Step 5.13) |  |  | 1 | 1 | 1 | 1
289004640 ( step 1.1 to Step 5.13) |  | Split 2 | 2 | 0 | 1 | 1
289004640 ( step 5.14 to Step 8.8) |  |  | 3 | 1 | 2 | 2
289004640 ( Full ) |  | Full | 1 | 1 | 2 | 2
Training Observations :
2890004640 - Page 3 NOTE :
Atlas Copco Micro torques system used by Waters Milford.
Need to observe difficulty in using manual torque. If long cycle time or consistency required, recommend to use "Atlas Copco Micor torque system "
2890004640 - Step 3.1
Hard to assemble the four(4) parts on the table with torque driver. It slips easily and hard to align and tight. Flex recommended to add fixture base to hold sample needle mounting base(415001487) to align needle retainer with curved spring and shoulder screw.
289004640 - Step 6.1
WAT410320 - M3 Washer - Shiny side must be down . Not in ASP. Update required.
2890004640 - Step 6.22
Fixture 999038335 - Fixture to be modifed to route the Flex cable based on Milford picture ( refer photo ).
Points to remember :
Step 1.3 : Leave screw loose.
Recommended to tight full and loose.
Step 3.1 For Easy Assemble ( trainer recommended )
Insert screw(4) in to the driver, insert 3 & 2 with driver then assemble with mounting base(1)
Step 1.1 to Step 8.8 - Expected operator mistake - Most of the screws wil look similar and easy to make mistake of using wrong screws.
Step 5.10 for easy assembly ( trainer recommended )
Use screwdriver to slide in drive pulley to insert with carriage.
Multiple action required to assemble and align the idler pulley ( Trainee to focus this step )
Step 6.1 ( trainer recommended )
M3 Washer ( WAT410320 ) shiny side down . Not spelled in ASP.
Step 6.11
Flex cable required visual check before assy. Any small scratch/damage make functional failure.
Step 6.21
Flex cable connector , black dot on the top . Wrong insertion damage the flex cable connector.
Step 6.28 ( critical process )
( Refer picture : Check actuation of carriage assembly. The loop should not extend out in this manner as it can get caught on other components in the assembly. )

## 3321
CUSTOMER : WATERS
PRODUCT : Arc SM-FTN |  |  |  | Trainer : | Vy Levan
3321 Assy | Pairing 4640 | Date | 279003321 - ASSY, RZZ MECH, CPSM 2, FTN-R |  |  |  | 3321 Identifier
|  |  | Gary | David | Jernail | Sylvester
Assy 1 | Assy 1 | Start | 2021-03-16 |  |  |  | 11MN1002
|  | End | 2021-03-16
Assy 2 | Assy 4 | Start |  |  | 2021-03-16 |  | 11MN1003
|  | End |  |  | 2021-03-16
Assy 3 | Assy 2 | Start |  |  |  | 2021-03-16 | 11MN1004
|  | End |  |  |  | 2021-03-16
Assy 4 | Assy 5 | Start | 2021-03-16 |  |  |  | 11MN1005
|  | End | 2021-03-16
Assy 5 | Assy 7 | Start |  |  | 2021-03-16 |  | 11MN1007
|  | End |  |  | 2021-03-16
Assy 6 | Assy 3 | Start |  | 2021-03-16 |  |  | 11MN1006
|  | End |  | 2021-03-16
Assy 7 | Assy 6 | Start |  |  |  | 2021-03-16 | 11MN1009
|  | End |  |  |  | 2021-03-16
Assy 8 | Assy 8 | Start | 2021-03-16 |  |  |  | 11MN1008
|  | End | 2021-03-16
Assy 9 | Assy 14 | Start |  | 2021-03-16
|  | End |  | 2021-03-16
Assy 10 | Assy 10 | Start |  |  | 2021-03-17 |  | 11MN1011
|  | End |  |  | 2021-03-17
Assy 11 | Assy 9 | Start |  |  |  | 2021-03-17
|  | End |  |  |  | 2021-03-17
Assy 12 | Assy 13 | Start |  |  | 2021-03-17 |  | 11MN1013
|  | End |  |  | 2021-03-17
Assy 13 | Assy 12 | Start |  |  |  | 2021-03-17 | 11MN1012
|  | End |  |  |  | 2021-03-17
Assy 14 | Assy 11 | Start | 2021-03-17 |  |  |  | 11MN1010
|  | End | 2021-03-17
Trainee Vs Assy trained
Assy's |  | Gary | David | Sylvester | Jernail
Assy : 279003321 |  | 4 | 2 | 4 | 4
Training Observations :
279003321 - Step 1.1
Item P/N & Picture Identification are not matching. ASP Update required
Points to remember :
Page 20 - Step 4.1 & 4.2
For easy assembly , Recommend to install 4.2 , then 4.1
Page 11 - Step 1.2
Firmly push the spline shaft in during screwing.
Page 14 - Step 2.3
Firmly push the leadscrew in during screwing.
Re-Work done :
Assy 3 - Needle carriage mounted reverse. Wrong mount located during RTZ testing. Re-work done.
Note on B12: 4640 mounted reversly in to 3321. Re-work done

## RTZ Testing
CUSTOMER : WATERS
PRODUCT : Arc SM-FTN | [5 empty] | Trainer : | Vy Levan
Test Seq | Assy combination used |  |  | Tested by | 735002271 - ACQUITY CPSM RTZ ASSEMBLY TESTER TE82995-5 | [7 empty] | Test Result | Tested on
| 1991 Identifier | 3321 Identifier | 4640 Identifier |  | Belt tension | Needle seal position | Needle seal force | 1st move error | encoder error | Zp axis | Z axis | R axis
1 | Assy 1 10MK1003 | Assy 1 11MN1002 | Assy 1 11MM1002 | David | 1.377 | Not captured. Unit Passed but clicked "FAIL" to see Dialog box result. Refer Test Observations | [6 empty] | Pass | 2021-03-17
|  |  |  | Sylvester | 1.354 | 52.15 | 56.43 | 0.13 deg | -0.23 deg | -1 | 1 | 0 | Pass | 2021-03-17
|  |  |  | Jernail | 1.35 | 52.24 | 56.48 | 0.13 deg | -0.18 deg | -1 | 1 | -1 | Pass | 2021-03-17
|  |  |  | Gary | 1.476 | 48.09 | 56.5 | 0.13 deg | 0.18 deg | -1 | 1 | -4 | Pass | 2021-03-17
2 | Assy 7 10MK1016 | Assy 2 11MN1003 | Assy 4 11MM1003 | Jernail | 1.098 | 53.51 | 56.61 | 0.13 deg | 0.36deg | 0 | 1 | 8 | Pass | 2021-03-18
3 | Assy 6 10MK1002 | Assy 3 11MN1004 | Assy 2 11MM1004 | Sylvester | 1.143 | 54.53 | 56.34 | 0.13 deg | 0.13 deg | -1 | -12 | 0 | Pass | 2021-03-18
4 | Assy 5 10MK1005 | Assy 4 11MN1005 | Assy 5 11MM1005 | Gary | 1.521 | 55.71 | 56.81 | 0.13 deg | 0.18 deg | -1 | 4 | -2 | Pass | 2021-03-18
6 | Assy 8 10MK1006 | Assy 5 11MN1007 | Assy 7 11MM1007 | Jernail | 1.377 | 49.53 | 57.11 | 0.13 deg | 0.18 deg | -1 | -12 | 9 | Pass | 2021-03-18
5 | Assy 11 10MK1004 | Assy 6 11MN1006 | Assy 3 11MM1006 | David | 1.539 | 48.43 | 56.72 | 0.13 deg | 0.14 deg | -1 | -2 | 0 | Pass | 2021-03-18
8 | Assy 4 10MK1010 | Assy 7 11MN1009 | Assy 6 11MM1009 | Sylvester | 1.287 | 52.24 | 56.4 | 0.13 deg | -23 deg | 0 | -13 | 1 | Pass | 2021-03-18
7 | Assy 13 1310MK1007 | Assy 8 11MN1008 | Assy 8 11MM1008 | David | 1.377 | 49.53 | 57.11 | 0.13 deg | 0.18 deg | 1 | -12 | 9 | Pass | 2021-03-18
|  | Assy 9 | Assy 14
10 | Assy 9 10MK1011 | Assy 10 11MN1011 | Assy 10 11MM1011 | Jernail | 1.606 | 50.88 | 56.49 | 0.13 deg | 0.18 deg | -1 | 2 | 2 | Pass | 2021-03-18
|  | Assy 11 | Assy 9
12 | Assy 14 10MK1014 | Assy 12 11MN1013 | Assy 13 11MM1013 | David | 1.31 | 51.65 | 56.57 | 0.13deg | 0.23 deg | 1 | -2 | 3 | Pass | 2021-03-18
11 | Assy 2 10MK1009 | Assy 13 11MN1012 | Assy 12 11MN1012 | Sylvester | 1.422 | 53.17 | 56.27 | 0.13 deg | 0.32 deg | 0 | -11 | 0 | Pass | 2021-03-18
9 | Assy 12 10MK1008 | Assy 14 11MN1010 | Assy 11 11MM1010 | Gary | 1.112 | 51.56 | 56.41 | 0.13 deg | 0.32 deg | 1 | 5 | -1 | Pass | 2021-03-18
Trainer inputs :
- Purpose of the RTZ tester is to make sure belt tension and needle carriage movements are good. Not for accuracy. Assy Passed will be further calibrated with Full unit testing.
- If the needle not center aligned, system will trigger failure. It wont allow to pass.
- Need seal change required every quarterly or Half yearly due to spring life cycle.
- Perfect belt tension is between 1.2 to 1.6 deg. If belt tension read <1.1 or >1.6, Please redo belt tension adjustment.
- If the belt tension is too tight, it fall below 1.0 deg
- if the belt tension is too loose, it fall above 1.6deg.
- Before the test, RTZ assy , needle carriage assy must be moved to right end of the assy. If assy not moved to right end, test may fail.
- During needle calibration, needle should not go inside the Teach pin. Must be above the calibration pin surface.
- If the test fail, Please "power off the Switch" first . Then follow dialogue box instruction to continue or retest.
Testing Matrix
- 279001991 manufactured assy to be tested with 279001993 master assy.
- 279003321 manufactured assy to be tested with 279001991 master assy.
- 279001991 manufactured assy and 279003321 manufactured unit can be tested together.
- Once tested, both PASS assy can be used in any chassis assy.
- Not an requirement to use both as pair in same assy.
Tester observations
- Final Dialogue box showing Pass or Fail to choose .. If choose "Fail" result showing "Passed". Test software correction required.
- During well plate Piercing test, Needle not piercing aluminium foil but Milford well plate aluminium foils are pierced with needle .
Note on C14: Loose Belt tension detected ( 1.72 deg ). Re-work done, belt tension detect 1.521.
Note on C16: Failed at Theta Belt tension test. Rework done. Belt tension detect 1.539

## 3341
CUSTOMER : WATERS
PRODUCT : Arc SM-FTN |  |  |  | Trainer :
3341 Assy | 289010714 Identifier no. | Date | 279003341 - ASSY, DRIVE, SYRINGE 1/4-20 ( * Build with decoupled 289010714 )
|  |  | Gary | David | Jernail | Sylvester
Assy 1 |  | Start
End
Assy 2 |  | Start
End
Assy 3 |  | Start
End
Assy 4 |  | Start
End
Assy 5 |  | Start
End
Assy 6 |  | Start
End
Assy 7 |  | Start
End
Assy 8 |  | Start
End
Assy 9 |  | Start
End
Assy 10 |  | Start
End
Assy 11 |  | Start
End
Assy 12 |  | Start
End
Assy 13 |  | Start
End
Assy 14 |  | Start
End
Trainee Vs Assy trained
Assy's |  | Gary | David | Sylvester | Jernail
Assy : 279003341
Training Observations :
Points to remember :