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1. [bookmark: _Toc527967776]Preparing Needle Wash and Blank Sample
1.1. [bookmark: _Toc527967777]Needle Wash Solution
1.1.1. Discard the Needle Wash at the end of the work week. 
1.1.2. Do not let Needle Wash bottles with Needle Wash solution sit out over holidays and/or weekends.
1.1.3. [bookmark: _GoBack]Protect the BSM lines over holidays and/or weekends by placing them in a sealed, clean, empty bottle or in a clean bag.
1.1.4. Prepare fresh Needle Wash solution at the beginning of each work week.
1.1.5. Clean the Needle Wash bottles and caps by rinsing with 100% MEOH.
1.1.6. Clean the graduated cylinders and funnels used to prepare Needle Wash each time the solution is made by rinsing with 100% MEOH.
1.1.7. Do not make more Needle Wash than can be used in the work week.

1.2. [bookmark: _Toc527967778]Preparing Blank Sample
1.2.1. Always use new vials and vial caps. Do not reuse vials or vial caps.
1.2.2. Prepare the Blank Sample when the Needle Wash solution is made. Do not open the Needle Wash bottle at the test system to aspirate more Blank Solution.
1.2.3. Using a new pipette aspirate some of the Needle Wash solution and dispense in a new vial. It is not required to fill the vial. About ¼ full is adequate. If the pipette cannot reach the Needle Wash solution in its bottle, pour some of the Needle Wash into a clean flask and aspirate from there.
1.2.4. Cap Blank vials using a white vial cap and store in a well plate.
1.2.5. Prepare as many vials as there are instruments to test that week.
1.2.6. Use a new blank vial on each UUT.

1.3. [bookmark: _Toc527967779]Sample Handling
1.3.1. Do not allow well plates with sample vials to sit out on benches. Store in a cabinet or drawer.
1.3.2. Do not touch the top of the vial caps.
1.3.3. Discard Blank sample vial after each UUT is tested. Do not reuse on the next UUT.
1.3.4. [bookmark: _Toc526422537][bookmark: _Toc527967780]Discard all unused sample vials at the end of the work week.
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