# Preparing Needle Wash and Blank Sample

### Needle Wash Solution

    - Discard the Needle Wash at the end of the work week.

    - Do not let Needle Wash bottles with Needle Wash solution sit out over holidays and/or weekends.

    - Protect the BSM lines over holidays and/or weekends by placing them in a sealed, clean, empty bottle or in a clean bag.

    - Prepare fresh Needle Wash solution at the beginning of each work week.

    - Clean the Needle Wash bottles and caps by rinsing with 100% MEOH.

    - Clean the graduated cylinders and funnels used to prepare Needle Wash each time the solution is made by rinsing with 100% MEOH.

    - Do not make more Needle Wash than can be used in the work week.

### Preparing Blank Sample

    - Always use new vials and vial caps. Do not reuse vials or vial caps.

    - Prepare the Blank Sample when the Needle Wash solution is made. Do not open the Needle Wash bottle at the test system to aspirate more Blank Solution.

    - Using a new pipette aspirate some of the Needle Wash solution and dispense in a new vial. It is not required to fill the vial. About ¼ full is adequate. If the pipette cannot reach the Needle Wash solution in its bottle, pour some of the Needle Wash into a clean flask and aspirate from there.

    - Cap Blank vials using a white vial cap and store in a well plate.

    - Prepare as many vials as there are instruments to test that week.

    - Use a new blank vial on each UUT.

### Sample Handling

    - Do not allow well plates with sample vials to sit out on benches. Store in a cabinet or drawer.

    - Do not touch the top of the vial caps.

    - Discard Blank sample vial after each UUT is tested. Do not reuse on the next UUT.

    - Discard all unused sample vials at the end of the work week.