MSV Test Method Description

Prime System (as required)

- Home both Port Select Valve and Dead End Valve

- Connect system fluidic tubing to 3-way connector

- Set pump to 100uL/min

- Wait for fluid to drip from tubing at Dead End Valve

- Click “Next Path”

- Step Port Select Valve

- Set pump to 100uL/min

- Wait for fluid to drip from tubing at Dead End Valve

- Click “Done Prime”

- Stop pump

- Click “Exit”

Test Valve

- Test Setup

  - Align slit on MSV

  - Place MSV on test fixture

  - Connect Lever

  - Connect tubing

- MSV Ports 2 to 3

  - User Prompt: Set MSV flow to ports 2 to 3

  - Move lever to test port 2 to 3

  - Set port select and dead valves to Prime ports 2 to 3

  - Set pump to flow 50uL/min

  - Wait for pump flow rate to reach 50uL/min

    - Check Transducer Pressure – Verify Not Exceeding Max Pressure

  - Set Pressure 210 PSI

  - Set valves to Test ports 2 to 3 (dead end)

    - Check Transducer Pressure – Verify Not Exceeding Max Pressure

  - Wait for transducer pressure to be 210 PSI +/- 3

  - Wait for flow meter to stabilize

    - 20 consecutive readings +/- 0.02 uL/min of previous reading

  - Take measurement as the average of the last 5 readings

  - Set pressure 0

- MSV Ports 1 to 3

  - User Prompt: Set MSV flow to ports 1 to 3

  - Move lever to test port 1 to 3

  - Set port select and dead valves to Prime ports 1 to 3

  - Set pump to flow 50uL/min

  - Wait for pump flow rate to reach 50uL/min

    - Check Transducer Pressure – Verify Not Exceeding Max Pressure

  - Set Pressure 210 PSI

  - Set valves to Test ports 1 to 3 (dead end)

    - Check Transducer Pressure – Verify Not Exceeding Max Pressure

  - Wait for transducer pressure to be 210 PSI +/- 3

  - Wait for flow meter to stabilize

    - 20 consecutive readings +/- 0.02 uL/min of previous reading

  - Take measurement as the average of the last 5 readings

  - Set pressure 0.

- MSV Ports 1 to 2

  - User Prompt: Set MSV flow to ports 1 to 2

  - Move lever to test port 1 to 2

  - Set Pressure 210 PSI

  - Set valves to Test ports 1 to 2

  - Wait for transducer pressure to be 210 PSI +/- 3

  - Set valves to Test ports 1 to 2 in opposite direction

  - Wait for transducer pressure to be 210 PSI +/-3

  - Wait for flow meter to stabilize

    - 20 consecutive readings +/- 0.02 uL/min of previous reading

  - Take measurement as the average of the last 5 readings.

  - Set pressure 0.

- Clean Up / Preparation for Next MSV

  - Dead End the Dead End Valve

  - Set Pressure 0.

  - User Prompt: Disconnect the MSV