## Slide 1: Qualification Binder Setup

## Slide 2: Binder Sections
Qualification Overview
Installation Qualification
Maintenance Documentation
Operational Qualification
Supporting Documentation

## Slide 3: Qualification Overview
  Includes both the ‘Qualification Approval Form’ and ‘Identification Of The Performer and The Reviewer’ form.
  Release notes, change notes, or revision history
    For QWB only: Change notes document any changes since the previous revision
Compliance Specialist Training Certificate
  Optional: Include additional training certificates (for software and instruments)
  Customer purchase order (PO), if available or requested by the customer
  (The document ‘Recommended Procedures Following Repair’ is included
  in the binder with the Qualification Overview document)

## Slide 4: Installation Qualification (IQ)
Place IQ documentation in order of:
  Software
  Detector(s)
  Pump
  Injector
Optional: Include Product Certificates, such as:
  Certificate of Structural Integrity
  Certificate of Performance or Certificate of Conformity

## Slide 5: Maintenance Procedures (MP)
Forms or procedures for recording the recommended maintenance procedures
Place MP documentation in order of:
  Software
  Detector(s)
  Pump
  Injector
Logs for recording software backups
Logs for recording module maintenance, repair, and system qualification

## Slide 6: Operational Qualification (OQ)
Place OQ documentation in order of:
  Software
  Detector(s)
  Pump
  Injector
Insert Certificate(s) of Analysis (COA) for any chemistry kit or column
Insert certificates of calibration of the tools used
Qualification Approval form

## Slide 7: Supporting Documentation
Documentation pertinent to the qualification
Does not have a discreet binder section
Discuss placement with the customer prior to performing the qualification service
May contain the System Setup document, Test Descriptions and Specifications, etc.