DOCUMENT TITLE: ACQUITY PREMIER TUV FIRMWARE DOWNLOAD/PRETEST

SYSTEM/ASSEMBLY NAME: Acquity premier tuv detector

PART NUMBER AFFECTED 186018006, 800000534

TEST EQUIPMENT NAME: TUV DOWNLOAD / PRETEST STATION

TEST EQUIPMENT NUMBER: te82964

TEST SOFTWARE: *N/A*

| AUTHOR(S): eric turner | DATE: June 12, 2020 |
| --- | --- |
| REVISED BY: n/a | DATE: N/A |

# Equipment Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| 1 | TE82964 | TUV Download /Pretest Station (Empower Work Station with system suitable option) |
| 1 | NA | RS232 Serial Cable |
| 1 | NA | Ethernet Cable |
| 1 | NA | Appropriately grounded anti-static wrist strap |
| 1 | 289002450 | TUV Medium Shunt |
| 1 | TE82989 | Leak Detector Simulator |

# reference documentation

| Document # | Description |
| --- | --- |
| WAT046538 | GENERIC HI-POT & GND CONT Test PROC. |

# Definitions and Abbreviations

| Term | Definition |
| --- | --- |
| ∆ Production Support Checks (PSC) | Tests performed in support of product design specifications to assure product quality, and/or verify pre/post test conditions. |

# Test SPecifications and Traceability Matrix

| Test # | Test Specification & Condition | Acceptance criteria | Reference |
| --- | --- | --- | --- |
| 1 | Visual Inspection | Fully assembled with no obvious defects | Production Support Check (PSC) |
| 2 | Electrical Safety | Pass | WAT046538 |
| 3 | Calibration | Unit calibrates without errors | Production Support Check (PSC) |
| 4 | Fans | Exhausts out rear and inside of unit | Production Support Check (PSC) |

- Visual inspection

  - Complete header information for Instrument Test Record.

  - Visually inspect unit for obvious defects.

  - Check that all cables are seated properly into their perspective connectors.

  - Check for proper ground connections.

  - Record Pass/Fail results for the Visual Inspection test onto Instrument Test Record.

- electrical safety tests

  - Perform the Electrical Safety Tests per High Pot and Ground Continuity Test Procedure, document WAT046538.

  - Record Pass/Fail results for Electrical Safety Test onto Instrument Test Record.

- Firmware loading

  - Connect power cord, RS232 cable and Ethernet cable to detector.

  - Switch detector power on.

  - From the desktop, double click on the LOADER icon.

  - In the firmware Loader Dialog, select the “Scan” box then click on the “Next” button.

  - Loader will show the Ethernet port on instrument available.

  - Select Ethernet download by clicking on it then click the “OK” button.

  - Select the appropriate firmware file (ACQUITY TUVII Detector.sft) for the instrument type then click OK.

  - Click on the “Next” button to start downloading the software.

  - Click on the “Finish” button when the download completes.

  - Switch instrument power OFF.

  - Record firmware version onto header of Instrument Test Record.

- power on / calibration

  - Install analytical cell 289002450 Medium Shunt into bench.

  - Connect flowcell ID cable to shunt

  - Plug TE82989 Leak Detect Simulator into panel of Acquity TUV. Set leak switch OFF

  - Switch unit power ON.

  - Select Acquity Terminal icon on desktop.  Type setserialnumber xxxxxxxxx (xxxxxxxxx is unit serial #), enter.

  - Double click the Stop Acquity Processes icon on desktop.

  - Double click the Acquity Console icon on desktop.

  - Select TUV Detectors.

  - Click on Control, and then click Reset TUV. (This will calibrate unit)

  - Verify that unit calibrates. (Lamp LED will be solid green when successfully calibrated.)

  - Record Pass/Fail results for Calibration onto Instrument Test Record.

- leak detect test

  - Select / launch Acquity Terminal on desktop.

  - TE82989 Leak Detect Simulator should be plugged in and leak switch OFF.

  - Verify that Leak Detect LED on simulator is illuminated.

  - Type dump leakdetector. Press Enter. Verify response:

* PRESENT -> * NO Leak

  - Set Leak switch ON.

  - Type dump leakdetector. Press Enter. Verify response:

* PRESENT -> * Leak detected

  - Record Pass/Fail results for Leak Detect onto Instrument Test Record.

- Fan tests

  - Verify that the rear panel fan is turning, and that air is blowing out the back of unit.

  - Verify that the two internal partition fans are turning, and that air is blowing towards the front of the unit.

  - Record the Pass/Fail results for the Fan Test onto Instrument Test Record.

- lamp id communications test

  - Enable ibutton data dump.

    - Type edit ibutton. Press Enter.

    - Type true. Press Enter.

  - Type dumplampmemory. Press Enter. Verify response:

Lamp Device Found

  - Record the Pass/Fail results for Lamp ID Communications Test onto Instrument Test Record.

- flow cell id communications test

  - Type dumpflowcellmemory. Press Enter. Verify response:

Flow Cell Device Found

  - Record Pass/Fail results for Flow cell ID Communications Test onto Instrument Test Record.

  - Disable ibutton data dump.

    - Type edit ibutton. Press Enter.

  - Type false. Press Enter.

#### Waters Corporation - TEST RECORD

| Part Number: | Product: |
| --- | --- |
| Operator: | Start Date/Time: |
| Serial Number: | End Date/Time: |
| Test System: | Order: |
| Test Software: | Location: |

TMD:

Firmware:

| Test Name | LSL | Measurement | USL | Unit |  | Outcome |
| --- | --- | --- | --- | --- | --- | --- |
| Visual Inspection | — | — | — | — |  | PASS FAIL See Event Log |
| Electrical Safety | — | — | — | — |  | PASS FAIL See Event Log |
| Calibration | — | — | — | — |  | PASS FAIL See Event Log |
| Leak Detect Test | — | — | — | — |  | PASS FAIL See Event Log |
| Fans | — | — | — | — |  | PASS FAIL See Event Log |
| Lamp ID | — | — | — | — |  | PASS FAIL See Event Log |
| Flowcell ID | — | — | — | — |  | PASS FAIL See Event Log |
| Note: Fill in 'N/A' if not applicable. |  |  |  |  |  |  |
| Instrument Overall Result - PASS FAIL |  |  |  |  |  |  |
| Signed: |  |  |  |  | Date: |  |

END OF DATA