# INSTALLATION QUALIFICATION REPORT

For the PCM II Final Test Station (TE82817-1)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| James Stors | Senior Test Validation Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Senior Operations Quality Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

The “Reviewed By” signature[s] denotes agreement with the description of the qualification activities.

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Emma Pierce | First document for approval |

# Purpose

##### The PCM II Final Test Station (TE82817-1) is located in the NPI laboratory in Milford, MA.  The station is used to perform the final instrument test of the following part numbers: 186002449 and 279001290.  The Test Station has two versions of the test software which are identical in functionality.  The PCM II Software tests the PCM II (Part #186002449. Refurb. Part #800000178) and the SFO Software tests the PCM II sub-assembly (Part #279001290, Refurb. Part #800000224).

##### In the final test, the UUT is connected to the test station and the performance of the rear panel I/O, the pump status, pump pressure, and pump control is verified to operate within design specifications.  When complete, a hard copy of the final test record is printed.

##### This Installation Qualification is being performed as required by the Quality Plan #2017-002.

# Scope of Report

This Installation Qualification Report details the qualification activities undertaken for the PCM II Final Test Station (TE82817-1) based on the requirements of the Installation Qualification Procedure, 730004135 Ver. 00.

This Installation Qualification Report defines the documents used to qualify the PCM II Final Test Station (TE82817-1).

# Summary of Results - Statement of Fitness for Intended Use

The PCM II Final Test Station was tested in accordance with the Installation Qualification Procedure, 730004135 Ver. 00.  All Test Scripts were performed and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the PCM II Final Test Station complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Qualification Report, the PCM II Final Test Station (TE82817-1) can be considered fully qualified and fit for intended use.

# Glossary

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| TSP | Test Specification and Procedure |
| CAL/PM | Calibration/Preventive Maintenance |
| NPI | New Product Introduction |
| UUT | Unit Under Test |
| S/N | Serial Number |
| P/N | Part Number |
| SFO | System Fluidics Organizer |
| PCM II | Pump Control Module II |
| N/A | Not Applicable |

# Training

The operator of the PCM II Final Test Station (TE82817-1) is the executor of this qualification, and was trained according to the PCM II Final Test Procedure, 735000887, Ver. 06 and the Installation Qualification Procedure, 730004135, Ver. 00.  Refer to Attachments 7.16 and 7.17 for the training records.

# Qualification Results

## Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.1.

#### The PCM II Final Test Procedure, 735000887, Ver. 06, was confirmed to be in SAP with a status of Release General, as evidenced below.

#### The TE82817-1 PM Procedure, 735002828, Ver. 00, was confirmed to be in SAP with a status of Release General, as evidenced below.

## Test Script IQ-02: Site Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.2.

#### The ambient temperature and humidity were measured using an Extech 445703 Hygro-thermometer, ID#95648, and was determined to be within specification as defined by Performance Specification, PCB, Pump Control WAT200454, Ver. 00.  The calibration record is provided as Attachment 7.3.

#### The electrical supply was measured using a Fluke 189 Multimeter, ID#98209, and was determined to be within specification as defined by Performance Specification, PCB, Pump Control WAT200454, Ver. 00.  The calibration record is provided as Attachment 7.4.

#### The Test Station was confirmed to have more than the three (3) dedicated electrical outlets required.

#### The space and test bench of the PCM II Final Test Station (TE82817-1) were confirmed to have adequate space for testing and service.  The bench was confirmed to be large and strong enough to support all test equipment.  A picture of the Test Station work area is provided as Attachment 7.5.

#### The Test Station was confirmed to have one (1) network available.

## Test Script IQ-03: Instrumentation Calibration Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed script is provided as Attachment 7.6.

#### The PCM II Final Test Station (TE82817-1) was confirmed to be in the Calibration Manager Database with a twelve (12) month maintenance cycle.

#### The Test Station was confirmed to have had preventative maintenance within this time frame.  A screen capture of the PCM II Final Test Station (TE82817-1) from the Equipment Master Database is provided as Attachment 7.7.

## Test Script IQ-04: Test Software and Instrument Firmware Version was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.8.

#### The Test Equipment Software version listed in PCM II Final Test Procedure 735000887, Ver. 06 is Test Software 0.0 and was verified to match the version listed in the Test Equipment program displays of the PCM II Test Program and the SFO Test Program.  A screen capture of the program displays are provided as Attachments 7.9 and 7.10, respectively.

#### The PCM II Firmware version listed in Production Master PCMII 361000331, Ver. 00 is Firmware v4.00 and was verified to match the version listed in the Loader Dialog Box.  A screen capture of the program is provided as Attachment 7.11.

## Test Script IQ-05: General Safety and Alarms was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.12.

#### It was confirmed that there were no mechanical safety pinch points or sharp edges.

#### The PCM II Final Test Station was confirmed to pass an Electrical Safety Hi-Pot/Ground Continuity test.  The calibration record is provided as Attachment 7.13.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.14.

#### One (1) UUT Part #186003765 (Serial #F17APS843M) containing the sub-assembly Part #279001290 was tested using the PCM II Final Test Station (TE82817-1) SFO program with all tests passing.  The final test record is provided as attachment 7.15.

## This table provides a summary of each test script as performed according to 7300004135 Ver. 00 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachment 7.1 | None |
| IQ-02 | Pass | Attachments 7.2, 7.3, 7.4, and 7.5 | None |
| IQ-03 | Pass | Attachments 7.6 and 7.7 | None |
| IQ-04 | Pass | Attachments 7.8, 7.9, 7.10, and 7.11 | None |
| IQ-05 | Pass | Attachments 7.12 and 7.13 | None |
| IQ-06 | Pass | Attachments 7.14 and 7.15 | None |

# Attachments

## Completed Test Script IQ-01: Associated Document Verification

## Completed Test Script IQ-02: Site Verification

## Objective Evidence IQ-02: Extech 445703 Hygo-thermometer (ID#95648) Calibration Record

## Objective Evidence IQ-02: Fluke 189 Multimeter (ID#98209) Calibration Record

## Objective Evidence IQ-02: Photo of Test Station Work Area

## Completed Test Script IQ-03: Instrument Calibration Verification

## Objective Evidence IQ-03: PCM II Final Test Station (TE82817-1) Calibration Record

## Completed Test Script IQ-04: Test Software and Instrument Firmware Version

## Objective Evidence IQ-04: Screen Capture of PCM II Test Equipment Software

## Objective Evidence IQ-04: Screen Capture of SFO Test Equipment Software

## Objective Evidence IQ-04: Screen Capture of Test Equipment Firmware

## Completed Test Script IQ-05: General Safety and Alarms

## Objective Evidence IQ-05: Hi-Pot Ground Continuity Tester (ID#90265) Calibration Record

## Completed Test Script IQ-06: Functional Verification of UUT Part #186003765 (279001290)

## Objective Evidence IQ-06: Instrument Final Test Record for UUT Part #186003765

## Completed Operator Training Record 735000887 Ver. 06

## Completed Executor Training Record 730004135 Ver. 00

# Deviation from Test Equipment Qualification

##### There were no deviations generated during the execution of the Installation Qualification Procedure 730004135 Ver. 00.

# References

##### The current approved versions of the following documents are referenced within this document:

| Document Number | Document title |
| --- | --- |
| 730004135 Ver. 00 | Installation Qualification Procedure |
| 735000887 Ver. 06 | PCM II Final Test Procedure |
| 735002828 Ver. 00 | TE82817-1 PM Procedure |
| WAT200454 Ver. 00 | Performance Spec, PCB, PMP Control |
| WAT046538 Ver. 09 | Generic Electrical Safety Test Procedure |
| 735001477 Ver. 05 | Serial Number Log, SFO |
| 361000331 Ver. 00 | Pump Control Module II Recreate Procedure v4.00.12 |
| 2017-002 | Legacy Products Relocation Quality Plan |