DOCUMENT TITLE: TUV detectors, final test procedure

SYSTEM/ASSEMBLY NAME: 2487 tuv detector, Acquity UPLC tuv detector, 2489 tuv detector, Optics Bench Spares, acquity tuv tc detector

PART NUMBER AFFECTED Wat 081110, 186015028, *186002487, WAS081147, 700002880, 700003550, 800000239, 700005245, 186015031, 800000303, 800000204, 800000102, 186015022, 700002554

TEST EQUIPMENT NAME: tuv final test station

TEST EQUIPMENT NUMBER: TE82903, TE82591

TEST SOFTWARE: Wat 081110, *186002487, WAS081147, 800000102,186015022, 700002554 *0.0* 186015028, 700002880, 700003550, 800000239, 700005245, 186015031, 800000303, 800000204 *0.1*

| AUTHOR(S): mark jussaume | DATE: November 06, 2006 |
| --- | --- |
| REVISED BY: mark jussaume | DATE: January 02, 2018 |

*To be tested as part number 289024872, instrument level for 186002487

# Equipment Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| 1 | TE82903 | TUV Final Test Station |
| AR | TE82591 | 2487 Optics Bench Test Station |
| 1 | 700004155 | Waters Two Cuvette Set |
| 1 | WAT042885 | Waters Wavelength Accuracy Calibration Solution (*see note 1 below) |
| 1 | 700004156 | Waters Six Cuvette Set (only 5 cuvettes used) |
| 1 | WAT042881 | Waters Absorbance Detector 10mm Cell Linearity Test Package (* see notes 1 and 2 below) |
| AR | 289002450 | TUV Med. Shunt |
| AR | 289001358 | Flow Cell Shunt |

NOTES:

1. There is a shelf life for both the wavelength accuracy and the linearity kits.

Check the expiration date on the kits before use.

2. Label Linearity cuvettes and fill with solutions from Test Package as follows:

Cuvette #1 fill with Ampule 1	100% Methanol

Cuvette #2 fill with Ampule 2	5.0 mg/L Propyl Paraben

Cuvette #3 fill with Ampule 4	15.0 mg/L Propyl Paraben

Cuvette #4 fill with Ampule 6	25.0 mg/L Propyl Paraben

Cuvette #5 fill with Ampule 7	30.0 mg/L Propyl Paraben

# reference documentation

| Document # | Description |
| --- | --- |
| WAT046538 | Electrical Safety Test Procedure |
| WAT08125 | 2487 Performance Specification |
| 721013304 | FSD-ACQUITY TUV Detector TC |
| 721007908 | 2489 Functional Specification |

# Definitions and Abbreviations

| Term | Definition |
| --- | --- |
| AR | As Required |
| mg | milligram |
| L | Liter |
| LSL | Lower Spec Limit |
| USL | Upper Spec Limit |
| RSD | Standard Deviation |
| nm | Nano Meter |
| AU | Absorbance Unit |
| UUT | Unit Under Test |
| ∆ Production Support Checks (PSC) | Tests performed in support of product design specifications to assure product quality, and/or verify pre/post test conditions. |

# Test SPecifications and Traceability Matrix (2487)

| Test # | Test Specification & Condition | Acceptance criteria | Reference |
| --- | --- | --- | --- |
| 1 | Drift Test | LSL = -0.0001 AU USL= 0.0001 AU | WAT081285 page 4 Detector Sensitivity: Drift: |
| 2 | Drift Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 3 | Noise Test | LSL = 0 AU USL = 0.000007 AU | WAT081285 page 4 Detector Sensitivity: Noise |
| 4 | Noise Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 5 | Dual Noise Test | LSL = 0 AU USL = 0.000050 AU | WAT081285 page 4 Detector Sensitivity: Dual Channel Noise |
| 6 | Dual Noise Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 7 | Linearity Test | LSL = 0 %RSD USL = 5 % RSD | WAT081285 page 4 Detector Sensitivity: Linearity |
| 8 | 255 nm Accuracy Test | LSL = 254 nm USL = 256 nm | WAT081285 page 4 Monochromator: Wavelength Accuracy |
| 9 | 378.8 nm Accuracy Test | LSL = 377.8 nm USL = 379.8 nm | WAT081285 page 4 Monochromator: Wavelength Accuracy |
| 10 | 522.4 nm Accuracy Test | LSL = 521.4 nm USL = 523.4 nm | WAT081285 page 4 Monochromator: Wavelength Accuracy |

# Test SPecifications and Traceability Matrix (Acquity TUV)

| Test # | Test Specification & Condition | Acceptance criteria | Reference |
| --- | --- | --- | --- |
| 1 | Drift Test | LSL = -0.000375 AU USL= 0.000375 AU | 721013304 Sec. 5.14.1.3 |
| 2 | Drift Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 3 | Noise Test | LSL = 0 USL = 0.000005 | 721013304 Sec. 5.14.1.1 |
| 4 | Noise Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 5 | Dual Noise Test | LSL = 0 AU USL = 0.000075 AU | 721013304 Sec. 5.3.1.8 |
| 6 | Dual Noise Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 7 | Thermal Control | PASS | PSC *See Note 1 |

NOTES:

*Note 1 To assure proper manufacture, each of the two PWM fans are switched off/on and the appropriate temperature response verified by reading the UUT thermocouple.

# Test SPecifications and Traceability Matrix(2489)

| Test # | Test Specification & Condition | Acceptance criteria | Reference |
| --- | --- | --- | --- |
| 1 | Drift Test | LSL = -0.0001 AU USL= 0.0001 AU | 721007908 Sec. 5.14.4 |
| 2 | Drift Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 3 | Noise Test | LSL = 0 USL = 0.000005 | 721007908 Sec. 5.15.2 |
| 4 | Noise Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 5 | Dual Noise Test | LSL = 0 AU USL = 0.000035 AU | 721007908 Sec.5.15.3 |
| 6 | Dual Noise Run Count | LSL = 0 EA USL = 99 EA | PSC |
| 7 | Linearity Test | LSL = 0 %RSD USL = 5 % RSD | 721007908 Sec.5.15.1 |
| 8 | 255 nm Accuracy Test | LSL = 254 nm USL = 256 nm | 721007908 Sec.5.15.5 |
| 9 | 378.8 nm Accuracy Test | LSL = 377.8 nm USL = 379.8 nm | 721007908 Sec.5.15.5 |
| 10 | 522.4 nm Accuracy Test | LSL = 521.4 nm USL = 523.4 nm | 721007908 Sec.5.15.5 |

- tester description

This test station allows the testing of up to 4 detectors at a time. All operator and instrument information may be entered via keyboard or barcode scanner. Operator may choose to run in Auto, Manual, or Diagnostics mode. Auto will execute each test sequentially. Manual will allow tests to be performed in any order. Diagnostics mode does not require any operator or unit information. Test data for each test, as well as the instruments final test results are saved automatically under a folder named by the instrument serial number. The INSTRUMENT FINAL TEST RECORD will print out upon request. A station overview showing station status and test progress is available for viewing or printing.

- tester setup

  - Plug Tester into a 100-240 Volts 50/60 HZ AC outlet.

  - Switch power on and allow pc to boot to logon screen. Enter “tester” for username and password.

  - Switch printer power on and make sure there is paper in tray.

- unit setup

When testing spare optics benches, place bench into appropriate test bed (an instrument chassis for 2489 & Acquity TUV benches, and TE82591 for 2487 benches). Perform instrument test. When testing TUVE bench spares, install flow cell shunt before starting test.

  - Place detector at Test Station and connect the appropriate RS232 cable and power cord. Use caution to not “Hot Plug” (instrument power switch should be off).

  - Switch detector power on.

  - (2487 & 2489 only) When the detector beeps and displays the power up diagnostics page, you must sequentially press the "Shift", ".", "1" keys. The unit will briefly display "RS232 Enabled".

  - Allow the unit to initialize and then warm up. 2487s and 2489s will display their home page, and Aquity TUVs LAMP status LED will turn solid green, when done.

  - Press "Shift Calibrate" to calibrate the 2487 or 2489.

NOTE: If the RS232 is not enabled, the detector will not communicate with the Test Station.

- test

  - Select the appropriate test program for unit type to be tested. Either the desktop icon or from the Start\Program Files\TE82903 menu.

  - Select the desired test mode, and follow the on screen instructions.

CAUTIONS:

Thermal Control Test: (TUV TC only) Units must be warmed up before performing test. The time it takes a unit to go from power on to initialization is sufficient.

Noise & Drift Tests: Units are very sensitive, so don't disturb units or benches while noise tests are running. Make sure cover is secure and front panel is on.

Linearity Test: This test uses five sealed cuvettes, marked one to five. They must be handled carefully. The smooth surfaces of these cuvettes must remain completely clean. Otherwise test results will be affected. These cuvettes are glass and may shatter under excessive pressures.

Wavelength Accuracy Test: This test uses two sealed cuvettes. They must be handled carefully. The smooth surfaces of these cuvettes must remain completely clean, otherwise test results will be affected. These cuvettes are glass and may shatter under excessive pressures.

- tests complete / removing units

  - Select “UUT DONE”. Then select YES to create test files and print report.

  - Clear lamp hours (2487 & 2489 only)

    - Press DIAG 4

    - Select lamp change, press “ENTER”.

    - Input 6 ones (1) for the serial number, press “ENTER”.

    - Enter date and year-press "ENTER" twice after the month and once after the day & year.

    - Press HOME.

    - Answer all questions with "ENTER"

Note:	For Acquity TUV TC Detectors the lamp history data/hours will automatically be cleared upon completion of test.

  - It is now safe to switch power off, disconnect cables, and remove from test station.