## Info by Mike Cottter Nov 06 14
ALLIANCE TEST EQUIPMENT - As Of November 06, 2014
Test | TE Number(s) | Test Procedure | Cal/PM Procedure | Test Equipment Name | Notes | Plant
Final Instrument | 82573 | WAT271129 | In Process | 2690 REAR PANEL TESTER | Labview (82573/Rear Panel Test) | Flextronics
| 82736 | WAT271129 | 735003457 | Alliance Final Test | Empower Project and Compression Test Config files text files
| 82733 | WAT271129 | In Process | VAC PURGE/PRIME SYSTEM | Basic Electro-Mechanical Assembly (Power Supply, Vacuum Pump)
Rheodyne Valve | 82585 | WAT050033 | WAT050032 | 2690/515 RHEODYNE VALVES TESTER | Z-World SBC
Injector/Seal Pack | 82503 | WAT043815 | WAT043816 | INJECTOR/SEAL PACK TESTER | Z-World SBC
Pump (individual) | 82519 | WAT069481 | WAT069004 | 2690 PUMP HEAD (MECHANICS) TESTER | Labview 5.1 Executable
Pump Tray | 82526 | WAT210700 | In Process | 2690 PUMP TRAY TESTER | Labview 5.1 Executable
CCD | 82520 | 735000431 | 735002998 | 2690 SAMPLE HANDLER TESTER | Labview Executable
Fluidics Tray | 82523 | WAT047159 | In Process | 2690 Fluidics Tray Tester | Z-World SBC
Check Valve | 82530 | WAT068000 | In Process | CARTRIDGE CHEK VALVE | May be in use?
Display/Keypad | 82546 | WAT242102 | In Process | 2690 KEYPAD/DISPLAY TESTER | Computer with Serial Port (or serial communication terminal, VT220, etc.)
HPMV | 82625 | 735000287 | 735000288 | HPMV Tester | Z-World SBC
Syringe Drive | 82656 | 735000300 | In Process | 2790/2690 Syringe Drive Tester | Basic Electronics with Stepper Driver IC (IC getting scarce)
Column Heater/Cooler | 82774 | 735000698 | 735000727 | Alliance Column Heater/Cooler Tester | Labview Executable, located at Venture | Venture

## Info Inspected for Milford
ALLIANCE TEST EQUIPMENT - As of November 12, 2014 | [10 empty] | 3 = Obsolete parts, impossible to locate | 3 = Only instance of this test equipment | 3 = Impossible or difficult to function if equipment breaks down | 3 = Difficult to repair or duplicate | 3 = High
[11 empty] | 2 = Older parts, and/or difficult to locate | 2 = Duplicate(s) testers in storage or elsewhere | 2 = Could test in instrument or otherwise improvise, w/some difficulty | 2 = Could be duplicated, with Some Effort | 2 = Moderate | AxBxCxDxE
[11 empty] | 1 = Parts easily attainable | 1 = Multiple instances at CM Facility | 1 = Could easily function (test in instrument or otherwise improvise) | 1 = Easily Duplicated | 1 = Low
Test | TE Number(s) | Test Procedure | TSP Rev | Test SW Rev in TSP | Test Software Rev displayed on screen | Cal/PM Procedure | Cal/PM Rev | Test Equipment Name | Notes | Plant | A - Component Scarcity | B - Quantity of This Tester | C - Other Alternatives? | D - Engineer's Assessment | E - Failure Risk | Relative Risk
Final Instrument | 82573 | WAT271129 | Extract from LEQ process |  | N / A | In Process | Extract from LEQ process | 2690 REAR PANEL TESTER | Labview (82573/Rear Panel Test). Flex claimed that OESR wasnot open to use newer test software revision. Old version was reinstalled after having problems with using newer version, and tester works normally again. | Flextronics | [5 empty] | 0
| 82736 | WAT271129 |  |  | N / A | 735003457 |  | Alliance Final Test | Empower Project and Compression Test Config files text files |  | 1 | 1 | 3 | 1 | 2 | 6
| 82733 | WAT271129 |  |  | N / A | In Process |  | VAC PURGE/PRIME SYSTEM | Basic Electro-Mechanical Assembly (Power Supply, Vacuum Pump) |  | 1 | 1 | 2 | 1 | 2 | 4
| N/A | N/A |  |  | N/A | N/A |  | Firmware Download Station | No TE number, but need to add to PC list to get support. |  | 1 | 1 | 1 | 1 | 1 | 1
Rheodyne Valve | 82585 | WAT050033 |  |  | N / A | WAT050032 |  | 2690/515 RHEODYNE VALVES TESTER | Z-World SBC. Flex didn't have this tester. | [6 empty] | 0
Injector/Seal Pack | 82503 | WAT043815 |  |  | N / A | WAT043816 |  | INJECTOR/SEAL PACK TESTER | Z-World SBC. Only -4 is being used. No Test Software. Concern on support of SBC in future. Use obsolete Alliance CPU Board, which requires old firmware and old PLD. Can breadboards and analog power supply be supported? | [6 empty] | 0
Pump (individual) | 82519 | WAT069481 |  |  | N / A | WAT069004 |  | 2690 PUMP HEAD (MECHANICS) TESTER | Labview 5.1 Executable. No Test Software Revision. M50 since 2006 being used. Use obsolete Alliance CPU Board, which requires old firmware and old PLD. That means it can't be replaced with new board. Require test engineering support from Matt. |  | 2 | 1 | 1 | 2 | 1 | 4
Pump Tray | 82526 | WAT210700 |  |  | N / A | In Process |  | 2690 PUMP TRAY TESTER | Labview 5.1 Executable. 3 Testers available with no dash, -1, and -4. No Test Software Revision. M50 is being used. -4 is not working because RS232 serial port failure. Concern on flow valve which needs replacement. Rheodyne valve is obsolete and no longer be supported. Replacement of COM Cable might be needed. |  | 3 | 1 | 2 | 2 | 1 | 12
CCD | 82520 | 735000431 |  |  | 1.4? Too smal to tell | 735002998 |  | 2690 SAMPLE HANDLER TESTER | Labview Executable with Test Software Rev 1.4. Indicator scratching issue is fixed. Contain both Tool number and TE Number and causing confusion. |  | 1 | 2 | 3 | 1 | 1 | 6
Fluidics Tray | 82523 | WAT047159 |  |  | N / A | In Process |  | 2690 Fluidics Tray Tester | Z-World SBC (Single Board Computer). -3 availabble. How long can we support kind of computer? Only use for spares on degasser. Tester's portion having 4 ports is not being used. Only 25% of tester is being used. | [6 empty] | 0
Check Valve | 82530 | WAT068000 |  |  | N / A | In Process |  | CARTRIDGE CHEK VALVE | May be in use? Flex didn't have tester, but has tool# 999029594 | [6 empty] | 0
Display/Keypad | 82546 | WAT242102 |  |  | N / A | In Process |  | 2690 KEYPAD/DISPLAY TESTER | Computer with Serial Port (or serial communication terminal, VT220, etc.). Flex currently uses its own PC, so Milford needs to provide PC and support it. No test software. Using windows hyperterminal. Low risk and reliable. |  | 1 | 1 | 1 | 1 | 1 | 1
HPMV | 82625 | 735000287 |  |  | N / A | 735000288 |  | HPMV Tester | Z-World SBC. -1 available. Capacity issue which might need duplication/upgrades because it currently tests 1 out 6 UUTs. How long can we continue to support Zylog SBC? | [6 empty] | 0
Syringe Drive | 82656 | 735000300 |  |  | N / A | In Process |  | 2790/2690 Syringe Drive Tester | Basic Electronics with Stepper Driver IC (IC getting scarce). No test software. Concern on obsolete parts such as IC and Power supply. It's been working well. Inside condition is so dusty and this is commonplace with panel misplaced on many testers. | [6 empty] | 0
| 82620 | 735000206 | [5 empty] | Heater/Chiller Tester | Only use for spares. It was assigned for Mark from Matt recently. | [6 empty] | 0
Column Heater/Cooler | 82774 | 735000698 | ??? | ??? | ??? | 735000727 |  | Alliance Column Heater/Cooler Tester | Labview Executable, located at Venture | Venture | [5 empty] | 0
Consideration | Description
1 | 4 PDA Detector Testers are not ready for production. Eric is working on final configuration because Milford doesn't have UUTs available. 2 units to be used until end of 2014.
2 | TE82882 - 11: Flex to connect 2 PDAs of same kind as discussed and agreed for final development by Eric. Tester should have switch connected and turned on. Serial Number to store in text file on desktop. Tester and all UUTs must be powered on as fully functional tester from 6 PM Singapore time everyday for Eric to work on.
3 | TE82526-4 : 2690 PUMP TRAY TESTER might need replacement of COM cabble
4 | TE82625-1: HPMV Tester need evaluation to meet capacity requirement. Inspect and issue Cal/PM as required. 100% testing on spares
5 | CM should refrain from tampering with critical components without proper authorization by notify WPPL to get Milford Design Engineer involved through OESR process.
6 | TE82800-3: Acutator Tester Consider using fingertight fitting with Peek ferrule to lessen frequent damage/change of pump head and tubing. Get part number and send samples to Joel. Matt Brule
7 | TE82520: Damage on indicator's plate and needle should be already fixed by raising needle higher with proper verification on 3 UUTs BEFORE and AFTER adjustment made. Pay attention to keep Master Carrier Plate flat and seating properly on Skid Plate during installation and removal of UUTs. Duplicate Tester is being made to meet capacity requirement.

## Info Inspected for WPPL
ALLIANCE TEST EQUIPMENT - As of November 12, 2014
Test | TE Number(s) | Test Procedure | Cal/PM Procedure | Test Equipment Name | Notes | Plant
Final Instrument | 82573 | WAT271129 | In Process | 2690 REAR PANEL TESTER | Labview (82573/Rear Panel Test). | Flextronics
| 82736 | WAT271129 | 735003457 | Alliance Final Test | Empower Project and Compression Test Config files text files
| 82733 | WAT271129 | In Process | VAC PURGE/PRIME SYSTEM | Basic Electro-Mechanical Assembly (Power Supply, Vacuum Pump)
| ????? |  |  | Firmware Download Station | Never has a TE number and/or can't be located. Milford will investigate and decide to assign TE#.
Rheodyne Valve | 82585 | WAT050033 | WAT050032 | 2690/515 RHEODYNE VALVES TESTER | Z-World SBC. Flex didn't have it.
Injector/Seal Pack | 82503 | WAT043815 | WAT043816 | INJECTOR/SEAL PACK TESTER | Z-World SBC. Only -4 is being used. Flex has 1 extra in storage with conditions unknown. No Test Software. Concern on support of SBC in future. Use obsolete Alliance CPU Board, which requires old firmware and old PLD. Can breadboards and analog power supply be supported?
Pump (individual) | 82519 | WAT069481 | WAT069004 | 2690 PUMP HEAD (MECHANICS) TESTER | Labview 5.1 Executable. No Test Software Revision. M50 since 2006 being used. Use obsolete Alliance CPU Board, which requires old firmware and old PLD. That means it can't be replaced with new board. Require test engineering support from Matt.
Pump Tray | 82526 | WAT210700 | In Process | 2690 PUMP TRAY TESTER | Labview 5.1 Executable. 3 Testers available with no dash, -1, and -4. No Test Software Revision. M50 is being used. -4 is not working because RS232 serial port failure. Concern on flow valve which needs replacement. Rheodyne valve is obsolete and no longer be supported. Replacement of COM Cable might be needed.
CCD | 82520 | 735000431 | 735002998 | 2690 SAMPLE HANDLER TESTER | Labview Executable with Test Software Rev 1.4. Indicator scratching issue is fixed. Contain both Tool number and TE Number and causing confusion.
Fluidics Tray | 82523 | WAT047159 | In Process | 2690 Fluidics Tray Tester | Z-World SBC (Single Board Computer). -3 availabble. How long can we support kind of computer? Only use for spares on degasser. Tester's portion having 4 ports is not being used. Only 25% of tester is being used.
Check Valve | 82530 | WAT068000 | In Process | CARTRIDGE CHEK VALVE | May be in use? Flex didn't have tester, but has tool# 999029594
Display/Keypad | 82546 | WAT242102 | In Process | 2690 KEYPAD/DISPLAY TESTER | Computer with Serial Port (or serial communication terminal, VT220, etc.). Flex currently uses its own PC, so Milford needs to support it. No test software. Using windows hyperterminal. Low risk and reliable.
HPMV | 82625 | 735000287 | 735000288 | HPMV Tester | Z-World SBC. -1 available. Capacity issue which might need duplication/upgrades because it currently tests 1 out 8 UUTs. How long can we continue to support Zylog SBC?
Syringe Drive | 82656 | 735000300 | In Process | 2790/2690 Syringe Drive Tester | Basic Electronics with Stepper Driver IC (IC getting scarce). No test software. Concern on obsolete parts such as IC and Power supply. It's been working well. Inside condition is so dusty and this is commonplace with panel misplaced on many testers.
| 82620 | 735000206 |  | Heater/Chiller Tester | Only use for spares. It was assigned for Mark from Matt recently.
Column Heater/Cooler | 82774 | 735000698 | 735000727 | Alliance Column Heater/Cooler Tester | Labview Executable, located at Venture. How come this tester located at Venture? Should WPPL take on this inspection? | Venture

## Sheet1
ALLIANCE TEST EQUIPMENT
Test | TE Number(s) | Test Procedure | Test Equipment Name | Plant
Final Instrument | 82573 | WAT271129 | 2690 REAR PANEL TESTER | Flextronics
| 82736 | WAT271129 | Alliance Final Test
| 82733 | WAT271129 | VAC PURGE/PRIME SYSTEM
Rheodyne Valve | 82585 | WAT050033 | 2690/515 RHEODYNE VALVES TESTER
Injector/Seal Pack | 82503 | WAT043815 | INJECTOR/SEAL PACK TESTER
Pump (individual) | 82519 | WAT069481 | 2690 PUMP HEAD (MECHANICS) TESTER
Pump Tray | 82526 | WAT210700 | 2690 PUMP TRAY TESTER
CCD | 82520 | 735000431 | 2690 SAMPLE HANDLER TESTER
Fluidics Tray | 82523 | WAT047159 | 2690 Fluidics Tray Tester
Check Valve | 82530 | WAT068000 | CARTRIDGE CHEK VALVE
Display/Keypad | 82546 | WAT242102 | 2690 KEYPAD/DISPLAY TESTER
HPMV | 82625 | 735000287 | HPMV Tester
Syringe Drive | 82656 | 735000300 | 2790/2690 Syringe Drive Tester
Heater/Cooler | 82620 | 735000206 | Heater/Chiller Tester
Column Heater/Cooler | 82774 | 735000698 | Alliance Column Heater/Cooler Tester | Venture