# SUMMARY

SUMMARY	1

1.	Testing the QDa Ion block & Pumping Block	2

1.1.	Requirement equipment for different Tests :	2

1.2.	Changing QDa block :	2

1.3.	TESTS	3

1.3.1.	Source pressure test	3

1.3.2.	Resolution, calibration check and recalibration	5

1.3.3.	Printing Tests reports	5

1.4.	Beams	6

2.	Storage of the Test Results	7

# Testing the QDa Ion block & Pumping Block

## Requirement equipment for different Tests :

- Allen key of 3 mm

- Gloves

- Safety glasses

## Changing QDa block :

- Switch off instrument with the power button

- Open the door

- Unscrew the 2 black plastic screws on the connector (red circles)

- Unscrew the 2 white plastic screws on the source (yellow circles)

- Remove the source

- Unscrew the four screws on the pumping block with an Allen key of 3.

- Take off delicately the Pumping block and ion block

- Put a new assembly in the instrument, align the two lines (one on the pumping block and one on the instrument)

- Tight the 4 screws on the pumping block, replace the source and tight the 4 plastic screws (black end white)

- Close the door

- Switch on the instrument.

## TESTS

### Source pressure test

- Open Hyperterminal

- Once the “Done executing startup script” message is displayed, open the Dashboard

- When the Turbo Speed Readback is at 97% (you will hear a click), the source pressure test begins, you can start a record to check the values of the probe, the heater and the pressure sensor

- Open Masslynx and MS Console

- When the SPT is completed, the critical readback checks start.

- When done, click on “Troubleshoot/Service Mode”

- Type pin : 1147 in the window

- Click on Troubleshoot/Show tune page

- Enable calibrant heater and wait for it to be at 150°C then enable High voltage supply and open the calibrant valve in negative mode (click Ion Mode/Electrospray -)

- Check the beams on the entire spector (Select one table, choose mass 640 and 1220 of span)

### Resolution, calibration check and recalibration

- Close the tune page

- Click on Maintain/Check Resolution

- Wait for it to be finished, the instrument will be “Ready”

- Open the tune page again

- Change the Low mass resolution of 0.20 in the engineers tab

- Close the tune page

- Click on Maintain/Check Resolution and wait for it to complete

- When the message “There is not a valid resolution for the instrument” is displayed, click on the “Resolve” Button on the right

- The calibration check will start automatically after.

- When done, click on Maintain/Recalibrate

### Printing Tests reports

- Click on “Logs”

- Save the 1st Resolution check diagnostic on a notepad file

- Save the resolution setup diagnostic on a notepad file

- Save the calibration check diagnostic on a notepad file

- Save the recalibration check diagnostic on a notepad file

- Make a screen capture of the Turbo Power page, in “QDA Detector” line on the left

## Beams

- Open the tune page, enable the calibrant heater, high voltage supply

- Print a PDF page of the entire spector in negative mode (click Ion Mode/Electrospray -) with the calibrant valve opened, call it  IC Neg FS

- Print a PDF page of the entire spector in negative mode with the calibrant valve closed, call it Neg Noise

- Print a PDF page of the entire spector in positive mode (click Ion Mode/Electrospray +) with the calibrant valve closed, call it Pos Noise

- Print a PDF page of the 4 specific masses in negative mode (135, 417,749 and 1081), call it Neg Beam

- Print a PDF page of the 4 specific masses in positive mode (169, 479, 761, 1159), call it Pos Beam

- Close Masslynx and switch off the instrument

# Storage of the Test Results

Stock results on network :

- Go on: ETSCSTORAGE01

- Ion block cleaning backup.

- QDA

- Save here

Test Report SQD2 / Xevo TQD

Lot number :

Serial number of Ion block :

Serial number of Hexapole :

Set-up solution :

| MS1 tests | m/z 59.05 | m/z 455.27 | m/z 1080.80 | m/z 2034.64 |
| --- | --- | --- | --- | --- |
| Measured mass (Da) |  |  |  |  |
| FWHM (Da) |  |  |  |  |
| Peak intensity |  |  |  |  |

Charging test :

| %Difference Low mass Pos | %Difference High mass Pos | %Difference Low mass Neg |
| --- | --- | --- |

Name of operator : ...

Signature :

Date : 30-Apr-19