# SUMMARY

SUMMARY	1

1.	TESTING THE XEVO TQS ION BLOCK & STEPWAVE	2

1.1.	Requirement equipment for different tests	2

1.2.	Changing ion block and its hexapole	2

1.3.	Tests	3

1.3.1.	Initial test	3

1.3.2.	API set-up test _ Data processing aquisition	4

1.3.3.	Charging test	5

1.3.4.	When the test is finished	11

2.	STORAGE OF THE TEST RESULTS	11

TESTING THE XEVO TQ-S ION BLOCK & STEPWAVE

## Requirement equipment for different tests

- Allen key of 2.5 mm

- Allen key of 3 mm

- Gloves

- Safety glasses

## Changing ion block and its hexapole

- Remove the vacuum.

- Open the isolation chamber, unscrew 4 screws which maintain ion block with its support, with an Allen key of 2.5 mm.

- Unscrew the four screws of the ion block support with an Allen key of 3 mm.

- Take off delicately the stepwave with the helper tool.

- Put a new stepwave with the same tool. Push until it stops.

- Screw support and ion block.

- Close the isolation chamber.

## Tests

### Initial test

- Start Mass spectrometer and when it has booted, start specific Masslynx for Xevo TQ-S Mass Spectrometer (SCN 905).

- File, open new project, ion block cleaning project.

- Project is on the following path :

"ETSC STORAGE 01"

ION BLOCK CLEANING

BACKUP\DATA Xevo TQS\

- Click on tab MS Tune, and Vacuum Pump ; The vacuum is for at least 2 hours before starting the test.

- Click on the LED on the bottom right, to turn green => Switches the instrument to operate.

### API set-up test _ Data processing acquisition

- In the MS Tune page.

- Check that vacuum is on and gas valve is open for the nebulisation.

- Analyser M52 Penning : 1.10-5

- Ion guide pinami : 5.10-3

- Collision cell pinami : 2.10-4

- Check the following parameters :

| Voltages | Capillary (kV) | 3.50 |
| --- | --- | --- |
|  | Cone (V) | 80 |
|  | Source offset (V) | 60 |
| Temperatures | Desolvation temp (°C) | 200 |
| Gas flow | Desolvation (L/Hr) | 550 |
|  | Cone (L/Hr) | 150 |
|  | Nebuliser (Bar) | 7 |

- For "Analyser parameters" :

| LM Resolution 1 | 3 |
| --- | --- |
| HM Resolution 1 | 15 |
| Ion Energy 1 | Adjust it (see after table) |

- Use an API set-up solution (#700002646-10).

- Infusion flow rate 10µL/min.

- Let flush the capillary.

- Adjust ion energy :

- Select only 1080.8 Da on one window.

- Increase ion energy at 2V.

- Adjust the peak to have it in 100% of the window.

- Reduce ion energy to have 5% of original peak.

- Add 5 units.

- You have optimal ion energy.

- Select all windows with all masses : 59 Da ; 455.3 Da ; 1080.8 Da ; 2034.6 Da.

- Check the mass parameters.

- Check the Beam.

- Start the acquisition in the sample set and use "Xevo TQS Cleaning Test".

- Go on next test.

### Charging test

CAUTION : Ensure that the instrument has been pumping for at least 2 hours before running the charging test.

### Positive ion beam test

Switch on the API gas, place the instrument into Operate, and then infuse the API set-up solution Xevo TQS at 10 µL/min.

- Set the Desolvation temperature and gas flow rate to the same values in the ESI positive and the ESI negative tunes pages (250 L/hr, 150°C).

- On the tune page, set the following parameters :

| Function | Set | Mass | Span | Gain |
| --- | --- | --- | --- | --- |
| MS1 Scan | N/A | 59 | 5 | 10 |
| MS1 Scan | N/A | 455.27 | 5 | 10 |
| MS1 scan | N/A | 1080.8 | 5 | 10 |
| MS1 Scan | N/A | 2034.62 | 5 | 15 |

- Optimize the capillary voltage.

- Set the following parameters :

- Sample cone voltage : 80V

- MS1 LM resolution 3

- HM resolution 15

- Set the ion energy to the optimum for the instrument.

- Click ion mode > ESI -ve ion and check for a beam in negative ion.

- Set the following masses :

- Window 1 59 Da

- Window 2 255.3 Da

- Window 3554.3 Da

- Optimize the capillary voltage.

- Set the following parameters :

- Sample cone voltage 40V

- MS1 LM resolution 3

- HM resolution 15

- Set the ion energy to the optimum for the instrument.

### Preparation

- Create the following MS method experiment file for positive, and after one for negative ion beam.

NOTE : Each function has the parameters MS1 mode, Continuum, Scan time of 0.5 seconds.

| Function | ES Mode | Start Time | End Time | Start Mass | End Mass |
| --- | --- | --- | --- | --- | --- |
| 1 2 3 4 5 6 7 8 | Positive | 0.50 5.50 0.30 1.00 0.50 5.50 0.30 1.00 | 1.00 6.00 0.40 5.50 1.00 6.00 0.40 5.50 | 55.00 55.00 55.00 55.00 2030.00 2030.00 2030.00 2030.00 | 65.00 65.00 65.00 65.00 2040.00 2040.00 2040.00 2040.00 |

| 9 10 11 12 13 14 15 16 17 18 | Negative | 6.10 6.30 0.00 11.80 6.80 11.80 6.30 6.10 0.00 6.80 | 6.20 6.80 0.30 12.30 11.80 12.30 6.80 6.20 0.30 11.80 | 55.00 55.00 55.00 55.00 55.00 2030.00 2030.00 2030.00 2030.00 2030.00 | 65.00 65.00 65.00 65.00 65.00 2040.00 2040.00 2040.00 2040.00 2040.00 |
| --- | --- | --- | --- | --- | --- |

MS method experiment

- Change the inter-scan delay to duplicate the tune page.

Figure 1 : Inter-scan setup dialog

CAUTION : Once the charging test is complete, the inter-scan delay must be reset to the state (automatic or manual) used by the customer.

- Use this experiment file in the sample list page.

- Ensure that the file name represents the serial number of the instrument.

- Save the sample list.

### Acquisition

- Set fluidics to infusion.

- Set a flow rate of 10µL/min and start the infusion.

- Start the acquisition.

### Processing

- Open the chromatogram TIC functions shown in table :

| Function Number | ESI Mode | Description |
| --- | --- | --- |
| 1 | Positive | Low mass intensity at start of run |
| 2 | Positive | Low mass intensity at end of run |
| 5 | Positive | High mass intensity at start of run |
| 6 | Positive | High mass intensity at end of run |
| 10 | Negative | Low mass intensity at start of run |
| 12 | Negative | Low mass intensity at end of run |

- Use Global / Replace and Smooth all the chromatograms using the setting in Figure 2 :

Figure 2 : Smooth chromatogram settings

- Calculate the % difference between the following functions, using the equation :

- Functions 1 and 2 (Low mass positive ion mode on positive chromatogram)

- Functions 5 and 6 (High mass positive ion mode on positive chromatogram)

- Functions 10 and 12 (Low mass negative ion mode on negative chromatogram)

% Difference = 100 X ((End of run smoothed TIC Intensity/Start of run smoothed TIC Intensity) - 1)

NOTE : Negative % results = loss of intensity over time.

Positive % results = a gain of intensity over time.

### Specification

If any of the traces are outside the following tolerances, then charging is present.

- Functions 1 and 2 must not exceed -18% to +13%

- Functions 5 and 6 must not exceed -23% to +27%

- Functions 10 and 12 must not exceed -19% to +15%

- If the charging test fails, clean the collision cell, check the quadrupole for ion burn, and clean if necessary.

- Repeat the charging test as described above.

### Report and printout

Screen capture and print the chromatogram in portrait configuration and label with the Ion block serial number and their respective % Difference values.

Figure 3 : Example of charging test chromatogram

CAUTION : Once the charging test is complete, the inter-scan delay must be reset to the state (automatic or manual) used by the customer.

### When the test is finished

- Switch the instrument to standby

- Vent the instrument : on MS tune, vacuum, vent

- Switch off the API gas

- Open source location once the pump stops working, take off ion block, its support and the stepwave

STORAGE OF THE TEST RESULTS

- Stock results on network :

- Go on : ETSCSTORAGE01

- Ion block cleaning backup

- Specific ion block of interests

- Copy here the project

Test Report

Lot number :

Serial number of Ion block :

Serial number of Stepwave :

Tested on Xevo TQS S/N : WAA139

Set-up solution :

| MS1 tests | m/z 59.05 | m/z 455.27 | m/z 1080.80 | m/z 2034.64 |
| --- | --- | --- | --- | --- |
| Measured mass (Da) |  |  |  |  |
| FWHM (Da) |  |  |  |  |
| Peak intensity |  |  |  |  |
| MS2 tests | m/z 59.05 | m/z 455.27 | m/z 1080.80 | m/z 2034.64 |
| Measured mass (Da) |  |  |  |  |
| FWHM (Da) |  |  |  |  |
| Peak intensity |  |  |  |  |

Charging test :

| %Difference 1-2 | %Difference 5-6 | %Difference 10-12 |
| --- | --- | --- |

Name of operator : ...

Signature :

Date : 16-Mar-17