## Summary of Actions
No | Optics Bench S/N | (New) Optic Bench Alignment Method | UUT S/N | TLT 1st | TLT 2nd | TLT 3rd | TLT 4th | TLT 5th | TLT 6th | TLT 7th | TLT 8th | (Old) Optic Bench Alignment Method | TLT 1st | TLT 2nd | TLT 3rd
1 | 2007110-01 | Pass | G20475228G | Rayleigh WL Fail Changed FC | Pass | NA | NA | NA | NA |  |  | NA | NA | NA
2 | 2007110-02 | Pass | G20475227G | Raman WL Fail Retest | Raman WL Fail Changed FC | Rayleigh WL Fail Purged & Retest | Pass | NA | NA |  |  | NA | NA | NA
3 | 2007110-03 | Fail | G20475225G | SNR Fail | Restart Fail | Pass | NA | NA | NA |  |  | Pass | Pass
4 | 2007110-04 | Fail | G20475224G | Rayleigh WL Fail Retest | Pass | NA | NA | NA | NA |  |  | Pass | Rayleigh WL Fail Retest | Pass
5 | 2007110-05 | Pass | G20475226G | Rayleigh WL Fail Retest | Rayleigh WL Fail Changed Flow Cell | Normalize Fail Retest | Restart Fail Retest with Original Flow Cell | Rayleigh WL Fail Changed Flow Cell | Normalize Fail Changed Flow Cell | Raman WL Fail Changed Flow Cell | Pass | NA | NA | NA
6 | 2007140-01 | Fail | G20475242G | TLT 1st: Pass | Pass | Rayleigh WL Fail Retest | Rayleigh WL Fail Retest optics bench | Pass
7 | 2007140-02 | Fail | G20475243G | TLT 1st: Pass | Pass | Pass
8 | 2007140-03 | Pass | G20475244G | TLT 1st: Pass
9 | 2007140-04 | Pass | G20475245G | TLT 1st: Pass
9 | 2007140-05 | Pass | G20475246G | TLT 1st: Raman WL Fail Retest; TLT 2nd: Raman WL Fail Changed FC; TLT 3rd: Pass
Note on E3: Mi Mi Soe Myint: Orginal FC was failed in other unit

## Summary of Data
|  |  | Optics Bench new method |  |  |  |  | Final Test | [7 empty] | Re-Align |  |  |  |  | Final Test
Optics Bench S/N | New method Pass/Fail | UUT S/N | Wavelength | Sample Energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Rayleigh WL | Rayleigh Energy | Raman WL | Raman Energy | Ram-Ray Valley Energy | Val-Ram | Ram-Ray | SNR | Wavelength | Sample Energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Rayleigh WL | Rayleigh Energy | Raman WL | Raman Energy | Ram-Ray Valley Energy | Val-Ram | Ram-Ray | SNR
2007110-01 | Pass | G20475228G | 397 | 23.8 | 1.18 | 39.21 | 1656.2 | 349.6 | 1946021 | 396.8 | 39621 | 11112 | 28.05 | 2.04 | 1618.61
2007110-02 | Pass | G20475227G | 396.2 | 76.5 | 1 | 42.54 | 1785.7 | 352.4 | 8040751 | 398.2 | 135609 | 50588 | 37.3 | 1.69 | 1998.47
2007110-03 | Fail | G20475225G | 396.5 | 16.2 | 1.83 | 35.66 | 1038.8 | 351.8 | 889460 | 397.5 | 28267 | 8469 | 29.96 | 3.18 | 1847.68 | 397.3 | 22.5 | 1.12 | 43.39 | 1766.7 | 351.3 | 2245349 | 397.8 | 39134 | 12535 | 32.03 | 1.74 | 2041.91
2007110-04 | Fail | G20475224G | 397.6 | 14.5 | 0.9 | 46.11 | 1401.1 | 350.7 | 1661670 | 396.3 | 21907 | 9185 | 41.93 | 1.32 | 1396.83 | 397 | 23.4 | 1.08 | 43.03 | 1947.5 | 352.2 | 2666283 | 397.8 | 41523 | 14210 | 34.22 | 1.56 | 2290.48
2007110-05 | Pass | G20475226G | 397.2 | 20.3 | 1.21 | 36.05 | 1731.9 | 351.6 | 2165106 | 398.8 | 37691 | 11907 | 31.59 | 1.74 | 1881.14
2007140-01 | Fail | G20475242G | 397.9 | 23.1 | 1.01 | 40.05 | 1581.4 | 351.5 | 2130464 | 398.1 | 39322 | 13657 | 34.73 | 1.85 | 1660.21 | 397.6 | 26.6 | 1.33 | 47.12 | 1856.6 | 350.8 | 1467975 | 397.3 | 43188 | 15010 | 34.76 | 2.94 | 1533.81
2007140-02 | Fail | G20475243G | 397.3 | 17.9 | 0.89 | 41.13 | 2051.3 | 352.3 | 1903046 | 398.7 | 27405 | 9925 | 36.22 | 1.44 | 1675.07 | 396.7 | 21.5 | 1.2 | 40.93 | 2082.3 | 352.3 | 2368280 | 398.7 | 39587 | 11995 | 30.3 | 1.67 | 2154.23
2007140-03 | Pass | G20475244G | 396.3 | 35.2 | 1.07 | 37.17 | 2005.5 | 350.4 | 3282258 | 397 | 62602 | 20100 | 32.11 | 1.91 | 1565.74
2007140-04 | Pass | G20475245G | 396.8 | 28.3 | 1.05 | 42.19 | 2118.4 | 350.1 | 4142586 | 398 | 55782 | 19509 | 34.97 | 1.35 | 2025.06
2007140-05 | Pass | G20475246G | 397.2 | 21.8 | 1.22 | 35.39 | 1615.8 | 351.4 | 3499639 | 399.4 | 36768 | 13832 | 37.62 | 1.05 | 1909.59

## SigmaSure export - Optics Bench
Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Photodiode Out (V) | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | M1 M2 Adj V | M3 Adjust Counts(x1M) | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy
2020-07-13 02:21 | New Align Tools | 279000400 | 2007110-01 | PASS | [18 empty] | 4.105 | 10.7229 | 396 | 8.9 | 1.14 | 49.16 | 713.7
2020-07-13 03:01 | New Align Tools | 279003225 | 2007110-01 | PASS | [18 empty] | 4.126 | 5.1546 | 397.9 | 14.6 | 1.14 | 45.53 | 903.1
2020-07-13 04:14 | New Align Tools | 279003225 | 2007110-01 | PASS | [18 empty] | 4.195 | 9.6388 | 397 | 23.8 | 1.18 | 39.21 | 1656.2
2020-07-13 06:05 | New Align Tools | 279003225 | 2007110-02 | PASS | [18 empty] | 4.131 | 10.5744 | 398 | 24.6 | 1.01 | 42.95 | 733.7
2020-07-13 06:32 | New Align Tools | 279003225 | 2007110-02 | FAIL | [18 empty] | 4.154 | 8.5561 | 398.4 | 76.4 | 0.96 | 39.1 | 1691.6
2020-07-13 07:35 | New Align Tools | 279003225 | 2007110-02 | PASS | [18 empty] | 4.175 | 7.6006 | 396.2 | 76.5 | 0.99 | 42.54 | 1785.7
2020-07-13 20:40 | New Align Tools | 279003225 | 2007110-02 | PASS | [18 empty] | 4.209 | 7.5122 | 396.9 | 85.3 | 0.96 | 41.28 | 1856.8
2020-07-13 21:08 | New Align Tools | 279003225 | 2007110-02 | PASS | [18 empty] | 4.189 | 6.0477 | 397.4 | 81.9 | 1.07 | 40.42 | 1806.7
2020-07-13 22:56 | New Align Tools | 279003225 | 2007110-03 | PASS | [18 empty] | 4.185 | 0.51496 | 396.6 | 8.4 | 0.76 | 57.99 | 1710.8
2020-07-13 23:30 | New Align Tools | 279003225 | 2007110-03 | PASS | [18 empty] | 4.165 | 5.9597 | 397.9 | 7.2 | 0.68 | 61.21 | 1824.5
2020-07-14 01:43 | New Align Tools | 279003225 | 2007110-04 | PASS | [18 empty] | 4.151 | 6.037 | 397.7 | 8.7 | 0.95 | 42.15 | 786.6
2020-07-14 06:38 | New Align Tools | 279003225 | 2007110-05 | PASS | [18 empty] | 4.21 | 0.84312 | 397.9 | 19.3 | 1.12 | 36.09 | 1776.9
2020-07-14 07:04 | New Align Tools | 279003225 | 2007110-05 | PASS | [18 empty] | 4.189 | 1.0358 | 397.3 | 20.9 | 1.22 | 35.75 | 1758.9
2020-07-14 20:30 | New Align Tools | 279003225 | 2007110-05 | PASS | [18 empty] | 4.214 | 1.0436 | 397.2 | 20.3 | 1.21 | 36.05 | 1731.9
2020-07-14 23:36 | New Align Tools | 279003225 | 2007110-05 | PASS | [18 empty] | 4.232 | 1.0443 | 397.4 | 19.5 | 1.22 | 35.8 | 1690.6
2020-07-15 00:04 | New Align Tools | 279003225 | 2007110-05 | FAIL | [18 empty] | 4.207 | 1.0185 | 398.4 | 20.8 | 1.23 | 35.53 | 1704.9
2020-07-15 01:27 | New Align Tools | 279003225 | 2007110-04 | PASS | [18 empty] | 3.526 | 6.0145 | 397.6 | 14.5 | 0.9 | 46.11 | 1401.1
2020-07-15 02:11 | New Align Tools | 279003225 | 2007110-04 | PASS | [18 empty] | 3.501 | 5.2227 | 397.5 | 14.2 | 1.27 | 39.89 | 1386.6
2020-07-15 02:55 | New Align Tools | 279003225 | 2007110-03 | PASS | [18 empty] | 3.501 | 5.2306 | 397.6 | 17.1 | 3.06 | 24.99 | 1014.3
2020-07-15 04:37 | New Align Tools | 279003225 | 2007110-03 | PASS | [18 empty] | 3.484 | 4.9916 | 396.5 | 16.2 | 1.83 | 35.66 | 1038.8
2020-07-17 15:50 | Old Align Tools | 279003225 | 2007110-03 | PASS | 402.6 | 28.4 | 1.16 | 37.71 | 1977.7 | 4.204 | 399.9 | 398.5 | 395.6 | 395.7 | 397.8 | 397.8 | 397.1 | 397.1
2020-07-17 18:27 | Old Align Tools | 279003225 | 2007110-04 | FAIL | 394.3 | 24.7 | 1.07 | 41.6 | 1844.9 | 4.153 | 0 | 0 | 396.3 | 0
2020-07-17 19:07 | Old Align Tools | 279003225 | 2007110-04 | PASS | 397 | 23.4 | 1.08 | 43.03 | 1947.5 | 4.17
2020-07-17 19:39 | Old Align Tools | 279003225 | 2007110-03 | PASS | 397 | 28.2 | 1.16 | 37.7 | 1936.6 | 4.228
2020-07-18 15:10 | Old Align Tools | 279003225 | 2007110-03 | PASS | 396.8 | 23 | 1.12 | 43.57 | 1772.5 | 4.117 | 397.2 | 397.2
2020-07-18 16:01 | Old Align Tools | 279003225 | 2007110-03 | PASS | 397.3 | 22.5 | 1.12 | 43.39 | 1766.7 | 4.171
2020-07-27 04:05 | New Align Tools | 279003225 | 2007140-01 | FAIL | [18 empty] | 4.122 | 6.1205 | 404.3 | 65.3 | 30440.71 | 0.51 | 0
2020-07-27 05:50 | New Align Tools | 279003225 | 2007140-01 | PASS | [18 empty] | 4.15 | 6.2933 | 397.1 | 6.6 | 0.98 | 49.46 | 707.5
2020-07-27 06:20 | New Align Tools | 279003225 | 2007140-01 | FAIL | [18 empty] | 4.189 | 5.9154
2020-07-27 07:02 | New Align Tools | 279003225 | 2007140-01 | PASS | [18 empty] | 4.195 | 5.8964 | 397.4 | 14.7 | 0.97 | 50.08 | 1592.6
2020-07-27 20:21 | New Align Tools | 279003225 | 2007140-01 | PASS | [18 empty] | 4.219 | 6.1102 | 397.3 | 23.2 | 0.91 | 41.25 | 1587.8
2020-07-27 20:51 | New Align Tools | 279003225 | 2007140-01 | PASS | [18 empty] | 4.185 | 5.1639 | 397.9 | 23.1 | 1.01 | 40.05 | 1581.4
2020-07-27 22:09 | New Align Tools | 279003225 | 2007140-02 | PASS | [18 empty] | 3.511 | 5.1694 | 397.8 | 20.7 | 0.88 | 37.79 | 2118
2020-07-27 23:26 | New Align Tools | 279003225 | 2007140-02 | PASS | [18 empty] | 4.185 | 5.2738 | 397.3 | 17.9 | 0.89 | 41.13 | 2051.3
2020-07-28 00:02 | New Align Tools | 279003225 | 2007140-03 | PASS | [18 empty] | 4.122 | 8.1329 | 396.5 | 37.3 | 0.99 | 35.56 | 2046.6
2020-07-28 01:42 | New Align Tools | 279003225 | 2007140-03 | PASS | [18 empty] | 4.201 | 9.6697 | 397.4 | 35.7 | 1 | 37.23 | 2001.3
2020-07-28 02:14 | New Align Tools | 279003225 | 2007140-03 | PASS | [18 empty] | 4.168 | 8.928 | 396.3 | 35.2 | 1.07 | 37.17 | 2005.5
2020-07-28 03:39 | New Align Tools | 279003225 | 2007140-04 | PASS | [18 empty] | 4.217 | 8.2038 | 396.1 | 31.9 | 1.03 | 43.21 | 2226.4
2020-07-28 04:20 | New Align Tools | 279003225 | 2007140-04 | PASS | [18 empty] | 4.184 | 7.2966 | 396.8 | 31.5 | 1.06 | 42.61 | 2191.1
2020-07-29 01:27 | New Align Tools | 279003225 | 2007140-04 | PASS | [18 empty] | 4.165 | 6.863 | 396.8 | 28.3 | 1.05 | 42.19 | 2118.4
2020-07-29 02:14 | New Align Tools | 279003225 | 2007140-05 | PASS | [18 empty] | 4.195 | 5.5233 | 396.1 | 19.4 | 1.22 | 35.73 | 1654.8
2020-07-29 02:43 | New Align Tools | 279003225 | 2007140-05 | PASS | [18 empty] | 4.15 | 6.1735 | 396.2 | 17.7 | 1.13 | 36.66 | 1650.2
2020-07-29 03:33 | New Align Tools | 279003225 | 2007140-05 | PASS | [18 empty] | 4.155 | 6.1384 | 398 | 20.9 | 1.09 | 37.65 | 1550
2020-07-29 04:05 | New Align Tools | 279003225 | 2007140-05 | PASS | [18 empty] | 4.16 | 6.1959 | 397.9 | 22.2 | 1.22 | 35.06 | 1629.7
2020-07-29 21:09 | New Align Tools | 279003225 | 2007140-05 | PASS | [18 empty] | 4.175 | 6.1753 | 397.2 | 21.8 | 1.22 | 35.39 | 1615.8
2020-07-30 10:50 | Old Align Tools | 279003225 | 2007140-01 | PASS | 397.6 | 25.2 | 1.03 | 45.38 | 1606.4 | 4.23
2020-07-30 11:09 | Old Align Tools | 279003225 | 2007140-01 | FAIL | 0 | 27.1 | 1.34 | 41.75 | 1873.3 | 4.205
2020-07-30 11:47 | Old Align Tools | 279003225 | 2007140-01 | PASS | 397.6 | 26.6 | 1.33 | 47.12 | 1856.6 | 4.192
2020-07-30 12:23 | Old Align Tools | 279003225 | 2007140-02 | PASS | 396.7 | 21.9 | 1.21 | 40.97 | 1852.6 | 4.187 | 396.9 | 396.9
2020-07-30 13:41 | Old Align Tools | 279003225 | 2007140-02 | PASS | 396.7 | 21.5 | 1.2 | 40.93 | 2082.3 | 4.177
2020-08-03 10:22 | Old Align Tools | 279003225 | 2007140-01 | PASS | 0 | 23.8 | 1.37 | 49.92 | 1863.1 | 4.181 | 397.7
2020-08-03 11:30 | Old Align Tools | 279003225 | 2007140-01 | PASS | 397.7 | 23 | 1.44 | 46.81 | 1881.7 | 4.195 | 397.2 | 396.7 | 396.7 | 396.7
2020-08-03 12:04 | Old Align Tools | 279003225 | 2007140-01 | PASS | 396.7 | 22 | 1.45 | 46.22 | 1875 | 4.19

## SigmaSure export - Unit Level
Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
2020-07-17 16:09 | 186247500 | G20475224G | FAIL | 353.4 | 1711338
2020-07-17 16:45 | 186247500 | G20475224G | PASS | 350.7 | 1661670 | 396.3 | 21907 | 9185 | 41.93 | 1.32 | 1396.83 | 0.99981
2020-07-18 09:34 | 186247500 | G20475225G | FAIL | 349.6 | 918493 | 394.7 | 30612 | 9582 | 31.3 | 3.33 | 0
2020-07-18 10:38 | 186247500 | G20475225G | FAIL
2020-07-18 11:18 | 186247500 | G20475224G | FAIL | 353.9 | 2668542
2020-07-18 13:47 | 186247500 | G20475224G | PASS | 352.2 | 2666283 | 397.8 | 41523 | 14210 | 34.22 | 1.56 | 2290.48 | 0.99978
2020-07-18 13:47 | 186247500 | G20475225G | PASS | 351.8 | 889460 | 397.5 | 28267 | 8469 | 29.96 | 3.18 | 1847.68 | 0.99978
2020-07-20 10:30 | 186247500 | G20475225G | PASS | 351.3 | 2245349 | 397.8 | 39134 | 12535 | 32.03 | 1.74 | 2041.91 | 0.99977
2020-07-20 12:19 | 186247500 | G20475226G | FAIL | 353 | 1818935
2020-07-20 13:47 | 186247500 | G20475226G | FAIL | 353 | 1727151
2020-07-20 14:47 | 186247500 | G20475226G | FAIL
2020-07-20 16:03 | 186247500 | G20475227G | FAIL | 351.5 | 16204491 | 380.7 | 28368
2020-07-20 16:04 | 186247500 | G20475226G | FAIL
2020-07-20 18:00 | 186247500 | G20475227G | FAIL | 351.4 | 16186631 | 380.6 | 28736
2020-07-20 18:17 | 186247500 | G20475228G | FAIL | 353.8 | 1922490
2020-07-20 18:49 | 186247500 | G20475227G | FAIL | 359.8 | 21
2020-07-21 09:36 | 186247500 | G20475228G | PASS | 349.6 | 1946021 | 396.8 | 39621 | 11112 | 28.05 | 2.04 | 1618.61 | 0.99984
2020-07-21 09:38 | 186247500 | G20475227G | PASS | 352.4 | 8040751 | 398.2 | 135609 | 50588 | 37.3 | 1.69 | 1998.47 | 0.99967
2020-07-21 10:28 | 186247500 | G20475226G | FAIL | 353.7 | 1710046
2020-07-21 11:09 | 186247500 | G20475226G | FAIL
2020-07-21 13:36 | 186247500 | G20475226G | FAIL | 352.5 | 1608596 | 400.4 | 35306
2020-07-21 14:00 | 186247500 | G20475226G | PASS | 351.6 | 2165106 | 398.8 | 37691 | 11907 | 31.59 | 1.74 | 1881.14 | 0.99985
2020-07-29 15:34 | 186247500 | G20475242G | PASS | 351.5 | 2130464 | 398.1 | 39322 | 13657 | 34.73 | 1.85 | 1660.21 | 0.99982
2020-07-29 16:57 | 186247500 | G20475243G | PASS | 352.3 | 1903046 | 398.7 | 27405 | 9925 | 36.22 | 1.44 | 1675.07 | 0.99983
2020-07-30 10:19 | 186247500 | G20475244G | PASS | 350.4 | 3282258 | 397 | 62602 | 20100 | 32.11 | 1.91 | 1565.74 | 0.99988
2020-07-30 12:22 | 186247500 | G20475245G | PASS | 349.5 | 3958529 | 396.8 | 55502 | 19996 | 36.03 | 1.4 | 1828.52 | 0.99978
2020-07-30 15:17 | 186247500 | G20475242G | FAIL | 352.9 | 2105063
2020-07-30 15:48 | 186247500 | G20475246G | FAIL | 352.2 | 2022311 | 400.9 | 34476
2020-07-30 16:59 | 186247500 | G20475242G | FAIL | 352.9 | 1930032
2020-07-30 17:30 | 186247500 | G20475246G | FAIL | 351.6 | 3615135 | 399.6 | 36828
2020-07-30 19:14 | 186247500 | G20475246G | PASS | 351.4 | 3499639 | 399.4 | 36768 | 13832 | 37.62 | 1.05 | 1909.59 | 0.99979
2020-08-01 07:33 | 186247500 | G20475242G | FAIL | 353 | 2210255
2020-08-01 07:45 | 186247500 | G20475243G | PASS | 352.3 | 2368280 | 398.7 | 39587 | 11995 | 30.3 | 1.67 | 2154.23 | 0.99979
2020-08-01 07:54 | 186247500 | G20475245G | PASS | 350.1 | 4142586 | 398 | 55782 | 19509 | 34.97 | 1.35 | 2025.06 | 0.99976
2020-08-02 09:25 | 186247500 | G20475242G | FAIL | 352.7 | 3320956

## 2007110-01
| Date | Align Method | part-no | serial-no | Outcome | M1 M2 Adj V | M3 Adjust Counts(x1M) | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy
Optics Bench | 2020-07-13 02:21 | New Align Tools | 279000400 | 2007110-01 | PASS | 4.105 | 10.7229 | 396 | 8.9 | 1.14 | 49.16 | 713.7
| 2020-07-13 03:01 | New Align Tools | 279003225 | 2007110-01 | PASS | 4.126 | 5.1546 | 397.9 | 14.6 | 1.14 | 45.53 | 903.1
| 2020-07-13 04:14 | New Align Tools | 279003225 | 2007110-01 | PASS | 4.195 | 9.6388 | 397 | 23.8 | 1.18 | 39.21 | 1656.2 |  | Or asked differently: are these results of the M3 being maximized to the LED flow cell?
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-20 18:17 | 186247500 | G20475228G | FAIL | 353.8 | 1922490
****** Changed Flow Cell ******
| 2020-07-21 09:36 | 186247500 | G20475228G | PASS | 349.6 | 1946021 | 396.8 | 39621 | 11112 | 28.05 | 2.04 | 1618.61 | 0.99984
Number highlighed in light grey are the numbers reported in "Summary of Data" tab

## 2007110-02
| Date | Align Method | part-no | serial-no | Outcome | M1 M2 Adj V | M3 Adjust Counts(x1M) | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy
Optics Bench | 2020-07-13 06:05 | New Align Tools | 279003225 | 2007110-02 | PASS | 4.131 | 10.5744 | 398 | 24.6 | 1.01 | 42.95 | 733.7
| 2020-07-13 06:32 | New Align Tools | 279003225 | 2007110-02 | FAIL | 4.154 | 8.5561 | 398.4 | 76.4 | 0.96 | 39.1 | 1691.6
| 2020-07-13 07:35 | New Align Tools | 279003225 | 2007110-02 | PASS | 4.175 | 7.6006 | 396.2 | 76.5 | 0.99 | 42.54 | 1785.7
| 2020-07-13 20:40 | New Align Tools | 279003225 | 2007110-02 | PASS | 4.209 | 7.5122 | 396.9 | 85.3 | 0.96 | 41.28 | 1856.8
| 2020-07-13 21:08 | New Align Tools | 279003225 | 2007110-02 | PASS | 4.189 | 6.0477 | 397.4 | 81.9 | 1.07 | 40.42 | 1806.7
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-20 16:03 | 186247500 | G20475227G | FAIL | 351.5 | 16204491 | 380.7 | 28368
****** Re-test ******
| 2020-07-20 18:00 | 186247500 | G20475227G | FAIL | 351.4 | 16186631 | 380.6 | 28736
******Changed Flow Cell******
| 2020-07-20 18:49 | 186247500 | G20475227G | FAIL | 359.8 | 21
****** Purge and Re-test ******
| 2020-07-21 09:38 | 186247500 | G20475227G | PASS | 352.4 | 8040751 | 398.2 | 135609 | 50588 | 37.3 | 1.69 | 1998.47 | 0.99967
Number highlighed in light grey are the numbers reported in "Summary of Data" tab
Ram WL is way off. Maybe an air bubble was in the cell.

## 2007110-03
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Photodiode Out (V) | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | Norm WL | M1 M2 Adj V | M3 Adjust Counts(x1M)
Optics Bench | 2020-07-13 22:56 | New Align Tools | 279003225 | 2007110-03 | PASS | 396.6 | 8.4 | 0.76 | 57.99 | 1710.8 | [9 empty] | 4.185 | 0.51496
| 2020-07-13 23:30 | New Align Tools | 279003225 | 2007110-03 | PASS | 397.9 | 7.2 | 0.68 | 61.21 | 1824.5 | [9 empty] | 4.165 | 5.9597
| 2020-07-15 02:55 | New Align Tools | 279003225 | 2007110-03 | PASS | 397.6 | 17.1 | 3.06 | 24.99 | 1014.3 | [9 empty] | 3.501 | 5.2306
| 2020-07-15 04:37 | New Align Tools | 279003225 | 2007110-03 | PASS | 396.5 | 16.2 | 1.83 | 35.66 | 1038.8 | [9 empty] | 3.484 | 4.9916
| 2020-07-17 15:50 | Old Align Tools | 279003225 | 2007110-03 | PASS | 402.6 | 28.4 | 1.16 | 37.71 | 1977.7 | 4.204 | 399.9 | 398.5 | 395.6 | 395.7 | 397.8 | 397.8 | 397.1 | 397.1
| 2020-07-17 19:39 | Old Align Tools | 279003225 | 2007110-03 | PASS | 397 | 28.2 | 1.16 | 37.7 | 1936.6 | 4.228
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-18 09:34 | 186247500 | G20475225G | FAIL | 349.6 | 918493 | 394.7 | 30612 | 9582 | 31.3 | 3.33 | 0
****** Re-test ******
| 2020-07-18 10:38 | 186247500 | G20475225G | FAIL | Restrat failed
****** Re-test ******
| 2020-07-18 13:47 | 186247500 | G20475225G | PASS | 351.8 | 889460 | 397.5 | 28267 | 8469 | 29.96 | 3.18 | 1847.68 | 0.99978
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Photodiode Out (V) | Norm WL | Norm WL
Optics Bench | 2020-07-18 15:10 | Old Align Tools | 279003225 | 2007110-03 | PASS | 396.8 | 23 | 1.12 | 43.57 | 1772.5 | 4.117 | 397.2 | 397.2
| 2020-07-18 16:01 | Old Align Tools | 279003225 | 2007110-03 | PASS | 397.3 | 22.5 | 1.12 | 43.39 | 1766.7 | 4.171
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-20 10:30 | 186247500 | G20475225G | PASS | 351.3 | 2245349 | 397.8 | 39134 | 12535 | 32.03 | 1.74 | 2041.91 | 0.99977
Number highlighed in light grey are the numbers reported in "Summary of Data" tab
Slight improvement to Signal/Noise using Old alignment method

## Sheet1
Align using LED flow cell (do not adjust mirrors after M3 is maximized but use golden flow cell to capture data)
| Optics Bench serial number: ??????? |  | Unit serial number: ?????????
Norm WL |  | Ray Wl
Sample Energy |  | Ray energy
Reference Energy |  | Ram WL
Ram-Ray Ratio |  | Ram energy
Val-Ram Ratio |  | Valley Energy
Capture Hyperterminal scan |  | Optical resolution
Signal to Noise 1
Hyperterminal scan: |  | Signal to Noise 2 |  | After final test can the signal to noise be run 2 more times using the front screen
1) use front screen to set excitation (EX) wavelength to 350 nm |  | Signal to Noise 3
2) capture text and enter "scan egrating 300 450 wl"
Align using golden flow cell (make any mirror adjustments that normally would be made)
| Optics Bench serial number: [same as above] |  | Unit serial number: [same as above]
Norm WL |  | Ray Wl
Sample Energy |  | Ray energy
Reference Energy |  | Ram WL
Ram-Ray Ratio |  | Ram energy
Val-Ram Ratio |  | Valley Energy
Capture Hyperterminal scan |  | Optical resolution
Signal to Noise 1
Hyperterminal scan: |  | Signal to Noise 2 |  | After final test can the signal to noise be run 2 more times using the front screen
1) use front screen to set excitation (EX) wavelength to 350 nm |  | Signal to Noise 3
2) capture text and enter "scan egrating 300 450 wl"

## 2007110-04
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | M1 M2 Adj V | M3 Adjust Counts(x1M)
Optics Bench | 2020-07-14 01:43 | New Align Tools | 279003225 | 2007110-04 | PASS | 397.7 | 8.7 | 0.95 | 42.15 | 786.6 | 4.151 | 6.037
| 2020-07-15 01:27 | New Align Tools | 279003225 | 2007110-04 | PASS | 397.6 | 14.5 | 0.9 | 46.11 | 1401.1 | 3.526 | 6.0145
| 2020-07-15 02:11 | New Align Tools | 279003225 | 2007110-04 | PASS | 397.5 | 14.2 | 1.27 | 39.89 | 1386.6 | 3.501 | 5.2227
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-17 16:09 | 186247500 | G20475224G | FAIL | 353.4 | 1711338
****** Re-test ******
| 2020-07-17 16:45 | 186247500 | G20475224G | PASS | 350.7 | 1661670 | 396.3 | 21907 | 9185 | 41.93 | 1.32 | 1396.83 | 0.99981
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Photodiode Out (V) | Norm WL | Norm WL | Norm WL | Norm WL
Optics Bench | 2020-07-17 18:27 | Old Align Tools | 279003225 | 2007110-04 | FAIL | 394.3 | 24.7 | 1.07 | 41.6 | 1844.9 | 4.153 | 0 | 0 | 396.3 | 0
| 2020-07-17 19:07 | Old Align Tools | 279003225 | 2007110-04 | PASS | 397 | 23.4 | 1.08 | 43.03 | 1947.5 | 4.17
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-18 11:18 | 186247500 | G20475224G | FAIL | 353.9 | 2668542
****** Re-test ******
| 2020-07-18 13:47 | 186247500 | G20475224G | PASS | 352.2 | 2666283 | 397.8 | 41523 | 14210 | 34.22 | 1.56 | 2290.48 | 0.99978
Number highlighed in light grey are the numbers reported in "Summary of Data" tab
Major improvement to Signal/Noise using Old alignment method

## 2007110-05
| Date | Align Method | part-no | serial-no | Outcome | M1 M2 Adj V | M3 Adjust Counts(x1M) | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy
Optics Bench | 2020-07-14 06:38 | New Align Tools | 279003225 | 2007110-05 | PASS | 4.21 | 0.84312 | 397.9 | 19.3 | 1.12 | 36.09 | 1776.9
| 2020-07-14 07:04 | New Align Tools | 279003225 | 2007110-05 | PASS | 4.189 | 1.0358 | 397.3 | 20.9 | 1.22 | 35.75 | 1758.9
| 2020-07-14 20:30 | New Align Tools | 279003225 | 2007110-05 | PASS | 4.214 | 1.0436 | 397.2 | 20.3 | 1.21 | 36.05 | 1731.9
| 2020-07-14 23:36 | New Align Tools | 279003225 | 2007110-05 | PASS | 4.232 | 1.0443 | 397.4 | 19.5 | 1.22 | 35.8 | 1690.6
| 2020-07-15 00:04 | New Align Tools | 279003225 | 2007110-05 | FAIL | 4.207 | 1.0185 | 398.4 | 20.8 | 1.23 | 35.53 | 1704.9
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-20 12:19 | 186247500 | G20475226G | FAIL | 353 | 1818935
****** Re-test ******
| 2020-07-20 13:47 | 186247500 | G20475226G | FAIL | 353 | 1727151
****** Change Flow Cell ******
| 2020-07-21 10:28 | 186247500 | G20475226G | FAIL | 353.7 | 1710046
****** Retry original Flow cell ******
| 2020-07-21 13:36 | 186247500 | G20475226G | FAIL | 352.5 | 1608596 | 400.4 | 35306
****** Change Flow Cell ******
| 2020-07-21 14:00 | 186247500 | G20475226G | PASS | 351.6 | 2165106 | 398.8 | 37691 | 11907 | 31.59 | 1.74 | 1881.14 | 0.99985
Number highlighed in light grey are the numbers reported in "Summary of Data" tab

## 2007140-01
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | M1 M2 Adj V | M3 Adjust Counts(x1M)
Optics Bench | 2020-07-27 04:05 | New Align Tools | 279003225 | 2007140-01 | FAIL | 404.3 | 65.3 | 30440.71 | 0.51 | 0 | 4.122 | 6.1205
| 2020-07-27 05:50 | New Align Tools | 279003225 | 2007140-01 | PASS | 397.1 | 6.6 | 0.98 | 49.46 | 707.5 | 4.15 | 6.2933
| 2020-07-27 06:20 | New Align Tools | 279003225 | 2007140-01 | FAIL | [5 empty] | 4.189 | 5.9154
| 2020-07-27 07:02 | New Align Tools | 279003225 | 2007140-01 | PASS | 397.4 | 14.7 | 0.97 | 50.08 | 1592.6 | 4.195 | 5.8964
| 2020-07-27 20:21 | New Align Tools | 279003225 | 2007140-01 | PASS | 397.3 | 23.2 | 0.91 | 41.25 | 1587.8 | 4.219 | 6.1102
| 2020-07-27 20:51 | New Align Tools | 279003225 | 2007140-01 | PASS | 397.9 | 23.1 | 1.01 | 40.05 | 1581.4 | 4.185 | 5.1639
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-29 15:34 | 186247500 | G20475242G | PASS | 351.5 | 2130464 | 398.1 | 39322 | 13657 | 34.73 | 1.85 | 1660.21 | 0.99982
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Photodiode Out (V) | Norm WL | Norm WL | Norm WL | Norm WL
Optics Bench | 2020-07-30 10:50 | Old Align Tools | 279003225 | 2007140-01 | PASS | 397.6 | 25.2 | 1.03 | 45.38 | 1606.4 | 4.23
| 2020-07-30 11:09 | Old Align Tools | 279003225 | 2007140-01 | FAIL | 0 | 27.1 | 1.34 | 41.75 | 1873.3 | 4.205
| 2020-07-30 11:47 | Old Align Tools | 279003225 | 2007140-01 | PASS | 397.6 | 26.6 | 1.33 | 47.12 | 1856.6 | 4.192
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-30 15:17 | 186247500 | G20475242G | FAIL | 352.9 | 2105063
| 2020-07-30 16:59 | 186247500 | G20475242G | FAIL | 352.9 | 1930032
| 2020-08-01 07:33 | 186247500 | G20475242G | FAIL | 353 | 2210255
| 2020-08-02 09:25 | 186247500 | G20475242G | FAIL | 352.7 | 3320956
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Photodiode Out (V) | Norm WL | Norm WL | Norm WL | Norm WL
Optics Bench | 2020-08-03 10:22 | Old Align Tools | 279003225 | 2007140-01 | PASS | 0 | 23.8 | 1.37 | 49.92 | 1863.1 | 4.181 | 397.7
| 2020-08-03 11:30 | Old Align Tools | 279003225 | 2007140-01 | PASS | 397.7 | 23 | 1.44 | 46.81 | 1881.7 | 4.195 | 397.2 | 396.7 | 396.7 | 396.7
| 2020-08-03 12:04 | Old Align Tools | 279003225 | 2007140-01 | PASS | 396.7 | 22 | 1.45 | 46.22 | 1875 | 4.19
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | ??????? | ???????????? | ??????? | ??????? | 350.8 | 1467975 | 397.3 | 43188 | 15010 | 2.94 | ?????? | 1533.81 | ???????
Number highlighed in light grey are the numbers reported in "Summary of Data" tab
Could not find this data in SigmaSure under serial number G20475242G.
Very slight decrease to Signal/Noise using Old alignment method

## 2007140-02
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | M1 M2 Adj V | M3 Adjust Counts(x1M)
Optics Bench | 2020-07-27 22:09 | New Align Tools | 279003225 | 2007140-02 | PASS | 397.8 | 20.7 | 0.88 | 37.79 | 2118 | 3.511 | 5.1694
| 2020-07-27 23:26 | New Align Tools | 279003225 | 2007140-02 | PASS | 397.3 | 17.9 | 0.89 | 41.13 | 2051.3 | 4.185 | 5.2738
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-29 16:57 | 186247500 | G20475243G | PASS | 352.3 | 1903046 | 398.7 | 27405 | 9925 | 36.22 | 1.44 | 1675.07 | 0.99983
| Date | Align Method | part-no | serial-no | Outcome | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy | Photodiode Out (V) | Norm WL | Norm WL
Optics Bench | 2020-07-30 12:23 | Old Align Tools | 279003225 | 2007140-02 | PASS | 396.7 | 21.9 | 1.21 | 40.97 | 1852.6 | 4.187 | 396.9 | 396.9
| 2020-07-30 13:41 | Old Align Tools | 279003225 | 2007140-02 | PASS | 396.7 | 21.5 | 1.2 | 40.93 | 2082.3 | 4.177
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-08-01 07:45 | 186247500 | G20475243G | PASS | 352.3 | 2368280 | 398.7 | 39587 | 11995 | 30.3 | 1.67 | 2154.23 | 0.99979
Number highlighed in light grey are the numbers reported in "Summary of Data" tab
Major improvement to Signal/Noise using Old alignment method

## 2007140-03
| Date | Align Method | part-no | serial-no | Outcome | M1 M2 Adj V | M3 Adjust Counts(x1M) | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy
Optics Bench | 2020-07-28 00:02 | New Align Tools | 279003225 | 2007140-03 | PASS | 4.122 | 8.1329 | 396.5 | 37.3 | 0.99 | 35.56 | 2046.6
| 2020-07-28 01:42 | New Align Tools | 279003225 | 2007140-03 | PASS | 4.201 | 9.6697 | 397.4 | 35.7 | 1 | 37.23 | 2001.3
| 2020-07-28 02:14 | New Align Tools | 279003225 | 2007140-03 | PASS | 4.168 | 8.928 | 396.3 | 35.2 | 1.07 | 37.17 | 2005.5
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-30 10:19 | 186247500 | G20475244G | PASS | 350.4 | 3282258 | 397 | 62602 | 20100 | 32.11 | 1.91 | 1565.74 | 0.99988
Number highlighed in light grey are the numbers reported in "Summary of Data" tab

## 2007140-04
| Date | Align Method | part-no | serial-no | Outcome | M1 M2 Adj V | M3 Adjust Counts(x1M) | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy
Optics Bench | 2020-07-28 03:39 | New Align Tools | 279003225 | 2007140-04 | PASS | 4.217 | 8.2038 | 396.1 | 31.9 | 1.03 | 43.21 | 2226.4
| 2020-07-28 04:20 | New Align Tools | 279003225 | 2007140-04 | PASS | 4.184 | 7.2966 | 396.8 | 31.5 | 1.06 | 42.61 | 2191.1
| 2020-07-29 01:27 | New Align Tools | 279003225 | 2007140-04 | PASS | 4.165 | 6.863 | 396.8 | 28.3 | 1.05 | 42.19 | 2118.4
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-30 12:22 | 186247500 | G20475245G | PASS | 349.5 | 3958529 | 396.8 | 55502 | 19996 | 36.03 | 1.4 | 1828.52 | 0.99978
| 2020-08-01 07:54 | 186247500 | G20475245G | PASS | 350.1 | 4142586 | 398 | 55782 | 19509 | 34.97 | 1.35 | 2025.06 | 0.99976
Number highlighed in light grey are the numbers reported in "Summary of Data" tab

## 2007140-05
| Date | Align Method | part-no | serial-no | Outcome | M1 M2 Adj V | M3 Adjust Counts(x1M) | Norm WL | Sample energy | Ram-Ray Ratio | Val-Ram Ratio | Ref Energy
Optics Bench | 2020-07-29 02:14 | New Align Tools | 279003225 | 2007140-05 | PASS | 4.195 | 5.5233 | 396.1 | 19.4 | 1.22 | 35.73 | 1654.8
| 2020-07-29 02:43 | New Align Tools | 279003225 | 2007140-05 | PASS | 4.15 | 6.1735 | 396.2 | 17.7 | 1.13 | 36.66 | 1650.2
| 2020-07-29 03:33 | New Align Tools | 279003225 | 2007140-05 | PASS | 4.155 | 6.1384 | 398 | 20.9 | 1.09 | 37.65 | 1550
| 2020-07-29 04:05 | New Align Tools | 279003225 | 2007140-05 | PASS | 4.16 | 6.1959 | 397.9 | 22.2 | 1.22 | 35.06 | 1629.7
| 2020-07-29 21:09 | New Align Tools | 279003225 | 2007140-05 | PASS | 4.175 | 6.1753 | 397.2 | 21.8 | 1.22 | 35.39 | 1615.8
| Date | part-no | serial-no | Outcome | RayWL | Ray Energy | Ram WL | Ram Energy | Valley Energy | Optical Res | Stray Light Ratio | Signal/Noise | PMT Coef
Unit | 2020-07-30 15:48 | 186247500 | G20475246G | FAIL | 352.2 | 2022311 | 400.9 | 34476
| 2020-07-30 17:30 | 186247500 | G20475246G | FAIL | 351.6 | 3615135 | 399.6 | 36828
****** Changed Flow Cell ******
| 2020-07-30 19:14 | 186247500 | G20475246G | PASS | 351.4 | 3499639 | 399.4 | 36768 | 13832 | 37.62 | 1.05 | 1909.59 | 0.99979
Number highlighed in light grey are the numbers reported in "Summary of Data" tab