## Sheet1
Row Labels | Count of Unit
38 | 1
39 | 79
42 | 168
43 | 219
44 | 2
46 | 499
48 | 436
49 | 212
50 | 94
Grand Total | 1710

## Pass
Alliance CarryOver Test Monitoring
TSP WAT271129 Rev 33 | [5 empty] | Alliance Test Tesmplate : 1.8
UUT Firmware : v 3.04 | [7 empty] | Injector Adjustment Test
Unit | Date | S/N | Test Station (station. Test port) | Carryover % (<0.002%) | Spec | Remarks | Evaluation | Measurement (Top) 38 to 50 Steps | Measurement (Width) 60 to 90 Steps
1 | 2018-04-06 | C18SM7917A | 4.4 | 0.0003 | 0.002 | Pass on first test |  | High Carryover
2 | 2018-04-06 | D18SM7918A | 5.1 | 0.0003 | 0.002 | Pass on first test |  | Abnormal Peak
3 | 2018-04-07 | C18SM7909A | 2.3 | 0.0004 | 0.002 | Pass on first test
4 | 2018-04-06 | D18SM4964A | 4.2 | 0.0003 | 0.002 | Pass on first test
5 | 2018-04-06 | D18SM4965A | 5.1 | 0.0003 | 0.002 | Pass on first test
6 | 2018-04-07 | D18SM4968A | 3.1 | 0.0004 | 0.002 | Pass on first test
7 | 2018-04-07 | D18SM4974A | 2.2 | 0.0002 | 0.002 | Pass on first test
8 | 2018-04-07 | C18SM4897A | 4.2 | 0.0003 | 0.002 | Pass on first test
9 | 2018-04-09 | D18SM4962A | 5.2 | 0.0012 | 0.002 | Resolve unit
10 | 2018-04-07 | D18SM4966A | 5.4 | 0.0002 | 0.002 | Resolve unit
11 | 2018-04-07 | D18SM4967A | 5.1 | 0.0002 | 0.002 | Pass on first test
12 | 2018-04-07 | D18SM4969A | 3.2 | 0.0005 | 0.002 | Pass on first test
13 | 2018-04-07 | D18SM4970A | 4.4 | 0.0008 | 0.002 | Pass on first test |  |  |  | Final Test Station | Total | Remarks
14 | 2018-04-09 | D18SM4972A | 1.3 | 0.0003 | 0.002 | Resolve unit |  |  |  | 1.1 | 97
15 | 2018-04-07 | D18SM4973A | 2.1 | 0.0003 | 0.002 | Pass on first test |  |  |  | 1.2 | 123
16 | 2018-04-07 | D18SM4975A | 4.1 | 0.0003 | 0.002 | Pass on first test |  |  |  | 1.3 | 143
17 | 2018-04-09 | D18SM4976A | 3.2 | 0.0003 | 0.002 | Pass on first test |  |  |  | 1.4 | 148
18 | 2018-04-09 | D18SM4977A | 2.3 | 0.0003 | 0.002 | Pass on first test |  |  |  | 2.1 | 148
19 | 2018-04-09 | D18SM4978A | 3.1 | 0.0003 | 0.002 | Pass on first test |  |  |  | 2.2 | 149
20 | 2018-04-09 | D18SM4979A | 4.2 | 0.0002 | 0.002 | Pass on first test |  |  |  | 2.3 | 151
21 | 2018-04-09 | D18SM4981A | 4.1 | 0.0003 | 0.002 | Pass on first test |  |  |  | 2.4 | 159
22 | 2018-04-09 | D18SM4983A | 3.4 | 0.0003 | 0.002 | Pass on first test |  |  |  | 3.1 | 173
23 | 2018-04-06 | C18SM7897A | 1.4 | 0.0004 | 0.002 | Pass on first test |  |  |  | 3.2 | 184
24 | 2018-04-06 | C18SM7898A | 3.1 | 0.0012 | 0.002 | Resolve unit |  |  |  | 3.3 | 156
25 | 2018-04-06 | C18SM7903A | 2.1 | 0.0003 | 0.002 | Resolve unit |  |  |  | 3.4 | 151
26 | 2018-04-06 | C18SM7904A | 1.1 | 0.0004 | 0.002 | Resolve unit |  |  |  | 4.1 | 166
27 | 2018-04-06 | C18SM7906A | 1.4 | 0.0002 | 0.002 | Resolve unit |  |  |  | 4.2 | 166
28 | 2018-04-06 | C18SM7908A | 3.2 | 0.0004 | 0.002 | Resolve unit |  |  |  | 4.3 | 164
29 | 2018-04-09 | C18SM7910A | 2.4 | 0.0003 | 0.002 | Pass on first test |  |  |  | 4.4 | 128
30 | 2018-04-06 | C18SM7911A | 3.2 | 0.0015 | 0.002 | Pass on first test |  |  |  | 5.1 | 126
31 | 2018-04-06 | C18SM7912A | 3.3 | 0.0005 | 0.002 | Resolve unit |  |  |  | 5.2 | 103
32 | 2018-04-06 | C18SM7914A | 4.2 | 0.0005 | 0.002 | Pass on first test |  |  |  | 5.3 | 80
33 | 2018-04-06 | C18SM7915A | 4.1 | 0.0006 | 0.002 | Pass on first test |  |  |  | 5.4 | 53
34 | 2018-04-06 | C18SM7916A | 4.3 | 0.0002 | 0.002 | Resolve unit |  |  |  | 6.1 | 12 | Released 19 Jun 18
35 | 2018-04-09 | C18SM7913A | 5.1 | 0.0003 | 0.002 | Resolve unit |  |  |  | 6.2 | 10 | Released 19 Jun 18
36 | 2018-04-09 | D18SM4982A | 4.4 | 0.0002 | 0.002 | Pass on first test |  |  |  | 6.3 | 4 | Released 19 Jun 18
37 | 2018-04-10 | D18SM4984A | 2.2 | 0.0002 | 0.002 | Pass on first test |  |  |  | 6.4 | 0 | Released 24 Jul 18
38 | 2018-04-10 | D18SM4985A | 4.1 | 0.0002 | 0.002 | Pass on first test |  |  |  | Total | 2794
39 | 2018-04-09 | D18SM4986A | 3.1 | 0.0003 | 0.002 | Pass on first test
40 | 2018-04-10 | D18SM4988A | 1.4 | 0.001 | 0.002 | Pass on first test
41 | 2018-04-10 | D18SM4989A | 1.3 | 0.0002 | 0.002 | Pass on first test
42 | 2018-04-10 | D18SM4990A | 2.3 | 0.0002 | 0.002 | Pass on first test
43 | 2018-04-10 | D18SM4991A | 2.4 | 0.0003 | 0.002 | Pass on first test
44 | 2018-04-10 | D18SM4993A | 3.2 | 0.0001 | 0.002 | Pass on first test
45 | 2018-04-11 | C18SM7907A | 4.3 | 0.0012 | 0.002 | Resolve unit
46 | 2018-04-11 | D18SM4015A | 2.4 | 0.0004 | 0.002 | Pass on first test
47 | 2018-04-11 | D18SM4007A | 4.1 | 0.0006 | 0.002 | Pass on first test
48 | 2018-04-11 | D18SM4005A | 5.1 | 0.0002 | 0.002 | Pass on first test
49 | 2018-04-11 | D18SM4004A | 3.2 | 0.0002 | 0.002 | Resolve unit
50 | 2018-04-10 | D18SM4001A | 1.4 | 0.0002 | 0.002 | Pass on first test
51 | 2018-04-11 | D18SM4999A | 1.3 | 0.0005 | 0.002 | Resolve unit
52 | 2018-04-11 | D18SM4998A | 2.4 | 0.0003 | 0.002 | Pass on first test
53 | 2018-04-11 | D18SM4996A | 5.1 | 0.0004 | 0.002 | Pass on first test
54 | 2018-04-10 | D18SM4994A | 4.2 | 0.0002 | 0.002 | Pass on first test
55 | 2018-04-11 | D18SM4980A | 3.4 | 0.0001 | 0.002 | Resolve unit
56 | 2018-04-09 | D18SM4971A | 1.4 | 0.0002 | 0.002 | Resolve unit
57 | 2018-04-07 | D18SM4963A | 5.3 | 0.0008 | 0.002 | Resolve unit
58 | 2018-04-11 | D18SM4857A | 1.3 | 0.0003 | 0.002 | Pass on first test
59 | 2018-04-12 | D18SM4987A | 2.1 | 0.0003 | 0.002 | Pass on first test
60 | 2018-04-12 | D18SM4003A | 1.2 | 0.0004 | 0.002 | Resolve unit
61 | 2018-04-12 | D18SM4006A | 3.2 | 0.0001 | 0.002 | Pass on first test
62 | 2018-04-12 | D18SM4009A | 1.4 | 0.0009 | 0.002 | Pass on first test
63 | 2018-04-12 | D18SM4014A | 2.4 | 0.0004 | 0.002 | Pass on first test
64 | 2018-04-12 | D18SM4016A | 1.3 | 0.0001 | 0.002 | Pass on first test
65 | 2018-04-12 | D18SM4018A | 4.1 | 0.0008 | 0.002 | Pass on first test
66 | 2018-04-12 | D18SM4022A | 3.1 | 0.0001 | 0.002 | Pass on first test
67 | 2018-04-11 | D18SM4992A | 3.1 | NA | 0.002 | Resolve unit
68 | 2018-04-14 | D18SM4008A | 2.1 | 0.0003 | 0.002 | Pass on first test
69 | 2018-04-13 | D18SM4010A | 1.2 | 0.0003 | 0.002 | Pass on first test
70 | 2018-04-13 | D18SM4011A | 2.2 | NA | 0.002 | Pass on first test
71 | 2018-04-13 | D18SM4012A | 1.4 | 0.0019 | 0.002 | Pass on first test
72 | 2018-04-13 | D18SM4025A | 4.3 | NA | 0.002 | Pass on first test
73 | 2018-04-13 | D18SM4027A | 3.4 | 0.0003 | 0.002 | Pass on first test
74 | 2018-04-13 | D18SM4029A | 3.1 | NA | 0.002 | Pass on first test
75 | 2018-04-13 | D18SM4031A | 2.4 | 0.0005 | 0.002 | Pass on first test
76 | 2018-04-13 | D18SM4032A | 2.1 | 0.0002 | 0.002 | Pass on first test
77 | 2018-04-13 | D18SM4033A | 3.2 | 0.0004 | 0.002 | Pass on first test
78 | 2018-04-13 | D18SM4034A | 3.3 | 0 | 0.002 | Pass on first test
79 | 2018-04-13 | D18SM4037A | 4.2 | NA | 0.002 | Pass on first test
80 | 2018-04-13 | D18SM4038A | 4.4 | NA | 0.002 | Pass on first test
81 | 2018-04-14 | D18SM7920A | 3.1 | 0.0004 | 0.002 | Pass on first test
82 | 2018-04-14 | D18SM7921A | 3.2 | NA | 0.002 | Pass on first test
83 | 2018-04-14 | D18SM7922A | 4.3 | NA | 0.002 | Pass on first test
84 | 2018-04-14 | D18SM7925A | 4.4 | 0.0001 | 0.002 | Pass on first test
85 | 2018-04-14 | D18SM7926A | 1.4 | 0.0002 | 0.002 | Pass on first test
86 | 2018-04-14 | D18SM4021A | 2.3 | 0.0003 | 0.002 | Resolve unit
87 | 2018-04-13 | D18SM4036A | 4.1 | NA | 0.002 | Pass on first test
88 | 2018-04-14 | D18SM4039A | 2.2 | NA | 0.002 | Pass on first test
89 | 2018-04-13 | D18SM4040A | 2.1 | 0.0004 | 0.002 | Pass on first test
90 | 2018-04-14 | D18SM4041A | 2.4 | 0.0003 | 0.002 | Pass on first test
91 | 2018-04-14 | D18SM4043A | 2.4 | 0.0001 | 0.002 | Pass on first test
92 | 2018-04-14 | D18SM4044A | 1.4 | 0.0005 | 0.002 | Pass on first test
93 | 2018-04-12 | C18SM7905A | 3.3 | NA | 0.002 | Resolve unit
94 | 2018-04-16 | D18SM7928A | 1.3 | 0.0002 | 0.002 | Pass on first test
95 | 2018-04-16 | D18SM7930A | 3.2 | 0.0001 | 0.002 | Pass on first test
96 | 2018-04-16 | D18SM7933A | 4.1 | NA | 0.002 | Resolve unit
97 | 2018-04-16 | D18SM7936A | 2.2 | NA | 0.002 | Pass on first test
98 | 2018-04-16 | D18SM7937A | 1.4 | NA | 0.002 | Pass on first test
99 | 2018-04-16 | D18SM7938A | 4.3 | 0.0006 | 0.002 | Pass on first test
100 | 2018-04-16 | D18SM7939A | 3.2 | 0.0003 | 0.002 | Pass on first test
101 | 2018-04-16 | D18SM7941A | 1.3 | 0.0006 | 0.002 | Pass on first test
102 | 2018-04-16 | D18SM4995A | 5.2 | NA | 0.002 | Pass on first test
103 | 2018-04-16 | D18SM4997A | 5.2 | NA | 0.002 | Pass on first test
104 | 2018-04-10 | D18SM4000A | 5.2 | NA | 0.002 | Pass on first test
105 | 2018-04-16 | D18SM40002 | 5.3 | 0.0013 | 0.002 | Pass on first test
106 | 2018-04-16 | D18SM4013A | 5.3 | 0.0017 | 0.002 | Pass on first test
107 | 2018-04-16 | D18SM4017A | 5.2 | 0.0003 | 0.002 | Pass on first test
108 | 2018-04-16 | D18SM4019A | 5.2 | 0.0003 | 0.002 | Pass on first test
109 | 2018-04-16 | D18SM4020A | 5.3 | 0.0013 | 0.002 | Pass on first test
110 | 2018-04-16 | D18SM4023A | 5.3 | 0.0013 | 0.002 | Pass on first test
111 | 2018-04-16 | D18SM4026A | 5.2 | NA | 0.002 | Pass on first test
112 | 2018-04-16 | D18SM4028A | 5.1 | 0.0003 | 0.002 | Pass on first test
113 | 2018-04-16 | D18SM4030A | 5.3 | 0.0012 | 0.002 | Pass on first test
114 | 2018-04-14 | D18SM7923A | 4.2 | 0.0004 | 0.002 | Resolve unit
115 | 2018-04-16 | D18SM7924A | 1.2 | NA | 0.002 | Resolve unit
116 | 2018-04-14 | D18SM7927A | 2.1 | 0.0002 | 0.002 | Resolve unit
117 | 2018-04-16 | D18SM7929A | 2.4 | 0.0006 | 0.002 | Resolve unit
118 | 2018-04-16 | D18SM7931A | 4.4 | NA | 0.002 | Resolve unit
119 | 2018-04-16 | D18SM7935A | 3.4 | 0 | 0.002 | Resolve unit
120 | 2018-04-16 | D18SM7940A | 1.4 | NA | 0.002 | Resolve unit
121 | 2018-04-16 | D18SM7942A | 5.2 | 0.0017 | 0.002 | Pass on first test
122 | 2018-04-16 | D18SM7944A | 5.3 | 0.0014 | 0.002 | Resolve unit
123 | 2018-04-16 | D18SM7946A | 2.2 | NA | 0.002 | Pass on first test
124 | 2018-04-16 | D18SM7024A | 5.4 | 0.0007 | 0.002 | Pass on first test
125 | 2018-04-16 | D18SM4035A | 5.1 | 0.0003 | 0.002 | Pass on first test
126 | 2018-04-16 | D18SM4045A | 4.3 | 0.0004 | 0.002 | Pass on first test
127 | 2018-04-16 | D18SM4046A | 4.3 | 0.0002 | 0.002 | Pass on first test
128 | 2018-04-17 | D18SM4048A | 2.2 | 0.0013 | 0.002 | Resolve unit
129 | 2018-04-17 | D18SM4051A | 3.3 | 0.0011 | 0.002 | Pass on first test
130 | 2018-04-17 | D18SM4052A | 2.1 | 0.0002 | 0.002 | Pass on first test
131 | 2018-04-17 | D18SM4054A | 4.4 | 0.0008 | 0.002 | Pass on first test
132 | 2018-04-18 | D18SM7919A | 3.4 | 0 | 0.002 | Resolve unit
133 | 2018-04-17 | D18SM7934A | 2.3 | 0.0006 | 0.002 | Resolve unit
134 | 2018-04-18 | D18SM7943A | 3.1 | 0.0002 | 0.002 | Resolve unit
135 | 2018-04-18 | D18SM7945A | 1.4 | 0.0002 | 0.002 | Resolve unit
136 | 2018-04-19 | D18SM4047A | 5.1 | NA | 0.002 | Resolve unit
137 | 2018-04-18 | D18SM4049A | 4.1 | 0.0001 | 0.002 | Resolve unit
138 | 2018-04-17 | D18SM4050A | 3.2 | NA | 0.002 | Resolve unit
139 | 2018-04-18 | D18SM4055A | 5.2 | NA | 0.002 | Pass on first test
140 | 2018-04-18 | D18SM4056A | 2.4 | 0.0002 | 0.002 | Resolve unit
141 | 2018-04-18 | D18SM4060A | 4.3 | 0.0004 | 0.002 | Pass on first test
142 | 2018-04-18 | D18SM4061A | 2.2 | 0.0006 | 0.002 | Pass on first test
143 | 2018-04-18 | D18SM4063A | 3.2 | NA | 0.002 | Resolve unit
144 | 2018-04-18 | D18SM4064A | 1.1 | 0.0002 | 0.002 | Pass on first test
145 | 2018-04-18 | D18SM4065A | 2.3 | 0.0017 | 0.002 | Resolve unit
146 | 2018-04-18 | D18SM4066A | 1.2 | NA | 0.002 | Pass on first test
147 | 2018-04-18 | D18SM4067A | 4.1 | NA | 0.002 | Pass on first test
148 | 2018-04-18 | D18SM4068A | 3.3 | 0.001 | 0.002 | Pass on first test
149 | 2018-04-18 | D18SM4071A | 1.4 | 0.0015 | 0.002 | Pass on first test
150 | 2018-04-19 | D18SM4053A | 4.2 | NA | 0.002 | Resolve unit
151 | 2018-04-20 | D18SM4057A | 2.1 | 0.0015 | 0.002 | Resolve unit
152 | 2018-04-19 | D18SM4059A | 4.4 | 0.0013 | 0.002 | Resolve unit
153 | 2018-04-19 | D18SM4069A | 5.2 | NA | 0.002 | Resolve unit
154 | 2018-04-19 | D18SM4070A | 2.4 | NA | 0.002 | Pass on first test
155 | 2018-04-19 | D18SM4073A | 3.1 | 0.0003 | 0.002 | Pass on first test
156 | 2018-04-19 | D18SM4075A | 2.3 | NA | 0.002 | Resolve unit
157 | 2018-04-19 | D18SM4076A | 3.2 | 0.0001 | 0.002 | Resolve unit
158 | 2018-04-19 | D18SM4077A | 3.3 | 0.0011 | 0.002 | Pass on first test
159 | 2018-04-19 | D18SM4078A | 3.4 | NA | 0.002 | Resolve unit
160 | 2018-04-19 | D18SM4080A | 4.3 | NA | 0.002 | Pass on first test
161 | 2018-04-19 | D18SM4081A | 3.1 | 0.0007 | 0.002 | Resolve unit
162 | 2018-04-19 | D18SM4085A | 4.2 | NA | 0.002 | Pass on first test
163 | 2018-04-19 | D18SM4087A | 3.2 | 0.0007 | 0.002 | Pass on first test
164 | 2018-04-20 | D18SM4042A | 1.3 | 0.0003 | 0.002 | Resolve unit | ASP execution eval
165 | 2018-04-20 | D18SM4072A | 1.2 | NA | 0.002 | Pass on first test | ASP execution eval
166 | 2018-04-20 | D18SM4074A | 2.2 | NA | 0.002 | Resolve unit
167 | 2018-04-21 | D18SM4079A | 4.1 | 0.0001 | 0.002 | Resolve unit
168 | 2018-04-20 | D18SM4082A | 5.1 | 0.0004 | 0.002 | Pass on first test
169 | 2018-04-20 | D18SM4084A | 2.4 | 0.0002 | 0.002 | Resolve unit
170 | 2018-04-20 | D18SM4088A | 3.1 | 0.0004 | 0.002 | Pass on first test
171 | 2018-04-20 | D18SM4090A | 4.2 | 0.0015 | 0.002 | Resolve unit
172 | 2018-04-20 | D18SM4091A | 5.2 | 0.0003 | 0.002 | Pass on first test
173 | 2018-04-20 | D18SM4092A | 2.3 | 0 | 0.002 | Resolve unit
174 | 2018-04-20 | D18SM4093A | 3.2 | 0.0003 | 0.002 | Pass on first test
175 | 2018-04-21 | D18SM4094A | 3.4 | NA | 0.002 | Resolve unit
176 | 2018-04-20 | D18SM4098A | 2.1 | 0.0003 | 0.002 | Pass on first test
177 | 2018-04-20 | D18SM4099A | 1.3 | 0.0005 | 0.002 | Pass on first test | ASP execution eval
178 | 2018-04-20 | D18SM4100A | 3.1 | NA | 0.002 | Pass on first test | ASP execution eval
179 | 2018-04-20 | D18SM4102A | 3.2 | 0.0002 | 0.002 | Pass on first test | ASP execution eval
180 | 2018-04-21 | D18SM4104A | 5.1 | 0.0003 | 0.002 | Pass on first test | ASP execution eval
181 | 2018-04-21 | D18SM4106A | 2.1 | 0.001 | 0.002 | Pass on first test | ASP execution eval
182 | 2018-04-23 | D18SM4086A | 4.4 | NA | 0.002 | Pass on first test
183 | 2018-04-23 | D18SM4089A | 3.3 | 0.0003 | 0.002 | Resolve unit
184 | 2018-04-20 | D18SM4095A | 5.2 | NA | 0.002 | Pass on first test
185 | 2018-04-23 | D18SM4096A | 3.2 | 0.0005 | 0.002 | Pass on first test
186 | 2018-04-23 | D18SM4097A | 1.2 | NA | 0.002 | Pass on first test | ASP execution eval
187 | 2018-04-23 | D18SM4103A | 2.4 | NA | 0.002 | Pass on first test
188 | 2018-04-23 | D18SM4105A | 2.2 | NA | 0.002 | Pass on first test | ASP execution eval
189 | 2018-04-23 | D18SM4107A | 2.1 | 0.0008 | 0.002 | Pass on first test | ASP execution eval
190 | 2018-04-23 | D18SM4111A | 1.3 | 0.0002 | 0.002 | Pass on first test
191 | 2018-04-23 | D18SM4112A | 2.1 | 0.0002 | 0.002 | Pass on first test
192 | 2018-04-23 | D18SM4116A | 5.1 | 0.0008 | 0.002 | Pass on first test | ASP execution eval
193 | 2018-04-24 | D18SM7947A | 3.4 | 0.0006 | 0.002 | Resolve unit
194 | 2018-04-24 | D18SM7952A | 3.1 | 0.0002 | 0.002 | Pass on first test
195 | 2018-04-18 | C18SM7896A | 4.3 | NA | 0.002 | Pass on first test
196 | 2018-04-24 | C18SM7900A | 4.3 | NA | 0.002 | Resolve unit | ASP execution eval
197 | 2018-04-23 | C18SM7901A | 4.3 | 0.0011 | 0.002 | Pass on first test
198 | 2018-04-17 | C18SM7902A | 2.2 | NA | 0.002 | Pass on first test
199 | 2018-04-24 | D18SM4058A | 4.2 | 0.0018 | 0.002 | Pass on first test
200 | 2018-04-24 | D18SM4101A | 2.4 | 0.0001 | 0.002 | Resolve unit
201 | 2018-04-24 | D18SM4108A | 3.1 | 0.0007 | 0.002 | Resolve unit | ASP execution eval
202 | 2018-04-24 | D18SM4110A | 4.2 | 0.001 | 0.002 | Resolve unit
203 | 2018-04-24 | D18SM4115A | 2.2 | 0.0006 | 0.002 | Pass on first test | ASP execution eval
204 | 2018-04-24 | D18SM4118A | 4.1 | 0.0001 | 0.002 | Pass on first test | ASP execution eval
205 | 2018-04-24 | D18SM4119A | 1.3 | 0.0016 | 0.002 | Resolve unit | ASP execution eval
206 | 2018-04-24 | D18SM4120A | 2.3 | 0.0002 | 0.002 | Resolve unit | ASP execution eval
207 | 2018-04-24 | D18SM4121A | 3.4 | 0.0004 | 0.002 | Pass on first test | ASP execution eval
208 | 2018-04-24 | D18SM4122A | 3.2 | 0.0005 | 0.002 | Pass on first test | ASP execution eval
209 | 2018-04-25 | D18SM4083A | 1.2 | NA | 0.002 | Resolve unit | ASP execution eval
210 | 2018-04-25 | D18SM4113A | 4.4 | 0.0013 | 0.002 | Resolve unit | ASP execution eval
211 | 2018-04-25 | D18SM4117A | 5.2 | NA | 0.002 | Resolve unit | ASP execution eval
212 | 2018-04-25 | D18SM4123A | 1.4 | 0.0011 | 0.002 | Resolve unit
213 | 2018-04-25 | D18SM4125A | 2.2 | 0.0003 | 0.002 | Pass on first test
214 | 2018-04-26 | D18SM4126A | 4.1 | 0.0003 | 0.002 | Resolve unit
215 | 2018-04-25 | D18SM4127A | 5.3 | 0.0003 | 0.002 | Resolve unit
216 | 2018-04-25 | D18SM4128A | 1.4 | NA | 0.002 | Pass on first test
217 | 2018-04-25 | D18SM4129A | 5.4 | 0.0005 | 0.002 | Pass on first test
218 | 2018-04-25 | D18SM4131A | 1.3 | 0.0002 | 0.002 | Pass on first test
219 | 2018-04-25 | D18SM4132A | 2.2 | 0.0014 | 0.002 | Pass on first test
220 | 2018-04-26 | D18SM4133A | 3.1 | 0.0009 | 0.002 | Pass on first test
221 | 2018-04-26 | D18SM4134A | 1.4 | 0.0001 | 0.002 | Resolve unit
222 | 2018-04-26 | D18SM4136A | 4.3 | 0.0002 | 0.002 | Pass on first test
223 | 2018-04-26 | D18SM4137A | 1.2 | NA | 0.002 | Resolve unit
224 | 2018-04-26 | D18SM4140A | 1.3 | 0.0006 | 0.002 | Resolve unit
225 | 2018-04-25 | D18SM7948A | 3.2 | 0.0004 | 0.002 | Pass on first test
226 | 2018-04-25 | D18SM7949A | 3.1 | 0.0011 | 0.002 | Pass on first test
227 | 2018-04-25 | D18SM7950A | 2.3 | 0.0003 | 0.002 | Pass on first test
228 | 2018-04-25 | D18SM7951A | 1.3 | 0.001 | 0.002 | Pass on first test
229 | 2018-04-26 | D18SM7953A | 2.4 | 0.001 | 0.002 | Resolve unit
230 | 2018-04-25 | D18SM7955A | 4.3 | NA | 0.002 | Pass on first test
231 | 2018-04-26 | D18SM7956A | 3.4 | 0.001 | 0.002 | Resolve unit
232 | 2018-04-26 | D18SM4109A | 3.2 | 0.0018 | 0.002 | Resolve unit | ASP execution eval
233 | 2018-04-26 | D18SM4114A | 2.1 | NA | 0.002 | Resolve unit | ASP execution eval
234 | 2018-04-26 | D18SM4124A | 5.1 | 0.0008 | 0.002 | Resolve unit
235 | 2018-04-26 | D18SM4141A | 1.3 | 0.0001 | 0.002 | Pass on first test
236 | 2018-04-27 | D18SM4146A | 5.3 | 0.002 | 0.002 | Resolve unit
237 | 2018-04-26 | D18SM4147A | 3.4 | 0.0001 | 0.002 | Pass on first test
238 | 2018-04-26 | D18SM4149A | 3.1 | 0.0012 | 0.002 | Pass on first test
239 | 2018-04-26 | D18SM4150A | 1.2 | 0.0002 | 0.002 | Pass on first test
240 | 2018-04-27 | D18SM4155A | 4.2 | 0.0005 | 0.002 | Resolve unit
241 | 2018-04-26 | D18SM4142A | 2.2 | 0.0018 | 0.002 | Resolve unit
242 | 2018-04-26 | D18SM4143A | 1.4 | 0.001 | 0.002 | Resolve unit
243 | 2018-04-26 | D18SM4144A | 4.3 | 0.0003 | 0.002 | Pass on first test
244 | 2018-04-26 | D18SM4135A | 5.4 | 0.0004 | 0.002 | Resolve unit
245 | 2018-04-27 | D18SM4145A | 4.1 | 0.0009 | 0.002 | Resolve unit
246 | 2018-04-27 | D18SM4148A | 2.4 | 0.0009 | 0.002 | Resolve unit
247 | 2018-04-26 | D18SM4151A | 4.4 | 0.0019 | 0.002 | Pass on first test
248 | 2018-04-27 | D18SM4152A | 5.2 | NA | 0.002 | Resolve unit
249 | 2018-04-27 | D18SM4154A | 4.3 | NA | 0.002 | Resolve unit
250 | 2018-04-27 | D18SM4156A | 1.2 | 0.0004 | 0.002 | Pass on first test
251 | 2018-04-27 | D18SM4158A | 1.4 | 0.0003 | 0.002 | Pass on first test
252 | 2018-04-27 | D18SM4159A | 2.2 | 0.0013 | 0.002 | Pass on first test
253 | 2018-04-27 | D18SM4160A | 3.1 | 0.0002 | 0.002 | Resolve unit
254 | 2018-04-27 | D18SM4165A | 5.1 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
255 | 2018-04-28 | D18SM4166A | 5.2 | 0.0014 | 0.002 | Pass on first test | Double Press Seal Eval
256 | 2018-04-28 | D18SM4167A | 5.3 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
257 | 2018-04-28 | D18SM4169A | 1.1 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
258 | 2018-04-28 | D18SM4170A | 1.2 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
259 | 2018-04-28 | D18SM4171A | 2.4 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
260 | 2018-04-27 | D18SM4162A | 1.1 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
261 | 2018-04-27 | D18SM4138A | 3.3 | 0.0008 | 0.002 | Resolve unit
262 | 2018-04-26 | D18SM7954A | 4.2 | 0.0012 | 0.002 | Resolve unit
263 | 2018-04-28 | D18SM7957A | 2.1 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
264 | 2018-04-28 | D18SM7959A | 2.2 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
265 | 2018-04-28 | D18SM7960A | 4.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
266 | 2018-04-28 | D18SM7962A | 2.4 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
267 | 2018-04-28 | D18SM7963A | 5.1 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
268 | 2018-04-28 | D18SM7965A | 5.3 | 0.0014 | 0.002 | Pass on first test | Double Press Seal Eval
269 | 2018-04-28 | D18SM7966A | 1.3 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
270 | 2018-04-30 | D18SM7964A | 5.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
271 | 2018-04-30 | D18SM7967A | 2.2 | 0.0016 | 0.002 | Resolve unit | Double Press Seal Eval
272 | 2018-04-30 | D18SM7968A | 2.1 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
273 | 2018-04-30 | D18SM7969A | 1.2 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
274 | 2018-04-30 | D18SM7970A | 4.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
275 | 2018-04-30 | D18SM7971A | 3.3 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
276 | 2018-04-30 | D18SM7972A | 1.3 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
277 | 2018-04-30 | D18SM7974A | 5.1 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
278 | 2018-04-25 | D18SM4130A | 2.3 | 0.0003 | 0.002 | Pass on first test
279 | 2018-04-30 | D18SM4139A | 2.3 | 0.0012 | 0.002 | Resolve unit
280 | 2018-04-30 | D18SM4157A | 1.1 | 0.001 | 0.002 | Pass on first test
281 | 2018-04-30 | D18SM4161A | 3.2 | 0.0002 | 0.002 | Resolve unit
282 | 2018-04-30 | D18SM4163A | 4.2 | 0.0013 | 0.002 | Resolve unit | Double Press Seal Eval
283 | 2018-04-30 | D18SM4164A | 4.4 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
284 | 2018-04-30 | D18SM4172A | 5.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
285 | 2018-04-30 | D18SM4173A | 2.4 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
286 | 2018-04-30 | D18SM4174A | 3.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
287 | 2018-04-30 | D18SM4175A | 2.3 | 0.0018 | 0.002 | Resolve unit | Double Press Seal Eval
288 | 2018-04-30 | D18SM4177A | 3.3 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
289 | 2018-05-02 | D18SM4178A | 4.3 | NA | 0.002 | Resolve unit | Double Press Seal Eval
290 | 2018-04-30 | D18SM4182A | 2.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
291 | 2018-05-02 | D18SM4176A | 2.1 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
292 | 2018-05-02 | D18SM4179A | 3.2 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
293 | 2018-05-02 | D18SM4181A | 4.2 | NA | 0.002 | Resolve unit | Double Press Seal Eval
294 | 2018-05-02 | D18SM4183A | 2.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
295 | 2018-05-02 | D18SM4184A | 1.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
296 | 2018-05-02 | D18SM4185A | 1.2 | 0.0017 | 0.002 | Resolve unit | Double Press Seal Eval
297 | 2018-05-02 | D18SM4186A | 1.1 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
298 | 2018-05-02 | D18SM4187A | 3.3 | 0.0006 | 0.002 | Pass on first test | Double Press Seal Eval
299 | 2018-05-02 | D18SM4188A | 2.4 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
300 | 2018-05-03 | D18SM4189A | 2.2 | 0.0016 | 0.002 | Resolve unit | Double Press Seal Eval
301 | 2018-05-02 | D18SM4190A | 5.2 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
302 | 2018-05-02 | D18SM4191A | 4.4 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
303 | 2018-05-02 | D18SM4192A | 4.2 | 0.0013 | 0.002 | Resolve unit | Double Press Seal Eval
304 | 2018-05-02 | D18SM4194A | 4.3 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
305 | 2018-05-03 | D18SM4196A | 1.3 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
306 | 2018-05-03 | D18SM4197A | 4.1 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
307 | 2018-05-03 | D18SM4199A | 3.3 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
308 | 2018-05-02 | D18SM4201A | 3.2 | 0.0014 | 0.002 | Resolve unit | Double Press Seal Eval
309 | 2018-05-02 | D18SM4205A | 1.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
310 | 2018-05-03 | D18SM4168A | 5.4 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
311 | 2018-05-03 | D18SM4195A | 2.4 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
312 | 2018-05-03 | D18SM4198A | 1.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
313 | 2018-05-03 | D18SM4200A | 4.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
314 | 2018-05-03 | D18SM4202A | 3.2 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
315 | 2018-05-03 | D18SM4203A | 1.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
316 | 2018-05-03 | D18SM4204A | 5.2 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
317 | 2018-05-02 | D18SM4206A | 2.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
318 | 2018-05-03 | D18SM4208A | 4.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
319 | 2018-05-03 | D18SM4209A | 3.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
320 | 2018-05-03 | D18SM4210A | 5.1 | 0.0011 | 0.002 | Pass on first test | Double Press Seal Eval
321 | 2018-05-03 | D18SM4211A | 4.1 | 0.0019 | 0.002 | Resolve unit | Double Press Seal Eval
322 | 2018-05-03 | D18SM4212A | 1.1 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
323 | 2018-05-03 | D18SM4213A | 2.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
324 | 2018-05-03 | D18SM4214A | 1.3 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
325 | 2018-05-03 | D18SM4215A | 2.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
326 | 2018-05-03 | D18SM4217A | 2.4 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
327 | 2018-05-03 | D18SM4219A | 1.4 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
328 | 2018-05-03 | D18SM4220A | 3.2 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
329 | 2018-04-30 | D18SM7958A | 3.1 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
330 | 2018-04-28 | D18SM7961A | 4.1 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
331 | 2018-05-02 | D18SM7973A | 5.3 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
332 | 2018-05-04 | E18SM7975A | 5.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
333 | 2018-05-04 | E18SM7976A | 2.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
334 | 2018-05-04 | E18SM7977A | 2.2 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
335 | 2018-05-04 | E18SM7978A | 1.4 | 0.0016 | 0.002 | Resolve unit | Double Press Seal Eval
336 | 2018-05-04 | E18SM7979A | 1.3 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
337 | 2018-05-04 | E18SM7981A | 4.2 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
338 | 2018-05-04 | E18SM7982A | 1.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
339 | 2018-05-04 | E18SM7983A | 1.1 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
340 | 2018-05-04 | E18SM7984A | 4.4 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
341 | 2018-05-05 | E18SM7985A | 3.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
342 | 2018-05-04 | E18SM7986A | 2.4 | NA | 0.002 | Pass on first test | Double Press Seal Eval
343 | 2018-05-04 | E18SM7987A | 1.3 | 0.0009 | 0.002 | Resolve unit | Double Press Seal Eval
344 | 2018-05-05 | E18SM7989A | 2.4 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
345 | 2018-05-03 | E18SM4207A | 4.3 | 0.0012 | 0.002 | Resolve unit | Double Press Seal Eval
346 | 2018-05-03 | E18SM4209A | 3.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
347 | 2018-05-03 | E18SM4210A | 5.1 | 0.0011 | 0.002 | Pass on first test | Double Press Seal Eval
348 | 2018-04-30 | D18SM4180A | 5.3 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
349 | 2018-05-04 | E18SM4216A | 3.3 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
350 | 2018-05-03 | E18SM4218A | 3.1 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
351 | 2018-05-04 | E18SM4222A | 4.1 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
352 | 2018-05-04 | E18SM4223A | 4.4 | 0.0012 | 0.002 | Resolve unit | Double Press Seal Eval
353 | 2018-05-05 | E18SM4225A | 5.1 | 0.0009 | 0.002 | Resolve unit | Double Press Seal Eval
354 | 2018-05-04 | E18SM4226A | 2.3 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
355 | 2018-05-05 | E18SM$227A | 4.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
356 | 2018-05-05 | E18SM4228A | 4.3 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
357 | 2018-05-05 | E18SM4229A | 4.1 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
358 | 2018-05-05 | E18SM4230A | 3.2 | NA | 0.002 | Resolve unit | Double Press Seal Eval
359 | 2018-05-05 | E18SM4234A | 3.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
360 | 2018-05-05 | E18SM$235A | 1.4 | NA | 0.002 | Resolve unit | Double Press Seal Eval
361 | 2018-05-05 | E18SM4238A | 4.1 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
362 | 2018-05-07 | E18SM4232A | 2.3 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
363 | 2018-05-07 | E18SM4233A | 5.2 | NA | 0.002 | Resolve unit | Double Press Seal Eval
364 | 2018-05-07 | E18SM4236A | 1.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
365 | 2018-05-07 | E18SM4237A | 1.3 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
366 | 2018-05-07 | E18SM4239A | 1.3 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
367 | 2018-05-07 | E18SM4240A | 5.1 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
368 | 2018-05-07 | E18SM4241A | 4.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
369 | 2018-05-07 | E18SM4242A | 4.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
370 | 2018-05-07 | E18SM4243A | 3.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
371 | 2018-05-07 | E18SM4244A | 1.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
372 | 2018-05-07 | E18SM4245A | 4.4 | NA | 0.002 | Pass on first test | Double Press Seal Eval
373 | 2018-05-07 | E18SM4246A | 3.2 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
374 | 2018-05-07 | E18SM4247A | 3.3 | NA | 0.002 | Resolve unit | Double Press Seal Eval
375 | 2018-05-07 | E18SM4248A | 3.4 | 0.0018 | 0.002 | Resolve unit | Double Press Seal Eval
376 | 2018-05-07 | E18SM4249A | 2.4 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
377 | 2018-05-07 | E18SM4250A | 1.4 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
378 | 2018-05-07 | E18SM4251A | 4.1 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
379 | 2018-05-07 | E18SM4252A | 5.2 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
380 | 2018-05-07 | E18SM4253A | 5.3 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
381 | 2018-05-07 | E18SM4254A | 4.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
382 | 2018-05-07 | E18SM4256A | 4.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
383 | 2018-05-07 | E18SM4257A | 3.1 | 0.0011 | 0.002 | Resolve unit | Double Press Seal Eval
384 | 2018-05-07 | E18SM4260A | 3.3 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
385 | 2018-05-07 | E18SM4261A | 1.2 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
386 | 2018-05-07 | E18SM4262A | 1.1 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
387 | 2018-05-08 | E18SM7980A | 3.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
388 | 2018-05-08 | E18SM4255A | 2.3 | NA | 0.002 | Resolve unit | Double Press Seal Eval
389 | 2018-05-08 | E18SM4258A | 3.2 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
390 | 2018-05-07 | E18SM4259A | 4.4 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
391 | 2018-05-07 | E18SM4263A | 1.3 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
392 | 2018-05-08 | E18SM4264A | 4.1 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
393 | 2018-05-08 | E18SM4265A | 5.1 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
394 | 2018-05-08 | E18SM4266A | 3.4 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
395 | 2018-05-08 | E18SM4267A | 2.4 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
396 | 2018-05-08 | E18SM4268A | 5.2 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
397 | 2018-05-08 | E18SM4269A | 5.3 | 0.0011 | 0.002 | Resolve unit | Double Press Seal Eval
398 | 2018-05-08 | E18SM4270A | 4.2 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
399 | 2018-05-08 | E18SM4271A | 4.4 | 0.0006 | 0.002 | Pass on first test | Double Press Seal Eval
400 | 2018-05-08 | E18SM4272A | 4.3 | 0.0015 | 0.002 | Pass on first test | Double Press Seal Eval
401 | 2018-05-08 | E18SM4273A | 3.3 | 0.0019 | 0.002 | Resolve unit | Double Press Seal Eval
402 | 2018-05-08 | E18SM4274A | 3.1 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
403 | 2018-05-08 | E18SM4278A | 1.1 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
404 | 2018-05-08 | E18SM4275A | 1.4 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
405 | 2018-05-08 | E18SM4280A | 2.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
406 | 2018-05-08 | E18SM4281A | 2.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
407 | 2018-05-09 | E18SM4283A | 5.1 | 0.0001 | 0.002 | Resolve unit | Double Press Seal Eval
408 | 2018-05-09 | E18SM4284A | 2.4 | 0.0015 | 0.002 | Resolve unit | Double Press Seal Eval
409 | 2018-05-09 | E18SM4285A | 5.3 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
410 | 2018-05-09 | E18SM4289A | 4.2 | 0.0002 | 0.002 | Resolve unit | Double Press Seal Eval
411 | 2018-05-09 | E18SM4290A | 3.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
412 | 2018-05-09 | E18SM4291A | 1.3 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
413 | 2018-05-09 | E18SM4292A | 2.3 | NA | 0.002 | Pass on first test | Double Press Seal Eval
414 | 2018-05-09 | E18SM4293A | 2.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
415 | 2018-05-09 | E18SM4294A | 2.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
416 | 2018-05-09 | E18SM4295A | 1.4 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
417 | 2018-05-09 | E18SM4297A | 3.1 | NA | 0.002 | Pass on first test | Double Press Seal Eval
418 | 2018-05-09 | E18SM4298A | 1.1 | NA | 0.002 | Resolve unit | Double Press Seal Eval
419 | 2018-05-10 | E18SM4276A | 1.2 | 0.0001 | 0.002 | Resolve unit | Double Press Seal Eval
420 | 2018-05-10 | E18SM4277A | 5.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
421 | 2018-05-10 | E18SM4288A | 3.3 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
422 | 2018-05-09 | E18SM4299A | 5.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
423 | 2018-05-09 | E18SM4300A | 5.3 | 0.0012 | 0.002 | Pass on first test | Double Press Seal Eval
424 | 2018-05-09 | E18SM4301A | 3.2 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
425 | 2018-05-10 | E18SM4302A | 2.4 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
426 | 2018-05-10 | E18SM4303A | 2.3 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
427 | 2018-05-10 | E18SM4304A | 1.4 | 0.0012 | 0.002 | Pass on first test | Double Press Seal Eval
428 | 2018-05-10 | E18SM4305A | 1.3 | 0.0018 | 0.002 | Pass on first test | Double Press Seal Eval
429 | 2018-05-10 | E18SM4306A | 3.1 | 0.0006 | 0.002 | Resolve unit | Double Press Seal Eval
430 | 2018-05-10 | E18SM4311A | 5.2 | 0.0009 | 0.002 | Resolve unit | Double Press Seal Eval
431 | 2018-05-10 | E18SM4312A | 1.1 | 0.0001 | 0.002 | Resolve unit | Double Press Seal Eval
432 | 2018-05-10 | E18SM4313A | 4.4 | 0.0012 | 0.002 | Pass on first test | Double Press Seal Eval
433 | 2018-05-10 | E18SM4315A | 5.3 | 0.0008 | 0.002 | Resolve unit | Double Press Seal Eval
434 | 2018-05-10 | E18SM4319A | 2.3 | 0.0017 | 0.002 | Pass on first test | Double Press Seal Eval
435 | 2018-05-10 | E18SM4320A | 2.4 | 0.0014 | 0.002 | Pass on first test | Double Press Seal Eval
436 | 2018-05-10 | E18SM4323A | 3.1 | 0.0019 | 0.002 | Pass on first test | Double Press Seal Eval
437 | 2018-05-11 | E18SM4231A | 3.3 | 0.0015 | 0.002 | Resolve unit | Double Press Seal Eval
438 | 2018-05-11 | E18SM4307A | 2.1 | 0.0012 | 0.002 | Resolve unit | Double Press Seal Eval
439 | 2018-05-11 | E18SM4309A | 4.1 | NA | 0.002 | Resolve unit | Double Press Seal Eval
440 | 2018-05-11 | E18SM4310A | 4.2 | 0.0014 | 0.002 | Resolve unit | Double Press Seal Eval
441 | 2018-05-11 | E18SM4314A | 5.1 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
442 | 2018-05-11 | E18SM4316A | 1.2 | 0.0004 | 0.002 | Resolve unit | Double Press Seal Eval
443 | 2018-05-10 | E18SM4318A | 3.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
444 | 2018-05-11 | E18SM4321A | 1.4 | 0.0001 | 0.002 | Pass on first test | Double Press Seal Eval
445 | 2018-05-11 | E18SM4322A | 1.3 | 0.0019 | 0.002 | Resolve unit | Double Press Seal Eval
446 | 2018-05-11 | E18SM4324A | 1.1 | 0.0006 | 0.002 | Pass on first test | Double Press Seal Eval
447 | 2018-05-11 | E18SM4327A | 2.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
448 | 2018-05-11 | E18SM4328A | 5.3 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
449 | 2018-05-11 | E18SM4329A | 3.2 | 0.0013 | 0.002 | Resolve unit | Double Press Seal Eval
450 | 2018-05-11 | E18SM4330A | 4.2 | 0.0017 | 0.002 | Pass on first test | Double Press Seal Eval
451 | 2018-05-12 | E18SM4331A | 1.4 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
452 | 2018-05-11 | E18SM4332A | 1.3 | 0.0015 | 0.002 | Pass on first test | Double Press Seal Eval
453 | 2018-05-12 | E18SM4333A | 3.3 | 0.0014 | 0.002 | Resolve unit | Double Press Seal Eval
454 | 2018-05-12 | E18SM4286A | 4.3 | 0.0011 | 0.002 | Resolve unit | Double Press Seal Eval
455 | 2018-05-12 | E18SM4325A | 3.1 | 0.0019 | 0.002 | Resolve unit | Double Press Seal Eval
456 | 2018-05-12 | E18SM4334A | 2.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
457 | 2018-05-12 | E18SM4335A | 1.1 | 0.0011 | 0.002 | Resolve unit | Double Press Seal Eval
458 | 2018-05-12 | E18SM4336A | 4.2 | 0.0012 | 0.002 | Resolve unit | Double Press Seal Eval
459 | 2018-05-14 | E18SM4337A | 2.4 | 0.0011 | 0.002 | Resolve unit | Double Press Seal Eval
460 | 2018-05-12 | E18SM4338A | 3.2 | 0.0012 | 0.002 | Resolve unit | Double Press Seal Eval
461 | 2018-05-14 | E18SM4339A | 2.1 | 0.0009 | 0.002 | Resolve unit | Double Press Seal Eval
462 | 2018-05-12 | E18SM4340A | 4.1 | 0.0016 | 0.002 | Pass on first test | Double Press Seal Eval
463 | 2018-05-12 | E18SM4341A | 5.1 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
464 | 2018-05-12 | E18SM4342A | 1.2 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
465 | 2018-05-14 | E18SM4343A | 5.3 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
466 | 2018-05-14 | E18SM4344A | 1.3 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
467 | 2018-05-14 | E18SM4345A | 3.3 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
468 | 2018-05-14 | E18SM4346A | 5.1 | 0.0019 | 0.002 | Pass on first test | Double Press Seal Eval
469 | 2018-05-14 | E18SM4347A | 3.1 | 0.0011 | 0.002 | Pass on first test | Double Press Seal Eval
470 | 2018-05-12 | E18SM4348A | 1.4 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
471 | 2018-05-14 | E18SM4349A | 1.1 | 0.0013 | 0.002 | Resolve unit | Double Press Seal Eval
472 | 2018-05-14 | E18SM4350A | 4.1 | 0.0016 | 0.002 | Pass on first test | Double Press Seal Eval
473 | 2018-05-14 | E18SM4351A | 4.2 | 0.0004 | 0.002 | Pass on first test | Double Press Seal Eval
474 | 2018-05-12 | E18SM4352A | 1.2 | 0.0003 | 0.002 | Pass on first test | Double Press Seal Eval
475 | 2018-05-14 | E18SM4353A | 3.2 | 0.001 | 0.002 | Resolve unit | Double Press Seal Eval
476 | 2018-05-14 | E18SM4355A | 2.4 | 0.0008 | 0.002 | Resolve unit | Double Press Seal Eval
477 | 2018-05-14 | E18SM4356A | 5.4 | 0.0006 | 0.002 | Pass on first test | Double Press Seal Eval
478 | 2018-05-14 | E18SM4357A | 4.3 | 0.0016 | 0.002 | Resolve unit | Double Press Seal Eval
479 | 2018-05-14 | E18SM4358A | 3.4 | 0.0016 | 0.002 | Pass on first test | Double Press Seal Eval
480 | 2018-05-14 | E18SM4360A | 3.1 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
481 | 2018-05-14 | E18SM4363A | 4.2 | NA | 0.002 | Pass on first test | Double Press Seal Eval
482 | 2018-05-14 | E18SM4365A | 1.1 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
483 | 2018-05-14 | E18SM4224A | 2.3 | 0.0005 | 0.002 | Resolve unit | Double Press Seal Eval
484 | 2018-05-12 | E18SM4287A | 3.4 | 0.0015 | 0.002 | Resolve unit | Double Press Seal Eval
485 | 2018-05-14 | E18SM4326A | 3.3 | 0.0002 | 0.002 | Pass on first test | Double Press Seal Eval
486 | 2018-05-14 | E18SM7988A | 1.2 | 0.0003 | 0.002 | Resolve unit | Double Press Seal Eval
487 | 2018-05-15 | E18SM4361A | 5.2 | 0.0019 | 0.002 | Resolve unit | Double Press Seal Eval
488 | 2018-05-15 | E18SM4364A | 3.2 | 0.0014 | 0.002 | Resolve unit | Double Press Seal Eval
489 | 2018-05-15 | E18SM4296A | 2.3 | 0.0008 | 0.002 | Pass on first test | Double Press Seal Eval
490 | 2018-05-15 | E18SM4308A | 2.2 | 0.001 | 0.002 | Pass on first test | Double Press Seal Eval
491 | 2018-05-15 | E18SM4362A | 3.4 | 0.0011 | 0.002 | Pass on first test | Double Press Seal Eval
492 | 2018-05-15 | E18SM4367A | 2.4 | 0.0007 | 0.002 | Resolve unit | Double Press Seal Eval
493 | 2018-05-15 | E18SM4368A | 3.3 | 0.0018 | 0.002 | Pass on first test | Double Press Seal Eval
494 | 2018-05-15 | E18SM4369A | 1.1 | 0.0006 | 0.002 | Pass on first test | Double Press Seal Eval
495 | 2018-05-15 | E18SM4373A | 1.2 | 0.0009 | 0.002 | Pass on first test | Double Press Seal Eval
496 | 2018-05-15 | E18SM4376A | 2.3 | 0.0005 | 0.002 | Pass on first test | Double Press Seal Eval
497 | 2018-05-15 | E18SM4371A | 5.1 | 0.0007 | 0.002 | Pass on first test | Double Press Seal Eval
498 | 2018-05-15 | E18SM4279A | 4.1 | 0.0017 | 0.002 | Resolve unit | Flex 20:80 Mix
499 | 2018-05-15 | E18SM4317A | 4.4 | 0.0009 | 0.002 | Resolve unit | Flex 20:80 Mix
500 | 2018-05-15 | E18SM4366A | 5.4 | 0.0012 | 0.002 | Pass on first test | Flex 20:80 Mix
501 | 2018-05-15 | E18SM4375A | 4.2 | 0.0009 | 0.002 | Pass on first test | Flex 20:80 Mix
502 | 2018-05-15 | E18SM4379A | 3.4 | 0.0005 | 0.002 | Pass on first test | Flex 20:80 Mix
503 | 2018-05-15 | E18SM4382A | 3.3 | 0.0004 | 0.002 | Pass on first test | Flex 20:80 Mix
504 | 2018-05-16 | E18SM4354A | 1.4 | 0.0005 | 0.002 | Resolve unit | MC 20:80 Mix
505 | 2018-05-16 | E18SM4359A | 2.1 | 0.0009 | 0.002 | Resolve unit | MC 20:80 Mix
506 | 2018-05-16 | E18SM4372A | 3.1 | 0.0012 | 0.002 | Resolve unit | MC 20:80 Mix
507 | 2018-05-16 | E18SM4377A | 3.2 | 0.0002 | 0.002 | Resolve unit | MC 20:80 Mix
508 | 2018-05-16 | E18SM4378A | 4.3 | NA | 0.002 | Resolve unit | MC 20:80 Mix
509 | 2018-05-16 | E18SM4380A | 5.1 | 0.0012 | 0.002 | Pass on first test | MC 20:80 Mix
510 | 2018-05-16 | E18SM4383A | 2.3 | NA | 0.002 | Resolve unit | MC 20:80 Mix
511 | 2018-05-16 | E18SM4385A | 2.2 | 0.0009 | 0.002 | Resolve unit | MC 20:80 Mix
512 | 2018-05-16 | E18SM4386A | 1.2 | 0.001 | 0.002 | Pass on first test | MC 20:80 Mix
513 | 2018-05-16 | E18SM4390A | 4.4 | 0.0007 | 0.002 | Pass on first test | MC 20:80 Mix
514 | 2018-05-16 | E18SM7991A | 3.3 | 0.0019 | 0.002 | Resolve unit | MC 20:80 Mix
515 | 2018-05-16 | E18SM4384A | 3.4 | 0.0019 | 0.002 | Resolve unit | MC 20:80 Mix
516 | 2018-05-16 | E18SM4389A | 4.2 | 0.0006 | 0.002 | Resolve unit | MC 20:80 Mix
517 | 2018-05-16 | E18SM7990A | 4.3 | 0.0006 | 0.002 | Pass on first test | MC 20:80 Mix
518 | 2018-05-16 | E18SM7993A | 5.4 | 0.0008 | 0.002 | Pass on first test | MC 20:80 Mix
519 | 2018-05-16 | E18SM7000A | 4.4 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
520 | 2018-05-16 | E18SM7001A | 4.1 | 0.0014 | 0.002 | Resolve unit | MC 20:80 Mix
521 | 2018-05-16 | E18SM7994A | 3.2 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
522 | 2018-05-16 | E18SM7995A | 2.1 | 0.0007 | 0.002 | Pass on first test | MC 20:80 Mix
523 | 2018-05-16 | E18SM7996A | 1.4 | NA | 0.002 | Pass on first test | MC 20:80 Mix
524 | 2018-05-16 | E18SM7997A | 2.2 | 0.0015 | 0.002 | Resolve unit | MC 20:80 Mix
525 | 2018-05-17 | E18SM4374A | 5.2 | 0.0014 | 0.002 | Resolve unit | MC 20:80 Mix
526 | 2018-05-17 | E18SM4381A | 2.4 | 0.0011 | 0.002 | Resolve unit | MC 20:80 Mix
527 | 2018-05-17 | E18SM7998A | 3.1 | 0.0003 | 0.002 | Resolve unit | MC 20:80 Mix
528 | 2018-05-17 | E18SM7999A | 1.2 | 0.0009 | 0.002 | Resolve unit | MC 20:80 Mix
529 | 2018-05-17 | E18SM7002A | 5.1 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
530 | 2018-05-17 | E18SM7003A | 2.3 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
531 | 2018-05-17 | E18SM7004A | 4.2 | 0.001 | 0.002 | Resolve unit | MC 20:80 Mix
532 | 2018-05-17 | E18SM7005A | 3.2 | 0.001 | 0.002 | Resolve unit | MC 20:80 Mix
533 | 2018-05-17 | E18SM7006A | 4.3 | 0.0014 | 0.002 | Pass on first test | MC 20:80 Mix
534 | 2018-05-17 | E18SM7007A | 2.1 | 0.0006 | 0.002 | Resolve unit | MC 20:80 Mix
535 | 2018-05-17 | E18SM7008A | 3.3 | 0.0003 | 0.002 | Pass on first test | MC 20:80 Mix
536 | 2018-05-17 | E18SM7009A | 3.4 | 0.0003 | 0.002 | Resolve unit | MC 20:80 Mix
537 | 2018-05-17 | E18SM7011A | 4.4 | 0.0019 | 0.002 | Pass on first test | MC 20:80 Mix
538 | 2018-05-17 | E18SM7012A | 1.4 | 0.0014 | 0.002 | Pass on first test | MC 20:80 Mix
539 | 2018-05-17 | E18SM7013A | 5.2 | 0.0013 | 0.002 | Pass on first test | MC 20:80 Mix
540 | 2018-05-17 | E18SM7018A | 2.4 | NA | 0.002 | Pass on first test | MC 20:80 Mix
541 | 2018-05-17 | E18SM7992A | 5.3 | 0.0012 | 0.002 | Resolve unit | MC 20:80 Mix
542 | 2018-05-18 | E18SM7010A | 4.1 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
543 | 2018-05-18 | E18SM7016A | 4.2 | 0.0004 | 0.002 | Pass on first test | MC 20:80 Mix
544 | 2018-05-18 | E18SM7019A | 1.3 | 0.0019 | 0.002 | Resolve unit | MC 20:80 Mix
545 | 2018-05-18 | E18SM7020A | 1.4 | 0.0019 | 0.002 | Resolve unit | MC 20:80 Mix
546 | 2018-05-18 | E18SM7021A | 4.3 | 0.0008 | 0.002 | Pass on first test | MC 20:80 Mix
547 | 2018-05-19 | E18SM7022A | 3.2 | 0.0008 | 0.002 | Resolve unit | MC 20:80 Mix
548 | 2018-05-19 | E18SM7023A | 3.1 | NA | 0.002 | Resolve unit | MC 20:80 Mix
549 | 2018-05-18 | E18SM7026A | 3.4 | 0.0009 | 0.002 | Pass on first test | MC 20:80 Mix
550 | 2018-05-18 | E18SM7029A | 3.3 | 0.0009 | 0.002 | Pass on first test | MC 20:80 Mix
551 | 2018-05-18 | E18SM7030A | 2.3 | 0.0008 | 0.002 | Pass on first test | MC 20:80 Mix
552 | 2018-05-18 | E18SM7032A | 4.1 | 0.0014 | 0.002 | Pass on first test | MC 20:80 Mix
553 | 2018-05-19 | E18SM7033A | 5.1 | NA | 0.002 | Pass on first test | MC 20:80 Mix
554 | 2018-05-19 | E18SM7034A | 4.3 | NA | 0.002 | Pass on first test | MC 20:80 Mix
555 | 2018-05-19 | E18SM7036A | 4.4 | NA | 0.002 | Pass on first test | MC 20:80 Mix
556 | 2018-05-19 | E18SM7044A | 4.1 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
557 | 2018-05-19 | E18SM7037A | 4.2 | NA | 0.002 | Resolve unit | MC 20:80 Mix
558 | 2018-05-19 | E18SM7039A | 2.3 | 0.0001 | 0.002 | Pass on first test | MC 20:80 Mix
559 | 2018-05-19 | E18SM7043A | 1.4 | 0.0003 | 0.002 | Pass on first test | MC 20:80 Mix
560 | 2018-05-19 | E18SM7046A | 5.1 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
561 | 2018-05-18 | E18SM4387A | 5.4 | 0.0005 | 0.002 | Pass on first test | MC 20:80 Mix
562 | 2018-05-19 | E18SM4388A | 1.1 | 0.0009 | 0.002 | Resolve unit | MC 20:80 Mix
563 | 2018-05-21 | E18SM7038A | 3.3 | 0.0004 | 0.002 | Pass on first test | MC 20:80 Mix
564 | 2018-05-21 | E18SM7041A | 5.4 | 0.0001 | 0.002 | Pass on first test | MC 20:80 Mix
565 | 2018-05-21 | E18SM7042A | 1.3 | 0.0001 | 0.002 | Resolve unit | MC 20:80 Mix
566 | 2018-05-21 | E18SM7045A | 3.1 | 0.0017 | 0.002 | Resolve unit | MC 20:80 Mix
567 | 2018-05-21 | E18SM7047A | 3.2 | 0.0006 | 0.002 | Resolve unit | MC 20:80 Mix
568 | 2018-05-21 | E18SM4397A | 4.1 | 0.0006 | 0.002 | Pass on first test | MC 20:80 Mix
569 | 2018-05-21 | E18SM4399A | 5.1 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
570 | 2018-05-21 | E18SM4401A | 4.2 | NA | 0.002 | Pass on first test | MC 20:80 Mix
571 | 2018-05-21 | E18SM4392A | 4.3 | 0.0018 | 0.002 | Pass on first test | MC 20:80 Mix
572 | 2018-05-22 | E18SM4404A | 4.1 | 0.0003 | 0.002 | Resolve unit | MC 20:80 Mix
573 | 2018-05-22 | E18SM4409A | 4.4 | 0.0001 | 0.002 | Resolve unit | MC 20:80 Mix
574 | 2018-05-22 | E18SM4411A | 3.3 | 0.001 | 0.002 | Pass on first test | MC 20:80 Mix
575 | 2018-05-22 | E18SM4412A | 1.3 | 0.0001 | 0.002 | Pass on first test | MC 20:80 Mix
576 | 2018-05-22 | E18SM4414A | 2.4 | 0.0004 | 0.002 | Pass on first test | MC 20:80 Mix
577 | 2018-05-22 | E18SM4416A | 3.2 | 0.0012 | 0.002 | Pass on first test | MC 20:80 Mix
578 | 2018-05-22 | E18SM4418A | 3.4 | 0.0007 | 0.002 | Resolve unit | MC 20:80 Mix
579 | 2018-05-22 | E18SM4421A | 5.3 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix
580 | 2018-05-22 | E18SM4423A | 4.2 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix
581 | 2018-05-22 | E18SM4425A | 2.4 | 0.0004 | 0.002 | Pass on first test | MC 20:80 Mix
582 | 2018-05-22 | E18SM4426A | 4.3 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix
583 | 2018-05-22 | E18SM4432A | 2.3 | 0.0003 | 0.002 | Pass on first test | MC 20:80 Mix
584 | 2018-05-22 | E18SM7014A | 2.2 | 0.0013 | 0.002 | Resolve unit | MC 20:80 Mix
585 | 2018-05-22 | E18SM7017A | 1.2 | 0.0003 | 0.002 | Resolve unit | MC 20:80 Mix
586 | 2018-05-22 | E18SM7024A | 3.1 | 0.0017 | 0.002 | Pass on first test | MC 20:80 Mix
587 | 2018-05-22 | E18SM7031A | 5.2 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix
588 | 2018-05-22 | E18SM4413A | 5.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix
589 | 2018-05-23 | E18SM4434A | 3.1 | NA | 0.002 | Pass on first test | MC 20:80 Mix
590 | 2018-05-23 | E18SM4429A | 4.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
591 | 2018-05-23 | E18SM4431A | 4.4 | 0.0018 | 0.002 | Resolve unit | WPPL 20:80 Mix
592 | 2018-05-23 | E18SM4437A | 4.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
593 | 2018-05-23 | E18SM4427A | 3.2 | 0.001 | 0.002 | Pass on first test | MC 20:80 Mix
594 | 2018-05-23 | E18SM4407A | 2.1 | 0.0004 | 0.002 | Pass on first test | MC 20:80 Mix
595 | 2018-05-23 | E18SM4420A | 5.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix
596 | 2018-05-23 | E18SM4435A | 4.2 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix
597 | 2018-05-23 | E18SM4419A | 3.3 | 0.0007 | 0.002 | Pass on first test | MC 20:80 Mix
598 | 2018-05-23 | E18SM4422A | 3.2 | 0.001 | 0.002 | Pass on first test | MC 20:80 Mix
599 | 2018-05-23 | E18SM4433A | 2.3 | 0.0007 | 0.002 | Resolve unit | MC 20:80 Mix
600 | 2018-05-23 | E18SM4436A | 3.4 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
601 | 2018-05-23 | E18SM4441A | 4.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix
602 | 2018-05-23 | E18SM4442A | 3.2 | 0.0004 | 0.002 | Resolve unit | MC 20:80 Mix
603 | 2018-05-23 | E18SM4443A | 1.3 | 0.001 | 0.002 | Pass on first test | MC 20:80 Mix
604 | 2018-05-23 | E18SM4449A | 4.3 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix
605 | 2018-05-23 | E18SM4450A | 3.4 | 0.0008 | 0.002 | Pass on first test | MC 20:80 Mix
606 | 2018-05-23 | E18SM7025A | 1.3 | 0.0012 | 0.002 | Resolve unit | MC 20:80 Mix
607 | 2018-05-23 | E18SM7035A | 2.2 | 0.0002 | 0.002 | Pass on first test | MC 20:80 Mix
608 | 2018-05-24 | E18SM4445A | 1.2 | 0.0018 | 0.002 | Pass on first test | WPPL 20:80 Mix
609 | 2018-05-24 | E18SM4440A | 5.1 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix
610 | 2018-05-24 | E18SM4455A | 4.1 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix
611 | 2018-05-24 | E18SM4451A | 5.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
612 | 2018-05-24 | E18SM4438A | 1.1 | NA | 0.002 | Pass on first test | MC 20:80 Mix
613 | 2018-05-24 | E18SM4463A | 1.1 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix
614 | 2018-05-24 | E18SM4415A | 1.4 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix
615 | 2018-05-24 | E18SM4461A | 1.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix
616 | 2018-05-24 | E18SM4447A | 2.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix
617 | 2018-05-24 | E18SM4462A | 2.4 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix
618 | 2018-05-24 | E18SM4458A | 3.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
619 | 2018-05-24 | E18SM4430A | 3.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
620 | 2018-05-24 | E18SM4471A | 3.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
621 | 2018-05-24 | E18SM4405A | 4.2 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix
622 | 2018-05-24 | E18SM4474A | 4.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix
623 | 2018-05-24 | E18SM4468A | 4.2 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix
624 | 2018-05-24 | E18SM4469A | 5.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix
625 | 2018-05-24 | E18SM4453A | 5.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix
626 | 2018-05-25 | E18SM4448A | 3.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix
627 | 2018-05-25 | E18SM4452A | 2.2 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix
628 | 2018-05-25 | E18SM4454A | 2.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
629 | 2018-05-25 | E18SM4457A | 4.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
630 | 2018-05-25 | E18SM4459A | 3.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
631 | 2018-05-25 | E18SM4465A | 3.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
632 | 2018-05-25 | E18SM4466A | 4.4 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix
633 | 2018-05-25 | E18SM4467A | 1.2 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix
634 | 2018-05-25 | E18SM4470A | 5.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
635 | 2018-05-25 | E18SM4472A | 4.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
636 | 2018-05-25 | E18SM4473A | 1.3 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix
637 | 2018-05-25 | E18SM4475A | 5.1 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix
638 | 2018-05-25 | E18SM4476A | 1.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
639 | 2018-05-25 | E18SM4478A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
640 | 2018-05-25 | E18SM4479A | 5.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
641 | 2018-05-25 | E18SM4480A | 2.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
642 | 2018-05-25 | E18SM4481A | 3.4 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix
643 | 2018-05-25 | E18SM4482A | 2.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
644 | 2018-05-25 | E18SM4485A | 4.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix
645 | 2018-05-25 | E18SM4486A | 1.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix
646 | 2018-05-25 | E18SM4487A | 3.2 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix
647 | 2018-05-25 | E18SM4488A | 1.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix
648 | 2018-05-26 | E18SM4456A | 1.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
649 | 2018-05-26 | E18SM4398A | 4.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
650 | 2018-05-25 | E18SM4439A | 5.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix
651 | 2018-05-25 | E18SM4464A | 2.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix
652 | 2018-05-26 | E18SM4477A | 3.2 | 0.0001 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
653 | 2018-05-26 | E18SM4489A | 1.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
654 | 2018-05-26 | E18SM4490A | 2.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
655 | 2018-05-26 | E18SM4491A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
656 | 2018-05-26 | E18SM4493A | 4.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
657 | 2018-05-26 | E18SM4494A | 3.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
658 | 2018-05-26 | E18SM4495A | 1.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
659 | 2018-05-26 | E18SM4496A | 1.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
660 | 2018-05-26 | E18SM4497A | 2.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
661 | 2018-05-26 | E18SM4499A | 1.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
662 | 2018-05-26 | E18SM4500A | 5.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
663 | 2018-05-26 | E18SM4504A | 1.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
664 | 2018-05-26 | E18SM4483A | 2.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
665 | 2018-05-28 | E18SM4424A | 2.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
666 | 2018-05-28 | E18SM4498A | 4.2 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
667 | 2018-05-28 | E18SM4501A | 5.2 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
668 | 2018-05-28 | E18SM4502A | 5.4 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
669 | 2018-05-28 | E18SM4505A | 3.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
670 | 2018-05-28 | E18SM4509A | 5.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
671 | 2018-05-28 | E18SM4510A | 2.3 | 0.0018 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
672 | 2018-05-28 | E18SM4511A | 2.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
673 | 2018-05-28 | E18SM4512A | 2.2 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
674 | 2018-05-28 | E18SM4513A | 1.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
675 | 2018-05-28 | E18SM4514A | 4.3 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
676 | 2018-05-28 | E18SM4516A | 3.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
677 | 2018-05-28 | E18SM4520A | 1.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
678 | 2018-05-28 | E18SM4525A | 3.1 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
679 | 2018-05-30 | E18SM4522A | 5.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
680 | 2018-05-25 | E18SM4484A | 4.4 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
681 | 2018-05-30 | E18SM4517A | 4.2 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
682 | 2018-05-30 | E18SM4521A | 5.2 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
683 | 2018-05-30 | E18SM4523A | 2.4 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
684 | 2018-05-30 | E18SM4526A | 3.4 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
685 | 2018-05-30 | E18SM4528A | 2.1 | 0.0018 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
686 | 2018-05-30 | E18SM4530A | 1.4 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
687 | 2018-05-30 | E18SM4531A | 4.4 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
688 | 2018-05-30 | E18SM4533A | 2.3 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
689 | 2018-05-30 | E18SM4536A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
690 | 2018-05-30 | E18SM4537A | 2.2 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
691 | 2018-05-30 | E18SM4538A | 2.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
692 | 2018-05-30 | E18SM4539A | 1.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
693 | 2018-05-30 | E18SM4540A | 3.2 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
694 | 2018-05-31 | E18SM4393A | 5.4 | 0.0013 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
695 | 2018-05-31 | E18SM4444A | 4.1 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
696 | 2018-05-31 | E18SM4446A | 2.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
697 | 2018-05-31 | E18SM4460A | 1.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
698 | 2018-05-31 | E18SM4508A | 5.1 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
699 | 2018-05-31 | E18SM4518A | 5.1 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
700 | 2018-05-31 | E18SM4532A | 1.2 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
701 | 2018-05-31 | E18SM4535A | 3.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
702 | 2018-05-31 | E18SM4541A | 5.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
703 | 2018-05-31 | E18SM4542A | 5.3 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
704 | 2018-05-31 | E18SM4543A | 2.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
705 | 2018-05-31 | E18SM4544A | 3.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
706 | 2018-05-31 | E18SM4545A | 2.4 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
707 | 2018-05-31 | E18SM4548A | 2.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
708 | 2018-05-31 | E18SM4551A | 1.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
709 | 2018-05-31 | E18SM4552A | 2.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
710 | 2018-05-31 | E18SM4553A | 1.3 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
711 | 2018-05-31 | E18SM4519A | 1.3 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
712 | 2018-05-31 | E18SM7040A | 5.3 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
713 | 2018-05-31 | E18SM7049A | 5.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
714 | 2018-05-31 | E18SM7052A | 3.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
715 | 2018-05-31 | E18SM7053A | 1.1 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
716 | 2018-05-31 | E18SM7056A | 2.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
717 | 2018-05-31 | E18SM4527A | 3.3 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
718 | 2018-05-31 | E18SM4524A | 4.4 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
719 | 2018-06-01 | E18SM7015A | 5.4 | 0.0019 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
720 | 2018-06-01 | E18SM7055A | 4.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
721 | 2018-06-01 | E18SM7060A | 2.3 | 0.0007 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
722 | 2018-06-01 | E18SM7061A | 1.3 | 0.0018 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
723 | 2018-06-01 | E18SM7070A | 1.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
724 | 2018-06-01 | E18SM7051A | 5.1 | 0.0019 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
725 | 2018-06-01 | E18SM7058A | 2.1 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
726 | 2018-06-01 | E18SM7054A | 3.4 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
727 | 2018-06-01 | E18SM7050A | 3.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
728 | 2018-06-01 | E18SM7064A | 4.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
729 | 2018-06-01 | E18SM7066A | 2.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
730 | 2018-06-02 | E18SM7057A | 5.2 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
731 | 2018-06-02 | E18SM7067A | 1.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
732 | 2018-06-02 | E18SM7068A | 1.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
733 | 2018-06-02 | E18SM7059A | 1.2 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
734 | 2018-06-02 | E18SM7071A | 4.4 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
735 | 2018-06-02 | E18SM7074A | 5.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
736 | 2018-06-02 | E18SM7048A | 1.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
737 | 2018-06-02 | E18SM4062A | 3.2 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
738 | 2018-06-02 | E18SM4063A | 3.1 | 0.0009 | 0.002 | Resolve Unit | WPPL 20:80 Mix (3rd bottle)
739 | 2018-06-02 | E18SM4549A | 4.2 | 0.0013 | 0.002 | Resolve Unit | WPPL 20:80 Mix (3rd bottle)
740 | 2018-06-01 | E18SM4534A | 2.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
741 | 2018-06-02 | E18SM4550A | 4.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
742 | 2018-06-02 | E18SM4554A | 3.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
743 | 2018-06-02 | E18SM4555A | 5.1 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
744 | 2018-06-02 | E18SM4556A | 2.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
745 | 2018-06-02 | E18SM4557A | 2.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
746 | 2018-06-02 | F18SM4559A | 2.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
747 | 2018-05-28 | E18SM4507A | 3.3 | N.A | 0.002 | Pass on first test | WPPL 20:80 Mix (2nd bottle)
748 | 2018-06-04 | E18SM4515A | 2.2 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
749 | 2018-06-04 | E18SM4529A | 4.1 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
750 | 2018-06-04 | F18SM4561A | 1.4 | 0.0018 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
751 | 2018-06-04 | F18SM4562A | 1.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
752 | 2018-06-04 | F18SM4563A | 5.2 | N.A | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
753 | 2018-06-04 | F18SM4564A | 2.3 | N.A | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
754 | 2018-06-04 | F18SM4566A | 1.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
755 | 2018-06-04 | F18SM4567A | 3.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
756 | 2018-06-04 | F18SM4568A | 1.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
757 | 2018-06-04 | F18SM4569A | 2.1 | N.A | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
758 | 2018-06-04 | F18SM4571A | 5.1 | N.A | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
759 | 2018-06-04 | F18SM4572A | 4.4 | N.A | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
760 | 2018-06-04 | F18SM4573A | 5.3 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
761 | 2018-06-04 | F18SM4578A | 2.2 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
762 | 2018-06-04 | F18SM4579A | 2.3 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
763 | 2018-06-04 | F18SM4585A | 4.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
764 | 2018-06-04 | E18SM7028A | 1.4 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
765 | 2018-06-04 | E18SM7065A | 4.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
766 | 2018-06-04 | E18SM7069A | 4.2 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
767 | 2018-06-04 | E18SM7072A | 3.4 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
768 | 2018-06-04 | E18SM7073A | 5.4 | N.A | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
769 | 2018-06-04 | E18SM7075A | 4.1 | N.A | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
770 | 2018-06-05 | E18SM4370A | 3.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
771 | 2018-06-04 | E18SM4395A | 3.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
772 | 2018-06-04 | E18SM4506A | 1.2 | 0.0016 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
773 | 2018-06-04 | E18SM4558A | 3.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
774 | 2018-06-05 | E18SM4560A | 2.5 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
775 | 2018-06-04 | E18SM4565A | 3.3 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
776 | 2018-06-05 | E18SM4574A | 1.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
777 | 2018-06-04 | E18SM4575A | 3.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
778 | 2018-06-05 | E18SM4576A | 5.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
779 | 2018-06-05 | E18SM4577A | 5.3 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
780 | 2018-06-05 | E18SM4582A | 2.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
781 | 2018-06-05 | E18SM4583A | 5.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
782 | 2018-06-04 | E18SM4584A | 5.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
783 | 2018-06-04 | E18SM4587A | 4.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
784 | 2018-06-05 | E18SM4588A | 2.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
785 | 2018-06-05 | E18SM4589A | 1.3 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
786 | 2018-06-05 | E18SM4592A | 3.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
787 | 2018-06-05 | E18SM4593A | 1.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
788 | 2018-06-05 | E18SM4594A | 3.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
789 | 2018-06-05 | E18SM4597A | 2.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
790 | 2018-06-05 | F18SM4581A | 4.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
791 | 2018-06-05 | F18SM4586A | 1.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (3rd bottle)
792 | 2018-06-05 | F18SM4595A | 4.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
793 | 2018-06-05 | F18SM4599A | 3.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
794 | 2018-06-05 | F18SM4602A | 2.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
795 | 2018-06-05 | F18SM4604A | 1.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
796 | 2018-06-05 | F18SM4608A | 2.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
797 | 2018-06-06 | E18SM4428A | 5.3 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
798 | 2018-06-06 | F18SM4570A | 4.1 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
799 | 2018-06-06 | F18SM4590A | 1.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
800 | 2018-06-06 | F18SM4591A | 3.1 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
801 | 2018-06-06 | F18SM4596A | 4.3 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
802 | 2018-06-06 | F18SM4598A | 5.1 | 0.0019 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
803 | 2018-06-06 | F18SM4601A | 1.2 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
804 | 2018-06-05 | F18SM4603A | 5.2 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
805 | 2018-06-05 | F18SM4606A | 2.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
806 | 2018-06-05 | F18SM4607A | 5.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (3rd bottle)
807 | 2018-06-06 | F18SM4609A | 2.2 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
808 | 2018-06-06 | F18SM4610A | 1.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
809 | 2018-06-06 | F18SM4611A | 1.4 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
810 | 2018-06-06 | F18SM4614A | 2.3 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
811 | 2018-06-06 | F18SM4616A | 2.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
812 | 2018-06-06 | F18SM4618A | 5.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
813 | 2018-06-06 | F18SM4619A | 1.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
814 | 2018-06-06 | F18SM4620A | 4.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
815 | 2018-06-06 | F18SM4621A | 1.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
816 | 2018-06-06 | F18SM4622A | 2.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
817 | 2018-06-06 | F18SM4623A | 5.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
818 | 2018-06-06 | F18SM4625A | 1.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
819 | 2018-06-06 | F18SM4627A | 3.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
820 | 2018-06-07 | F18SM4417A | 2.3 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
821 | 2018-06-07 | F18SM4580A | 4.4 | 0.0013 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
822 | 2018-06-06 | F18SM4612A | 3.3 | 0.0016 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
823 | 2018-06-07 | F18SM4613A | 4.2 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
824 | 2018-06-07 | F18SM4617A | 5.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
825 | 2018-06-07 | F18SM4624A | 1.3 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
826 | 2018-06-07 | F18SM4626A | 3.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
827 | 2018-06-07 | F18SM4628A | 2.1 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (4th bottle)
828 | 2018-06-07 | F18SM4629A | 1.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
829 | 2018-06-07 | F18SM4631A | 1.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
830 | 2018-06-07 | F18SM4633A | 2.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
831 | 2018-06-07 | F18SM4634A | 5.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
832 | 2018-06-07 | F18SM4635A | 4.3 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
833 | 2018-06-07 | F18SM4636A | 3.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
834 | 2018-06-07 | F18SM4637A | 4.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
835 | 2018-06-07 | F18SM4630A | 3.4 | 0.0007 | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
836 | 2018-06-08 | F18SM4639A | 3.1 | NA | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
837 | 2018-06-07 | F18SM7076A | 5.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
838 | 2018-06-08 | F18SM7077A | 5.4 | 0.0011 | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
839 | 2018-06-07 | F18SM7078A | 1.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
840 | 2018-06-07 | F18SM7079A | 2.3 | 0 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
841 | 2018-06-08 | F18SM7080A | 2.1 | 0.0005 | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
842 | 2018-06-07 | F18SM7081A | 2.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
843 | 2018-06-07 | F18SM7083A | 4.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
844 | 2018-06-08 | F18SM7085A | 4.1 | 0.0015 | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
845 | 2018-06-08 | F18SM7086A | 3.2 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
846 | 2018-06-08 | F18SM7087A | 5.1 | 0.0003 | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
847 | 2018-06-08 | F18SM7088A | 4.4 | 0.001 | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
848 | 2018-06-08 | F18SM7090A | 1.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
849 | 2018-06-08 | F18SM7094A | 4.3 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
850 | 2018-06-08 | F18SM7095A | 3.4 | 0 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
851 | 2018-06-08 | F18SM7096A | 5.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
852 | 2018-06-08 | F18SM7097A | 5.3 | 0.0005 | 0.002 | Resolve Unit | WPPL 20:80 Mix (4th bottle)
853 | 2018-06-08 | F18SM7098A | 2.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
854 | 2018-06-08 | F18SM7099A | 3.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
855 | 2018-06-08 | F18SM7100A | 3.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
856 | 2018-06-08 | F18SM7101A | 5.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
857 | 2018-06-08 | F18SM7104A | 4.3 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
858 | 2018-06-07 | F18SM7082A | 5.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
859 | 2018-06-09 | F18SM7092A | 2.3 | 0.0007 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
860 | 2018-06-09 | F18SM7102A | 3.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
861 | 2018-06-09 | F18SM7105A | 1.4 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
862 | 2018-06-08 | F18SM4640A | 4.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle)
863 | 2018-06-09 | F18SM4641A | 1.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
864 | 2018-06-09 | F18SM4642A | 2.4 | 0.0001 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
865 | 2018-06-09 | F18SM4643A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
866 | 2018-06-09 | F18SM4644A | 3.2 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
867 | 2018-06-09 | F18SM4646A | 4.3 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
868 | 2018-06-09 | F18SM4647A | 5.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
869 | 2018-06-09 | F18SM4648A | 4.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
870 | 2018-06-09 | F18SM4651A | 2.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
871 | 2018-06-11 | E18SM4547A | 4.4 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
872 | 2018-06-11 | F18SM4600A | 3.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
873 | 2018-06-11 | F18SM4638A | 5.1 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
874 | 2018-06-11 | F18SM4645A | 4.1 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
875 | 2018-06-11 | F18SM4649A | 5.2 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
876 | 2018-06-11 | F18SM4650A | 5.3 | 0.0017 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
877 | 2018-06-11 | F18SM4652A | 2.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
878 | 2018-06-11 | F18SM4653A | 4.4 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
879 | 2018-06-11 | F18SM4654A | 3.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
880 | 2018-06-11 | F18SM4655A | 3.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
881 | 2018-06-11 | F18SM4656A | 4.2 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
882 | 2018-06-11 | F18SM4657A | 1.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
883 | 2018-06-11 | F18SM4658A | 4.3 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
884 | 2018-06-11 | F18SM4659A | 3.2 | 0.0016 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
885 | 2018-06-11 | F18SM4660A | 2.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
886 | 2018-06-11 | F18SM4661A | 1.2 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
887 | 2018-06-11 | F18SM4663A | 2.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
888 | 2018-06-11 | F18SM4664A | 5.4 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
889 | 2018-06-11 | F18SM4665A | 4.1 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
890 | 2018-06-11 | F18SM4666A | 3.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
891 | 2018-06-11 | F18SM4667A | 1.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
892 | 2018-06-11 | F18SM4668A | 1.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
893 | 2018-06-11 | F18SM4670A | 1.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
894 | 2018-06-11 | F18SM4671A | 5.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
895 | 2018-06-11 | F18SM4673A | 4.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
896 | 2018-06-11 | F18SM4679A | 4.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
897 | 2018-06-11 | F18SM4682A | 3.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
898 | 2018-06-12 | F18SM7089A | 1.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
899 | 2018-06-12 | F18SM7093A | 2.2 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
900 | 2018-06-12 | F18SM7103A | 4.1 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
901 | 2018-06-12 | E18SM4492A | 3.4 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
902 | 2018-06-12 | F18SM4662A | 3.3 | 0.0009 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
903 | 2018-06-12 | F18SM4672A | 5.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
904 | 2018-06-12 | F18SM4674A | 2.4 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
905 | 2018-06-12 | F18SM4675A | 2.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
906 | 2018-06-12 | F18SM4676A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
907 | 2018-06-12 | F18SM4677A | 2.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
908 | 2018-06-12 | F18SM4678A | 2.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
909 | 2018-06-12 | F18SM4680A | 1.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
910 | 2018-06-12 | F18SM4681A | 3.4 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
911 | 2018-06-12 | F18SM4684A | 1.4 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
912 | 2018-06-12 | F18SM4685A | 3.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
913 | 2018-06-12 | F18SM4686A | 2.1 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
914 | 2018-06-12 | F18SM4687A | 2.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
915 | 2018-06-12 | F18SM4689A | 2.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
916 | 2018-06-12 | F18SM4691A | 4.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
917 | 2018-06-12 | F18SM4695A | 3.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
918 | 2018-06-12 | F18SM4698A | 2.1 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
919 | 2018-06-12 | E18SM4408A | 4.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
920 | 2018-06-12 | E18SM4546A | 5.4 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (2nd bottle)
921 | 2018-06-12 | F18SM4688A | 2.2 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
922 | 2018-06-12 | F18SM4692A | 3.3 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (5th bottle)
923 | 2018-06-13 | F18SM4693A | 3.1 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
924 | 2018-06-12 | F18SM4696A | 1.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
925 | 2018-06-12 | F18SM4697A | 2.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
926 | 2018-06-12 | F18SM4703A | 3.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
927 | 2018-06-13 | E18SM4503A | 4.4 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
928 | 2018-06-13 | F18SM4669A | 1.2 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
929 | 2018-06-13 | F18SM4683A | 1.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
930 | 2018-06-13 | F18SM4694A | 5.3 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
931 | 2018-06-13 | F18SM4700A | 4.3 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
932 | 2018-06-13 | F18SM4701A | 2.1 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
933 | 2018-06-13 | F18SM4702A | 2.3 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
934 | 2018-06-13 | F18SM4704A | 2.4 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
935 | 2018-06-13 | F18SM4705A | 3.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
936 | 2018-06-13 | F18SM4707A | 3.3 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
937 | 2018-06-13 | F18SM4708A | 4.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
938 | 2018-06-13 | F18SM4709A | 1.4 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
939 | 2018-06-13 | F18SM4710A | 1.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
940 | 2018-06-13 | F18SM4711A | 3.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
941 | 2018-06-12 | F18SM7091A | 5.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (5th bottle)
942 | 2018-06-13 | F18SM7106A | 5.4 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
943 | 2018-06-13 | F18SM7108A | 3.1 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
944 | 2018-06-13 | F18SM7110A | 4.3 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
945 | 2018-06-13 | F18SM7111A | 2.4 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
946 | 2018-06-13 | F18SM7113A | 3.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
947 | 2018-06-13 | F18SM7116A | 1.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
948 | 2018-06-14 | F18SM4690A | 5.4 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
949 | 2018-06-14 | F18SM4699A | 4.3 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
950 | 2018-06-14 | F18SM4706A | 2.2 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
951 | 2018-06-14 | F18SM7107A | 5.3 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
952 | 2018-06-14 | F18SM7109A | 4.1 | 0.0018 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
953 | 2018-06-14 | F18SM7112A | 4.2 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
954 | 2018-06-14 | F18SM7114A | 2.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
955 | 2018-06-14 | F18SM7115A | 1.3 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
956 | 2018-06-14 | F18SM7117A | 1.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
957 | 2018-06-14 | F18SM7118A | 1.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
958 | 2018-06-14 | F18SM7120A | 2.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
959 | 2018-06-14 | F18SM7121A | 2.4 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
960 | 2018-06-14 | F18SM7124A | 3.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
961 | 2018-06-14 | F18SM7126A | 4.4 | 0.0007 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
962 | 2018-06-14 | F18SM7127A | 5.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
963 | 2018-06-14 | F18SM7129A | 2.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
964 | 2018-06-14 | F18SM7130A | 1.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
965 | 2018-06-14 | F18SM7131A | 4.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
966 | 2018-06-14 | F18SM7133A | 4.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
967 | 2018-06-14 | F18SM7084A | 5.2 | 0.0001 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
968 | 2018-06-18 | F18SM7123A | 3.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
969 | 2018-06-18 | F18SM7132A | 1.3 | 0.0018 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
970 | 2018-06-18 | F18SM7119A | 5.1 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
971 | 2018-06-18 | F18SM7122A | 3.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle)
972 | 2018-06-18 | F18SM7128A | 5.1 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
973 | 2018-06-18 | F18SM7134A | 4.1 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
974 | 2018-06-18 | F18SM7135A | 4.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
975 | 2018-06-18 | F18SM7136A | 4.2 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
976 | 2018-06-18 | F18SM7137A | 4.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
977 | 2018-06-18 | F18SM7139A | 3.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
978 | 2018-06-18 | F18SM7141A | 2.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
979 | 2018-06-18 | F18SM7142A | 2.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
980 | 2018-06-18 | F18SM4712A | 2.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
981 | 2018-06-18 | F18SM4715A | 1.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
982 | 2018-06-18 | F18SM4716A | 2.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
983 | 2018-06-18 | F18SM4717A | 1.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
984 | 2018-06-18 | F18SM4718A | 1.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
985 | 2018-06-18 | F18SM4720A | 4.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle) Seal Pack Evaluation
986 | 2018-06-18 | F18SM4722A | 4.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
987 | 2018-06-18 | F18SM4724A | 1.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (6th bottle)
988 | 2018-06-19 | F18SM7138A | 5.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle) Seal Pack Evaluation
989 | 2018-06-19 | F18SM4713A | 2.1 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle) Seal Pack Evaluation
990 | 2018-06-19 | F18SM4714A | 3.2 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
991 | 2018-06-19 | F18SM4719A | 4.4 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
992 | 2018-06-19 | F18SM4721A | 4.2 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
993 | 2018-06-19 | F18SM4723A | 3.4 | 0.0009 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
994 | 2018-06-19 | F18SM4725A | 2.4 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
995 | 2018-06-19 | F18SM4726A | 5.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
996 | 2018-06-19 | F18SM4727A | 3.1 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
997 | 2018-06-19 | F18SM4729A | 4.3 | 0.0013 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
998 | 2018-06-19 | F18SM4730A | 2.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
999 | 2018-06-19 | F18SM4731A | 4.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1000 | 2018-06-19 | F18SM4732A | 1.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1001 | 2018-06-19 | F18SM4733A | 2.3 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1002 | 2018-06-19 | F18SM4734A | 1.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1003 | 2018-06-19 | F18SM4735A | 1.2 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1004 | 2018-06-19 | F18SM4737A | 1.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1005 | 2018-06-19 | F18SM4738A | 2.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1006 | 2018-06-19 | F18SM4736A | 5.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1007 | 2018-06-19 | F18SM4739A | 4.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1008 | 2018-06-19 | F18SM4740A | 3.2 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle) Enlargement of Seal Pack (WAT033276)
1009 | 2018-06-19 | F18SM4741A | 4.4 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle) Enlargement of Seal Pack (WAT033276)
1010 | 2018-06-19 | F18SM4742A | 5.2 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle) Enlargement of Seal Pack (WAT033276)
1011 | 2018-06-19 | F18SM4743A | 3.4 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle) Enlargement of Seal Pack (WAT033276)
1012 | 2018-06-19 | F18SM4744A | 2.3 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle) Enlargement of Seal Pack (WAT033276)
1013 | 2018-06-20 | F18SM4745A | 1.4 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1014 | 2018-06-19 | F18SM4746A | 2.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1015 | 2018-06-19 | F18SM4747A | 2.4 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1016 | 2018-06-20 | F18SM4748A | 1.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1017 | 2018-06-20 | F18SM4749A | 1.1 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1018 | 2018-06-20 | F18SM4750A | 2.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1019 | 2018-06-20 | F18SM4751A | 4.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1020 | 2018-06-20 | F18SM4752A | 1.2 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1021 | 2018-06-20 | F18SM4753A | 2.1 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1022 | 2018-06-20 | F18SM4754A | 5.3 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1023 | 2018-06-20 | F18SM4755A | 5.2 | 0.0018 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1024 | 2018-06-20 | F18SM4756A | 5.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1025 | 2018-06-20 | F18SM4757A | 2.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1026 | 2018-06-20 | F18SM4759A | 4.3 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1027 | 2018-06-20 | F18SM4761A | 4.2 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1028 | 2018-06-20 | F18SM4762A | 2.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1029 | 2018-06-20 | F18SM4764A | 3.2 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1030 | 2018-06-20 | F18SM4765A | 3.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1031 | 2018-06-20 | F18SM4766A | 1.4 | 0.0018 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1032 | 2018-06-21 | F18SM7140A | 1.1 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle) Seal Pack Evaluation
1033 | 2018-06-20 | F18SM4605A | 3.3 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1034 | 2018-06-20 | F18SM4767A | 5.2 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1035 | 2018-06-21 | F18SM4768A | 1.2 | 0.0016 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1036 | 2018-06-21 | F18SM4769A | 5.1 | 0.0017 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1037 | 2018-06-20 | F18SM4770A | 4.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1038 | 2018-06-21 | F18SM4772A | 2.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1039 | 2018-06-20 | F18SM4773A | 2.4 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1040 | 2018-06-21 | F18SM4774A | 2.1 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1041 | 2018-06-21 | F18SM4775A | 1.3 | 0.0013 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1042 | 2018-06-21 | F18SM4776A | 4.4 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1043 | 2018-06-21 | F18SM4728A | 5.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1044 | 2018-06-21 | F18SM4763A | 3.1 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1045 | 2018-06-21 | F18SM4771A | 4.2 | 0.0007 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1046 | 2018-06-21 | F18SM4778A | 5.3 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1047 | 2018-06-21 | F18SM4779A | 4.3 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1048 | 2018-06-21 | F18SM4780A | 5.2 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1049 | 2018-06-21 | F18SM4782A | 3.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1050 | 2018-06-21 | F18SM4783A | 3.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1051 | 2018-06-21 | F18SM4784A | 3.2 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1052 | 2018-06-21 | F18SM4785A | 2.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1053 | 2018-06-21 | F18SM4786A | 2.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1054 | 2018-06-22 | F18SM4787A | 2.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1055 | 2018-06-21 | F18SM4788A | 1.4 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1056 | 2018-06-21 | F18SM4789A | 1.3 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1057 | 2018-06-21 | F18SM4790A | 1.2 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1058 | 2018-06-20 | F18SM4758A | 4.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1059 | 2018-06-21 | F18SM4760A | 2.2 | 0.0013 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1060 | 2018-06-22 | F18SM4632A | 6.1 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1061 | 2018-06-21 | F18SM4791A | 5.1 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1062 | 2018-06-21 | F18SM7125A | 6.2 | 0.0009 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1063 | 2018-06-22 | F18SM7143A | 4.1 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1064 | 2018-06-21 | F18SM7144A | 3.2 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1065 | 2018-06-21 | F18SM7145A | 5.3 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1066 | 2018-06-22 | F18SM7146A | 3.3 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (7th bottle)
1067 | 2018-06-22 | F18SM7147A | 4.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1068 | 2018-06-21 | F18SM7149A | 3.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1069 | 2018-06-22 | F18SM7150A | 1.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1070 | 2018-06-22 | F18SM7151A | 1.2 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1071 | 2018-06-22 | F18SM7152A | 1.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1072 | 2018-06-22 | F18SM7153A | 2.1 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1073 | 2018-06-22 | F18SM7154A | 2.2 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1074 | 2018-06-22 | F18SM7155A | 2.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1075 | 2018-06-22 | F18SM7156A | 3.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1076 | 2018-06-22 | F18SM7157A | 2.3 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1077 | 2018-06-22 | F18SM7158A | 3.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1078 | 2018-06-22 | F18SM7159A | 3.4 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1079 | 2018-06-22 | F18SM7160A | 4.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1080 | 2018-06-22 | F18SM7161A | 4.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1081 | 2018-06-22 | F18SM7162A | 4.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1082 | 2018-06-22 | F18SM7166A | 5.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1083 | 2018-06-22 | F18SM7167A | 1.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1084 | 2018-06-22 | F18SM7168A | 1.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1085 | 2018-06-22 | F18SM7169A | 1.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle)
1086 | 2018-06-25 | F18SM7186A | 6.2 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 46 | 75
1087 | 2018-06-26 | F18SM4798A | 4.2 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 76
1088 | 2018-06-25 | F18SM4803A | 6.1 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1089 | 2018-06-26 | F18SM4805A | 4.2 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 77
1090 | 2018-06-26 | F18SM4811A | 1.2 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 46 | 72
1091 | 2018-06-26 | F18SM4812A | 1.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 70
1092 | 2018-06-26 | F18SM4813A | 2.3 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 48 | 75
1093 | 2018-06-26 | F18SM4816A | 5.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1094 | 2018-06-26 | F18SM4817A | 4.4 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 76
1095 | 2018-06-26 | F18SM4818A | 4.3 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 78
1096 | 2018-06-26 | F18SM4819A | 2.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1097 | 2018-06-26 | F18SM4821A | 1.4 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 74
1098 | 2018-06-26 | F18SM4824A | 1.3 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 72
1099 | 2018-06-26 | F18SM4827A | 3.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1100 | 2018-06-21 | F18SM4781A | 5.4 | 0.0018 | 0.002 | Pass on first test | WPPL 20:80 Mix (7th bottle) | 48 | 75
1101 | 2018-06-25 | F18SM4792A | 1.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1102 | 2018-06-25 | F18SM4793A | 1.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1103 | 2018-06-25 | F18SM4794A | 1.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 70
1104 | 2018-06-25 | F18SM4795A | 3.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1105 | 2018-06-25 | F18SM4796A | 2.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1106 | 2018-06-25 | F18SM4797A | 2.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1107 | 2018-06-25 | F18SM4800A | 2.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1108 | 2018-06-25 | F18SM4801A | 3.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1109 | 2018-06-25 | F18SM4802A | 3.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1110 | 2018-06-25 | F18SM4804A | 4.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 70
1111 | 2018-06-25 | F18SM4806A | 2.1 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 43 | 75
1112 | 2018-06-25 | F18SM4807A | 4.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1113 | 2018-06-25 | F18SM4808A | 1.4 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1114 | 2018-06-25 | F18SM4809A | 4.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1115 | 2018-06-25 | F18SM4810A | 1.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 78
1116 | 2018-06-27 | F18SM4814A | 5.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1117 | 2018-06-27 | F18SM4820A | 2.2 | 0.0007 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 48 | 75
1118 | 2018-06-26 | F18SM4822A | 1.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1119 | 2018-06-26 | F18SM4823A | 4.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 76
1120 | 2018-06-26 | F18SM4825A | 1.2 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 39 | 72
1121 | 2018-06-26 | F18SM4828A | 3.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1122 | 2018-06-26 | F18SM4829A | 4.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 77
1123 | 2018-06-26 | F18SM4830A | 3.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 74
1124 | 2018-06-26 | F18SM4831A | 3.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1125 | 2018-06-26 | F18SM4832A | 1.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1126 | 2018-06-26 | F18SM4833A | 4.2 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 43 | 76
1127 | 2018-06-26 | F18SM4834A | 2.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1128 | 2018-06-26 | F18SM4835A | 1.3 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 43 | 75
1129 | 2018-06-26 | F18SM4837A | 2.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1130 | 2018-06-26 | F18SM4838A | 5.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1131 | 2018-06-26 | F18SM4839A | 3.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 76
1132 | 2018-06-26 | F18SM4842A | 1.2 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1133 | 2018-06-26 | F18SM7171A | 6.3 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 42 | 77
1134 | 2018-06-27 | F18SM7193A | 1.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 76
1135 | 2018-06-27 | F18SM7195A | 2.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 77
1136 | 2018-06-27 | F18SM7196A | 4.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 74
1137 | 2018-06-27 | F18SM7197A | 5.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1138 | 2018-06-27 | F18SM7199A | 5.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1139 | 2018-06-27 | F18SM7200A | 3.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1140 | 2018-06-27 | F18SM7203A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 70
1141 | 2018-06-27 | F18SM7205A | 3.4 | 0 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1142 | 2018-06-27 | F18SM7208A | 4.1 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1143 | 2018-06-27 | F18SM4799A | 6.2 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 48 | 75
1144 | 2018-06-27 | F18SM4815A | 5.1 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 43 | 78
1145 | 2018-06-27 | F18SM4826A | 2.4 | NA | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 46 | 73
1146 | 2018-06-27 | F18SM4836A | 4.3 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 46 | 73
1147 | 2018-06-27 | F18SM4841A | 6.1 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1148 | 2018-06-27 | F18SM4843A | 3.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 78
1149 | 2018-06-27 | F18SM4844A | 3.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1150 | 2018-06-27 | F18SM4845A | 3.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1151 | 2018-06-27 | F18SM4846A | 4.1 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 76
1152 | 2018-06-27 | F18SM4847A | 4.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1153 | 2018-06-27 | F18SM4848A | 3.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1154 | 2018-06-27 | F18SM4849A | 1.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1155 | 2018-06-27 | F18SM4850A | 1.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 76
1156 | 2018-06-27 | F18SM4851A | 1.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1157 | 2018-06-06 | F18SM4615A | 2.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (4th bottle) | 50 | 76
1158 | 2018-06-28 | F18SM7214A | 5.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1159 | 2018-06-28 | F18SM7215A | 5.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 72
1160 | 2018-06-28 | F18SM7229A | 5.1 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1161 | 2018-06-28 | F18SM7218A | 2.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 72
1162 | 2018-06-28 | F18SM7206A | 2.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1163 | 2018-06-28 | F18SM7220A | 2.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1164 | 2018-06-28 | F18SM7219A | 2.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1165 | 2018-06-28 | F18SM7222A | 2.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1166 | 2018-06-27 | F18SM7194A | 2.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1167 | 2018-06-28 | F18SM7198A | 3.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1168 | 2018-06-28 | F18SM7204A | 3.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1169 | 2018-06-28 | F18SM7207A | 2.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1170 | 2018-06-28 | F18SM7209A | 3.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1171 | 2018-06-28 | F18SM7210A | 4.1 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1172 | 2018-06-28 | F18SM7211A | 4.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 76
1173 | 2018-06-28 | F18SM7212A | 4.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1174 | 2018-06-28 | F18SM7213A | 4.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1175 | 2018-06-28 | F18SM7216A | 4.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1176 | 2018-06-28 | F18SM7217A | 1.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1177 | 2018-06-28 | F18SM7221A | 1.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1178 | 2018-06-28 | F18SM7223A | 1.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1179 | 2018-06-28 | F18SM7224A | 1.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1180 | 2018-06-28 | F18SM7225A | 3.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1181 | 2018-06-28 | F18SM7227A | 4.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1182 | 2018-06-28 | F18SM7228A | 4.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1183 | 2018-06-28 | F18SM7230A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1184 | 2018-06-27 | F18SM4840A | 2.2 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1185 | 2018-06-28 | F18SM7201A | 3.3 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 72
1186 | 2018-06-28 | F18SM7202A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1187 | 2018-06-28 | F18SM7226A | 3.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1188 | 2018-06-29 | F18SM7231A | 5.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 72
1189 | 2018-06-29 | F18SM7232A | 3.3 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 48 | 72
1190 | 2018-06-30 | F18SM7233A | 2.3 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 78
1191 | 2018-06-30 | F18SM7235A | 5.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1192 | 2018-06-30 | F18SM7237A | 2.4 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 76
1193 | 2018-06-29 | F18SM7238A | 5.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1194 | 2018-06-29 | F18SM7240A | 4.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 77
1195 | 2018-06-29 | F18SM7241A | 4.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 77
1196 | 2018-06-29 | F18SM7242A | 4.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 76
1197 | 2018-06-29 | F18SM7243A | 4.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 76
1198 | 2018-06-29 | F18SM7244A | 3.4 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1199 | 2018-06-29 | F18SM7245A | 3.3 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 74
1200 | 2018-06-30 | F18SM7246A | 3.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1201 | 2018-06-30 | F18SM4852A | 3.3 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 73
1202 | 2018-06-30 | F18SM4853A | 3.4 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1203 | 2018-06-30 | F18SM4854A | 4.2 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 72
1204 | 2018-06-30 | F18SM4856A | 4.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1205 | 2018-06-30 | F18SM4857A | 3.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1206 | 2018-06-30 | F18SM4858A | 5.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1207 | 2018-07-02 | F18SM4859A | 2.4 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 46 | 73
1208 | 2018-07-02 | F18SM4860A | 5.2 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1209 | 2018-07-02 | F18SM4861A | 3.1 | 0.0001 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 43 | 73
1210 | 2018-07-02 | F18SM4862A | 1.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1211 | 2018-06-30 | F18SM4863A | 3.4 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 75
1212 | 2018-07-02 | F18SM4864A | 2.3 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1213 | 2018-07-02 | F18SM4865A | 4.3 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1214 | 2018-07-02 | F18SM4867A | 4.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 75
1215 | 2018-07-02 | F18SM4868A | 3.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 76
1216 | 2018-07-02 | F18SM4875A | 5.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 77
1217 | 2018-07-02 | F18SM4876A | 5.4 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1218 | 2018-07-02 | F18SM4877A | 6.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 76
1219 | 2018-07-02 | F18SM4878A | 6.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1220 | 2018-07-03 | G18SM4882A | 3.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1221 | 2018-07-02 | F18SM7165A | 3.4 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 42 | 74
1222 | 2018-07-02 | F18SM7234A | 3.2 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1223 | 2018-07-02 | F18SM7236A | 3.2 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 78
1224 | 2018-06-30 | F18SM7247A | 4.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1225 | 2018-07-03 | F18SM7248A | 2.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1226 | 2018-07-03 | F18SM4870A | 1.3 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1227 | 2018-07-03 | F18SM4871A | 5.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1228 | 2018-07-03 | F18SM4872A | 1.2 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1229 | 2018-07-03 | F18SM4873A | 1.1 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1230 | 2018-07-02 | F18SM4874A | 2.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1231 | 2018-07-02 | F18SM4879A | 6.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1232 | 2018-07-03 | F18SM4880A | 2.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 76
1233 | 2018-07-03 | F18SM4881A | 1.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1234 | 2018-07-03 | G18SM4883A | 3.2 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 75
1235 | 2018-07-03 | G18SM4887A | 4.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 74
1236 | 2018-07-03 | G18SM4890A | 5.2 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1237 | 2018-07-03 | G18SM4891A | 3.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1238 | 2018-07-03 | G18SM4893A | 5.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 77
1239 | 2018-07-03 | G18SM4884A | 2.3 | 0.0012 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 48 | 72
1240 | 2018-07-03 | G18SM4885A | 3.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 75
1241 | 2018-07-03 | G18SM4892A | 3.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 50 | 75
1242 | 2018-07-03 | G18SM4894A | 2.4 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 49 | 76
1243 | 2018-07-03 | G18SM4898A | 3.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 74
1244 | 2018-07-03 | G18SM4899A | 3.2 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1245 | 2018-07-03 | G18SM4900A | 1.4 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 46 | 75
1246 | 2018-07-03 | G18SM4901A | 4.2 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1247 | 2018-07-03 | G18SM4903A | 5.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 42 | 76
1248 | 2018-07-03 | G18SM4907A | 2.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 72
1249 | 2018-07-03 | G18SM4908A | 3.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 49 | 76
1250 | 2018-07-03 | G18SM4909A | 5.3 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (8th bottle) | 42 | 76
1251 | 2018-07-03 | G18SM4910A | 5.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 76
1252 | 2018-07-03 | G18SM4911A | 4.2 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 46 | 73
1253 | 2018-06-29 | F18SM7239A | 5.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 77
1254 | 2018-07-04 | F18SM4866A | 4.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 78
1255 | 2018-07-03 | G18SM4886A | 3.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 48 | 75
1256 | 2018-07-03 | G18SM4889A | 4.3 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (8th bottle) | 43 | 73
1257 | 2018-07-04 | G18SM4896A | 1.3 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1258 | 2018-07-04 | G18SM4897A | 5.4 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 48 | 75
1259 | 2018-07-04 | G18SM4902A | 2.2 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 49 | 76
1260 | 2018-07-04 | G18SM4904A | 1.2 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 48 | 75
1261 | 2018-07-04 | G18SM4905A | 5.2 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 43 | 75
1262 | 2018-07-04 | G18SM4912A | 3.1 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 75
1263 | 2018-07-04 | G18SM4913A | 2.3 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 72
1264 | 2018-07-04 | G18SM4914A | 2.4 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1265 | 2018-07-04 | G18SM4915A | 2.1 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 75
1266 | 2018-07-04 | G18SM4916A | 1.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 72
1267 | 2018-07-04 | G18SM4917A | 3.2 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1268 | 2018-07-04 | G18SM4918A | 3.3 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1269 | 2018-07-04 | G18SM4919A | 3.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1270 | 2018-07-04 | G18SM4921A | 4.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1271 | 2018-07-05 | G18SM4922A | 5.1 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 73
1272 | 2018-07-04 | G18SM4923A | 2.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 74
1273 | 2018-07-04 | G18SM4926A | 2.4 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1274 | 2018-07-04 | G18SM4927A | 2.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1275 | 2018-07-05 | G18SM4928A | 1.3 | 0.0015 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 73
1276 | 2018-07-05 | F18SM4855A | 4.4 | 0.0001 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 72
1277 | 2018-07-05 | G18SM4888A | 3.4 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1278 | 2018-07-05 | G18SM4895A | 3.2 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 72
1279 | 2018-07-05 | G18SM4920A | 4.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 76
1280 | 2018-07-05 | G18SM4924A | 4.3 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 74
1281 | 2018-07-05 | G18SM4929A | 2.3 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1282 | 2018-07-04 | G18SM4930A | 1.2 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1283 | 2018-07-05 | G18SM4931A | 3.3 | 0.0009 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 75
1284 | 2018-07-05 | G18SM4932A | 5.2 | 0.0013 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 48 | 72
1285 | 2018-07-05 | G18SM4933A | 5.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 72
1286 | 2018-07-05 | G18SM4934A | 6.1 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 76
1287 | 2018-07-05 | G18SM4935A | 5.3 | 0.0011 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 39 | 76
1288 | 2018-07-05 | G18SM4936A | 6.2 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1289 | 2018-07-05 | G18SM4937A | 6.3 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 76
1290 | 2018-07-05 | G18SM4938A | 5.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 78
1291 | 2018-07-05 | G18SM4939A | 2.2 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1292 | 2018-07-05 | G18SM4944A | 1.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 79
1293 | 2018-07-07 | G18SM4960A | 2.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1294 | 2018-07-07 | G18SM4963A | 3.4 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1295 | 2018-07-07 | G18SM4964A | 1.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1296 | 2018-07-07 | G18SM4966A | 1.3 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 74
1297 | 2018-07-07 | G18SM4967A | 4.2 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 43 | 75
1298 | 2018-07-07 | G18SM4970A | 4.1 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 43 | 73
1299 | 2018-07-07 | G18SM4972A | 3.2 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 75
1300 | 2018-07-07 | G18SM4976A | 1.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 74
1301 | 2018-07-07 | G18SM4977A | 5.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1302 | 2018-07-07 | G18SM4978A | 5.2 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1303 | 2018-07-04 | G18SM4925A | 3.1 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1304 | 2018-07-05 | G18SM4940A | 2.1 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1305 | 2018-07-06 | G18SM4941A | 1.4 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 44 | 72
1306 | 2018-07-06 | G18SM4942A | 1.3 | 0.0008 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 75
1307 | 2018-07-05 | G18SM4945A | 2.4 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1308 | 2018-07-05 | G18SM4946A | 2.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1309 | 2018-07-06 | G18SM4947A | 3.3 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 74
1310 | 2018-07-05 | G18SM4948A | 3.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 76
1311 | 2018-07-05 | G18SM4949A | 3.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1312 | 2018-07-06 | G18SM4950A | 3.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 75
1313 | 2018-07-06 | G18SM4951A | 4.1 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 42 | 76
1314 | 2018-07-06 | G18SM4952A | 4.2 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 39 | 77
1315 | 2018-07-06 | G18SM4954A | 4.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1316 | 2018-07-06 | G18SM4955A | 5.1 | 0.0018 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 48 | 75
1317 | 2018-07-06 | G18SM4956A | 5.2 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 73
1318 | 2018-07-06 | G18SM4957A | 2.2 | 0.0013 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 48 | 70
1319 | 2018-07-06 | G18SM4958A | 2.1 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1320 | 2018-07-06 | G18SM4959A | 2.3 | 0.0017 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 72
1321 | 2018-07-06 | G18SM4961A | 3.1 | 0.0015 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1322 | 2018-07-06 | G18SM4962A | 3.2 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 73
1323 | 2018-07-09 | G18SM4943A | 1.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1324 | 2018-07-07 | G18SM4953A | 4.3 | 0.0017 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 48 | 70
1325 | 2018-07-07 | G18SM4965A | 2.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 49 | 74
1326 | 2018-07-07 | G18SM4971A | 3.1 | 0.0014 | 0.002 | Resolve unit | WPPL 20:80 Mix (9th bottle) | 46 | 70
1327 | 2018-07-09 | G18SM4973A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 72
1328 | 2018-07-09 | G18SM4979A | 3.2 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1329 | 2018-07-07 | G18SM4980A | 3.3 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 72
1330 | 2018-07-09 | G18SM4984A | 6.2 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 42 | 74
1331 | 2018-07-09 | G18SM4985A | 5.4 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 76
1332 | 2018-07-07 | G18SM4986A | 2.3 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 48 | 77
1333 | 2018-07-09 | G18SM4987A | 2.4 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 76
1334 | 2018-07-07 | G18SM4988A | 1.1 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 46 | 75
1335 | 2018-07-09 | G18SM4990A | 3.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 74
1336 | 2018-07-09 | G18SM4991A | 2.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1337 | 2018-07-09 | G18SM4992A | 3.3 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 77
1338 | 2018-07-09 | G18SM4993A | 1.4 | 0.0007 | 0.002 | Resolve unit | WPPL 20:80 Mix (10th bottle) | 48 | 75
1339 | 2018-07-09 | G18SM4994A | 1.3 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1340 | 2018-07-09 | G18SM4995A | 5.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1341 | 2018-07-09 | G18SM4996A | 5.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 42 | 76
1342 | 2018-07-09 | G18SM4997A | 3.2 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 72
1343 | 2018-07-09 | G18SM4998A | 4.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 76
1344 | 2018-07-09 | G18SM4002A | 5.1 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1345 | 2018-07-10 | G18SM7249A | 2.4 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 76
1346 | 2018-07-10 | G18SM7251A | 3.4 | 0.0011 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1347 | 2018-07-10 | G18SM7252A | 4.4 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 76
1348 | 2018-07-10 | G18SM7254A | 1.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 74
1349 | 2018-07-10 | G18SM7258A | 2.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1350 | 2018-07-11 | G18SM4968A | 4.3 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 73
1351 | 2018-07-10 | G18SM4983A | 4.2 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1352 | 2018-07-10 | G18SM4999A | 1.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 50 | 76
1353 | 2018-07-10 | G18SM4004A | 1.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1354 | 2018-07-10 | G18SM4009A | 1.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 77
1355 | 2018-07-10 | F18SM4869A | 5.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 50 | 73
1356 | 2018-07-09 | G18SM4906A | 5.3 | 0.0013 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 74
1357 | 2018-07-10 | G18SM4974A | 2.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 72
1358 | 2018-07-10 | G18SM4975A | 4.4 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1359 | 2018-07-09 | G18SM4989A | 3.4 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1360 | 2018-07-10 | G18SM4000A | 3.4 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 76
1361 | 2018-07-09 | G18SM4003A | 5.2 | 0.0002 | 0.002 | Resolve unit | WPPL 20:80 Mix (10th bottle) | 42 | 76
1362 | 2018-07-09 | G18SM4005A | 1.2 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1363 | 2018-07-10 | G18SM4006A | 4.3 | 0.0019 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 42 | 74
1364 | 2018-07-10 | G18SM4011A | 3.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 77
1365 | 2018-07-10 | G18SM4012A | 3.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 70
1366 | 2018-07-10 | G18SM4013A | 4.1 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1367 | 2018-07-10 | G18SM4982A | 6.1 | 0.0004 | 0.002 | Resolve unit | WPPL 20:80 Mix (10th bottle) | 43 | 73
1368 | 2018-07-10 | G18SM4001A | 3.3 | 0.001 | 0.002 | Resolve unit | WPPL 20:80 Mix (10th bottle) | 48 | 78
1369 | 2018-07-10 | G18SM4007A | 2.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 72
1370 | 2018-07-11 | G18SM4008A | 5.4 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 72
1371 | 2018-07-10 | G18SM4010A | 2.1 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1372 | 2018-07-11 | G18SM7250A | 2.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1373 | 2018-07-11 | G18SM7253A | 5.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 77
1374 | 2018-07-11 | G18SM7256A | 2.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 42 | 79
1375 | 2018-07-11 | G18SM7260A | 2.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 72
1376 | 2018-07-11 | G18SM7261A | 4.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1377 | 2018-07-11 | G18SM7262A | 5.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 76
1378 | 2018-07-11 | G18SM7263A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1379 | 2018-07-11 | G18SM7264A | 4.1 | 0.0014 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 39 | 77
1380 | 2018-07-11 | G18SM7265A | 4.2 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1381 | 2018-07-11 | G18SM7266A | 3.4 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 77
1382 | 2018-07-11 | G18SM7267A | 1.4 | 0.0006 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 76
1383 | 2018-07-12 | G18SM7268A | 2.1 | 0.0009 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 77
1384 | 2018-07-11 | G18SM7270A | 1.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1385 | 2018-07-11 | G18SM7271A | 4.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 42 | 76
1386 | 2018-07-11 | G18SM7276A | 6.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 42 | 76
1386 | 2018-07-11 | G18SM7277A | 2.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 50 | 75
1387 | 2018-07-11 | G18SM7279A | 3.4 | 0.0003 | 0.002 | Resolve unit | WPPL 20:80 Mix (10th bottle) | 48 | 77
1388 | 2018-07-11 | G18SM7282A | 4.4 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 73
1389 | 2018-07-11 | G18SM7283A | 5.1 | 0.0012 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 39 | 74
1390 | 2018-07-12 | G18SM7259A | 3.2 | 0.0005 | 0.002 | Resolve unit | WPPL 20:80 Mix (10th bottle) | 48 | 75
1391 | 2018-07-11 | G18SM7273A | 6.2 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1392 | 2018-07-12 | G18SM7274A | 2.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1393 | 2018-07-11 | G18SM7275A | 5.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 76
1394 | 2018-07-12 | G18SM7278A | 3.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1395 | 2018-07-12 | G18SM7280A | 4.1 | 0.0006 | 0.002 | Resolve unit | WPPL 20:80 Mix (10th bottle) | 48 | 75
1396 | 2018-07-12 | G18SM7284A | 1.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1397 | 2018-07-12 | G18SM7286A | 4.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1398 | 2018-07-12 | G18SM7287A | 4.3 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 77
1399 | 2018-07-12 | G18SM7288A | 1.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1400 | 2018-07-12 | G18SM7289A | 1.4 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1401 | 2018-07-07 | G18SM4969A | 2.3 | 0.0007 | 0.002 | Pass on first test | WPPL 20:80 Mix (9th bottle) | 39 | 74
1402 | 2018-07-14 | G18SM4015A | 4.4 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
1403 | 2018-07-14 | G18SM4034A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1404 | 2018-07-13 | G18SM4036A | 3.4 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 76
1405 | 2018-07-14 | G18SM4037A | 1.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1406 | 2018-07-14 | G18SM4038A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1407 | 2018-07-13 | G18SM4039A | 4.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 42 | 72
1408 | 2018-07-14 | G18SM4041A | 5.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1409 | 2018-07-14 | G18SM4042A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1410 | 2018-07-13 | G18SM4044A | 5.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 74
1411 | 2018-07-14 | G18SM4046A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1412 | 2018-07-14 | G18SM4049A | 3.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1413 | 2018-07-14 | G18SM4050A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1414 | 2018-07-14 | G18SM4052A | 3.4 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1415 | 2018-07-14 | G18SM4054A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1416 | 2018-07-12 | G18SM7257A | 5.2 | 0.0016 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1417 | 2018-07-14 | G18SM7269A | 5.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1418 | 2018-07-14 | G18SM7281A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1419 | 2018-07-12 | G18SM7285A | 2.2 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 49 | 74
1420 | 2018-07-13 | G18SM4014A | 2.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 72
1421 | 2018-07-13 | G18SM4016A | 3.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 72
1422 | 2018-07-13 | G18SM4017A | 5.1 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 76
1423 | 2018-07-12 | G18SM4018A | 2.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1424 | 2018-07-12 | G18SM4019A | 2.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1425 | 2018-07-13 | G18SM4020A | 3.1 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1426 | 2018-07-12 | G18SM4021A | 3.2 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1427 | 2018-07-13 | G18SM4022A | 2.4 | 0.0002 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1428 | 2018-07-12 | G18SM4023A | 1.2 | 0.001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1429 | 2018-07-12 | G18SM4024A | 1.3 | 0.0004 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1430 | 2018-07-13 | G18SM4026A | 4.1 | NA | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1431 | 2018-07-13 | G18SM4027A | 4.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 73
1432 | 2018-07-13 | G18SM4028A | 1.4 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 75
1433 | 2018-07-13 | G18SM4029A | 1.3 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 43 | 76
1434 | 2018-07-13 | G18SM4030A | 1.2 | 0.0008 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1435 | 2018-07-13 | G18SM4031A | 2.1 | 0.0001 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 72
1436 | 2018-07-13 | G18SM4032A | 2.2 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 48 | 75
1437 | 2018-07-13 | G18SM4033A | 2.3 | 0.0005 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 75
1438 | 2018-07-14 | G18SM4035A | 3.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1439 | 2018-07-13 | G18SM4043A | 4.3 | 0.0003 | 0.002 | Pass on first test | WPPL 20:80 Mix (10th bottle) | 46 | 78
1440 | 2018-07-14 | G18SM4025A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
1441 | 2018-07-16 | G18SM4040A | 5.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1442 | 2018-07-14 | G18SM4045A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1443 | 2018-07-16 | G18SM4047A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1444 | 2018-07-14 | G18SM4048A | 4.3 | 0.0004 | 0.002 | Resolve unit | FLEX 30:70 Mix | 42 | 76
1445 | 2018-07-16 | G18SM4051A | 2.3 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1446 | 2018-07-14 | G18SM4053A | 4.4 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1447 | 2018-07-17 | G18SM4056A | 5.2 | 0.0001 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 78
1448 | 2018-07-16 | G18SM4057A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1449 | 2018-07-16 | G18SM4058A | 5.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1450 | 2018-07-16 | G18SM4059A | 5.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
1451 | 2018-07-16 | G18SM4060A | 2.2 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 69
1452 | 2018-07-16 | G18SM4061A | 2.4 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
1453 | 2018-07-16 | G18SM4064A | 1.4 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1454 | 2018-07-16 | G18SM4065A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
1455 | 2018-07-16 | G18SM4066A | 3.2 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 69
1456 | 2018-07-16 | G18SM4067A | 3.4 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1457 | 2018-07-17 | G18SM4068A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1458 | 2018-07-16 | G18SM4071A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
1459 | 2018-07-16 | G18SM4072A | 5.1 | 0.0006 | 0.002 | Resolve unit | FLEX 30:70 Mix | 39 | 72
1460 | 2018-07-18 | G18SM7255A | 3.3 | 0.0012 | 0.002 | Resolve unit | FLEX 30:70 Mix | 43 | 76
1461 | 2018-07-17 | G18SM4062A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1462 | 2018-07-17 | G18SM4063A | 6.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1463 | 2018-07-17 | G18SM4069A | 4.4 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
1464 | 2018-07-17 | G18SM4070A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1465 | 2018-07-16 | G18SM4073A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1466 | 2018-07-17 | G18SM4074A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1467 | 2018-07-17 | G18SM4075A | 2.3 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1468 | 2018-07-17 | G18SM4076A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
1469 | 2018-07-17 | G18SM4077A | 1.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1470 | 2018-07-17 | G18SM4078A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1471 | 2018-07-17 | G18SM4080A | 3.4 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1472 | 2018-07-17 | G18SM4081A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1473 | 2018-07-17 | G18SM4083A | 4.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1474 | 2018-07-17 | G18SM4084A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1475 | 2018-07-17 | G18SM4085A | 5.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1476 | 2018-07-17 | G18SM4086A | 5.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 67
1477 | 2018-07-17 | G18SM4087A | 5.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
1478 | 2018-07-17 | G18SM4088A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1479 | 2018-07-17 | G18SM4089A | 1.3 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1480 | 2018-07-17 | G18SM4090A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1481 | 2018-07-18 | G18SM4095A | 4.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1482 | 2018-07-18 | G18SM7272A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1483 | 2018-07-17 | G18SM4079A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
1484 | 2018-07-18 | G18SM4082A | 4.1 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 72
1485 | 2018-07-18 | G18SM4091A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1486 | 2018-07-18 | G18SM4092A | 6.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1487 | 2018-07-18 | G18SM4093A | 4.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1488 | 2018-07-18 | G18SM4094A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1489 | 2018-07-18 | G18SM4096A | 1.1 | 0.0003 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 75
1490 | 2018-07-18 | G18SM4097A | 1.4 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1491 | 2018-07-18 | G18SM4098A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1492 | 2018-07-18 | G18SM4099A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1493 | 2018-07-18 | G18SM4100A | 1.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1494 | 2018-07-18 | G18SM4101A | 3.4 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1495 | 2018-07-18 | G18SM4102A | 4.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1496 | 2018-07-18 | G18SM4103A | 4.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1497 | 2018-07-18 | G18SM4104A | 3.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1498 | 2018-07-18 | G18SM4105A | 5.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1499 | 2018-07-18 | G18SM4109A | 2.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1500 | 2018-07-18 | G18SM4110A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1501 | 2018-07-18 | G18SM4111A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1502 | 2018-07-18 | G18SM4113A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1503 | 2018-07-18 | G18SM4114A | 1.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
1504 | 2018-07-18 | G18SM4106A | 5.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1505 | 2018-07-19 | G18SM4107A | 5.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1506 | 2018-07-18 | G18SM4108A | 5.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1507 | 2018-07-18 | G18SM4112A | 3.2 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1508 | 2018-07-18 | G18SM4115A | 3.3 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1509 | 2018-07-19 | G18SM4116A | 3.4 | 0.0005 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 74
1510 | 2018-07-19 | G18SM4117A | 1.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1511 | 2018-07-19 | G18SM4118A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1512 | 2018-07-19 | G18SM4119A | 1.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1513 | 2018-07-19 | G18SM4120A | 4.1 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1514 | 2018-07-19 | G18SM4121A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1515 | 2018-07-19 | G18SM4122A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1516 | 2018-07-19 | G18SM4125A | 1.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1517 | 2018-07-19 | G18SM4127A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1518 | 2018-07-19 | G18SM4135A | 1.2 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1519 | 2018-07-19 | G18SM4137A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1520 | 2018-07-20 | G18SM4140A | 5.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1521 | 2018-07-20 | G18SM4143A | 4.1 | 0.0003 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 75
1522 | 2018-07-20 | G18SM4146A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1523 | 2018-07-20 | G18SM4148A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1524 | 2018-07-20 | G18SM4150A | 1.4 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1525 | 2018-07-20 | G18SM4153A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1526 | 2018-07-20 | G18SM4155A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1527 | 2018-07-20 | G18SM4055A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1528 | 2018-07-19 | G18SM4123A | 4.3 | 0.0001 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
1529 | 2018-07-19 | G18SM4126A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1530 | 2018-07-20 | G18SM4128A | 3.1 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1531 | 2018-07-19 | G18SM4129A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1532 | 2018-07-20 | G18SM4130A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1533 | 2018-07-19 | G18SM4132A | 5.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1534 | 2018-07-19 | G18SM4133A | 5.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1535 | 2018-07-20 | G18SM4134A | 1.3 | 0.0008 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 75
1536 | 2018-07-20 | G18SM4136A | 1.4 | 0.0009 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 74
1537 | 2018-07-20 | G18SM4138A | 3.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1538 | 2018-07-20 | G18SM4139A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1539 | 2018-07-20 | G18SM4141A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1540 | 2018-07-20 | G18SM4142A | 5.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1541 | 2018-07-20 | G18SM4144A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1542 | 2018-07-20 | G18SM4147A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
1543 | 2018-07-23 | G18SM4145A | 3.4 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 42 | 74
1544 | 2018-07-20 | G18SM4149A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1545 | 2018-07-20 | G18SM4151A | 1.3 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 67
1546 | 2018-07-23 | G18SM4152A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1547 | 2018-07-20 | G18SM4154A | 1.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1548 | 2018-07-20 | G18SM4156A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
1549 | 2018-07-23 | G18SM4157A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1550 | 2018-07-20 | G18SM4158A | 4.2 | 0.0008 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 72
1551 | 2018-07-23 | G18SM4159A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1552 | 2018-07-23 | G18SM4160A | 4.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1553 | 2018-07-23 | G18SM4161A | 3.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
1554 | 2018-07-23 | G18SM4163A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
1555 | 2018-07-23 | G18SM4164A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1556 | 2018-07-23 | G18SM4166A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1557 | 2018-07-23 | G18SM4167A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1558 | 2018-07-23 | G18SM4168A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1559 | 2018-07-23 | G18SM4170A | 1.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1560 | 2018-07-23 | G18SM4165A | 2.1 | 0.0019 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 69
1561 | 2018-07-23 | G18SM4169A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1562 | 2018-07-23 | G18SM4171A | 1.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
1563 | 2018-07-23 | G18SM4172A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
1564 | 2018-07-23 | G18SM4173A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 70
1565 | 2018-07-23 | G18SM4174A | 4.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1566 | 2018-07-23 | G18SM4175A | 5.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
1567 | 2018-07-23 | G18SM4176A | 5.3 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 70
1568 | 2018-07-23 | G18SM4177A | 5.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1569 | 2018-07-23 | G18SM4178A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1570 | 2018-07-23 | G18SM7290A | 2.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
1571 | 2018-07-24 | G18SM7291A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1572 | 2018-07-23 | G18SM7292A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1573 | 2018-07-23 | G18SM7294A | 2.1 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1574 | 2018-07-23 | G18SM7295A | 1.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1575 | 2018-07-23 | G18SM7296A | 1.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
1576 | 2018-07-24 | G18SM7298A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1577 | 2018-07-24 | G18SM7299A | 3.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1578 | 2018-07-24 | G18SM7300A | 4.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1579 | 2018-07-24 | G18SM7302A | 4.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1580 | 2018-07-24 | G18SM7306A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1581 | 2018-07-24 | G18SM7308A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1582 | 2018-07-24 | G18SM7309A | 3.3 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
1583 | 2018-07-24 | G18SM7310A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
1584 | 2018-07-24 | G18SM7311A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1585 | 2018-07-24 | G18SM7313A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
1586 | 2018-07-24 | G18SM7293A | 2.2 | 0.0014 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 72
1587 | 2018-07-24 | G18SM7297A | 1.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1588 | 2018-07-25 | G18SM7301A | 4.2 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 70
1589 | 2018-07-24 | G18SM7304A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1590 | 2018-07-24 | G18SM7305A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1591 | 2018-07-24 | G18SM7312A | 4.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1592 | 2018-07-25 | G18SM7314A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1593 | 2018-07-25 | G18SM7315A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
1594 | 2018-07-25 | G18SM7316A | 3.1 | 0.0001 | 0.002 | Resolve unit | FLEX 30:70 Mix | 42 | 70
1595 | 2018-07-25 | G18SM7317A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1596 | 2018-07-25 | G18SM7318A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1597 | 2018-07-25 | G18SM7319A | 2.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1598 | 2018-07-25 | G18SM7320A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1599 | 2018-07-25 | G18SM7321A | 5.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1600 | 2018-07-25 | G18SM7322A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1601 | 2018-07-25 | G18SM7323A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1602 | 2018-07-25 | G18SM7324A | 4.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1603 | 2018-07-25 | G18SM7325A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1604 | 2018-07-25 | G18SM7326A | 1.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1605 | 2018-07-25 | G18SM7327A | 4.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1606 | 2018-07-25 | G18SM7328A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
1607 | 2018-07-25 | G18SM7330A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1608 | 2018-07-24 | G18SM4124A | 3.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1609 | 2018-07-20 | G18SM4131A | 4.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1610 | 2018-07-25 | G18SM4179A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1611 | 2018-07-25 | G18SM4180A | 5.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1612 | 2018-07-25 | G18SM4182A | 2.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1613 | 2018-07-25 | G18SM4184A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1614 | 2018-07-26 | G18SM4185A | 3.2 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1615 | 2018-07-25 | G18SM4186A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1616 | 2018-07-26 | G18SM4187A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
1617 | 2018-07-26 | G18SM4188A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1618 | 2018-07-26 | G18SM4189A | 4.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1619 | 2018-07-26 | G18SM4190A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1620 | 2018-07-26 | G18SM4191A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1621 | 2018-07-26 | G18SM4192A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1622 | 2018-07-26 | G18SM4193A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1623 | 2018-07-26 | G18SM4194A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1624 | 2018-07-26 | G18SM4195A | 1.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1625 | 2018-07-26 | G18SM4198A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1626 | 2018-07-26 | G18SM4199A | 2.21 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1627 | 2018-07-26 | G18SM4200A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1628 | 2018-07-26 | G18SM7303A | 1.3 | 0.0001 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 70
1629 | 2018-07-25 | G18SM7329A | 5.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
1630 | 2018-07-25 | G18SM7331A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1631 | 2018-07-25 | G18SM4181A | 2.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1632 | 2018-07-25 | G18SM4183A | 1.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1633 | 2018-07-26 | G18SM4196A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1634 | 2018-07-26 | G18SM4197A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 78
1635 | 2018-07-26 | G18SM4201A | 2.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1636 | 2018-07-26 | G18SM4202A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1637 | 2018-07-26 | G18SM4203A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1638 | 2018-07-26 | G18SM4204A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1639 | 2018-07-26 | G18SM4205A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1640 | 2018-07-27 | G18SM4207A | 5.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1641 | 2018-07-26 | G18SM4208A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1642 | 2018-07-26 | G18SM4209A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1643 | 2018-07-26 | G18SM4210A | 4.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1644 | 2018-07-27 | G18SM4211A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1645 | 2018-07-27 | G18SM4214A | 4.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1646 | 2018-07-27 | G18SM4215A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1647 | 2018-07-27 | G18SM4216A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1648 | 2018-07-27 | G18SM4206A | 1.4 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 78
1649 | 2018-07-27 | G18SM4213A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
1650 | 2018-07-27 | G18SM4217A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1651 | 2018-07-27 | G18SM4218A | 4.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1652 | 2018-07-27 | G18SM4220A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 79
1653 | 2018-07-27 | G18SM4227A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1654 | 2018-07-30 | G18SM4212A | 3.4 | 0.0006 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 77
1655 | 2018-07-30 | G18SM4219A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
1656 | 2018-07-30 | G18SM4222A | 5.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1657 | 2018-07-30 | G18SM4223A | 2.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1658 | 2018-07-30 | G18SM4225A | 1.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1659 | 2018-07-27 | G18SM4228A | 1.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1660 | 2018-07-30 | G18SM4229A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 80
1661 | 2018-07-30 | G18SM4230A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1662 | 2018-07-30 | G18SM4231A | 3.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1663 | 2018-07-30 | G18SM4232A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1664 | 2018-07-30 | G18SM4233A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1665 | 2018-07-30 | G18SM4234A | 4.3 | 0.0006 | 0.002 | Resolve unit | FLEX 30:70 Mix | 39 | 80
1666 | 2018-07-30 | G18SM4235A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1667 | 2018-07-30 | G18SM4236A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1668 | 2018-07-30 | G18SM4237A | 1.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1669 | 2018-07-30 | G18SM4238A | 5.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1670 | 2018-07-30 | G18SM4239A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1671 | 2018-07-30 | G18SM4240A | 5.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1672 | 2018-07-30 | G18SM4244A | 5.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1673 | 2018-07-30 | G18SM4224A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1674 | 2018-07-30 | G18SM4226A | 1.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1675 | 2018-07-30 | G18SM4241A | 6.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1676 | 2018-07-30 | G18SM4242A | 4.2 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1677 | 2018-07-30 | G18SM4243A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1678 | 2018-07-30 | G18SM4245A | 3.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1679 | 2018-07-31 | G18SM4246A | 4.4 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 75
1680 | 2018-07-30 | G18SM4247A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 79
1681 | 2018-07-31 | G18SM4248A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1682 | 2018-07-31 | G18SM4249A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1683 | 2018-07-31 | G18SM4250A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1684 | 2018-07-31 | G18SM4252A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 81
1685 | 2018-07-31 | G18SM4253A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1686 | 2018-07-31 | G18SM4254A | 3.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1687 | 2018-07-31 | G18SM4256A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1688 | 2018-07-31 | G18SM4258A | 5.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1689 | 2018-07-31 | G18SM4259A | 5.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 80
1690 | 2018-07-31 | G18SM4262A | 5.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 80
1691 | 2018-07-31 | G18SM4221A | 2.2 | 0.0001 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 80
1692 | 2018-07-31 | G18SM4255A | 4.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1693 | 2018-07-31 | G18SM4257A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1694 | 2018-07-31 | G18SM4260A | 2.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1695 | 2018-07-31 | G18SM4261A | 5.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
1696 | 2018-07-31 | G18SM4263A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
1697 | 2018-08-01 | G18SM4264A | 1.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1698 | 2018-08-01 | G18SM4265A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1699 | 2018-08-01 | G18SM4266A | 1.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1700 | 2018-07-31 | G18SM4267A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 78
1701 | 2018-07-31 | G18SM4268A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
1702 | 2018-08-01 | G18SM4269A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1703 | 2018-08-01 | G18SM4270A | 2.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1704 | 2018-08-01 | G18SM4271A | 3.3 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
1705 | 2018-08-01 | G18SM4272A | 4.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
1706 | 2018-08-01 | G18SM4275A | 3.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1707 | 2018-08-01 | G18SM4276A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1708 | 2018-08-01 | G18SM4279A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
1709 | 2018-08-01 | G18SM4283A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1710 | 2018-08-01 | G18SM4284A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1711 | 2018-08-01 | G18SM4286A | 1.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1712 | 2018-08-01 | G18SM4287A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1713 | 2018-08-01 | G18SM4251A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 79
1714 | 2018-08-01 | G18SM4273A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
1715 | 2018-08-02 | G18SM4274A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1716 | 2018-08-01 | G18SM4277A | 4.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1717 | 2018-08-02 | G18SM4278A | 2.3 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1718 | 2018-08-02 | G18SM4280A | 5.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1719 | 2018-08-02 | G18SM4281A | 5.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1720 | 2018-08-02 | G18SM4285A | 2.2 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 78
1721 | 2018-08-02 | G18SM4288A | 5.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1722 | 2018-08-02 | G18SM4289A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1723 | 2018-08-02 | G18SM4290A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1724 | 2018-08-02 | G18SM4291A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1725 | 2018-08-02 | G18SM4292A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1726 | 2018-08-02 | G18SM4294A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1727 | 2018-08-02 | G18SM4293A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1728 | 2018-08-02 | G18SM4295A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1729 | 2018-08-02 | G18SM4296A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1730 | 2018-08-02 | G18SM4297A | 2.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1731 | 2018-08-02 | G18SM4298A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1732 | 2018-08-02 | G18SM4299A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1733 | 2018-08-02 | G18SM4300A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1734 | 2018-08-02 | G18SM4302A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
1735 | 2018-08-02 | G18SM4303A | 1.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1736 | 2018-08-02 | G18SM4305A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1737 | 2018-08-03 | G18SM4282A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1738 | 2018-08-03 | G18SM4301A | 1.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1739 | 2018-08-03 | G18SM4304A | 3.3 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 77
1740 | 2018-08-03 | G18SM4306A | 4.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1741 | 2018-08-03 | G18SM4308A | 4.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1742 | 2018-08-06 | G18SM4310A | 3.2 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 76
1743 | 2018-08-03 | G18SM4311A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1744 | 2018-08-03 | G18SM4312A | 3.1 | 0.0006 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 77
1745 | 2018-08-06 | G18SM4315A | 2.2 | 0.0003 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 72
1746 | 2018-08-03 | G18SM4316A | 1.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1747 | 2018-08-03 | G18SM4317A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1748 | 2018-08-03 | G18SM4318A | 1.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1749 | 2018-08-04 | G18SM4319A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1750 | 2018-08-04 | G18SM4320A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1751 | 2018-08-04 | G18SM4321A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1752 | 2018-08-04 | G18SM4322A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1753 | 2018-08-06 | G18SM4307A | 3.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1754 | 2018-08-04 | H18SM4323A | 4.3 | 0.0004 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 72
1755 | 2018-08-04 | H18SM4324A | 5.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1756 | 2018-08-04 | H18SM4325A | 5.1 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 75
1757 | 2018-08-04 | H18SM4326A | 5.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1758 | 2018-08-06 | H18SM4327A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1759 | 2018-08-06 | H18SM4329A | 1.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1760 | 2018-08-06 | H18SM4331A | 1.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1761 | 2018-08-06 | H18SM4332A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1762 | 2018-08-06 | H18SM4333A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1763 | 2018-08-06 | H18SM4334A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1764 | 2018-08-06 | H18SM4335A | 3.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1765 | 2018-08-06 | H18SM4337A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1766 | 2018-08-06 | H18SM4338A | 4.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1767 | 2018-08-06 | H18SM4340A | 4.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
1768 | 2018-08-07 | H18SM4314A | 2.3 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 78
1769 | 2018-08-07 | H18SM4339A | 4.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1770 | 2018-08-07 | H18SM7332A | 5.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1771 | 2018-08-06 | H18SM7333A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1772 | 2018-08-07 | H18SM7334A | 2.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1773 | 2018-08-06 | H18SM7335A | 3.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1774 | 2018-08-07 | H18SM7336A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 80
1775 | 2018-08-07 | H18SM7338A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1776 | 2018-08-07 | H18SM7339A | 1.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1777 | 2018-08-07 | H18SM7340A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1778 | 2018-08-07 | H18SM7341A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1779 | 2018-08-07 | H18SM7342A | 3.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1780 | 2018-08-07 | H18SM7343A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1781 | 2018-08-07 | H18SM7344A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1782 | 2018-08-07 | H18SM7346A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1783 | 2018-08-07 | H18SM7347A | 4.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1784 | 2018-08-07 | H18SM7345A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1785 | 2018-08-07 | H18SM7348A | 5.1 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1786 | 2018-08-07 | H18SM7349A | 2.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1787 | 2018-08-07 | H18SM7373A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1788 | 2018-08-07 | H18SM7374A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1789 | 2018-08-07 | H18SM7375A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1790 | 2018-08-08 | H18SM4330A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1791 | 2018-08-08 | H18SM4342A | 1.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1792 | 2018-08-07 | H18SM7351A | 2.4 | 0.0019 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1793 | 2018-08-07 | H18SM7352A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1794 | 2018-08-07 | H18SM7353A | 1.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1795 | 2018-08-07 | H18SM7355A | 4.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1796 | 2018-08-08 | H18SM7357A | 4.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1797 | 2018-08-08 | H18SM7360A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1798 | 2018-08-08 | H18SM7361A | 2.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1799 | 2018-08-08 | H18SM7363A | 5.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1800 | 2018-08-08 | H18SM7364A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1801 | 2018-08-08 | H18SM7365A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1802 | 2018-08-08 | H18SM7366A | 1.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1803 | 2018-08-08 | H18SM7367A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 90
1804 | 2018-08-08 | H18SM7369A | 3.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1805 | 2018-08-08 | H18SM7371A | 3.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1806 | 2018-08-08 | H18SM7372A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
1807 | 2018-08-08 | H18SM7337A | 1.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1808 | 2018-08-08 | H18SM7356A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1809 | 2018-08-08 | H18SM7359A | 1.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1810 | 2018-08-08 | H18SM7362A | 2.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1811 | 2018-08-08 | H18SM7368A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1812 | 2018-08-08 | H18SM4328A | 1.3 | 0.0007 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
1813 | 2018-08-10 | H18SM4343A | 2.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1814 | 2018-08-08 | H18SM4344A | 3.4 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1815 | 2018-08-08 | H18SM4345A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1816 | 2018-08-08 | H18SM4346A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 73
1817 | 2018-08-10 | H18SM4347A | 2.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1818 | 2018-08-10 | H18SM4348A | 3.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1819 | 2018-08-10 | H18SM4349A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1820 | 2018-08-10 | H18SM4350A | 2.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1821 | 2018-08-10 | H18SM4351A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1822 | 2018-08-10 | H18SM4353A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1823 | 2018-08-10 | H18SM4355A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1824 | 2018-08-13 | H18SM4356A | 4.2 | 0.0004 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
1825 | 2018-08-10 | H18SM4357A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 79
1826 | 2018-08-13 | H18SM7350A | 1.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1827 | 2018-08-08 | H18SM7354A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1828 | 2018-08-08 | H18SM7358A | 3.4 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1829 | 2018-08-10 | H18SM7370A | 5.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1830 | 2018-08-06 | H18SM4313A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1831 | 2018-08-10 | H18SM4352A | 5.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1832 | 2018-08-10 | H18SM4354A | 3.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1833 | 2018-08-10 | H18SM4358A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1834 | 2018-08-13 | H18SM4359A | 2.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1835 | 2018-08-13 | H18SM4360A | 1.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1836 | 2018-08-13 | H18SM4361A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 84
1837 | 2018-08-13 | H18SM4362A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1838 | 2018-08-13 | H18SM4363A | 2.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1839 | 2018-08-13 | H18SM4364A | 3.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1840 | 2018-08-13 | H18SM4365A | 3.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1841 | 2018-08-13 | H18SM4366A | 3.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1842 | 2018-08-13 | H18SM4367A | 3.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1843 | 2018-08-13 | H18SM4368A | 4.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1844 | 2018-08-13 | H18SM4369A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1845 | 2018-08-13 | H18SM4370A | 4.3 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1846 | 2018-08-13 | H18SM4371A | 4.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 79
1847 | 2018-08-13 | H18SM4373A | 2.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1848 | 2018-08-13 | H18SM4375A | 1.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1849 | 2018-08-13 | H18SM4376A | 1.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1850 | 2018-08-14 | H18SM4336A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1851 | 2018-08-13 | H18SM4341A | 2.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1852 | 2018-08-14 | H18SM4372A | 4.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1853 | 2018-08-13 | H18SM4374A | 1.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1854 | 2018-08-13 | H18SM4377A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1855 | 2018-08-13 | H18SM4378A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1856 | 2018-08-13 | H18SM4379A | 3.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1857 | 2018-08-13 | H18SM4380A | 3.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1858 | 2018-08-13 | H18SM4381A | 3.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1859 | 2018-08-13 | H18SM4382A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
1860 | 2018-08-13 | H18SM4383A | 4.1 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1861 | 2018-08-14 | H18SM4385A | 4.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
1862 | 2018-08-14 | H18SM4386A | 5.1 | 0.0015 | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 76
1863 | 2018-08-14 | H18SM4387A | 5.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1864 | 2018-08-14 | H18SM4388A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1865 | 2018-08-14 | H18SM4389A | 3.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1866 | 2018-08-14 | H18SM4390A | 3.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1867 | 2018-08-14 | H18SM4391A | 3.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
1868 | 2018-08-14 | H18SM4393A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1869 | 2018-08-14 | H18SM4394A | 2.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1870 | 2018-08-14 | H18SM4396A | 4.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1871 | 2018-08-15 | H18SM4981A | 5.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1872 | 2018-08-15 | H18SM4392A | 3.1 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 76
1873 | 2018-08-15 | H18SM4395A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1874 | 2018-08-14 | H18SM4397A | 4.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1875 | 2018-08-14 | H18SM4398A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1876 | 2018-08-14 | H18SM4399A | 3.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1877 | 2018-08-15 | H18SM4400A | 3.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
1878 | 2018-08-14 | H18SM4401A | 4.1 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1879 | 2018-08-15 | H18SM4402A | 2.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1880 | 2018-08-15 | H18SM4403A | 2.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1881 | 2018-08-15 | H18SM4404A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1882 | 2018-08-15 | H18SM4405A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1883 | 2018-08-15 | H18SM4406A | 1.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1884 | 2018-08-15 | H18SM4407A | 2.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
1885 | 2018-08-15 | H18SM4408A | 3.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1886 | 2018-08-15 | H18SM4409A | 4.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1887 | 2018-08-15 | H18SM4410A | 4.4 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1888 | 2018-08-15 | H18SM4411A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1889 | 2018-08-15 | H18SM4412A | 4.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1890 | 2018-08-15 | H18SM4413A | 4.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1891 | 2018-08-15 | H18SM4414A | 3.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1892 | 2018-08-15 | H18SM4415A | 2.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
1893 | 2018-08-16 | H18SM4425A | 3.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 81
1894 | 2018-08-16 | H18SM4429A | 4.2 | 0.0018 | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 70
1895 | 2018-08-16 | H18SM4431A | 2.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1896 | 2018-08-16 | H18SM4435A | 3.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 79
1897 | 2018-08-16 | H18SM4437A | 1.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1898 | 2018-08-16 | H18SM4440A | 1.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1899 | 2018-08-16 | H18SM4444A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1900 | 2018-08-16 | H18SM4445A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1901 | 2018-08-16 | H18SM4447A | 4.2 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1902 | 2018-08-16 | H18SM4384A | 5.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1903 | 2018-08-16 | H18SM4416A | 5.2 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1904 | 2018-08-15 | H18SM4417A | 1.3 | 0.0019 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1905 | 2018-08-15 | H18SM4418A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
1906 | 2018-08-15 | H18SM4419A | 1.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1907 | 2018-08-15 | H18SM4420A | 2.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1908 | 2018-08-15 | H18SM4421A | 1.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
1909 | 2018-08-15 | H18SM4422A | 2.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1910 | 2018-08-16 | H18SM4424A | 3.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
1911 | 2018-08-15 | H18SM4426A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1912 | 2018-08-16 | H18SM4427A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1913 | 2018-08-16 | H18SM4428A | 5.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1914 | 2018-08-16 | H18SM4430A | 4.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1915 | 2018-08-16 | H18SM4432A | 4.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1916 | 2018-08-16 | H18SM4433A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1917 | 2018-08-16 | H18SM4434A | 2.1 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
1918 | 2018-08-16 | H18SM4436A | 1.4 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1919 | 2018-08-16 | H18SM4438A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1920 | 2018-08-16 | H18SM4439A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
1921 | 2018-08-20 | H18SM7388A | 3.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1922 | 2018-08-20 | H18SM7391A | 5.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1923 | 2018-08-20 | H18SM7393A | 5.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
1924 | 2018-08-20 | H18SM7395A | 6.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1925 | 2018-08-20 | H18SM7396A | 6.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1926 | 2018-08-20 | H18SM7397A | 2.4 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 43 | 76
1927 | 2018-08-20 | H18SM7402A | 2.1 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1928 | 2018-08-20 | H18SM7404A | 1.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
1929 | 2018-08-20 | H18SM7405A | 3.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1930 | 2018-08-21 | H18SM7385A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 78
1931 | 2018-08-20 | H18SM7376A | 4.4 | 0.0011 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 77
1932 | 2018-08-20 | H18SM7377A | 5.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 81
1933 | 2018-08-20 | H18SM7378A | 2.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 78
1934 | 2018-08-20 | H18SM7379A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
1935 | 2018-08-20 | H18SM7380A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1936 | 2018-08-20 | H18SM7381A | 1.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1937 | 2018-08-20 | H18SM7382A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1938 | 2018-08-20 | H18SM7383A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1939 | 2018-08-20 | H18SM7384A | 1.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1940 | 2018-08-20 | H18SM7386A | 3.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1941 | 2018-08-20 | H18SM7387A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
1942 | 2018-08-20 | H18SM7389A | 4.1 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
1943 | 2018-08-20 | H18SM7390A | 3.2 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1944 | 2018-08-20 | H18SM7392A | 5.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1945 | 2018-08-20 | H18SM7398A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1946 | 2018-08-20 | H18SM7399A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 80
1947 | 2018-08-21 | H18SM7400A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1948 | 2018-08-17 | H18SM4423A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
1949 | 2018-08-20 | H18SM4443A | 5.3 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1950 | 2018-08-20 | H18SM4480A | 3.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 38 | 78
1951 | 2018-08-21 | H18SM7406A | 4.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1952 | 2018-08-21 | H18SM7407A | 2.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1953 | 2018-08-21 | H18SM7408A | 3.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1954 | 2018-08-21 | H18SM7409A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 79
1955 | 2018-08-21 | H18SM7410A | 3.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1956 | 2018-08-21 | H18SM7411A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 81
1957 | 2018-08-21 | H18SM7412A | 4.4 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1958 | 2018-08-21 | H18SM7413A | 4.2 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1959 | 2018-08-21 | H18SM7414A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
1960 | 2018-08-21 | H18SM7415A | 5.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
1961 | 2018-08-21 | H18SM7417A | 2.2 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1962 | 2018-08-21 | H18SM7418A | 2.1 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 81
1963 | 2018-08-21 | H18SM7420A | 5.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 79
1964 | 2018-08-21 | H18SM7421A | 5.4 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
1965 | 2018-08-23 | H18SM7422A | 1.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 78
1966 | 2018-08-23 | H18SM7423A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
1967 | 2018-08-23 | H18SM7394A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
1968 | 2018-08-23 | H18SM7401A | 5.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1969 | 2018-08-23 | H18SM7403A | 1.4 | 0.0017 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 70
1970 | 2018-08-23 | H18SM7416A | 5.2 | 0.0013 | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 78
1971 | 2018-08-23 | H18SM7419A | 1.3 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1972 | 2018-08-23 | H18SM7424A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
1973 | 2018-08-23 | H18SM7426A | 2.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1974 | 2018-08-23 | H18SM7427A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1975 | 2018-08-23 | H18SM7428A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1976 | 2018-08-23 | H18SM7429A | 3.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1977 | 2018-08-23 | H18SM7430A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1978 | 2018-08-23 | H18SM7431A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
1979 | 2018-08-23 | H18SM7432A | 4.1 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 80
1980 | 2018-08-23 | H18SM7433A | 4.2 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1981 | 2018-08-23 | H18SM7434A | 4.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1982 | 2018-08-23 | H18SM7435A | 4.3 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1983 | 2018-08-25 | H18SM4496A | 4.3 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
1984 | 2018-08-24 | H18SM4506A | 5.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
1985 | 2018-08-24 | H18SM4511A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
1986 | 2018-08-25 | H18SM4513A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1987 | 2018-08-25 | H18SM4515A | 1.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
1988 | 2018-08-25 | H18SM4518A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
1989 | 2018-08-25 | H18SM4519A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1990 | 2018-08-25 | H18SM4520A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
1991 | 2018-08-25 | H18SM4521A | 4.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
1992 | 2018-08-25 | H18SM4522A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
1993 | 2018-08-23 | H18SM4482A | 5.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 79
1994 | 2018-08-23 | H18SM4483A | 1.1 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1995 | 2018-08-24 | H18SM4484A | 1.2 | 0.0004 | 0.002 | Resolve unit | FLEX 30:70 Mix | 43 | 70
1996 | 2018-08-23 | H18SM4485A | 2.2 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 70
1997 | 2018-08-24 | H18SM4486A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
1998 | 2018-08-23 | H18SM4487A | 2.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
1999 | 2018-08-23 | H18SM4488A | 3.2 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2000 | 2018-08-24 | H18SM4489A | 3.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2001 | 2018-08-24 | H18SM4490A | 2.4 | 0.0005 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 75
2002 | 2018-08-23 | H18SM4491A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 61
2003 | 2018-08-24 | H18SM4492A | 3.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2004 | 2018-08-24 | H18SM4493A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2005 | 2018-08-24 | H18SM4494A | 4.2 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2006 | 2018-08-24 | H18SM4495A | 5.1 | 0.0006 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 72
2007 | 2018-08-24 | H18SM4497A | 4.4 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2008 | 2018-08-24 | H18SM4498A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2009 | 2018-08-24 | H18SM4499A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2010 | 2018-08-24 | H18SM4501A | 1.2 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2011 | 2018-08-24 | H18SM4502A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2012 | 2018-08-24 | H18SM4503A | 2.4 | 0 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2013 | 2018-08-24 | H18SM4504A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2014 | 2018-08-25 | H18SM4505A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2015 | 2018-08-24 | H18SM4508A | 4.2 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2016 | 2018-08-24 | H18SM4509A | 3.3 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
2017 | 2018-08-24 | H18SM4510A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2018 | 2018-08-25 | H18SM4512A | 2.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2019 | 2018-08-25 | H18SM4516A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2020 | 2018-08-25 | H18SM4517A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2021 | 2018-08-27 | H18SM4523A | 4.4 | 0.0008 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 70
2022 | 2018-08-27 | H18SM4538A | 1.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2023 | 2018-08-27 | H18SM4542A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2024 | 2018-08-27 | H18SM4544A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2025 | 2018-08-27 | H18SM4545A | 4.2 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2026 | 2018-08-27 | H18SM4546A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2027 | 2018-08-27 | H18SM4547A | 4.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2028 | 2018-08-27 | H18SM4549A | 1.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2029 | 2018-08-27 | H18SM4550A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2030 | 2018-08-27 | H18SM4551A | 5.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2031 | 2018-08-27 | H18SM4553A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2032 | 2018-08-27 | H18SM4556A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2033 | 2018-08-27 | H18SM4558A | 4.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2034 | 2018-08-28 | H18SM4561A | 2.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2035 | 2018-08-28 | H18SM4562A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2036 | 2018-08-27 | H18SM4507A | 5.3 | 0.0005 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 70
2037 | 2018-08-27 | H18SM4514A | 2.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2038 | 2018-08-25 | H18SM4524A | 5.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2039 | 2018-08-25 | H18SM4526A | 1.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2040 | 2018-08-27 | H18SM4527A | 5.2 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
2041 | 2018-08-25 | H18SM4528A | 2.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2042 | 2018-08-25 | H18SM4529A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2043 | 2018-08-27 | H18SM4530A | 3.1 | 0.0005 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
2044 | 2018-08-27 | H18SM4531A | 1.2 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 70
2045 | 2018-08-27 | H18SM4532A | 1.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2046 | 2018-08-27 | H18SM4533A | 1.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2047 | 2018-08-27 | H18SM4534A | 4.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2048 | 2018-08-27 | H18SM4535A | 2.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2049 | 2018-08-27 | H18SM4536A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2050 | 2018-08-27 | H18SM4537A | 4.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2051 | 2018-08-27 | H18SM4539A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2052 | 2018-08-27 | H18SM4540A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2053 | 2018-08-27 | H18SM4541A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2054 | 2018-08-27 | H18SM4543A | 4.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2055 | 2018-08-27 | H18SM4548A | 3.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2056 | 2018-08-28 | H18SM4525A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2057 | 2018-08-28 | H18SM4552A | 1.3 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2058 | 2018-08-27 | H18SM4555A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
2059 | 2018-08-28 | H18SM4557A | 3.3 | 0.0011 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 70
2060 | 2018-08-28 | H18SM4559A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2061 | 2018-08-28 | H18SM4560A | 1.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2062 | 2018-08-28 | H18SM4563A | 3.1 | 0.0018 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2063 | 2018-08-28 | H18SM4564A | 5.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
2064 | 2018-08-28 | H18SM4565A | 4.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2065 | 2018-08-28 | H18SM4566A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2066 | 2018-08-28 | H18SM4568A | 4.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2067 | 2018-08-28 | H18SM4569A | 4.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2068 | 2018-08-28 | H18SM4570A | 5.2 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2069 | 2018-08-28 | H18SM4571A | 5.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2070 | 2018-08-28 | H18SM4573A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2071 | 2018-08-28 | H18SM4574A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2072 | 2018-08-28 | H18SM4576A | 1.2 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 79
2073 | 2018-08-28 | H18SM4577A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2074 | 2018-08-28 | H18SM4578A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2075 | 2018-08-29 | H18SM4554A | 1.4 | 0.0015 | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 75
2076 | 2018-08-28 | H18SM4572A | 2.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2077 | 2018-08-28 | H18SM4575A | 1.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2078 | 2018-08-29 | H18SM4579A | 4.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
2079 | 2018-08-29 | H18SM4580A | 3.1 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2080 | 2018-08-28 | H18SM4581A | 3.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2081 | 2018-08-29 | H18SM4582A | 5.2 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2082 | 2018-08-29 | H18SM4583A | 4.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2083 | 2018-08-29 | H18SM4584A | 4.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2084 | 2018-08-29 | H18SM4585A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2085 | 2018-08-29 | H18SM4586A | 5.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2086 | 2018-08-29 | H18SM4587A | 5.4 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2087 | 2018-08-29 | H18SM4588A | 4.3 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2088 | 2018-08-29 | H18SM4589A | 6.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2089 | 2018-08-29 | H18SM4591A | 1.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2090 | 2018-08-29 | H18SM4592A | 1.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2091 | 2018-08-29 | H18SM4593A | 3.3 | 0.0019 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2092 | 2018-08-29 | H18SM4594A | 4.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2093 | 2018-08-29 | H18SM4595A | 2.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
2094 | 2018-08-29 | H18SM4596A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2095 | 2018-08-29 | H18SM4597A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2096 | 2018-08-29 | H18SM4598A | 3.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
2097 | 2018-08-29 | H18SM4599A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 67
2098 | 2018-08-29 | H18SM4600A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2099 | 2018-08-29 | H18SM4602A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2100 | 2018-08-30 | H18SM7436A | 1.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2101 | 2018-08-30 | H18SM7437A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2102 | 2018-08-30 | H18SM7438A | 1.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2103 | 2018-08-30 | H18SM7439A | 1.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2104 | 2018-08-30 | H18SM7440A | 2.3 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2105 | 2018-08-30 | H18SM7441A | 2.2 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2106 | 2018-08-30 | H18SM7442A | 1.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2107 | 2018-08-30 | H18SM7443A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2108 | 2018-08-30 | H18SM7444A | 3.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2109 | 2018-08-30 | H18SM7445A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2110 | 2018-08-30 | H18SM7446A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2111 | 2018-08-30 | H18SM7448A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2112 | 2018-08-30 | H18SM7450A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2113 | 2018-08-30 | H18SM7452A | 1.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2114 | 2018-08-30 | H18SM7455A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2115 | 2018-08-30 | H18SM7456A | 5.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2116 | 2018-08-30 | H18SM4500A | 2.4 | 0.0001 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
2117 | 2018-08-28 | H18SM4567A | 3.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2118 | 2018-08-30 | H18SM4590A | 2.2 | 0.0018 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2119 | 2018-08-30 | H18SM4601A | 5.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2120 | 2018-08-30 | H18SM4603A | 5.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2121 | 2018-08-29 | H18SM4604A | 5.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2122 | 2018-08-31 | H18SM4605A | 5.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2123 | 2018-08-31 | H18SM4606A | 3.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2124 | 2018-08-30 | H18SM4607A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2125 | 2018-08-30 | H18SM4608A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2126 | 2018-08-30 | H18SM4609A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2127 | 2018-08-30 | H18SM4610A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2128 | 2018-08-30 | H18SM7447A | 4.1 | 0.0018 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2129 | 2018-08-30 | H18SM7451A | 5.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2130 | 2018-08-30 | H18SM7453A | 1.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2131 | 2018-08-30 | H18SM7457A | 1.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2132 | 2018-08-30 | H18SM7458A | 2.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2133 | 2018-08-30 | H18SM7454A | 2.2 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2134 | 2018-08-30 | H18SM7449A | 2.1 | 0.0017 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 72
2135 | 2018-08-31 | H18SM4621A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2136 | 2018-08-31 | H18SM4622A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2137 | 2018-08-31 | H18SM4626A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2138 | 2018-08-31 | H18SM4627A | 2.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2139 | 2018-08-31 | H18SM4630A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2140 | 2018-08-31 | H18SM4611A | 1.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2141 | 2018-08-31 | H18SM4612A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2142 | 2018-08-31 | H18SM4614A | 1.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2143 | 2018-08-31 | H18SM4615A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2144 | 2018-08-31 | H18SM4616A | 2.2 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2145 | 2018-08-31 | H18SM4617A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2146 | 2018-08-31 | H18SM4618A | 2.4 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2147 | 2018-08-31 | H18SM4619A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2148 | 2018-08-31 | H18SM4620A | 3.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2149 | 2018-09-03 | H18SM4613A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2150 | 2018-08-31 | H18SM4623A | 4.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2151 | 2018-08-31 | H18SM4624A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2152 | 2018-09-03 | H18SM4625A | 4.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2153 | 2018-09-03 | H18SM4628A | 2.3 | 0.0014 | 0.002 | Resolve unit | FLEX 30:70 Mix | 43 | 75
2154 | 2018-08-31 | H18SM4629A | 2.2 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2155 | 2018-09-03 | H18SM4631A | 3.2 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 78
2156 | 2018-09-03 | H18SM4632A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2157 | 2018-09-03 | H18SM4633A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2158 | 2018-09-03 | H18SM4634A | 1.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2159 | 2018-09-03 | H18SM4635A | 2.2 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 79
2160 | 2018-09-03 | H18SM4636A | 1.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 79
2161 | 2018-09-03 | H18SM4637A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2162 | 2018-09-03 | H18SM4638A | 3.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2163 | 2018-09-03 | H18SM4639A | 3.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2164 | 2018-09-03 | H18SM4640A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2165 | 2018-09-03 | H18SM4641A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2166 | 2018-09-03 | H18SM4643A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2167 | 2018-09-03 | H18SM4644A | 5.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2168 | 2018-09-03 | H18SM4645A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2169 | 2018-09-03 | H18SM4653A | 4.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2170 | 2018-09-03 | H18SM4655A | 4.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2171 | 2018-09-03 | H18SM4656A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2172 | 2018-09-03 | H18SM4642A | 3.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2173 | 2018-09-03 | J18SM4646A | 2.3 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2174 | 2018-09-03 | J18SM4647A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2175 | 2018-09-03 | J18SM4648A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2176 | 2018-09-03 | J18SM4649A | 2.2 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2177 | 2018-09-03 | J18SM4652A | 3.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2178 | 2018-09-03 | J18SM4651A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2179 | 2018-09-04 | J18SM4654A | 5.1 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2180 | 2018-09-04 | J18SM4658A | 5.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2181 | 2018-09-04 | J18SM4661A | 4.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2182 | 2018-09-04 | J18SM4662A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2183 | 2018-09-04 | J18SM4664A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2184 | 2018-09-04 | J18SM4665A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2185 | 2018-09-04 | J18SM4650A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
2186 | 2018-09-04 | J18SM4657A | 5.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2187 | 2018-09-04 | J18SM4660A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2188 | 2018-09-04 | J18SM4663A | 4.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2189 | 2018-09-04 | J18SM4666A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2190 | 2018-09-04 | J18SM4667A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2191 | 2018-09-04 | J18SM4668A | 2.4 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2192 | 2018-09-04 | J18SM4670A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2193 | 2018-09-04 | H18SM7459A | 2.2 | 0.0013 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 72
2194 | 2018-09-04 | H18SM7460A | 1.4 | 0.0018 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2195 | 2018-09-04 | J18SM7462A | 1.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2196 | 2018-09-04 | J18SM7463A | 1.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2197 | 2018-09-04 | J18SM7464A | 1.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2198 | 2018-09-04 | J18SM7465A | 3.3 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2199 | 2018-09-04 | J18SM7467A | 4.1 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2200 | 2018-09-04 | J18SM7468A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2201 | 2018-09-04 | J18SM7469A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2202 | 2018-09-04 | J18SM7470A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2203 | 2018-09-04 | J18SM7471A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
2204 | 2018-09-05 | J18SM7461A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2205 | 2018-09-04 | J18SM7466A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2206 | 2018-09-05 | J18SM7473A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2207 | 2018-09-05 | J18SM7474A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2208 | 2018-09-05 | J18SM7475A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2209 | 2018-09-05 | J18SM7476A | 2.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2210 | 2018-09-05 | J18SM7478A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2211 | 2018-09-05 | J18SM7479A | 2.4 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2212 | 2018-09-05 | J18SM7481A | 3.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2213 | 2018-09-05 | J18SM7483A | 3.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2214 | 2018-09-05 | J18SM7484A | 4.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2215 | 2018-09-05 | J18SM7486A | 4.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2216 | 2018-09-05 | J18SM7487A | 4.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2217 | 2018-09-05 | J18SM7488A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2218 | 2018-09-05 | J18SM7489A | 5.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2219 | 2018-09-05 | J18SM7490A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2220 | 2018-09-05 | J18SM7492A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
2221 | 2018-09-05 | J18SM7493A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2222 | 2018-09-06 | J18SM7472A | 1.2 | 0.0004 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 77
2223 | 2018-09-06 | J18SM7477A | 1.4 | 0.0001 | 0.002 | Resolve unit | FLEX 30:70 Mix | 44 | 72
2224 | 2018-09-06 | J18SM7480A | 3.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2225 | 2018-09-06 | J18SM7482A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2226 | 2018-09-05 | J18SM7485A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2227 | 2018-09-06 | J18SM7491A | 1.1 | 0.0012 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
2228 | 2018-09-06 | J18SM7494A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2229 | 2018-09-06 | J18SM7495A | 5.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2230 | 2018-09-06 | J18SM7496A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2231 | 2018-09-06 | J18SM7497A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2232 | 2018-09-06 | J18SM7498A | 2.3 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2233 | 2018-09-06 | J18SM7499A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2234 | 2018-09-06 | J18SM7500A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2235 | 2018-09-06 | J18SM7501A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2236 | 2018-09-06 | J18SM7502A | 3.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2237 | 2018-09-06 | J18SM7503A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2238 | 2018-09-06 | J18SM7504A | 4.2 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2239 | 2018-09-08 | J18SM4697A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2240 | 2018-09-08 | J18SM4699A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2241 | 2018-09-08 | J18SM4704A | 2.2 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2242 | 2018-09-08 | J18SM4705A | 3.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2243 | 2018-09-06 | J18SM4681A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2244 | 2018-09-06 | J18SM4682A | 4.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2245 | 2018-09-07 | J18SM4683A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2246 | 2018-09-07 | J18SM4684A | 2.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2247 | 2018-09-07 | J18SM4685A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2248 | 2018-09-07 | J18SM4686A | 1.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2249 | 2018-09-07 | J18SM4687A | 1.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2250 | 2018-09-07 | J18SM4688A | 1.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2251 | 2018-09-07 | J18SM4689A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2252 | 2018-09-07 | J18SM4691A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2253 | 2018-09-07 | J18SM4692A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2254 | 2018-09-07 | J18SM4693A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2255 | 2018-09-07 | J18SM4694A | 3.3 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2256 | 2018-09-07 | J18SM4695A | 3.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2257 | 2018-09-07 | J18SM4696A | 4.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2258 | 2018-09-07 | J18SM4698A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2259 | 2018-09-07 | J18SM4700A | 5.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2260 | 2018-09-07 | J18SM4701A | 5.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2261 | 2018-09-07 | J18SM4680A | 5.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2262 | 2018-09-10 | J18SM4702A | 5.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2263 | 2018-09-08 | J18SM4703A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2264 | 2018-09-08 | J18SM4706A | 2.4 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2265 | 2018-09-08 | J18SM4707A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2266 | 2018-09-08 | J18SM4708A | 3.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2267 | 2018-09-10 | J18SM4709A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2268 | 2018-09-08 | J18SM4710A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2269 | 2018-09-10 | J18SM4711A | 3.4 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2270 | 2018-09-10 | J18SM4712A | 4.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2271 | 2018-09-10 | J18SM4713A | 4.2 | 0.0011 | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 75
2272 | 2018-09-10 | J18SM4715A | 5.1 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2273 | 2018-09-10 | J18SM4716A | 3.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2274 | 2018-09-10 | J18SM4717A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2275 | 2018-09-10 | J18SM4718A | 3.1 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2276 | 2018-09-10 | J18SM4719A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2277 | 2018-09-10 | J18SM4720A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2278 | 2018-09-10 | J18SM4724A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2279 | 2018-09-10 | J18SM4725A | 5.1 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2280 | 2018-09-10 | J18SM4726A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2281 | 2018-09-10 | J18SM4727A | 3.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2282 | 2018-09-10 | J18SM4729A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
2283 | 2018-09-10 | J18SM4730A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2284 | 2018-09-10 | J18SM4714A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2285 | 2018-09-11 | J18SM4721A | 2.3 | 0.0003 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 72
2286 | 2018-09-10 | J18SM4723A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2287 | 2018-09-11 | J18SM4728A | 5.2 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2288 | 2018-09-10 | J18SM4731A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 80
2289 | 2018-09-11 | J18SM4732A | 1.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2290 | 2018-09-10 | J18SM4733A | 2.1 | 0.0019 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2291 | 2018-09-11 | J18SM4734A | 1.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2292 | 2018-09-11 | J18SM4735A | 1.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2293 | 2018-09-11 | J18SM4736A | 3.1 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2294 | 2018-09-11 | J18SM4737A | 1.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2295 | 2018-09-11 | J18SM4738A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2296 | 2018-09-11 | J18SM4739A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2297 | 2018-09-11 | J18SM4740A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2298 | 2018-09-11 | J18SM4741A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2299 | 2018-09-11 | J18SM4742A | 4.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2300 | 2018-09-11 | J18SM4690A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2301 | 2018-09-11 | J18SM4743A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2302 | 2018-09-11 | J18SM4744A | 5.1 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2303 | 2018-09-11 | J18SM4745A | 4.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2304 | 2018-09-11 | J18SM4746A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2305 | 2018-09-11 | J18SM4747A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2306 | 2018-09-11 | J18SM4748A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2307 | 2018-09-11 | J18SM4749A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2308 | 2018-09-11 | J18SM4750A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2309 | 2018-09-11 | J18SM4751A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2310 | 2018-09-11 | J18SM4753A | 3.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 78
2311 | 2018-09-11 | J18SM4754A | 2.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2312 | 2018-09-11 | J18SM4758A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2313 | 2018-09-12 | J18SM4722A | 1.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2314 | 2018-09-12 | J18SM4752A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2315 | 2018-09-12 | J18SM4755A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2316 | 2018-09-12 | J18SM4756A | 5.2 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 68
2317 | 2018-09-12 | J18SM4757A | 1.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2318 | 2018-09-12 | J18SM4759A | 1.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 66
2319 | 2018-09-12 | J18SM4760A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2320 | 2018-09-12 | J18SM4761A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2321 | 2018-09-12 | J18SM4762A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2322 | 2018-09-12 | J18SM4763A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2323 | 2018-09-12 | J18SM4764A | 2.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2324 | 2018-09-12 | J18SM4765A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2325 | 2018-09-12 | J18SM4766A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2326 | 2018-09-12 | J18SM4767A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2327 | 2018-09-12 | J18SM4768A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2328 | 2018-09-12 | J18SM4769A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2329 | 2018-09-12 | J18SM4770A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2330 | 2018-09-12 | J18SM4771A | 5.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2331 | 2018-09-12 | J18SM4772A | 5.2 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2332 | 2018-09-12 | J18SM4773A | 5.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
2333 | 2018-09-12 | J18SM4777A | 1.2 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 72
2334 | 2018-09-12 | J18SM47778A | 1.3 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 69
2335 | 2018-09-12 | J18SM47779A | 1.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2336 | 2018-09-14 | J18SM4806A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 76
2337 | 2018-09-14 | J18SM4810A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2338 | 2018-09-14 | J18SM4815A | 2.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2339 | 2018-09-14 | J18SM4816A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2340 | 2018-09-14 | J18SM4817A | 2.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2341 | 2018-09-14 | J18SM4818A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2342 | 2018-09-14 | J18SM4823A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2343 | 2018-09-14 | J18SM4825A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2344 | 2018-09-14 | J18SM4827A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2345 | 2018-09-14 | J18SM4828A | 1.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2346 | 2018-09-14 | J18SM4829A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2347 | 2018-09-14 | J18SM4830A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2348 | 2018-09-15 | J18SM4831A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2349 | 2018-09-15 | J18SM4832A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2350 | 2018-09-15 | J18SM4834A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2351 | 2018-09-15 | J18SM4837A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 80
2352 | 2018-09-13 | J18SM4798A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2353 | 2018-09-13 | J18SM4799A | 3.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2354 | 2018-09-13 | J18SM4800A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2355 | 2018-09-13 | J18SM4801A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2356 | 2018-09-13 | J18SM4802A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 73
2357 | 2018-09-13 | J18SM4803A | 2.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2358 | 2018-09-13 | J18SM4804A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2359 | 2018-09-14 | J18SM4805A | 1.4 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 72
2360 | 2018-09-14 | J18SM4807A | 1.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2361 | 2018-09-14 | J18SM4808A | 1.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2362 | 2018-09-14 | J18SM4809A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2363 | 2018-09-14 | J18SM4811A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2364 | 2018-09-14 | J18SM4812A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2365 | 2018-09-14 | J18SM4813A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2366 | 2018-09-14 | J18SM4814A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2367 | 2018-09-12 | J18SM4774A | 5.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2368 | 2018-09-12 | J18SM4775A | 1.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2369 | 2018-09-13 | J18SM4776A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2370 | 2018-09-13 | J18SM4780A | 2.1 | 0.0006 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 69
2371 | 2018-09-13 | J18SM4781A | 2.2 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 70
2372 | 2018-09-13 | J18SM4782A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 68
2373 | 2018-09-13 | J18SM4783A | 2.3 | 0 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 68
2374 | 2018-09-13 | J18SM4784A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2375 | 2018-09-13 | J18SM4785A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2376 | 2018-09-13 | J18SM4786A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 68
2377 | 2018-09-13 | J18SM4787A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2378 | 2018-09-13 | J18SM4788A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2379 | 2018-09-13 | J18SM4789A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2380 | 2018-09-13 | J18SM4790A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2381 | 2018-09-13 | J18SM4791A | 1.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2382 | 2018-09-13 | J18SM4792A | 4.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2383 | 2018-09-13 | J18SM4793A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2384 | 2018-09-13 | J18SM4794A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
2385 | 2018-09-13 | J18SM4795A | 5.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2386 | 2018-09-13 | J18SM4796A | 5.2 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2387 | 2018-09-13 | J18SM4797A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | . | 72
2388 | 2018-09-15 | J18SM4819A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2389 | 2018-09-14 | J18SM4820A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2390 | 2018-09-15 | J18SM4821A | 5.1 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2391 | 2018-09-14 | J18SM4822A | 4.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2392 | 2018-09-14 | J18SM4824A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2393 | 2018-09-14 | J18SM4826A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2394 | 2018-09-15 | J18SM4833A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2395 | 2018-09-15 | J18SM4835A | 4.2 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2396 | 2018-09-15 | J18SM4836A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2397 | 2018-09-17 | J18SM4838A | 1.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2398 | 2018-09-17 | J18SM4839A | 1.2 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 72
2399 | 2018-09-17 | J18SM4840A | 1.3 | 0.0004 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 74
2400 | 2018-09-17 | J18SM4841A | 1.4 | 0.0007 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
2401 | 2018-09-17 | J18SM4842A | 2.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2402 | 2018-09-17 | J18SM4843A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2403 | 2018-09-17 | J18SM4844A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2404 | 2018-09-17 | J18SM4845A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2405 | 2018-09-17 | J18SM7505A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2406 | 2018-09-17 | J18SM7506A | 1.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2407 | 2018-09-18 | J18SM7507A | 3.3 | 0.0009 | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 75
2408 | 2018-09-17 | J18SM7508A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2409 | 2018-09-17 | J18SM7509A | 3.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2410 | 2018-09-17 | J18SM7510A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2411 | 2018-09-17 | J18SM7511A | 4.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2412 | 2018-09-17 | J18SM7512A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2413 | 2018-09-17 | J18SM7513A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2414 | 2018-09-17 | J18SM7514A | 5.2 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2415 | 2018-09-17 | J18SM7516A | 5.3 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2416 | 2018-09-17 | J18SM7517A | 5.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2417 | 2018-09-17 | J18SM7520A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2418 | 2018-09-17 | J18SM7521A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2419 | 2018-09-17 | J18SM7523A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2420 | 2018-09-17 | J18SM7515A | 5.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2421 | 2018-09-18 | J18SM7518A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2422 | 2018-09-18 | J18SM7519A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2423 | 2018-09-18 | J18SM7522A | 2.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2424 | 2018-09-18 | J18SM7524A | 1.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2425 | 2018-09-18 | J18SM7525A | 1.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2426 | 2018-09-18 | J18SM7526A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2427 | 2018-09-18 | J18SM7527A | 1.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2428 | 2018-09-18 | J18SM7528A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2429 | 2018-09-18 | J18SM7529A | 4.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2430 | 2018-09-18 | J18SM7530A | 5.2 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2431 | 2018-09-18 | J18SM7532A | 5.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2432 | 2018-09-18 | J18SM7533A | 5.3 | 0.0014 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2433 | 2018-09-18 | J18SM7535A | 2.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2434 | 2018-09-18 | J18SM7538A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2435 | 2018-09-18 | J18SM7534A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2436 | 2018-09-18 | J18SM7537A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2437 | 2018-09-18 | J18SM7539A | 1.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2438 | 2018-09-18 | J18SM7543A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2439 | 2018-09-19 | J18SM7531A | 3.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2440 | 2018-09-19 | J18SM7536A | 2.1 | 0.0005 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 72
2441 | 2018-09-19 | J18SM7542A | 4.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2442 | 2018-09-19 | J18SM7544A | 4.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2443 | 2018-09-19 | J18SM7540A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2444 | 2018-09-19 | J18SM7541A | 4.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2445 | 2018-09-19 | J18SM7545A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2446 | 2018-09-19 | J18SM7546A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2447 | 2018-09-19 | J18SM7547A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2448 | 2018-09-19 | J18SM7548A | 2.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 75
2449 | 2018-09-19 | J18SM7549A | 3.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2450 | 2018-09-19 | J18SM7550A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2451 | 2018-09-19 | J18SM7551A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2452 | 2018-09-19 | J18SM7552A | 2.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2453 | 2018-09-19 | J18SM7553A | 4.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2454 | 2018-09-19 | J18SM7554A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2455 | 2018-09-19 | J18SM7555A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2456 | 2018-09-19 | J18SM7556A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2457 | 2018-09-19 | J18SM7557A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2458 | 2018-09-19 | J18SM7558A | 2.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2459 | 2018-09-19 | J18SM7559A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2460 | 2018-09-19 | J18SM7560A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2461 | 2018-09-19 | J18SM7562A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2462 | 2018-09-19 | J18SM7563A | 3.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2463 | 2018-09-20 | J18SM7561A | 3.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 79
2464 | 2018-09-20 | J18SM7564A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2465 | 2018-09-20 | J18SM7566A | 4.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2466 | 2018-09-20 | J18SM7567A | 4.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2467 | 2018-09-20 | J18SM7568A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2468 | 2018-09-20 | J18SM7569A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2469 | 2018-09-20 | J18SM7570A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2470 | 2018-09-20 | J18SM7571A | 1.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2471 | 2018-09-20 | J18SM7572A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2472 | 2018-09-20 | J18SM7573A | 1.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2473 | 2018-09-20 | J18SM7574A | 1.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2474 | 2018-09-20 | J18SM7575A | 2.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2475 | 2018-09-19 | J18SM4846A | 4.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2476 | 21 sp 2018 | J18SM4848A | 3.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2477 | 2018-09-20 | J18SM4849A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2478 | 2018-09-20 | J18SM4850A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2479 | 2018-09-21 | J18SM4851A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2480 | 2018-09-21 | J18SM4852A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2481 | 2018-09-21 | J18SM4852A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2482 | 2018-09-21 | J18SM4854A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2483 | 2018-09-22 | J18SM4855A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2484 | 2018-09-21 | J18SM4856A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2485 | 2018-09-21 | J18SM4857A | 3.3 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2486 | 2018-09-21 | J18SM4858A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2487 | 2018-09-21 | J18SM4859A | 2.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2488 | 2018-09-21 | J18SM4860A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2489 | 2018-09-21 | J18SM4861A | 2.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2490 | 2018-09-21 | J18SM4862A | 1.4 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2491 | 2018-09-22 | J18SM4867A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2492 | 2018-09-20 | J18SM4847A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2493 | 2018-09-21 | J18SM4863A | 1.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2494 | 2018-09-21 | J18SM4864A | 1.2 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2495 | 2018-09-21 | J18SM4865A | 1.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2496 | 2018-09-24 | J18SM4866A | 3.3 | 0.0004 | 0.002 | Resolve unit | FLEX 30:70 Mix | 49 | 74
2497 | 2018-09-21 | J18SM4868A | 3.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2498 | 2018-09-22 | J18SM4869A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2499 | 2018-09-24 | J18SM4870A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2500 | 2018-09-24 | J18SM4871A | 1.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 70
2501 | 2018-09-24 | J18SM4872A | 2.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2502 | 2018-09-24 | J18SM4873A | 1.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2503 | 2018-09-24 | J18SM4874A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2504 | 2018-09-24 | J18SM4875A | 1.2 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2505 | 2018-09-24 | J18SM4877A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 79
2506 | 2018-09-24 | J18SM4878A | 2.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2507 | 2018-09-24 | J18SM4879A | 3.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2508 | 2018-09-24 | J18SM4880A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2509 | 2018-09-24 | J18SM4881A | 3.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2510 | 2018-09-24 | J18SM4882A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2511 | 2018-09-24 | J18SM4883A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 70
2512 | 2018-09-24 | J18SM4884A | 1.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
2513 | 2018-09-24 | J18SM4885A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2514 | 2018-09-24 | J18SM4888A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2515 | 2018-09-24 | J18SM4889A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2516 | 2018-09-24 | J18SM4892A | 3.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2517 | 2018-09-25 | J18SM4908A | 6.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2518 | 2018-09-25 | J18SM4923A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2519 | 2018-09-26 | J18SM4926A | 3.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2520 | 2018-09-26 | J18SM4927A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2521 | 2018-09-25 | J18SM4906A | 1.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2522 | 2018-09-25 | J18SM4907A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 69
2523 | 2018-09-25 | J18SM4909A | 6.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2524 | 2018-09-25 | J18SM4910A | 6.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2525 | 2018-09-25 | J18SM4911A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2526 | 2018-09-25 | J18SM4912A | 3.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2527 | 2018-09-25 | J18SM4913A | 3.3 | 0 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2528 | 2018-09-25 | J18SM4914A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2529 | 2018-09-25 | J18SM4915A | 4.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2530 | 2018-09-25 | J18SM4916A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2531 | 2018-09-25 | J18SM4918A | 2.4 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2532 | 2018-09-25 | J18SM4919A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2533 | 2018-09-25 | J18SM4920A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2534 | 2018-09-25 | J18SM4921A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2536 | 2018-09-26 | J18SM4922A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2537 | 2018-09-25 | J18SM4924A | 1.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2538 | 2018-09-25 | J18SM4925A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
2539 | 2018-09-24 | J18SM4876A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2540 | 2018-09-24 | J18SM4886A | 3.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2541 | 2018-09-25 | J18SM4887A | 1.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 79
2542 | 2018-09-24 | J18SM4890A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 73
2543 | 2018-09-25 | J18SM4891A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2544 | 2018-09-25 | J18SM4893A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
2545 | 2018-09-25 | J18SM4894A | 1.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2546 | 2018-09-25 | J18SM4895A | 1.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2547 | 2018-09-25 | J18SM4896A | 4.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2548 | 2018-09-25 | J18SM4897A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2549 | 2018-09-25 | J18SM4898A | 4.2 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2550 | 2018-09-25 | J18SM4899A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2551 | 2018-09-25 | J18SM4900A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2552 | 2018-09-25 | J18SM4901A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2553 | 2018-09-25 | J18SM4902A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2554 | 2018-09-25 | J18SM4903A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2555 | 2018-09-25 | J18SM4904A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2556 | 2018-09-25 | J18SM4905A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2557 | 2018-09-26 | J18SM4917A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 72
2558 | 2018-09-26 | J18SM4928A | 4.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2559 | 2018-09-26 | J18SM4929A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2560 | 2018-09-26 | J18SM4931A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 70
2561 | 2018-09-26 | J18SM4932A | 5.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2562 | 2018-09-26 | J18SM4933A | 5.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2563 | 2018-09-26 | J18SM4934A | 5.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2564 | 2018-09-26 | J18SM4935A | 4.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2565 | 2018-09-26 | J18SM4936A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2566 | 2018-09-26 | J18SM4937A | 3.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2567 | 2018-09-26 | J18SM4938A | 3.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2568 | 2018-09-27 | J18SM4939A | 3.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2569 | 2018-09-27 | J18SM4940A | 2.4 | 0.0006 | 0.002 | Resolve unit | FLEX 30:70 Mix | 43 | 73
2570 | 2018-09-27 | J18SM4942A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 70
2571 | 2018-09-27 | J18SM4943A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2572 | 2018-09-26 | J18SM4930A | 5.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 73
2573 | 2018-09-27 | J18SM4941A | 2.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2574 | 2018-09-27 | J18SM4944A | 1.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2575 | 2018-09-27 | J18SM4945A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2576 | 2018-09-27 | J18SM4946A | 1.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2577 | 2018-09-27 | J18SM4947A | 1.2 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2578 | 2018-09-27 | J18SM4948A | 1.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 73
2579 | 2018-09-27 | J18SM4949A | 3.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2580 | 2018-09-27 | J18SM4950A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2581 | 2018-09-27 | J18SM4951A | 4.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2582 | 2018-09-27 | J18SM4952A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2583 | 2018-09-27 | J18SM4953A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2584 | 2018-09-27 | J18SM4954A | 3.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2585 | 2018-09-27 | J18SM4955A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2586 | 2018-09-27 | J18SM4961A | 2.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 70
2587 | 2018-09-27 | J18SM4962A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2588 | 2018-09-29 | J18SM4975A | 1.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2589 | 2018-09-29 | J18SM4979A | 4.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2590 | 2018-09-29 | J18SM4982A | 3.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2591 | 2018-09-29 | J18SM4984A | 1.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 49
2592 | 2018-09-29 | J18SM4987A | 5.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2593 | 2018-09-29 | J18SM4988A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2594 | 2018-09-29 | J18SM4990A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2595 | 2018-09-29 | J18SM4991A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2596 | 2018-09-29 | J18SM4996A | 2.3 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2597 | 2018-09-29 | J18SM4999A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2598 | 2018-09-29 | J18SM4000A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2599 | 2018-09-29 | J18SM4003A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2600 | 2018-09-28 | J18SM4956A | 2.4 | 0.0005 | 0.002 | Resolve unit | FLEX 30:70 Mix | 48 | 75
2601 | 2018-09-28 | J18SM4957A | 2.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2602 | 2018-09-28 | J18SM4958A | 2.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2603 | 2018-09-28 | J18SM4959A | 1.4 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 73
2604 | 2018-09-28 | J18SM4963A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2605 | 2018-09-28 | J18SM4964A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2606 | 2018-09-28 | J18SM4965A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 72
2607 | 2018-09-28 | J18SM4966A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2608 | 2018-09-28 | J18SM4967A | 3.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 73
2609 | 2018-09-28 | J18SM4968A | 4.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2610 | 2018-09-28 | J18SM4969A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2611 | 2018-09-28 | J18SM4970A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2612 | 2018-09-28 | J18SM4971A | 5.1 | 0.0011 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2613 | 2018-09-28 | J18SM4972A | 2.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
2614 | 2018-09-28 | J18SM4973A | 2.1 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2615 | 2018-09-28 | J18SM4974A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2616 | 2018-09-28 | J18SM4976A | 4.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2617 | 2018-09-28 | J18SM4977A | 3.3 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2618 | 2018-09-28 | J18SM4978A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
2619 | 2018-09-28 | J18SM4980A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2620 | 2018-09-28 | J18SM4981A | 4.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 78
2621 | 2018-09-28 | J18SM4983A | 4.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2622 | 2018-10-01 | J18SM7577A | 5.2 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2623 | 2018-10-01 | J18SM7578A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2624 | 2018-10-01 | J18SM7579A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2625 | 2018-10-01 | J18SM7580A | 4.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2626 | 2018-10-01 | J18SM7581A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2627 | 2018-10-01 | J18SM7582A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2628 | 2018-10-01 | J18SM7583A | 2.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2629 | 2018-10-01 | J18SM7585A | 2.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2630 | 2018-10-01 | J18SM7586A | 3.3 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2631 | 2018-10-01 | J18SM7587A | 5.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2632 | 2018-10-01 | J18SM7588A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2633 | 2018-10-01 | J18SM7589A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2634 | 2018-10-01 | J18SM7590A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2635 | 2018-10-01 | J18SM7594A | 1.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2636 | 2018-10-01 | J18SM7595A | 1.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2637 | 2018-10-01 | J18SM7596A | 2.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2638 | 2018-10-01 | J18SM7597A | 3.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2639 | 2018-10-01 | J18SM7598A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2640 | 2018-10-01 | J18SM4960A | 1.3 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 50 | 76
2641 | 2018-09-29 | J18SM4985A | 1.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2642 | 2018-10-01 | J18SM4986A | 5.1 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2643 | 2018-09-29 | J18SM4992A | 3.4 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 76
2644 | 2018-09-29 | J18SM4993A | 3.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 79
2645 | 2018-10-01 | J18SM4994A | 3.1 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
2646 | 2018-10-01 | J18SM4995A | 2.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2647 | 2018-09-29 | J18SM4997A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2648 | 2018-09-29 | J18SM4998A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2649 | 2018-10-01 | J18SM4001A | 4.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2650 | 2018-10-01 | J18SM4002A | 3.3 | NA | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 73
2651 | 2018-09-29 | J18SM4004A | 4.4 | 0.0012 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2652 | 2018-10-01 | J18SM4005A | 5.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2653 | 2018-10-01 | J18SM4006A | 1.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2654 | 2018-10-01 | J18SM4007A | 1.4 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2655 | 2018-10-01 | J18SM4008A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2656 | 2018-10-02 | J18SM7584A | 2.2 | 0.0016 | 0.002 | Resolve unit | FLEX 30:70 Mix | 46 | 78
2657 | 2018-10-02 | J18SM7591A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2658 | 2018-10-02 | J18SM7592A | 4.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2659 | 2018-10-02 | J18SM7593A | 1.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2660 | 2018-10-02 | J18SM7599A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2661 | 2018-10-02 | J18SM7600A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2662 | 2018-10-02 | J18SM7601A | 3.3 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2663 | 2018-10-02 | J18SM7603A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2664 | 2018-10-02 | J18SM7604A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2665 | 2018-10-02 | J18SM7605A | 5.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2666 | 2018-10-02 | J18SM7607A | 3.4 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2667 | 2018-10-02 | J18SM7608A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2668 | 2018-10-02 | J18SM7610A | 1.4 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2669 | 2018-10-02 | J18SM7612A | 1.1 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2670 | 2018-10-02 | J18SM7614A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2671 | 2018-10-02 | J18SM7606A | 5.2 | 0.0019 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2672 | 2018-10-02 | J18SM7609A | 2.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2673 | 2018-10-02 | J18SM7611A | 2.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2674 | 2018-10-02 | J18SM7613A | 2.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2675 | 2018-10-02 | J18SM7615A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2676 | 2018-10-02 | J18SM7616A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2677 | 2018-10-02 | J18SM7617A | 4.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2678 | 2018-10-02 | J18SM4989A | 3.2 | 0.0002 | 0.002 | Resolve unit | FLEX 30:70 Mix | 43 | 78
2679 | 2018-10-03 | K18SM4009A | 5.1 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2680 | 2018-10-03 | K18SM4010A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2681 | 2018-10-02 | K18SM4011A | 4.1 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2682 | 2018-10-03 | K18SM4012A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2683 | 2018-10-03 | K18SM4013A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2684 | 2018-10-03 | K18SM4014A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2685 | 2018-10-03 | K18SM4015A | 4.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2686 | 2018-10-03 | K18SM4016A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2687 | 2018-10-03 | K18SM4017A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2688 | 2018-10-03 | K18SM4018A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2689 | 2018-10-03 | K18SM4020A | 5.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2690 | 2018-10-03 | K18SM4022A | 3.4 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 80
2691 | 2018-10-03 | K18SM4024A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2692 | 2018-10-03 | K18SM4025A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 79
2693 | 2018-10-04 | K18SM4019A | 5.2 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2694 | 2018-10-03 | K18SM4023A | 4.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2695 | 2018-10-04 | K18SM4026A | 3.2 | 0.0009 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2696 | 2018-10-04 | K18SM4027A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2697 | 2018-10-04 | K18SM4028A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2698 | 2018-10-04 | K18SM4029A | 3.4 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2699 | 2018-10-04 | K18SM4030A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2700 | 2018-10-04 | K18SM4032A | 4.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2701 | 2018-10-04 | K18SM4033A | 2.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2702 | 2018-10-04 | K18SM4034A | 2.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2703 | 2018-10-04 | K18SM4035A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2704 | 2018-10-04 | K18SM4036A | 2.1 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2705 | 2018-10-04 | K18SM4038A | 1.4 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2706 | 2018-10-04 | K18SM4039A | 1.3 | 0.0018 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2707 | 2018-10-04 | K18SM4043A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2708 | 2018-10-04 | K18SM4021A | 5.1 | 0.0013 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 72
2709 | 2018-10-05 | K18SM4031A | 3.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2710 | 2018-10-05 | K18SM4037A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2711 | 2018-10-05 | K18SM4041A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2712 | 2018-10-05 | K18SM4042A | 3.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2713 | 2018-10-05 | K18SM4044A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2714 | 2018-10-05 | K18SM4045A | 2.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2715 | 2018-10-05 | K18SM4048A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2716 | 2018-10-05 | K18SM4049A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2717 | 2018-10-05 | K18SM4050A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2718 | 2018-10-05 | K18SM4051A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2719 | 2018-10-05 | K18SM4052A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2720 | 2018-10-05 | K18SM4053A | 3.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2721 | 2018-10-05 | K18SM4054A | 5.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 73
2722 | 2018-10-05 | K18SM4055A | 5.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 78
2723 | 2018-10-05 | K18SM4056A | 1.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2724 | 2018-10-05 | K18SM4057A | 1.3 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2725 | 2018-10-05 | K18SM4058A | 1.2 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2726 | 2018-10-05 | K18SM4059A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2727 | 2018-10-05 | K18SM4060A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2728 | 2018-10-05 | K18SM4061A | 2.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2729 | 2018-10-06 | K18SM4064A | 4.1 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2730 | 2018-10-06 | K18SM4066A | 1.4 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2731 | 2018-10-06 | K18SM4069A | 2.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2732 | 2018-10-06 | K18SM4071A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2733 | 2018-10-06 | K18SM4072A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2734 | 2018-10-06 | K18SM4046A | 3.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2735 | 2018-10-06 | K18SM4047A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2736 | 2018-10-06 | K18SM4062A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2737 | 2018-10-06 | K18SM4063A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2738 | 2018-10-06 | K18SM4065A | 1.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2739 | 2018-10-06 | K18SM4067A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2740 | 2018-10-06 | K18SM4070A | 2.4 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2741 | 2018-10-06 | K18SM4073A | 4.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2742 | 2018-10-06 | K18SM4074A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2743 | 2018-10-06 | K18SM4077A | 3.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 79
2744 | 2018-10-06 | K18SM4080A | 4.2 | 0.0007 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2745 | 2018-10-06 | K18SM4082A | 4.1 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2746 | 2018-10-08 | K18SM4040A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2747 | 2018-10-06 | K18SM4068A | 2.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2748 | 2018-10-06 | K18SM4075A | 4.4 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2749 | 2018-10-08 | K18SM4076A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 76
2750 | 2018-10-08 | K18SM4078A | 4.1 | 0.001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2751 | 2018-10-08 | K18SM4079A | 4.3 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 74
2752 | 2018-10-08 | K18SM4081A | 4.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2753 | 2018-10-08 | K18SM4083A | 3.3 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2754 | 2018-10-08 | K18SM4084A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2755 | 2018-10-08 | K18SM4085A | 3.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 73
2756 | 2018-10-08 | K18SM4087A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2757 | 2018-10-08 | K18SM4089A | 1.2 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 75
2758 | 2018-10-08 | K18SM4090A | 1.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 77
2759 | 2018-10-08 | K18SM4091A | 3.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2760 | 2018-10-08 | K18SM4092A | 3.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 76
2761 | 2018-10-08 | K18SM4093A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2762 | 2018-10-08 | K18SM4094A | 5.1 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2763 | 2018-10-08 | K18SM4096A | 2.2 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2764 | 2018-10-08 | K18SM4097A | 2.1 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2765 | 2018-10-08 | K18SM4098A | 2.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2766 | 2018-10-08 | K18SM4099A | 1.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2767 | 2018-10-08 | K18SM4100A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2768 | 2018-10-08 | K18SM4101A | 4.2 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2769 | 2018-10-08 | K18SM4102A | 2.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2770 | 2018-10-08 | K18SM4103A | 4.1 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2771 | 2018-10-08 | K18SM4104A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2772 | 2018-10-08 | K18SM4105A | 3.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2773 | 2018-10-08 | K18SM4110A | 4.3 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 75
2774 | 2018-10-09 | K18SM4106A | 3.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2775 | 2018-10-08 | K18SM4107A | 4.4 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2776 | 2018-10-09 | K18SM4108A | 4.1 | 0.0008 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 73
2777 | 2018-10-08 | K18SM4109A | 5.1 | 0.0017 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
2778 | 2018-10-09 | K18SM4111A | 4.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 72
2779 | 2018-10-09 | K18SM4112A | 4.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 48 | 75
2780 | 2018-10-09 | K18SM4113A | 3.3 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 50 | 76
2781 | 2018-10-09 | K18SM4114A | 3.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2782 | 2018-10-09 | K18SM4115A | 3.1 | 0.0018 | 0.002 | Pass on first test | FLEX 30:70 Mix | 46 | 78
2783 | 2018-10-09 | K18SM4116A | 2.4 | 0.0006 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 79
2784 | 2018-10-09 | K18SM4118A | 2.2 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 49 | 77
2785 | 2018-10-09 | K18SM4119A | 2.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2786 | 2018-10-09 | K18SM4120A | 1.3 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2787 | 2018-10-09 | K18SM4121A | 1.2 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2788 | 2018-10-09 | K18SM4123A | 4.4 | NA | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 75
2789 | 2018-10-09 | K18SM4124A | 5.1 | 0.0016 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 77
2790 | 2018-10-09 | K18SM4125A | 4.2 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 74
2791 | 2018-10-09 | K18SM4126A | 4.1 | 0.0015 | 0.002 | Pass on first test | FLEX 30:70 Mix | 42 | 77
2792 | 2018-10-09 | K18SM4127A | 4.3 | 0.0002 | 0.002 | Pass on first test | FLEX 30:70 Mix | 43 | 76
2793 | 2018-10-09 | K18SM4128A | 3.4 | 0.0001 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 72
2794 | 2018-10-09 | K18SM4129A | 3.3 | 0.0004 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2795 | 2018-10-09 | K18SM4130A | 3.2 | 0.0005 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74
2796 | 2018-10-09 | K18SM4131A | 3.1 | 0.0003 | 0.002 | Pass on first test | FLEX 30:70 Mix | 39 | 74

## Fail
Alliance CarryOver Test Monitoring
Unit | Date | S/N | Test Station (station. Test port) | Carryover % (<0.002%) | Current Status | Failure Mode | Action to Resolve | Retest | Replace Blank | Replace Injector | Failure Mode | Failure Summary | Total | Final Test Station | No. Of UUT's
Ex | 2018-04-01 |  | X.X | 0.0025 | Resolved | Abnormal Graph |  |  |  |  | High Blank | Abnormal Graph | 15 | 1.1 | 13
1 | 2018-04-07 | D18SM4962A | 5.2 | 0.0008 | Resolved | Abnormal Graph | REPLACE BLANK (9) |  | 1 |  | High Carryover % | Post Blank No Value | 4 | 1.2 | 25
2 | 2018-04-07 | D18SM4966A | 5.4 | 0.0002 | Resolved | Abnormal Graph | REPLACE BLANK (9) |  | 1 |  | Abnormal Graph | High Blank | 287 | 1.3 | 31
3 | 2018-04-09 | D18SM4972A | 1.3 | NA | Resolved | Abnormal Graph | REPLACE BLANK (9) |  | 1 |  | Post Blank no value | Retention Time Drifted | 1 | 1.4 | 31
4 | 2018-04-06 | C18SM7898A | 3.1 | 0.0021 | Resolved | Abnormal Graph | REPLACE BLANK (9) |  | 1 |  | Retention Time Drifted | High Carryover % | 221 | 2.1 | 27
5 | 2018-04-06 | C18SM7903A | 2.1 | NA | Resolved | Post Blank no value | REPLACE BLANK (9) |  | 1 |  | Status |  |  | 2.2 | 25
6 | 2018-04-06 | C18SM7904A | 1.1 | 0.0017 | Resolved | Abnormal Graph | REPLACE BLANK (9) |  | 1 |  | Resolved |  |  | 2.3 | 33
7 | 2018-04-06 | C18SM7906A | 1.4 | 0.0006 | Resolved | Abnormal Graph | 1. REPLACE BLANK (9) 2. Change tester from 1.2 to tester 1.4 | 1 | 1 |  | Troubleshooting |  |  | 2.4 | 27
8 | 2018-04-06 | C18SM7908A | 3.2 | 0.001 | Resolved | Abnormal Graph | 1. REPLACE BLANK (9) 2. Change tester from 2.2 to tester 3.2 | 1 | 1 |  |  |  |  | 3.1 | 31
9 | 2018-04-06 | C18SM7912A | 3.3 | 0.0032 | Resolved | High Carryover % | REPLACE BLANK (9) |  | 1 |  |  |  |  | 3.2 | 35
10 | 2018-04-06 | C18SM7916A | 4.3 | NA | Resolved | Post Blank no value | REPLACE BLANK (9) |  | 1 |  |  |  |  | 3.3 | 35
11 | 2018-04-07 | C18SM7913A | 3.3 | 0.0011 | Resolved | High Blank | 1. REPLACE BLANK (9), 2. Change tester from 3.4 to tester 3.3 3. Change tester from 3.3 to 5.1 4. Replace Seal (P/n: WAT 033276) | 1 | 2 | 1 |  |  |  | 3.4 | 24
12 | 2018-04-10 | C18SM7907A | 4.3 | 0.0032 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replaced Blank (9) 5. Replaced Blank (9) 6. Replaced Blank (9) 7. Replaced Seal (P/n: WAT 033276) |  | 6 | 1 |  |  |  | 4.1 | 29
13 | 2018-04-11 | D18SM4004A | 3.2 | NA | Resolved | Abnormal Graph | REPLACE BLANK (9) |  | 1 |  |  |  |  | 4.2 | 30
14 | 2018-04-11 | D18SM4999A | 1.3 | NA | Resolved | Post Blank no value | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Retest 4. Retest 5. Retest | 3 | 2 |  |  |  |  | 4.3 | 22
15 | 2018-04-10 | D18SM4980A | 3.4 | 0.0008 | Resolved | Abnormal Graph | 1. REPLACE BLANK (9), 2. Replace Seal (P/n: WAT 033276) |  | 1 | 1 |  |  |  | 4.4 | 25
16 | 2018-04-09 | D18SM4971A | 1.4 | NA | Resolved | Post Blank no value | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) |  | 3 |  |  |  |  | 5.1 | 24
17 | 2018-04-07 | D18SM4963A | 5.3 | 0.0006 | Resolved | Abnormal Graph | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2 |  |  |  |  | 5.2 | 27
18 | 2018-04-12 | D18SM4003A | 1.2 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2 |  |  |  |  | 5.3 | 18
19 | 2018-04-11 | D18SM4992A | 3.1 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2 |  |  |  |  | 5.4 | 11
20 | 2018-04-13 | D18SM4021A | 2.3 | NA | Resolved | Retention Time Drifted | 1. Replaced Blank (9) 2. Replaced GPV p/n: WAT270524 |  | 1 |  |  |  |  | 6.1 | 2
21 | 2018-04-09 | C18SM7905A | 3.3 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2 |  |  |  |  | 6.2 | 1
22 | 2018-04-16 | D18SM7933A | 4.1 | 0.0067 | Resolved | High Carryover % | 1. Replace Seal (P/n: WAT 033276) |  |  | 1 |  |  |  | 6.3 | 0
23 | 2018-04-14 | D18SM7923A | 4.2 | 0.0013 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Seal (P/n WAT033276) | 1 | 1 | 1 |  |  |  | 6.4 | 0
24 | 2018-04-16 | D18SM7924A | 1.2 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Seal (P/n WAT03276) 3. Replaced Blank (9) | 1 | 1 | 1 |  |  |  | Total | 526
25 | 2018-04-14 | D18SM7927A | 2.1 | 0.0003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Seal (P/n WAT03276) |  | 1 | 1
26 | 2018-04-16 | D18SM7929A | 2.4 | 0.003 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Seal (P/n WAT03276) |  | 1 | 1
27 | 2018-04-16 | D18SM7931A | 4.4 | 0.0003 | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276) | 1 |  | 1
28 | 2018-04-16 | D18SM7932A | 3.1 | 0.0007 | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276) | 1 |  | 1
29 | 2018-04-16 | D18SM7935A | 3.4 | 0.0005 | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276) | 1 |  | 1
30 | 2018-04-16 | D18SM7940A | 1.4 | NA | Resolved | High Blank | 1. Retest 2. Replaced Seal (P/n WAT03276) | 1 |  | 1
31 | 2018-04-16 | D18SM7944A | 5.3 | 0.005 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Retest | 1 | 1
32 | 2018-04-16 | D18SM4045A | 4.3 | 0.0147 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Retest | 1 | 1
33 | 2018-04-17 | D18SM4048A | 2.2 | 0.0035 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
34 | 2018-04-18 | D18SM7919A | 3.4 | 0.002 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) | 1 |  | 1
35 | 2018-04-17 | D18SM7934A | 2.3 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector seal (P/n:WAT033276) 4. Replaced injector (WAT279003915) from s/n:15J9D055 to 15J9D087 | 1 | 1 | 2
36 | 2018-04-17 | D18SM7943A | 3.1 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915) | 1 | 1 | 1
37 | 2018-04-18 | D18SM7945A | 1.4 | 0.0009 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915) | 1 | 1 | 1
38 | 2018-04-18 | D18SM4047A | 5.1 | NA | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
39 | 2018-04-17 | D18SM4049A | 4.1 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector seal (P/n:WAT033276) | 1 | 1 | 1
40 | 2018-04-17 | D18SM4050A | 3.2 | 0.0002 | Resolved | High Blank | 1. Retest 2. Realigned and retest Injector (P/n WAT279003915) | 1 |  | 1
41 | 2018-04-17 | D18SM4056A | 2.4 | NA | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915) | 1 | 1 | 1
42 | 2018-04-18 | D18SM4063A | 3.2 | 0.0002 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
43 | 2018-04-18 | D18SM4065A | 2.3 | 0.0424 | Resolved | High Carryover % | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
44 | 2018-04-18 | D18SM4053A | 4.2 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915) | 1 | 1 | 1
45 | 2018-04-19 | D18SM4057A | 2.1 | 0.0051 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
46 | 2018-04-18 | D18SM4059A | 4.4 | 0.012 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
47 | 2018-04-19 | D18SM4069A | 5.2 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
48 | 2018-04-19 | D18SM4075A | 2.3 | 0.0036 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915) | 1 | 1 | 1
49 | 2018-04-19 | D18SM4076A | 3.2 | 0.008 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
50 | 2018-04-19 | D18SM4078A | 3.4 | 0.0002 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
51 | 2018-04-19 | D18SM4081A | 3.1 | 0.0047 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
52 | 2018-04-16 | D18SM4042A | 3.3 | 0.0013 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Seal (P/n WAT03276) 4. Change tester from tester 3.3 to tester 1.3 | 2 | 1 | 1
53 | 2018-04-19 | D18SM4074A | 2.2 | 0.0063 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3. Replaced tubing p/n 430001336 as tubing deformed during realignment of injector | 1 |  | 1
54 | 2018-04-20 | D18SM4079A | 4.1 | 0.0063 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3. Replaced tubing p/n 430001336 as tubing deformed during realignment of injector |  | 1 | 1
55 | 2018-04-19 | D18SM4084A | 2.4 | 0.0481 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
56 | 2018-04-20 | D18SM4090A | 4.2 | 0.0088 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
57 | 2018-04-20 | D18SM4092A | 2.3 | 0.005 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
58 | 2018-04-20 | D18SM4094A | 3.4 | 0.0009 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
59 | 2018-04-21 | D18SM4089A | 3.3 | 0.0056 | Resolved | High Carryover % | 1. Retest 2. Realigned and retest Injector (P/n WAT279003915) | 1 |  | 1
60 | 2018-04-24 | D18SM7947A | 3.4 | 0.0006 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
61 | 2018-04-24 | C18SM7900A | 4.3 | NA | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) 2. Change tester from tester 4.1 to tester 4.3 3. Realigned and retest Injector (P/n WAT279003915) 4. Realigned and retest Injector (P/n WAT279003915) | 1 |  | 3
62 | 2018-04-23 | D18SM4101A | 2.4 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
63 | 2018-04-23 | D18SM4108A | 3.1 | 0.0014 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
64 | 2018-04-24 | D18SM4110A | 4.2 | 0.001 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
65 | 2018-04-24 | D18SM4119A | 1.3 | 0.0059 | Resolved | High Carryover % | 1. Retest | 1
66 | 2018-04-24 | D18SM4120A | 2.3 | 0.001 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
67 | 2018-04-25 | D18SM4083A | 1.2 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3.Realigned and retest Injector (P/n WAT279003915) |  | 1 | 2
68 | 2018-04-25 | D18SM4113A | 4.4 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3.Realigned and retest Injector (P/n WAT279003915) |  | 1 | 2
69 | 2018-04-24 | D18SM4117A | 5.2 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
70 | 2018-04-25 | D18SM4123A | 1.4 | 0.0011 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
71 | 2018-04-24 | D18SM4126A | 4.1 | 0.0003 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) |  |  | 1
72 | 2018-04-25 | D18SM4127A | 5.3 | 0.0003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
73 | 2018-04-25 | D18SM4134A | 1.4 | NA | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
74 | 2018-04-26 | D18SM4137A | 1.2 | NA | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
75 | 2018-04-26 | D18SM4140A | 1.3 | 0.0007 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
76 | 2018-04-25 | D18SM7953A | 2.4 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) 3.Realigned and retest Injector (P/n WAT279003915) |  | 1 | 2
77 | 2018-04-25 | D18SM7956A | 3.4 | 0.0006 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
78 | 2018-04-26 | D18SM4109A | 3.2 | 0.0027 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Realigned and retest Injector (P/n WAT279003915) |  | 1 | 1
79 | 2018-04-26 | D18SM4114A | 2.1 | NA | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and retest Injector (P/n WAT279003915) | 1 | 1 | 1
80 | 2018-04-25 | D18SM4124A | 5.1 | 0.0003 | Resolved | High Blank | 1. Realigned and retest Injector (P/n WAT279003915) 2. Realigned and retest Injector (P/n WAT279003915) |  |  | 2
81 | 2018-04-26 | D18SM4146A | 5.3 | 0.0073 | Resolved | High Carryover % | 1. Retest | 1
82 | 2018-04-26 | D18SM4155A | 4.2 | 0.001 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
83 | 2018-04-26 | D18SM4142A | 2.2 | 0.0021 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
84 | 2018-04-26 | D18SM4143A | 1.4 | 0.0013 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
85 | 2018-04-26 | D18SM4135A | 5.4 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and Retest Injector (279003915) |  | 1 | 1
86 | 2018-04-27 | D18SM4145A | 4.1 | 0.0009 | Resolved | High Blank | 1. Replaced Blank (9) 2. Realigned and Retest Injector (279003915) |  | 1 | 1
87 | 2018-04-27 | D18SM4148A | 2.4 | 0.0009 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915) |  | 2 | 1
88 | 2018-04-27 | D18SM4152A | 5.2 | 0.0021 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
89 | 2018-04-27 | D18SM4154A | 4.3 | NA | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
90 | 2018-04-27 | D18SM4160A | 3.1 | 0.0002 | Resolved | High Blank | 1. Retest | 1
91 | 2018-04-27 | D18SM4138A | 3.3 | 0.0055 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915) |  | 2 | 1
92 | 2018-04-25 | D18SM7954A | 4.2 | 0.0012 | Resolved | High Blank | 1. Replaced Blank (9) 3. Realigned and Retest Injector (279003915) |  | 1 | 1
93 | 2018-04-28 | D18SM7957A | 2.1 | 0.0084 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
94 | 2018-04-30 | D18SM7964A | 5.2 | 0.0027 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D075 to 16J9D105 | 1 | 1 | 1
95 | 2018-04-30 | D18SM7967A | 2.2 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915) | 1 | 1 | 1
96 | 2018-04-28 | D18SM7968A | 2.1 | 0.0005 | Resolved | High Blank | 1. Realigned and Retest Injector (279003915) |  |  | 1
97 | 2018-04-30 | D18SM7974A | 5.1 | 0.0003 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
98 | 2018-04-28 | D18SM4139A | 2.3 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915) | 1 | 1 | 1
99 | 2018-04-28 | D18SM4161A | 3.2 | 0.0001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Realigned and Retest Injector (279003915) | 1 | 2 | 1
100 | 2018-04-30 | D18SM4163A | 4.2 | NA | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Realigned and Retest Injector (279003915) | 1 | 1 | 1
101 | 2018-04-30 | D18SM4164A | 4.4 | 0.0046 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D053 to 16J9D104 | 1 | 1 | 1
102 | 2018-04-30 | D18SM4175A | 2.3 | 0.0041 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
103 | 2018-04-30 | D18SM4178A | 4.3 | 0.0174 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
104 | 2018-05-02 | D18SM4176A | 2.1 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced injector (WAT279003915) from s/n:16J9D093 to 17J9D006 |  | 1 | 1
105 | 2018-04-30 | D18SM4181A | 4.2 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
106 | 2018-05-02 | D18SM4183A | 2.2 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D093 to 17J9D006 | 1 | 1 | 1
107 | 2018-05-02 | D18SM4185A | 1.2 | 0.0006 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
108 | 2018-05-03 | D18SM4189A | 2.2 | 0.0018 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9) | 1 | 2
109 | 2018-05-02 | D18SM4192A | 4.2 | 0.0046 | Resolved | High Carryover % | 1. Retest | 1
110 | 2018-05-02 | D18SM4196A | 1.3 | 0.0006 | Resolved | High Blank | 1. Retest | 1
111 | 2018-05-02 | D18SM4197A | 4.1 | 0.0025 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
112 | 2018-05-02 | D18SM4199A | 3.3 | 0.001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
113 | 2018-05-02 | D18SM4201A | 3.2 | 0.0016 | Resolved | High Blank | 1. Retest | 1
114 | 2018-04-30 | D18SM4168A | 5.4 | 0.0004 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced injector (WAT279003915) from s/n:16J9D060 to 17J9D028 | 1 | 1 | 1
115 | 2018-05-03 | D18SM4195A | 2.4 | 0.0004 | Resolved | High Blank | 1. Retest | 1
116 | 2018-05-03 | D18SM4211A | 4.1 | 0.0018 | Resolved | Abnormal Graph | 1. Retest 2. Replaced Blank (9) | 1 | 1
117 | 2018-05-03 | D18SM4214A | 1.3 | 0.0057 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
118 | 2018-05-03 | D18SM4217A | 2.2 | 0.0002 | Resolved | High Blank | 1. Retest | 1
119 | 2018-04-28 | D18SM7961A | 4.1 | 0.001 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
120 | 2018-05-04 | E18SM7978A | 1.4 | 0.0033 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (() | 1 | 1
121 | 2018-05-04 | E18SM7984A | 4.4 | 0.0004 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
122 | 2018-05-04 | E18SM7987A | 1.3 | 0.0009 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
123 | 2018-05-05 | E18SM7989A | 2.4 | 0.0005 | Resolved | High Blank | 1. Retest | 1
124 | 2018-05-03 | E18SM4207A | 4.3 | 0.0012 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
125 | 2018-04-30 | D18SM4180A | 5.3 | 0.0109 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replace Injector (279003915) from SN 16J9D086 to J17J9D055 | 1 | 1 | 1
126 | 2018-05-03 | E18SM4216A | 3.3 | 0.0048 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) 3. Replace Injector (279003915) from SN 17J9D049 to J17J9D026 | 1 | 1 | 1
127 | 2018-05-03 | E18SM4218A | 3.1 | 0.0002 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replace Injector (279003915) from SN 17J9D027 to J17J9D034 | 1 | 2 | 1
128 | 2018-05-04 | E18SM4222A | 4.1 | 0.002 | Resolved | High Blank | 1. Retest | 1
129 | 2018-05-04 | E18SM4223A | 4.4 | 0.0027 | Resolved | High Carryover % | 1. Retest | 1
130 | 2018-05-04 | E18SM4225A | 5.1 | 0.0057 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1
131 | 2018-05-05 | E18SM4230A | 3.2 | 0.0001 | Resolved | High Blank | 1. Retest | 1
132 | 2018-05-05 | E18SM4235A | 1.4 | 0.0021 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
133 | 2018-05-05 | E18SM4238A | 4.1 | NA | Resolved | Abnormal Graph | 1. Retest 2. Replaced Blank (9) | 1 | 1
134 | 2018-05-05 | E18SM4232A | 2.3 | 0.0004 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replaced Blank (9) | 1 | 3
135 | 2018-05-05 | E18SM4233A | 5.2 | 0.0005 | Resolved | High Blank | 1. Retest 2. Replaced Blank (9) | 1 | 1
136 | 2018-05-05 | E18SM4236A | 3.3 | NA | Resolved | Abnormal Graph | 1. Retest | 1
137 | 2018-05-07 | E18SM4240A | 5.1 | 0.0003 | Resolved | High Blank | 1. Retest | 1
138 | 2018-05-07 | E18SM4247A | 3.3 | NA | Resolved | High Blank | 1. Retest | 1
139 | 2018-05-07 | E18SM4248A | 3.4 | 0.0032 | Resolved | High Carryover % | 1. Retest | 1
140 | 2018-05-07 | E18SM4249A | 2.4 | 0.0034 | Resolved | High Carryover % | 1. Retest 2. Replaced Blank (9) | 1 | 1
141 | 2018-05-07 | E18SM4250A | 1.4 | 0.006 | Resolved | High Carryover % | 1. Retest | 1
142 | 2018-05-07 | E18SM4252A | 5.2 | 0.0009 | Resolved | High Blank | 1. Retest | 1
143 | 2018-05-07 | E18SM4257A | 3.3 | 0.0049 | Resolved | High Carryover % | 1. Retest | 1
144 | 2018-05-07 | E18SM4255A | 2.3 | NA | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) 4. Replace Injector (279003915) from SN 17J9D083 to 17J9D028 |  | 3 | 1
145 | 2018-05-08 | E18SM4258A | 3.2 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
146 | 2018-05-08 | E18SM4263A | 1.3 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replace Injector (279003915) from SN 17J0D008 to 17JCD008 |  | 2 | 1
147 | 2018-05-08 | E18SM4266A | 3.4 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
148 | 2018-05-08 | E18SM4269A | 5.3 | 0.0026 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
149 | 2018-05-08 | E18SM4270A | 4.2 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
150 | 2018-05-08 | E18SM4273A | 3.3 | 0.0036 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
151 | 2018-05-08 | E18SM4274A | 3.1 | 0.0042 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
152 | 2018-05-08 | E18SM4275A | 1.4 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
153 | 2018-05-09 | E18SM4283A | 5.1 | 0.0001 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) |  | 3
154 | 2018-05-09 | E18SM4284A | 2.4 | 0.0052 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
155 | 2018-05-09 | E18SM4289A | 4.2 | 0.0044 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
156 | 2018-05-09 | E18SM4295A | 1.4 | 0.0023 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
157 | 2018-05-09 | E18SM4298A | 1.1 | 0.0039 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
158 | 2018-05-09 | E18SM4276A | 1.2 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
159 | 2018-05-10 | E18SM4288A | 3.3 | 0.0001 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
160 | 2018-05-10 | E18SM4301A | 3.2 | 0.0016 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
161 | 2018-05-09 | E18SM4302A | 2.4 | 0.0023 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) |  | 3
162 | 2018-05-10 | E18SM4303A | 2.3 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
163 | 2018-05-10 | E18SM4306A | 3.1 | 0.0007 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) |  | 3
164 | 2018-05-10 | E18SM4311A | 5.2 | 0.0033 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
165 | 2018-05-10 | E18SM4312A | 1.1 | 0.0001 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
166 | 2018-05-10 | E18SM4315A | 5.3 | 0.0095 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
167 | 2018-05-11 | E18SM4231A | 3.3 | 0.0085 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
168 | 2018-05-11 | E18SM4307A | 2.1 | 0.0036 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
169 | 2018-05-11 | E18SM4309A | 4.1 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
170 | 2018-05-11 | E18SM4310A | 4.2 | 0.0006 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) |  | 3
171 | 2018-05-11 | E18SM4314A | 5.1 | 0.0008 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) |  | 3
172 | 2018-05-11 | E18SM4316A | 1.2 | 0.0003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
173 | 2018-05-11 | E18SM4322A | 1.3 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
174 | 2018-05-11 | E18SM4329A | 3.2 | 0.0031 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
175 | 2018-05-12 | E18SM4331A | 1.4 | 0.001 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
176 | 2018-05-12 | E18SM4333A | 3.3 | 0.0032 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
177 | 2018-05-12 | E18SM4286A | 4.3 | 0.0003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
178 | 2018-05-12 | E18SM4325A | 3.1 | 0.0025 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
179 | 2018-05-12 | E18SM4335A | 1.1 | 0.0023 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
180 | 2018-05-12 | E18SM4336A | 4.2 | 0.0141 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
181 | 2018-05-12 | E18SM4338A | 3.2 | 0.001 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
182 | 2018-05-14 | E18SM4337A | 2.4 | 0.0035 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
183 | 2018-05-14 | E18SM4339A | 2.1 | 0.0054 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
184 | 2018-05-14 | E18SM4343A | 5.3 | 0.0024 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
185 | 2018-05-14 | E18SM4344A | 1.3 | 0.0012 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) (new Vial and New Cap) |  | 3
186 | 2018-05-14 | E18SM4345A | 3.3 | 0.0027 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
187 | 2018-05-14 | E18SM4349A | 1.1 | 0.0015 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) (new Vial and New Cap) |  | 3
188 | 2018-05-14 | E18SM4353A | 3.2 | 0.0025 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
189 | 2018-05-14 | E18SM4224A | 2.3 | 0.0047 | Resolved | High Carryover % | 1. Replaced Blank (9) 2. Replaced Blank (9) 3. Replaced Blank (9) (new Vial and New Cap) |  | 3
190 | 2018-05-12 | E18SM4287A | 3.4 | 0.0034 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
191 | 2018-05-14 | E18SM4355A | 2.4 | 0.0025 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
192 | 2018-05-14 | E18SM4361A | 5.2 | 0.0054 | Resolved | High Carryover % | 1. Replaced Blank (9) |  | 1
193 | 2018-05-14 | E18SM7988A | 1.2 | 0.0003 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9) |  | 2
194 | 2018-05-15 | E18SM4364A | 3.2 | 0.0014 | Resolved | High Blank | 1. Replaced Blank (9) |  | 1
195 | 2018-05-15 | E18SM4279A | 4.1 | 0.0024 | Resolved | High Carryover % | 1. Replaced New Cap |  | 1
196 | 2018-05-15 | E18SM4317A | 4.4 | 0.0095 | Resolved | High Carryover % | 1. Replaced New Cap 2. Replaced New Cap, vial and Blank (9) 3. Replaced Injector (279003915) from SN 1959D011 to SN 20J9D035 |  | 2 | 1
197 | 2018-05-15 | E18SM4367A | 2.4 | 0.0014 | Resolved | High Blank | 1. Replaced Blank (9) 2. Replaced Blank (9)and vial cap |  | 2
198 | 2018-05-15 | E18SM4378A | 4.3 | 0.0002 | Resolved | High Blank | 1. Replaced Cap, vial and Mix 20:80
199 | 2018-05-16 | E18SM4354A | 1.4 | 0.0005 | Resolved | High Blank | 1. Replaced Blank (9) Flex Mix 30:70 2. Replaced Blank (9) Flex Mix 30:70 3. Replaced Injector (279003915) From SN 19JD044 To : SN 20J90019 Flex Mix 30:70 4. Replaced Blank (9) MC Mix 20:80 |  | 3 | 1
200 | 2018-05-16 | E18SM4359A | 2.1 | 0.003 | Resolved | High Carryover % | 1. Replaced Blank (9) Flex Mix 30:70 2. Replaced Blank (9) Flex Mix 30:70 3. Replaced Injector (279003915) From SN 19JD061 To : SN 20J90020 Flex Mix 30:70 4. Replaced Blank (9) MC Mix 20:80 |  | 3 | 1
201 | 2018-05-16 | E18SM4372A | 3.1 | 0.0024 | Resolved | High Carryover % | 1. Replaced Vial, Cap and Blank (9) Flex Mix 30:70 2. Replaced New Cap Flex Mix 30:70 3. Replaced Injector (279003915) From SN 20JD007 To : SN 20J90030 Flex Mix 30:70 4. Replaced Blank (9) MC Mix 20:80 |  | 3 | 1
202 | 2018-05-16 | E18SM4377A | 3.2 | 0.0002 | Resolved | Abnormal Graph | 1. Replaced Blank (9) Flex Mix 20:80 2. Replaced Blank (9) Flex Mix 20:80 3. Replaced Blank (9) MC Mix 20:80 |  | 3
203 | 2018-05-16 | E18SM4383A | 2.3 | 0.0004 | Resolved | High Blank | 1. Replaced Blank (9) Flex Mix 20:80 2. Retest with Test Group 3 and 4 Replaced Blank (9) MC Mix 20:80 |  | 2
204 | 2018-05-16 | E18SM4385A | 2.2 | 0.0002 | Resolved | High Blank | 1. Replaced Blank (9) Flex Mix 20:80 2. Replaced New vial, Cap and Blank (9) MC Mix 20:80 |  | 2
205 | 2018-05-16 | E18SM4384A | 3.4 | 0.0035 | Resolved | High Carryover % | 1. Retest with Test Group 3 and 4 Replaced Vial, Cap and Blank (9) MC Mix 20:80 2. Retest with Test Group 3 and 4 Replaced New Vial, Cap and Blank (9) MC Mix 20:80 |  | 2
206 | 2018-05-16 | E18SM4389A | 4.2 | 0.0012 | Resolved | High Blank | 1. Retest with Test Group 3 and 4 Replaced Vial, Cap and Blank (9) MC Mix 20:80 |  | 1
207 | 2018-05-16 | E18SM7991A | 3.3 | 0.0039 | Resolved | High Carryover % | 1. Retest with Test Group 3 and 4 Replaced Vial, Cap and Blank (9) MC Mix 20:80 |  | 1
208 | 2018-05-16 | E18SM7001A | 4.1 | 0.001 | Resolved | High Carryover % | 1. Retest with Test Group 3 and 4 Replaced Vial, Cap and Blank (9) MC Mix 20:80 |  | 1
209 | 2018-05-16 | E18SM7997A | 2.2 | 0.0023 | Resolved | High Carryover % | 1. Retest with Test Group 3 and 4 Replaced Vial, Cap and Blank (9) MC Mix 20:80 |  | 1
210 | 2018-05-17 | E18SM4374A | 5.2 | 0.0031 | Resolved | High Carryover % | 1. Retest with Test Group 3 and 4 Replaced Blank (9) MC Mix 20:80 |  | 1
211 | 2018-05-17 | E18SM4381A | 2.4 | 0.0024 | Resolved | High Carryover % | 1. Replace BLANK (9) 2. Replace BLANK (9) 3. Replace New Vial,Cap and BLANK (9) |  | 3
212 | 2018-05-17 | E18SM7998A | 3.1 | 0.0308 | Resolved | High Carryover % | 1. Test Group 3 and 4 and Replace BLANK (9) 2. Test Group 4 and Replace vial, can and BLANK (9) 3. Replace Injector (279003915) from Serial Num 20J9D042 to 20J9D070 |  | 2 | 1
213 | 2018-05-17 | E18SM7999A | 1.2 | 0.0007 | Resolved | High Blank | 1. Test Group 3 and 4 and Replace BLANK (9) 2. Test Group 4 and Replace vial, can and BLANK (9) |  | 2
214 | 2018-05-17 | E18SM7004A | 4.2 | 0.0027 | Resolved | High Carryover % | 1. Test Group 3 and Replace BLANK (9) 2. Test Group 3,4 and Replace vial, cap and BLANK (9) |  | 2
215 | 2018-05-17 | E18SM7005A | 3.2 | 0.003 | Resolved | High Carryover % | 1. Replace BLANK (9) |  | 1
216 | 2018-05-17 | E18SM7007A | 2.1 | 0.0005 | Resolved | High Blank | 1. Replace BLANK (9) |  | 1
217 | 2018-05-17 | E18SM7009A | 3.4 | 0.0005 | Resolved | High Blank | 1. Replace BLANK (9) |  | 1
218 | 2018-05-17 | E18SM7992A | 5.3 | 0.0022 | Resolved | High Carryover % | 1. Replace Blank (9), Cap and Vial |  | 1
219 | 2018-05-18 | E18SM7019A | 1.3 | 0.0013 | Resolved | High Blank | 1. Replace Blank (9) 2. Test Group 3 and 4 Replaced Blank (9) New Vial and cap |  | 2
220 | 2018-05-18 | E18SM7020A | 1.4 | 0.0015 | Resolved | High Blank | 1. Replace BLANK (9) |  | 1
221 | 2018-05-19 | E18SM7022A | 3.2 | 0.0032 | Resolved | High Carryover % | 1. Replace Cap and Blank (9) Test Group 3 and 4 2. Replaced Blank (9) New Vial Test Group 3 and 4 |  | 2
222 | 2018-05-19 | E18SM7023A | 3.1 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Blank (9) Test Group 3 2. Replaced Blank (9) New Vial Test Group 3 and 4 |  | 2
223 | 2018-05-19 | E18SM7037A | 4.2 | NA | Resolved | High Blank | 1. Replace Cap and Blank (9) |  | 2
224 | 2018-05-19 | E18SM4388A | 1.1 | NA | Resolved | High Carryover % | 1. Replace Blank (9) Test Group 3 2. Retest Test Group 3 and 4 3. Replaced Injector (279003915) from SN 20J9D025 to SN 20J9D065 4. Test Group 3 and 4 Replace Vial, Cap and Blank (9) |  | 3 | 1
225 | 2018-05-19 | E18SM7042A | 1.3 | 0.0043 | Resolved | High Carryover % | 1. Replace Cap and Blank (9) Test Group 3 2. Replaced Blank (9) New Vial, new cap Test Group 3 and 4 |  | 2
226 | 2018-05-21 | E18SM7045A | 3.1 | 0.0024 | Resolved | High Carryover % | 1. Replace Cap and Blank (9) Test Group 3 and 4 |  | 1
227 | 2018-05-19 | E18SM7047A | 3.2 | 0.0066 | Resolved | High Carryover % | 1. Replace vial, Cap and Blank (9) Test Group 3 and 4 |  | 1
228 | 2018-05-22 | E18SM4404A | 4.1 | 0.0243 | Resolved | High Carryover % | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 2. Replaced Injector (279003915) from SN 19J9D044 to SN 20J9D0124 |  | 1 | 1
229 | 2018-05-22 | E18SM4409A | 4.4 | 0.0052 | Resolved | High Carryover % | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 2. Replaced Injector (279003915) from SN 20J9D106 to SN 20J9D0135 |  | 1 | 1
230 | 2018-05-22 | E18SM4418A | 3.4 | 0.0007 | Resolved | High Blank | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 |  | 1
231 | 2018-05-22 | E18SM7014A | 2.2 | 0.042 | Resolved | High Carryover % | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 2. Replaced Injector (279003915) from SN 20J9D011 to SN 20J9D0125 |  | 1 | 1
232 | 2018-05-22 | E18SM7017A | 1.2 | 0.0172 | Resolved | High Carryover % | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 2. Replaced Injector (279003915) from SN 20J9D058 to SN 20J9D0126 |  | 1 | 1
233 | 2018-05-23 | E18SM4429A | 4.1 | 0.0127 | Resolved | High Carryover % | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 2. Replaced Injector (279003915) from SN 20J9D131 to SN 20J9D004 |  | 1 | 1
234 | 2018-05-23 | E18SM4431A | 4.4 | 0.003 | Resolved | High Carryover % | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 |  | 1
235 | 2018-05-22 | E18SM7025A | 1.3 | 0.0007 | Resolved | High Blank | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 |  | 1
236 | 2018-05-23 | E18SM4433A | 2.3 | 0.0016 | Resolved | High Blank | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 |  | 1
237 | 2018-05-23 | E18SM4435A | 4.2 | 0.0015 | Resolved | High Blank | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 |  | 1
238 | 2018-05-23 | E18SM4442A | 3.2 | 0.0021 | Resolved | High Carryover % | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 |  | 1
239 | 2018-05-24 | E18SM4471A | 3.3 | NA | Resolved | Abnormal Graph | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
240 | 2018-05-24 | E18SM4440A | 5.1 | 0.0028 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
241 | 2018-05-24 | E18SM4462A | 2.4 | 0.0438 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
242 | 2018-05-24 | E18SM4415A | 1.4 | 0.0015 | Resolved | High Blank | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 2. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 20J9D113 to SN 20J9D018 |  | 2 | 1
243 | 2018-05-23 | E18SM4447A | 2.1 | 0.0005 | Resolved | High Blank | 1. Replace 2 Caps and 2 vials and Blank (9) Test Group 3 and 4 |  | 1
244 | 2018-05-24 | E18SM4454A | 2.3 | 0.0043 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 2. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 3. Replaced Injector (279003915) from SN 21J9D014 to SN 21J9D041 |  | 2 | 1
245 | 2018-05-25 | E18SM4456A | 1.4 | NA | Resolved | High Blank | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
246 | 2018-05-25 | E18SM4457A | 4.3 | 0.0069 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 2. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 3. Replaced Injector (279003915) from SN 21J9D017 to SN 21J9D054 |  | 2 | 1
247 | 2018-05-24 | E18SM4459A | 3.4 | 0.0003 | Resolved | High Blank | 1. Replace 2 Caps and 2 Blank (9) Test Group 3 and 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D020 to SN 21J9D040 |  | 2 | 1
248 | 2018-05-25 | E18SM4466A | 4.4 | 0.0017 | Resolved | High Blank | 1. Replace 2 Caps and 2 Blank (9) Test Group 3 and 4 |  | 1
249 | 2018-05-25 | E18SM4470A | 5.2 | 0.0001 | Resolved | High Blank | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
250 | 2018-05-25 | E18SM4472A | 4.1 | 0.0154 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
251 | 2018-05-25 | E18SM4473A | 1.3 | 0.1883 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
252 | 2018-05-25 | E18SM4475A | 5.1 | 0.0035 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
253 | 2018-05-25 | E18SM4479A | 5.4 | 0.0002 | Resolved | High Blank | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
254 | 2018-05-25 | E18SM4487A | 3.2 | NA | Resolved | High Blank | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 |  | 1
255 | 2018-05-26 | E18SM4398A | 2.3 | 0.0004 | Resolved | High Blank | 1. Replace Caps and Blank (9) Test Group 3 and 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 20J9D097 to SN 20J9D128 |  | 2 | 1
256 | 2018-05-25 | E18SM4464A | 2.4 | 0.0045 | Resolved | High Carryover % | 1. Replace Caps and Blank (9) Test Group 3 and 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D026 to SN 21J9D061 |  | 2 | 1
257 | 2018-05-25 | E18SM4477A | 3.2 | NA | Resolved | High Blank | 1. Replace Caps and Blank (9) Test Group 3 and 4 |  | 1
258 | 2018-05-26 | E18SM4489A | 1.1 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
259 | 2018-05-26 | E18SM4495A | 1.2 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
260 | 2018-05-25 | E18SM4483A | 2.1 | NA | Resolved | High Blank | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D048 to SN 20J9D087 |  | 2 | 1
261 | 2018-05-26 | E18SM4498A | 4.2 | 0.0119 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 2. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 3. Replaced Injector (279003915) from SN 21J9D088 to SN 20J9D096 |  | 2 | 1
262 | 2018-05-26 | E18SM4501A | 5.2 | 0.0026 | Resolved | High Carryover % | 1. Replaced Post Blank and cap Restest with Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D068 to SN 20J9D094 |  | 2 | 1
263 | 2018-05-28 | E18SM4510A | 2.3 | 0.0042 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
264 | 2018-05-28 | E18SM4514A | 4.3 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Blank (9) Test Group 3 and 4 |  | 1
265 | 2018-05-28 | E18SM4525A | 3.1 | 0.0033 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
266 | 2018-05-25 | E18SM4484A | 4.4 | 0.0001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D049 to SN 20J9D062 |  | 2 | 1
267 | 2018-05-28 | E18SM4521A | 5.2 | 0.004 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 2
268 | 2018-05-28 | E18SM4523A | 2.4 | 0.0025 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
269 | 2018-05-30 | E18SM4526A | 3.4 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
270 | 2018-05-30 | E18SM4530A | 1.4 | 0.001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
271 | 2018-05-30 | E18SM4533A | 2.3 | 0.0007 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
272 | 2018-05-30 | E18SM4540A | 3.2 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
273 | 2018-05-31 | E18SM4519A | 1.3 | 0.025 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
274 | 2018-05-31 | E18SM4532A | 1.2 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
275 | 2018-05-31 | E18SM4393A | 5.4 | 0.0069 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D090 to SN 20J9D122 |  | 2 | 1
276 | 2018-05-31 | E18SM4444A | 4.1 | 0.0008 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D002 to SN 20J9D012 |  | 2 | 1
277 | 2018-05-31 | E18SM4460A | 1.4 | 0.0022 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D021 to SN 20J9D042 |  | 2 | 1
278 | 2018-05-31 | E18SM4508A | 5.1 | 0.0042 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
279 | 2018-05-31 | E18SM4518A | 5.1 | 0.7805 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
280 | 2018-05-31 | E18SM4535A | 3.1 | 0.0044 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 2. Replace Cap and Post Blank (9) Test Group 3 3. Replaced Injector (279003915) from SN 21J9D011 to SN 20J9D017 |  | 1 | 1
281 | 2018-05-31 | E18SM4544A | 3.2 | 0.0163 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
282 | 2018-05-31 | E18SM7040A | 5.3 | 0.004 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D086 to SN 20J9D119 |  | 2 | 1
283 | 2018-05-31 | E18SM4549A | 4.2 | 0.0074 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D111 to SN 20J9D048 |  | 2 | 1
284 | 2018-05-31 | E18SM4524A | 4.4 | 0.0011 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
285 | 2018-05-30 | E18SM4527A | 3.3 | 0.0001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 2
286 | 2018-06-01 | E18SM7051A | 5.1 | 0.0036 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
287 | 2018-06-01 | E18SM7058A | 2.1 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
288 | 2018-06-01 | E18SM7060A | 2.3 | 0.0026 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
289 | 2018-06-01 | E18SM7062A | 3.1 | 0.0029 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
290 | 2018-06-01 | E18SM7059A | 1.2 | 0.0052 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 2. Replace Cap and Post Blank (9) Test Group 3 3. Replaced Injector (279003915) from SN 21J9D020 to SN 20J9D044 |  | 1 | 1
291 | 2018-05-18 | E18SM7015A | 5.1 | 0.0036 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 2. Replace Cap and Post Blank (9) Test Group 3 and |  | 1
292 | 2018-06-02 | E18SM7048A | 1.1 | 0.0014 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 2. Replace Cap and Post Blank (9) Test Group 3 3. Replaced Injector (279003915) from SN 21J9D005 to SN 20J9D057 |  | 2 | 1
293 | 2018-06-02 | E18SM7057A | 5.2 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 2
294 | 2018-06-02 | E18SM4062A | 3.2 | 0.0059 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D023 to SN 20J9D047 |  | 1 | 1
295 | 2018-06-02 | E18SM7067A | 1.3 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 1
296 | 2018-05-28 | E18SM4515A | 4.1 | 0.0042 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D078 to SN 20J9D112 |  | 2 | 1
297 | 2018-06-04 | E18SM4529A | 4.1 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 21J9D103 to SN 20J9D015 |  | 2 | 1
298 | 2018-06-04 | F18SM4573A | 5.3 | 0.0109 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and 4 2. Replace Cap and Post Blank (9) Test Group 3 and 4 |  | 2
299 | 2018-06-04 | E18SM7028A | 1.4 | 0.0021 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 20J9D071 to SN 20J9D141 |  | 2 | 1
300 | 2018-06-04 | E18SM7069A | 4.2 | 0.0028 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
301 | 2018-06-04 | E18SM7072A | 3.4 | 0.0048 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D033 to SN 22J9D065 |  | 2 | 1
302 | 2018-06-04 | E18SM7073A | 5.4 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 |  | 2
303 | 2018-06-04 | E18SM7075A | 4.1 | NA | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D034 to SN 22J9D060 |  | 2 | 1
304 | 2018-06-04 | E18SM4395A | 3.2 | 0.0078 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
305 | 2018-06-04 | E18SM4506A | 1.2 | 0.0022 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D073 to SN 22J9D104 |  | 2 | 1
306 | 2018-06-02 | E18SM4558A | 3.1 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D038 to SN 22J9D082 |  | 2 | 1
307 | 2018-06-05 | E18SM4560A | 2.5 | 0.0076 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D042 to SN 22J9D083 |  | 2 | 1
308 | 2018-06-04 | E18SM4565A | 3.3 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D051 to SN 22J9D078 |  | 2 | 1
309 | 2018-06-04 | E18SM4576A | 5.2 | 0.0031 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Post Blank (9) Test Group 3 and 4 |  | 2
310 | 2018-06-05 | E18SM4588A | 2.3 | NA | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
311 | 2018-06-05 | F18SM4581A | 4.3 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Post Blank (9) Test Group 3 and 4 |  | 2
312 | 2018-06-05 | F18SM4586A | 1.1 | NA | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 |  | 2
313 | 2018-06-05 | E18SM4427A | 2.3 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 |  | 2
314 | 2018-06-06 | F18SM4570A | 4.1 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 |  | 2
315 | 2018-06-06 | F18SM4591A | 3.1 | 0.0021 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D084 to SN 22J9D100 |  | 2 | 1
316 | 2018-06-06 | F18SM4596A | 4.3 | 0.0281 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Post Blank (9) Test Group 3 and 4 |  | 2
317 | 2018-06-06 | F18SM4598A | 5.1 | 0.0034 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Post Blank (9) Test Group 3 and 4 |  | 2
318 | 2018-06-06 | F18SM4601A | 1.2 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
319 | 2018-06-06 | F18SM4611A | 1.4 | 0.0035 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
320 | 2018-06-07 | F18SM4417A | 2.3 | 0.0013 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
321 | 2018-06-06 | F18SM4580A | 4.4 | 0.0023 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
322 | 2018-06-06 | F18SM4612A | 3.3 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
323 | 2018-06-06 | F18SM4613A | 4.2 | 0.0014 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
324 | 2018-06-07 | F18SM4624A | 1.3 | 0.0016 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
325 | 2018-06-06 | F18SM4628A | 2.1 | 0.01 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
326 | 2018-06-07 | F18SM4630A | 3.4 | 0.0028 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
327 | 2018-06-08 | F18SM4639A | 3.1 | 0.0008 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
328 | 2018-06-08 | F18SM7077A | 5.4 | 0.0013 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
329 | 2018-06-07 | F18SM7080A | 2.1 | 0.0001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Caps and Blank (9) Test Group 3 and 4 3. Replaced Injector (279003915) from SN 22J9D125 to SN 22J9D031 |  | 2 | 1
330 | 2018-06-07 | F18SM7085A | 4.1 | 0.0047 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4
331 | 2018-06-08 | F18SM7087A | 5.1 | 0.0001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
332 | 2018-06-08 | F18SM7088A | 4.4 | 0.008 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
333 | 2018-06-08 | F18SM7097A | 5.3 | 0.0018 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
334 | 2018-06-08 | F18SM7092A | 2.3 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
335 | 2018-06-09 | F18SM7102A | 3.3 | 0.0064 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
336 | 2018-06-09 | F18SM4644A | 3.2 | 0.0151 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
337 | 2018-06-11 | E18SM4547A | 4.4 | 0.0051 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
338 | 2018-06-08 | F18SM4638A | 1.4 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
339 | 2018-06-11 | F18SM4645A | 4.1 | 0.002 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
340 | 2018-06-11 | F18SM4649A | 5.2 | 0.002 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
341 | 2018-06-09 | F18SM4650A | 5.3 | 0.0063 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
342 | 2018-06-11 | F18SM4653A | 4.4 | NA | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
343 | 2018-06-11 | F18SM4658A | 4.3 | 0.0028 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
344 | 2018-06-11 | F18SM4659A | 3.2 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
345 | 2018-06-11 | F18SM7093A | 2.2 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
346 | 2018-05-26 | E18SM4492A | 3.4 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 3. Replaced Injector (279003915) from SN 21J9D058 to SN 21J9D080 |  | 2 | 1
347 | 2018-06-11 | F18SM4662A | 3.3 | 0.0151 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
348 | 2018-06-11 | F18SM4672A | 5.4 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
349 | 2018-06-12 | E18SM4546A | 5.4 | 0.0081 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
350 | 2018-06-12 | F18SM4688A | 2.2 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
351 | 2018-06-12 | F18SM4692A | 3.3 | 0.0013 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
352 | 2018-06-13 | F18SM4693A | 3.1 | 0.0028 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
353 | 2018-06-13 | E18SM4503A | 4.4 | 0.0024 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
354 | 2018-06-13 | F18SM4669A | 1.2 | 0.0023 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
355 | 2018-06-13 | F18SM4683A | 1.3 | 0.0028 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
356 | 2018-06-12 | F18SM4694A | 5.3 | 0.0032 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
357 | 2018-06-13 | F18SM4700A | 4.3 | 0.0023 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
358 | 2018-06-13 | F18SM4702A | 2.3 | 0.0016 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
359 | 2018-06-13 | F18SM4705A | 3.2 | 0.1507 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
360 | 2018-06-13 | F18SM4707A | 3.3 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
361 | 2018-06-12 | F18SM4690A | 4.1 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
362 | 2018-06-13 | F18SM4699A | 1.1 | 0.0139 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
363 | 2018-06-13 | F18SM4706A | 2.2 | 0.0045 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 3. Replaced Injector (279003915) from SN 23J9D086 to SN 23J9D112 |  | 2 | 1
364 | 2018-06-13 | F18SM7107A | 5.3 | 0.0053 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
365 | 2018-06-14 | F18SM7109A | 4.1 | 0.0063 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
366 | 2018-06-14 | F18SM7112A | 4.2 | 0.0063 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
367 | 2018-06-14 | F18SM7114A | 2.1 | 0.0028 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
368 | 2018-06-13 | F18SM7115A | 1.3 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
369 | 2018-06-14 | F18SM7126A | 4.4 | 0.0011 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
370 | 2018-06-14 | F18SM7084A | 5.2 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
371 | 2018-06-14 | F18SM7123A | 3.4 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
372 | 2018-06-14 | F18SM7132A | 1.3 | 0.0056 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
373 | 2018-06-16 | F18SM7122A | 3.2 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
374 | 2018-06-16 | F18SM7128A | 5.1 | 0.0014 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
375 | 2018-06-18 | F18SM7119A | 5.12 | 0.0062 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
376 | 2018-06-19 | F18SM7138A | 5.2 | 0.0149 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
377 | 2018-06-19 | F18SM4713A | 2.1 | 0.0001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
378 | 2018-06-18 | F18SM4714A | 3.2 | 0.0105 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
379 | 2018-06-19 | F18SM4719A | 4.4 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
380 | 2018-06-19 | F18SM4721A | 4.2 | 0.0017 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
381 | 2018-06-19 | F18SM4723A | 3.4 | 0.0099 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
382 | 2018-06-19 | F18SM4725A | 2.4 | 0.0021 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
383 | 2018-06-19 | F18SM4727A | 3.1 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
384 | 2018-06-19 | F18SM4729A | 4.3 | 0.0024 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
385 | 2018-06-19 | F18SM4741A | 4.4 | 0.0226 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
386 | 2018-06-19 | F18SM4744A | 2.3 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
387 | 2018-06-19 | F18SM4745A | 1.4 | 0.0039 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
388 | 2018-06-20 | F18SM4752A | 1.2 | 0.0007 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
389 | 2018-06-20 | F18SM4753A | 2.1 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
390 | 2018-06-20 | F18SM4755A | 5.2 | 0.001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
391 | 2018-06-21 | F18SM7140A | 1.1 | 0.0024 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
392 | 2018-06-20 | F18SM4767A | 5.2 | 0.0017 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
393 | 2018-06-20 | F18SM4768A | 1.2 | 0.0009 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
394 | 2018-06-21 | F18SM4769A | 5.1 | 0.0026 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
395 | 2018-06-20 | F18SM4773A | 2.4 | 0.0037 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
396 | 2018-06-20 | F18SM4774A | 2.1 | 0.0042 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
397 | 2018-06-21 | F18SM4775A | 1.3 | 0.0027 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
398 | 2018-06-21 | F18SM4763A | 3.1 | 0.0041 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 3. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 3
399 | 2018-06-21 | F18SM4771A | 4.2 | 0.0042 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
400 | 2018-06-21 | F18SM4780A | 5.2 | 0.001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
401 | 2018-06-21 | F18SM4787A | 2.3 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
402 | 2018-06-21 | F18SM4789A | 1.3 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
403 | 2018-06-21 | F18SM4760A | 2.2 | 0.0016 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
404 | 2018-06-22 | F18SM4632A | 6.1 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 3. Replaced Injector (279003915) from SN 22J9D114 to SN 25J9D024 |  | 2 | 1
405 | 2018-06-20 | F18SM7125A | 5.4 | 0.0025 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
406 | 2018-06-22 | F18SM7143A | 4.1 | 0.0008 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
407 | 2018-06-22 | F18SM7146A | 3.3 | 0.0023 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
408 | 2018-06-25 | F18SM7186A | 6.2 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
409 | 2018-06-26 | F18SM4811A | 1.2 | 0.0008 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
410 | 2018-06-26 | F18SM4813A | 2.3 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
411 | 2018-06-25 | F18SM4806A | 2.1 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
412 | 2018-06-26 | F18SM4820A | 2.2 | 0.0057 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
413 | 2018-06-26 | F18SM4833A | 4.2 | 0.0113 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
414 | 2018-06-26 | F18SM4833A | 1.3 | 0.0001 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
415 | 2018-06-23 | F18SM7171A | 2.1 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
416 | 2018-06-25 | F18SM4799A | 2.3 | 0.0002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
417 | 2018-06-27 | F18SM4815A | 5.1 | 0.0168 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
418 | 2018-06-26 | F18SM4826A | 2.4 | 0.0158 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
419 | 2018-06-27 | F18SM4836A | 4.3 | 0.0025 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
420 | 2018-06-29 | F18SM7232A | 3.3 | 0.0021 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
421 | 2018-06-30 | F18SM4859A | 2.4 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
422 | 2018-06-30 | F18SM4861A | 3.1 | 0.0003 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
423 | 2018-06-22 | F18SM7165A | 4.4 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
424 | 2018-07-03 | G18SM4884A | 2.3 | 0.0024 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
425 | 2018-07-03 | G18SM4894A | 2.4 | 0.002 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
426 | 2018-07-03 | G18SM4900A | 1.4 | 0.0132 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
427 | 2018-07-03 | G18SM4909A | 5.3 | NA | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
428 | 2018-07-03 | G18SM4897A | 5.4 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
429 | 2018-07-04 | G18SM4902A | 2.2 | 0.0007 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
430 | 2018-07-04 | G18SM4904A | 1.2 | 0.0043 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
431 | 2018-07-03 | G18SM4905A | 5.2 | 0.0031 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
432 | 2018-07-04 | G18SM4913A | 2.3 | 0.0022 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
433 | 2018-07-05 | G18SM4922A | 5.1 | 0.0031 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Retest Passed |  | 2
434 | 2018-07-04 | G18SM4928A | 1.3 | 0.0006 | Resolved | High Blank | 1. Replace Post Blank (9) Retest Passed |  | 1
435 | 2018-07-05 | F18SM4855A | 4.4 | 0.0011 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 2. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 2
436 | 2018-07-05 | G18SM4931A | 3.3 | 0.0037 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
437 | 2018-07-05 | G18SM4932A | 5.2 | 0.0022 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
438 | 2018-07-05 | G18SM4935A | 5.3 | 0.0035 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
439 | 2018-07-05 | G18SM4942A | 1.3 | 0.0005 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
440 | 2018-07-06 | G18SM4955A | 5.1 | 0.0008 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
441 | 2018-07-06 | G18SM4956A | 5.2 | 0.0004 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
442 | 2018-07-06 | G18SM4957A | 2.2 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
443 | 2018-07-07 | G18SM4953A | 4.3 | 0.0013 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
444 | 2018-07-07 | G18SM4971A | 3.1 | 0.0024 | Resolved | High Carryover % | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
445 | 2018-07-09 | G18SM4993A | 1.4 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
446 | 2018-07-09 | G18SM4003A | 5.2 | 0.0006 | Resolved | High Blank | 1. Replace Cap and Post Blank (9) Test Group 3 and Test Group 4 |  | 1
447 | 2018-07-09 | G18SM4982A | 6.1 | 0.0011 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
448 | 2018-07-10 | G18SM4001A | 3.3 | 0.0009 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
449 | 2018-07-11 | G18SM7279A | 3.4 | 0.0021 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
450 | 2018-07-12 | G18SM7259A | 3.2 | 0.0122 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
451 | 2018-07-11 | G18SM7280A | 4.1 | 0.0006 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
452 | 2018-07-14 | G18SM4048A | 4.3 | 0.0009 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
453 | 2018-07-16 | G18SM4056A | 5.2 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
454 | 2018-07-16 | G18SM4060A | 2.2 | 0.0001 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
455 | 2018-07-16 | G18SM4072A | 5.1 | 0.0183 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
456 | 2018-07-17 | G18SM7255A | 3.3 | 0.001 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
457 | 2018-07-18 | G18SM4082A | 4.1 | 0.0022 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
458 | 2018-07-18 | G18SM4096A | 1.1 | 0.0003 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
459 | 2018-07-19 | G18SM4116A | 3.4 | 0.0034 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
460 | 2018-07-20 | G18SM4143A | 4.1 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
461 | 2018-07-19 | G18SM4123A | 4.3 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
462 | 2018-07-19 | G18SM4134A | 1.3 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
463 | 2018-07-19 | G18SM4136A | 1.4 | 0.01 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
464 | 2018-07-20 | G18SM4145A | 3.4 | 0 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
465 | 2018-07-20 | G18SM4151A | 1.3 | 0.0001 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
466 | 2018-07-20 | G18SM4158A | 4.2 | 0.0026 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
467 | 2018-07-23 | G18SM4165A | 2.1 | 0.0102 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
468 | 2018-07-24 | G18SM7293A | 2.2 | 0.0062 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
469 | 2018-07-24 | G18SM7301A | 4.2 | 0.0009 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
470 | 2018-07-25 | G18SM7316A | 3.1 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
471 | 2018-07-26 | G18SM7303A | 1.3 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
472 | 2018-07-26 | G18SM4206A | 2.1 | 0.0003 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
473 | 2018-07-27 | G18SM4212A | 3.4 | NA | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
474 | 2018-07-30 | G18SM4234A | 4.3 | 0.0098 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
475 | 2018-07-30 | G18SM4246A | 4.4 | 0.0005 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
476 | 2018-07-31 | G18SM4221A | 2.2 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
477 | 2018-08-01 | G18SM4285A | 2.2 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
478 | 2018-08-03 | G18SM4304A | 3.3 | 0.0172 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
479 | 2018-08-03 | G18SM4310A | 3.2 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
480 | 2018-08-03 | G18SM4312A | 3.1 | 0.0005 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
481 | 2018-08-03 | G18SM4315A | 2.2 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
482 | 2018-08-04 | H18SM4323A | 4.3 | 0.0007 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
483 | 2018-08-04 | H18SM4325A | 5.1 | NA | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
484 | 2018-08-07 | H18SM4328A | 1.3 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
485 | 2018-08-10 | H18SM4356A | 4.2 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
486 | 2018-08-14 | G18SM4386A | 5.1 | 0.0036 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
487 | 2018-08-14 | H18SM4392A | 3.1 | NA | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
488 | 2018-08-15 | G18SM4386A | 4.2 | 0.0118 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
489 | 2018-08-20 | H18SM7397A | 2.4 | 0.0005 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
490 | 2018-08-20 | H18SM7376A | 4.4 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
491 | 2018-08-21 | H18SM7403A | 1.4 | 0.0062 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
492 | 2018-08-23 | H18SM7416A | 5.2 | 0.0009 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
493 | 2018-08-23 | H18SM4484A | 1.2 | 0.0017 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
494 | 2018-08-23 | H18SM4490A | 2.4 | 0.0003 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
495 | 2018-08-24 | H18SM4495A | 5.1 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
496 | 2018-08-27 | H18SM4523A | 4.4 | 0.0067 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
497 | 2018-08-25 | H18SM4507A | 5.3 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
498 | 2018-08-25 | H18SM4530A | 3.1 | 0.0014 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
499 | 2018-08-27 | H18SM4531A | 1.2 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
500 | 2018-08-28 | H18SM4557A | 3.3 | 0.0261 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
501 | 2018-08-29 | H18SM4554A | 1.4 | 0.0087 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
502 | 2018-08-29 | H18SM4500A | 2.3 | 0.0001 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
503 | 2018-08-30 | H18SM7449A | 2.1 | 0.0037 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
504 | 2018-08-31 | H18SM4628A | 2.3 | 0.0014 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
505 | 2018-09-03 | H18SM4631A | 3.2 | 0.0001 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
506 | 2018-09-04 | H18SM7459A | 2.2 | 0.0065 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
507 | 2018-09-06 | J18SM7472A | 1.2 | 0.0003 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
508 | 2018-09-05 | J18SM7477A | 1.4 | 0.0008 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
509 | 2018-09-06 | J18SM7491A | 1.1 | 0.0011 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
510 | 2018-09-10 | J18SM4713A | 4.2 | 0.0007 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
511 | 2018-09-10 | J18SM4721A | 2.3 | 0.0013 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
512 | 2018-09-12 | H18SM4777A | 1.2 | 0.0028 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
513 | 2018-09-13 | J18SM4805A | 1.4 | 0.0004 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
514 | 2018-09-13 | J18SM4780A | 2.1 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
515 | 2018-09-13 | J18SM4781A | 2.2 | 0.0059 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
516 | 2018-09-17 | J18SM4839A | 1.2 | 0.0003 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
517 | 2018-09-17 | J18SM4840A | 1.3 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
518 | 2018-09-17 | J18SM4841A | 1.4 | 0.0008 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
519 | 2018-09-18 | J18SM7507A | 3.3 | 0.0114 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
520 | 2018-09-18 | J18SM7536A | 2.1 | 0.0006 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
521 | 2018-09-21 | J18SM4866A | 3.3 | 0.0028 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
522 | 2018-09-26 | J18SM4940A | 2.4 | 0.0001 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
523 | 2018-09-28 | J18SM4956A | 2.4 | 0.0095 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
524 | 2018-09-29 | H18SM4960A | 1.3 | 0.0002 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
525 | 2018-10-01 | J18SM4994A | 3.1 | 0.0001 | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
526 | 2018-09-29 | J18SM4994A | 3.3 | NA | Resolved | High Blank | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
527 | 2018-10-01 | J18SM7584A | 2.2 | 0.0045 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 1
528 | 2018-10-01 | J18SM4989A | 3.2 | 0.003 | Resolved | High Carryover % | 1. Replace Post Blank (9) Test Group 3 and Test Group 4 2. Replace Post Blank (9) Test Group 3 and Test Group 4 |  | 2
Note on J2: 1.Replace Blank use same cap, same vial, change solution 2. Replace blank change the cap same vial change solution 3. Same cap replace vial change solution)
Note on D3: sgycha: Test station 1; Port 3
Note on G3: sgycha: drop downlist; state if otherwise

## FM Mon
Count of S/N | Column Labels
Row Labels | Abnormal Graph | High Blank | Post Blank no value | Retention Time Drifted | High Carryover % | Grand Total
2018-04-01
2018-04-06 | 4 |  | 2 |  | 1 | 7
2018-04-07 | 3 | 1 |  |  |  | 4
2018-04-09 | 1 | 1 | 1 |  |  | 3
2018-04-10 | 1 |  |  |  | 1 | 2
2018-04-11 | 1 | 1 | 1 |  |  | 3
2018-04-12 |  | 1 |  |  |  | 1
2018-04-13 |  |  |  | 1 |  | 1
2018-04-14 |  | 2 |  |  |  | 2
2018-04-16 |  | 6 |  |  | 4 | 10
2018-04-17 |  | 5 |  |  | 1 | 6
2018-04-18 |  | 5 |  |  | 2 | 7
2018-04-19 |  | 2 |  |  | 6 | 8
2018-04-20 |  | 1 |  |  | 3 | 4
2018-04-21 |  |  |  |  | 1 | 1
2018-04-23 |  | 2 |  |  |  | 2
2018-04-24 |  | 6 |  |  | 1 | 7
2018-04-25 |  | 9 |  |  |  | 9
2018-04-26 |  | 6 |  |  | 3 | 9
2018-04-27 |  | 4 |  |  | 2 | 6
2018-04-28 |  | 4 |  |  | 1 | 5
2018-04-30 |  | 5 |  |  | 5 | 10
2018-05-02 |  | 6 |  |  | 2 | 8
2018-05-03 | 1 | 5 |  |  | 2 | 8
2018-05-04 |  | 3 |  |  | 3 | 6
2018-05-05 | 2 | 4 |  |  | 1 | 7
2018-05-07 |  | 4 |  |  | 4 | 8
2018-05-08 |  | 5 |  |  | 3 | 8
2018-05-09 |  | 2 |  |  | 5 | 7
2018-05-10 |  | 5 |  |  | 2 | 7
2018-05-11 |  | 5 |  |  | 3 | 8
2018-05-12 |  | 3 |  |  | 5 | 8
2018-05-14 |  | 3 |  |  | 8 | 11
2018-05-15 |  | 3 |  |  | 2 | 5
2018-05-16 | 1 | 4 |  |  | 6 | 11
2018-05-17 |  | 3 |  |  | 6 | 9
2018-05-18 |  | 2 |  |  | 1 | 3
2018-05-19 |  | 2 |  |  | 4 | 6
2018-05-21 |  |  |  |  | 1 | 1
2018-05-22 |  | 2 |  |  | 4 | 6
2018-05-23 |  | 3 |  |  | 3 | 6
2018-05-24 | 1 | 2 |  |  | 3 | 6
2018-05-25 |  | 8 |  |  | 5 | 13
2018-05-26 |  | 4 |  |  | 2 | 6
2018-05-28 |  | 1 |  |  | 5 | 6
2018-05-30 |  | 5 |  |  |  | 5
2018-05-31 |  | 3 |  |  | 9 | 12
2018-06-01 |  | 1 |  |  | 4 | 5
2018-06-02 |  | 4 |  |  | 1 | 5
2018-06-04 |  | 4 |  |  | 7 | 11
2018-06-05 |  | 4 |  |  | 1 | 5
2018-06-06 |  | 4 |  |  | 6 | 10
2018-06-07 |  | 3 |  |  | 2 | 5
2018-06-08 |  | 6 |  |  | 1 | 7
2018-06-09 |  |  |  |  | 3 | 3
2018-06-11 |  | 4 |  |  | 5 | 9
2018-06-12 |  | 3 |  |  | 2 | 5
2018-06-13 |  | 5 |  |  | 7 | 12
2018-06-14 |  | 3 |  |  | 4 | 7
2018-06-16 |  | 2 |  |  |  | 2
2018-06-18 |  |  |  |  | 2 | 2
2018-06-19 |  | 5 |  |  | 6 | 11
2018-06-20 |  | 5 |  |  | 3 | 8
2018-06-21 |  | 4 |  |  | 5 | 9
2018-06-22 |  | 3 |  |  | 1 | 4
2018-06-23 |  | 1 |  |  |  | 1
2018-06-25 |  | 3 |  |  |  | 3
2018-06-26 |  | 4 |  |  | 2 | 6
2018-06-27 |  |  |  |  | 2 | 2
2018-06-29 |  |  |  |  | 1 | 1
2018-06-30 |  | 2 |  |  |  | 2
2018-07-03 |  | 3 |  |  | 3 | 6
2018-07-04 |  | 2 |  |  | 2 | 4
2018-07-05 |  | 2 |  |  | 4 | 6
2018-07-06 |  | 3 |  |  |  | 3
2018-07-07 |  | 1 |  |  | 1 | 2
2018-07-09 |  | 3 |  |  |  | 3
2018-07-10 |  | 1 |  |  |  | 1
2018-07-11 |  | 1 |  |  | 1 | 2
2018-07-12 |  |  |  |  | 1 | 1
2018-07-14 |  | 1 |  |  |  | 1
2018-07-16 |  | 2 |  |  | 1 | 3
2018-07-17 |  | 1 |  |  |  | 1
2018-07-18 |  | 1 |  |  | 1 | 2
2018-07-19 |  | 2 |  |  | 2 | 4
2018-07-20 |  | 3 |  |  | 1 | 4
2018-07-23 |  |  |  |  | 1 | 1
2018-07-24 |  | 1 |  |  | 1 | 2
2018-07-25 |  | 1 |  |  |  | 1
2018-07-26 |  | 2 |  |  |  | 2
2018-07-27 |  | 1 |  |  |  | 1
2018-07-30 |  | 1 |  |  | 1 | 2
2018-07-31 |  | 1 |  |  |  | 1
2018-08-01 |  | 1 |  |  |  | 1
2018-08-03 |  | 3 |  |  | 1 | 4
2018-08-04 |  | 2 |  |  |  | 2
2018-08-07 |  | 1 |  |  |  | 1
2018-08-10 |  | 1 |  |  |  | 1
2018-08-14 |  | 1 |  |  | 1 | 2
2018-08-15 |  |  |  |  | 1 | 1
2018-08-20 |  | 2 |  |  |  | 2
2018-08-21 |  |  |  |  | 1 | 1
2018-08-23 |  | 3 |  |  |  | 3
2018-08-24 |  | 1 |  |  |  | 1
2018-08-25 |  | 2 |  |  |  | 2
2018-08-27 |  | 1 |  |  | 1 | 2
2018-08-28 |  |  |  |  | 1 | 1
2018-08-29 |  | 1 |  |  | 1 | 2
2018-08-30 |  |  |  |  | 1 | 1
2018-08-31 |  | 1 |  |  |  | 1
2018-09-03 |  | 1 |  |  |  | 1
2018-09-04 |  |  |  |  | 1 | 1
2018-09-06 |  | 2 |  |  |  | 2
2018-09-05 |  | 1 |  |  |  | 1
2018-09-10 |  | 2 |  |  |  | 2
2018-09-12 |  |  |  |  | 1 | 1
2018-09-13 |  | 2 |  |  | 1 | 3
2018-09-17 |  | 3 |  |  |  | 3
2018-09-18 |  | 1 |  |  | 1 | 2
2018-09-21 |  |  |  |  | 1 | 1
2018-09-26 |  | 1 |  |  |  | 1
2018-09-28 |  |  |  |  | 1 | 1
2018-09-29 |  | 2 |  |  |  | 2
2018-10-01 |  | 1 |  |  | 2 | 3
Grand Total | 15 | 287 | 4 | 1 | 221 | 528

## Failure V.02
# | Failure Code | Example | Trouble shooting guide
1 | High Carryover % |  | 07 May 2018 Changing BLANKS (9) \ Retest 16 May 2018 Retest Matrix TG 3 and 4 \ Change BLANKS (9) \ Vial and Caps Michael Cotter 20% MeOH mix 80% Ultrapure Waters [Blanks (9)] 23 May 2018 Replaced Post Blank solution, Rinse vial with Ultrapure water (3x) Replaced new Cap, Retest with TG 3 and 4 (Repeat if FAIL Carryover Test) If PASS,Ship If FAIL, Replaced Injector PN (279003915) from SN XXXXXXXX to SN XXXXXXXX If after replace and FAIL again, on hold for Eng'g Review
2 | High Blank |  | 07 May 2018 Changing Blank \ Retest 16 May 2018 Retest Matrix TG 3 and 4 \ Change BLANKS (9) \ Vial and Caps Michael Cotter 20% MeOH mix 80% Ultrapure Waters [Blanks (9)] 23 May 2018 Replaced Post Blank solution, Rinse vial with Ultrapure water (3x) Replaced new Cap, Retest with TG 3 and 4 (Repeat if FAIL Carryover Test) If PASS,Ship If FAIL, Replaced Injector PN (279003915) from SN XXXXXXXX to SN XXXXXXXX
3 | Abnormal Peak |  | 07 May 2018 CHANGING BLANKS \ Retest 16 May 2018 Retest Matrix TG 3 and 4 \ Change BLANKS (9) \ Vial and Caps Michael Cotter 20% MeOH mix 80% Ultrapure Waters [Blanks (9)] 23 May 2018 Replaced Post Blank solution, Rinse vial with Ultrapure water (3x) Replaced new Cap, Retest with TG 3 and 4 (Repeat if FAIL Carryover Test) If PASS,Ship If FAIL, Replaced Injector PN (279003915) from SN XXXXXXXX to SN XXXXXXXX If Pass,Ship If fail, replace blank fail again, on hold for review
4 | Post Blank no data |  | 11 April 2018 Used modified test result with Specification : Post Blank <0.002 (or Blank Peak(s) mising/blank)

## Double Press of Seal
Double Press Evaluation Result
1 | D18SM4162A | Static Leak | Result | Qty
2 | D18SM4163A | High Blank | High Blank | 5
3 | D18SM4164A | High Carryover | High Carryover | 2
4 | D18SM4165A | Pass | Pass | 13
5 | D18SM4166A | Pass
6 | D18SM4167A | Pass
7 | D18SM4168A | High Blank
8 | D18SM4169A | Pass
9 | D18SM4170A | Pass
10 | D18SM4171A | Pass
11 | D18SM7959A | Pass
12 | D18SM7958A | Static Leak
13 | D18SM7957A | High Blank
14 | D18SM7960A | Pass
15 | D18SM7961A | High Blank
16 | D18SM7962A | Pass
17 | D18SM7963A | pass
18 | D18SM7964A | High Blank
19 | D18SM7965A | Pass
20 | D18SM7966A | Pass

## Eval Summary V.05
SN | Evaluation | Cut In Serial/ Control Lot Size | Date | Total Tested | First Pass | Yield
1 | ASP Execution Eval_Hand Tighthen | 24 | 2018-04-20 | 24 | 14 | 58.33%
2 | Double Press Seal | D18SM4162A D18SM7957A | 2018-04-27 | 238 | 132 | 55.46%
3 | Flex Mix 20:80 | E18SM4279A E18SM4382A | 2013-05-15 | 6 | 4 | 66.67%
4 | MC Mix 20:80 | E18SM4354A E18SM7018A | 2018-05-16 | 93 | 55 | 59.14%
5 | WPPL 20:80 Mix | refer to Pass Tab | 2018-05-22 | 53 | 35 | 66.04%
6 | WPPL 20:80 Mix (2nd bottle) | refer to Pass Tab | 2018-05-26 | 72 | 40 | 55.56%
7 | WPPL 20:80 Mix (3nd bottle) Flushing with methanol on seal body | refer to Pass Tab | 2018-06-01 | 81 | 51 | 62.96%
8 | WPPL 20:80 Mix (4th bottle) Flushing with methanol on seal body | refer to Pass Tab | 2018-06-06 | 60 | 39 | 65%
9 | WPPL 20:80 Mix (5th bottle) Flushing with methanol on seal body | refer to Pass Tab | 2018-06-09 | 66 | 48 | 72.73%
10 | WPPL 20:80 Mix (6th bottle) Flushing with methanol on seal body | refer to Pass Tab | 2018-06-13 | 61 | 37 | 60.66%
11 | WPPL 20:80 Mix (6th bottle) Enlargement of Seal Wash Evaluation (WAT033273) | refer to Pass Tab | 2018-06-18 | 17 | 14 | 82.35%
12 | WPPL 20:80 Mix (7th bottle) | refer to Pass Tab | 2018-06-19 | 100 | 67 | 67%
13 | WPPL 20:80 Mix (7th bottle) Enlargement of Seal Wash Evaluation (WAT033273) | refer to Pass Tab | 2018-06-19 | 8 | 3 | 37.5%
14 | WPPL 20:80 Mix (8th bottle) Injector Adjustment Test Recording | refer to Pass Tab | 2018-06-23 | 118 | 105 | 88.98%
15 | WPPL 20:80 Mix (9th bottle) Injector Adjustment Test Recording Stop Methanol Seal Wash rinsing (1st SN G18SM4929A) | refer to Pass Tab | 2018-07-04 | 74 | 57 | 77.03%
16 | WPPL 20:80 Mix (10th bottle) Injector Adjustment Test Recording Change of Post Blank Cap - every Monday | refer to Pass Tab | 2018-07-09 | 97 | 90 | 92.78%
17 | FLEX 30:70Mix Injector Adjustment Test Recording Change of Post Blank Cap - every Monday | refer to Pass Tab | 2018-07-14 | 1370 | 1293 | 94.38%

## FPY
Count of Remarks | Column Labels
Row Labels | Pass on first test | Resolve unit | Grand Total | Row Labels | Sum of Retest | Sum of Replace Blank | Sum of Replace Injector | Date | Pass on first test | Resolve unit | Grand Total | FPY | Target | Date | Retest | Replace Blank | Replace Injector | Resolve Unit
21 sp 2018 | 1 |  | 1 | 2018-04-01 |  |  |  | Apr-18 | 189 | 103 | 292 | 64.73% | 90% | Apr-18 | 52 | 100 | 72 | 103
2018-04-06 | 8 | 7 | 15 | 2018-04-06 | 2 | 7 |  | 2018-05-01 | 247 | 180 | 427 | 57.85% | 90% | 2018-05-02 | 31 | 126 | 73 | 180
2018-04-07 | 9 | 2 | 11 | 2018-04-07 | 1 | 6 | 1 | 2018-06-01 | 352 | 138 | 490 | 71.84% | 90% | 2018-06-01 | 0 | 207 | 14 | 138
2018-04-09 | 9 | 4 | 13 | 2018-04-09 |  | 6 |  | 2018-07-01 | 434 | 56 | 490 | 88.57% | 90% | 2018-07-01 | 0 | 64 | 0 | 56
2018-04-10 | 10 |  | 10 | 2018-04-10 |  | 7 | 2 | 2018-08-01 | 425 | 27 | 452 | 94.03% | 90% | 2018-08-01 | 0 | 32 | 0 | 27
2018-04-11 | 6 | 5 | 11 | 2018-04-11 | 3 | 5 |  | 2018-09-01 | 454 | 24 | 474 | 95.78% | 90% | 2018-09-01 | 0 | 24 | 0 | 23
2018-04-12 | 7 | 2 | 9 | 2018-04-12 |  | 2 |  | 2018-10-01 | 25 | 3 | 28 | 89.29% | 90% | 2018-10-01 | 0 | 4 | 0 | 5
2018-04-13 | 14 |  | 14 | 2018-04-13 |  | 1 |  | 2018-10-02 | 22 | 2 | 24 | 91.67% | 90%
2018-04-14 | 10 | 3 | 13 | 2018-04-14 | 1 | 2 | 2 | 2018-10-03 | 14 | 0 | 14 | 100% | 90%
2018-04-16 | 24 | 7 | 31 | 2018-04-16 | 9 | 5 | 8 | 2018-10-04 | 15 | 0 | 15 | 100% | 90%
2018-04-17 | 4 | 3 | 7 | 2018-04-17 | 5 | 6 | 6 | 2018-10-05 | 20 | 0 | 20 | 100% | 90%
2018-04-18 | 9 | 7 | 16 | 2018-04-18 | 3 | 3 | 7 | 2018-10-06 | 19 | 0 | 19 | 100% | 90%
2018-04-19 | 6 | 8 | 14 | 2018-04-19 | 5 | 6 | 4 | 2018-10-08 | 28 | 0 | 28 | 100% | 90%
2018-04-20 | 10 | 6 | 16 | 2018-04-20 | 1 | 3 | 3 | 2018-10-09 | 21 | 0 | 21 | 100% | 90%
2018-04-21 | 2 | 2 | 4 | 2018-04-21 | 1 |  | 1
2018-04-23 | 10 | 1 | 11 | 2018-04-23 |  | 1 | 2
2018-04-24 | 6 | 7 | 13 | 2018-04-24 | 2 | 3 | 7
2018-04-25 | 11 | 5 | 16 | 2018-04-25 |  | 8 | 12
2018-04-26 | 8 | 13 | 21 | 2018-04-26 | 2 | 9 | 3
2018-04-27 | 5 | 8 | 13 | 2018-04-27 | 1 | 7 | 3
2018-04-28 | 11 | 2 | 13 | 2018-04-28 | 3 | 5 | 3
2018-04-30 | 10 | 11 | 21 | 2018-04-30 | 9 | 10 | 6
2018-05-02 | 11 | 7 | 18 | 2018-05-02 | 6 | 5 | 2
2018-05-03 | 15 | 11 | 26 | 2018-05-03 | 8 | 8 | 2
2018-05-04 | 9 | 6 | 15 | 2018-05-04 | 6 | 3
2018-05-05 | 5 | 5 | 10 | 2018-05-05 | 7 | 6
2018-05-07 | 16 | 11 | 27 | 2018-05-07 | 7 | 4 | 1
2018-05-08 | 10 | 8 | 18 | 2018-05-08 |  | 12 | 1
2018-05-09 | 9 | 6 | 15 | 2018-05-09 |  | 14
2018-05-10 | 8 | 8 | 16 | 2018-05-10 |  | 12
2018-05-11 | 6 | 8 | 14 | 2018-05-11 |  | 15
2018-05-12 | 6 | 8 | 14 | 2018-05-12 |  | 11
2018-05-14 | 10 | 11 | 21 | 2018-05-14 |  | 20
2018-05-15 | 12 | 5 | 17 | 2018-05-15 |  | 6 | 1
2018-05-16 | 9 | 12 | 21 | 2018-05-16 |  | 22 | 3
2018-05-17 | 8 | 9 | 17 | 2018-05-17 |  | 14 | 1
2018-05-18 | 8 | 2 | 10 | 2018-05-18 |  | 4
2018-05-19 | 7 | 4 | 11 | 2018-05-19 |  | 12 | 1
2018-05-21 | 6 | 3 | 9 | 2018-05-21 |  | 1
2018-05-22 | 12 | 5 | 17 | 2018-05-22 |  | 6 | 4
2018-05-23 | 13 | 6 | 19 | 2018-05-23 |  | 6 | 1
2018-05-24 | 14 | 4 | 18 | 2018-05-24 |  | 9 | 3
2018-05-25 | 13 | 12 | 25 | 2018-05-25 |  | 17 | 4
2018-05-26 | 9 | 6 | 15 | 2018-05-26 |  | 10 | 4
2018-05-28 | 10 | 5 | 15 | 2018-05-28 |  | 8 | 1
2018-05-30 | 8 | 6 | 14 | 2018-05-30 |  | 6
2018-05-31 | 13 | 12 | 25 | 2018-05-31 |  | 17 | 6
2018-06-01 | 8 | 4 | 12 | 2018-06-01 |  | 5 | 1
2018-06-02 | 9 | 7 | 16 | 2018-06-02 |  | 8 | 3
2018-06-04 | 17 | 12 | 29 | 2018-06-04 |  | 20 | 6
2018-06-05 | 17 | 6 | 23 | 2018-06-05 |  | 9 | 1
2018-06-06 | 15 | 7 | 22 | 2018-06-06 |  | 17 | 1
2018-06-07 | 14 | 7 | 21 | 2018-06-07 |  | 5 | 1
2018-06-08 | 11 | 7 | 18 | 2018-06-08 |  | 12
2018-06-09 | 7 | 4 | 11 | 2018-06-09 |  | 4
2018-06-11 | 19 | 8 | 27 | 2018-06-11 |  | 15
2018-06-12 | 22 | 7 | 29 | 2018-06-12 |  | 9
2018-06-13 | 12 | 9 | 21 | 2018-06-13 |  | 19 | 1
2018-06-14 | 10 | 10 | 20 | 2018-06-14 |  | 14
2018-06-18 | 15 | 5 | 20 | 2018-06-16 |  | 4
2018-06-19 | 16 | 11 | 27 | 2018-06-18 |  | 3
2018-06-20 | 16 | 6 | 22 | 2018-06-19 |  | 16
2018-06-21 | 17 | 11 | 28 | 2018-06-20 |  | 9
2018-06-22 | 18 | 4 | 22 | 2018-06-21 |  | 15
2018-06-25 | 15 | 2 | 17 | 2018-06-22 |  | 7 | 1
2018-06-26 | 23 | 5 | 28 | 2018-06-23 |  | 1
2018-06-27 | 22 | 5 | 27 | 2018-06-25 |  | 3
2018-06-28 | 28 |  | 28 | 2018-06-26 |  | 7
2018-06-29 | 9 | 1 | 10 | 2018-06-27 |  | 4
2018-06-30 | 12 |  | 12 | 2018-06-29 |  | 2
2018-07-02 | 14 | 3 | 17 | 2018-06-30 |  | 2
2018-07-03 | 25 | 4 | 29 | 2018-07-03 |  | 7
2018-07-04 | 15 | 5 | 20 | 2018-07-04 |  | 6
2018-07-05 | 17 | 6 | 23 | 2018-07-05 |  | 8
2018-07-06 | 10 | 4 | 14 | 2018-07-06 |  | 3
2018-07-07 | 15 | 2 | 17 | 2018-07-07 |  | 2
2018-07-09 | 18 | 2 | 20 | 2018-07-09 |  | 4
2018-07-10 | 19 | 2 | 21 | 2018-07-10 |  | 1
2018-07-11 | 21 | 1 | 22 | 2018-07-11 |  | 3
2018-07-12 | 15 | 2 | 17 | 2018-07-12 |  | 2
2018-07-13 | 17 |  | 17 | 2018-07-14 |  | 1
2018-07-14 | 17 | 1 | 18 | 2018-07-16 |  | 3
2018-07-16 | 13 | 2 | 15 | 2018-07-17 |  | 1
2018-07-17 | 21 | 1 | 22 | 2018-07-18 |  | 3
2018-07-18 | 24 | 3 | 27 | 2018-07-19 |  | 4
2018-07-19 | 15 | 2 | 17 | 2018-07-20 |  | 4
2018-07-20 | 19 | 5 | 24 | 2018-07-23 |  | 1
2018-07-23 | 25 | 2 | 27 | 2018-07-24 |  | 3
2018-07-24 | 16 | 1 | 17 | 2018-07-25 |  | 1
2018-07-25 | 24 | 2 | 26 | 2018-07-26 |  | 2
2018-07-26 | 23 | 1 | 24 | 2018-07-27 |  | 1
2018-07-27 | 11 | 1 | 12 | 2018-07-30 |  | 2
2018-07-30 | 23 | 2 | 25 | 2018-07-31 |  | 2
2018-07-31 | 17 | 2 | 19 | 2018-08-01 |  | 1
2018-08-01 | 17 |  | 17 | 2018-08-03 |  | 4
2018-08-02 | 20 | 1 | 21 | 2018-08-04 |  | 2
2018-08-03 | 8 | 2 | 10 | 2018-08-07 |  | 2
2018-08-04 | 6 | 2 | 8 | 2018-08-10 |  | 1
2018-08-06 | 14 | 2 | 16 | 2018-08-14 |  | 3
2018-08-07 | 24 |  | 24 | 2018-08-15 |  | 1
2018-08-08 | 23 | 1 | 24 | 2018-08-20 |  | 2
2018-08-10 | 13 |  | 13 | 2018-08-21 |  | 1
2018-08-13 | 26 | 1 | 27 | 2018-08-23 |  | 3
2018-08-14 | 15 | 1 | 16 | 2018-08-24 |  | 1
2018-08-15 | 24 | 1 | 25 | 2018-08-25 |  | 3
2018-08-16 | 20 | 1 | 21 | 2018-08-27 |  | 2
2018-08-17 | 1 |  | 1 | 2018-08-28 |  | 1
2018-08-20 | 25 | 2 | 27 | 2018-08-29 |  | 4
2018-08-21 | 16 |  | 16 | 2018-08-30 |  | 1
2018-08-23 | 22 | 2 | 24 | 2018-08-31 |  | 1
2018-08-24 | 17 | 3 | 20 | 2018-09-03 |  | 1
2018-08-25 | 16 |  | 16 | 2018-09-04 |  | 1
2018-08-27 | 26 | 4 | 30 | 2018-09-05 |  | 1
2018-08-28 | 23 | 1 | 24 | 2018-09-06 |  | 3
2018-08-29 | 22 | 1 | 23 | 2018-09-10 |  | 2
2018-08-30 | 29 | 2 | 31 | 2018-09-12 |  | 1
2018-08-31 | 19 |  | 19 | 2018-09-13 |  | 3
2018-09-03 | 25 | 2 | 27 | 2018-09-17 |  | 3
2018-09-04 | 25 | 1 | 26 | 2018-09-18 |  | 2
2018-09-05 | 18 |  | 18 | 2018-09-21 |  | 1
2018-09-06 | 15 | 3 | 18 | 2018-09-26 |  | 1
2018-09-07 | 17 |  | 17 | 2018-09-28 |  | 1
2018-09-08 | 9 |  | 9 | 2018-09-29 |  | 3
2018-09-10 | 20 | 1 | 21 | 2018-10-01 |  | 4
2018-09-11 | 24 | 1 | 25 | Grand Total | 82 | 684 | 120
2018-09-12 | 24 | 1 | 25
2018-09-13 | 24 | 2 | 26
2018-09-14 | 23 | 1 | 24
2018-09-15 | 9 |  | 9
2018-09-17 | 20 | 3 | 23
2018-09-18 | 18 | 1 | 19
2018-09-19 | 24 | 1 | 25
2018-09-20 | 15 |  | 15
2018-09-21 | 15 |  | 15
2018-09-22 | 3 |  | 3
2018-09-24 | 21 | 1 | 22
2018-09-25 | 33 |  | 33
2018-09-26 | 15 |  | 15
2018-09-27 | 18 | 1 | 19
2018-09-28 | 21 | 1 | 22
2018-09-29 | 18 |  | 18
2018-10-01 | 25 | 3 | 28
2018-10-02 | 22 | 2 | 24
2018-10-03 | 14 |  | 14
2018-10-04 | 15 |  | 15
2018-10-05 | 20 |  | 20
2018-10-06 | 19 |  | 19
2018-10-08 | 28 |  | 28
2018-10-09 | 21 |  | 21
Grand Total | 2267 | 529 | 2796

## Sheet2
Top Count | Carryover
. | 0.0003
38 | 0.0006
39 | 0.0013
39 | 0.0011
39 | 0.0017
39 | 0.0014
39 | 0.0012
39 | 0.0007
39 | 0.0015
39 | 0
39 | 0
39 | 0.0006
39 | 0
39 | 0.0003
39 | 0.0001
39 | 0.0001
39 | 0.0012
39 | 0.0003
39 | 0.0003
39 | 0.0004
39 | 0.0006
39 | 0.0002
39 | 0
39 | 0.0004
39 | 0
39 | 0
39 | 0.0002
39 | 0.0003
39 | 0.001
39 | 0.0007
39 | 0.001
39 | 0.0001
39 | 0
39 | 0.0001
39 | 0
39 | 0.0003
39 | 0.0005
39 | 0.0004
39 | 0.0003
39 | 0.0003
39 | 0.0001
39 | 0.0001
39 | 0.0002
39 | 0
39 | 0.0017
39 | 0.0003
39 | 0.0002
39 | 0.0011
39 | 0
39 | 0.0001
39 | 0
39 | 0
39 | 0.0005
39 | 0.0002
39 | 0.0012
39 | 0.0003
39 | 0.0007
39 | 0
39 | 0
39 | 0.0009
39 | 0
39 | 0.0002
39 | 0.0003
39 | 0
39 | 0.0007
39 | 0.0017
39 | 0.0005
39 | 0.0005
39 | 0.0001
39 | 0
39 | 0.0015
39 | 0
39 | 0.0008
39 | 0
39 | 0.0008
39 | 0.0006
39 | 0.0016
39 | 0.0001
39 | 0.0004
39 | 0.0005
39 | 0.0003
42 | 0.0013
42 | 0.0015
42 | 0.0002
42 | 0.0002
42 | 0.0004
42 | 0
42 | 0.0015
42 | 0.001
42 | 0.0017
42 | 0.0004
42 | 0.0002
42 | 0.0006
42 | 0.0005
42 | 0.0005
42 | 0.0009
42 | 0
42 | 0.001
42 | 0.0003
42 | 0.0013
42 | 0.0005
42 | 0.001
42 | 0.0006
42 | 0.0003
42 | 0.0002
42 | 0.0019
42 | 0.0003
42 | 0.0004
42 | 0.0001
42 | 0.0002
42 | 0.0004
42 | 0
42 | 0.0007
42 | 0.0014
42 | 0.0009
42 | 0.0002
42 | 0.0006
42 | 0.0002
42 | 0.0014
42 | 0.0001
42 | 0.0004
42 | 0
42 | 0
42 | 0
42 | 0
42 | 0
42 | 0.0001
42 | 0.0006
42 | 0.0004
42 | 0.001
42 | 0
42 | 0.0002
42 | 0.0001
42 | 0.0004
42 | 0
42 | 0.0005
42 | 0.0005
42 | 0.0009
42 | 0.0014
42 | 0.0008
42 | 0.0004
42 | 0.0006
42 | 0.0008
42 | 0.0003
42 | 0.0001
42 | 0.0007
42 | 0
42 | 0
42 | 0.0011
42 | 0.0004
42 | 0.0004
42 | 0.0005
42 | 0.0002
42 | 0
42 | 0.0004
42 | 0.0011
42 | 0.0002
42 | 0.0009
42 | 0.0002
42 | 0.0007
42 | 0.0004
42 | 0.001
42 | 0.0007
42 | 0.0007
42 | 0.0008
42 | 0.0007
42 | 0.0003
42 | 0
42 | 0.0001
42 | 0.0005
42 | 0.0003
42 | 0
42 | 0.0017
42 | 0.0003
42 | 0.0001
42 | 0.0003
42 | 0.0004
42 | 0.0001
42 | 0
42 | 0.0001
42 | 0.0007
42 | 0
42 | 0.0005
42 | 0.0002
42 | 0.0005
42 | 0.0004
42 | 0.0008
42 | 0.0014
42 | 0
42 | 0.0003
42 | 0.0008
42 | 0.001
42 | 0.0002
42 | 0.0001
42 | 0.0002
42 | 0.0004
42 | 0.0013
42 | 0.0004
42 | 0.0004
42 | 0.0011
42 | 0.0003
42 | 0.0002
42 | 0.0002
42 | 0.0004
42 | 0
42 | 0.0001
42 | 0.0009
42 | 0
42 | 0.0003
42 | 0.0002
42 | 0
42 | 0
42 | 0.0005
42 | 0.0002
42 | 0
42 | 0.0006
42 | 0.0005
42 | 0.0002
42 | 0.0009
42 | 0.0012
42 | 0.0006
42 | 0.0001
42 | 0.0007
42 | 0.0003
42 | 0.0004
42 | 0.0001
42 | 0.0013
42 | 0.0007
42 | 0.0017
42 | 0.0002
42 | 0.0003
42 | 0.0017
42 | 0.0001
42 | 0
42 | 0.0002
42 | 0.0005
42 | 0
42 | 0.0003
42 | 0
42 | 0.001
42 | 0.0003
42 | 0.0002
42 | 0.0005
42 | 0
42 | 0
42 | 0.0017
42 | 0
42 | 0.0004
42 | 0.0015
43 | 0.0004
43 | 0.0002
43 | 0
43 | 0.0004
43 | 0.0003
43 | 0.0013
43 | 0
43 | 0
43 | 0
43 | 0.0011
43 | 0
43 | 0.0009
43 | 0.0009
43 | 0.0008
43 | 0.0002
43 | 0.0001
43 | 0.0003
43 | 0.0007
43 | 0.0013
43 | 0.0006
43 | 0.0003
43 | 0.0009
43 | 0.0019
43 | 0.0012
43 | 0.0013
43 | 0.0008
43 | 0.0004
43 | 0.0016
43 | 0
43 | 0.0002
43 | 0.0001
43 | 0.0003
43 | 0.0001
43 | 0.0001
43 | 0.0007
43 | 0.0012
43 | 0.0014
43 | 0.0012
43 | 0.0006
43 | 0.0006
43 | 0.0002
43 | 0.0003
43 | 0.0001
43 | 0
43 | 0.0003
43 | 0.0005
43 | 0
43 | 0.0001
43 | 0.0002
43 | 0.0008
43 | 0.0001
43 | 0.0007
43 | 0.001
43 | 0
43 | 0.0003
43 | 0.0002
43 | 0
43 | 0
43 | 0.0005
43 | 0.0002
43 | 0
43 | 0
43 | 0.0003
43 | 0
43 | 0.0016
43 | 0.0005
43 | 0.0001
43 | 0.0001
43 | 0.0008
43 | 0
43 | 0.0008
43 | 0.0002
43 | 0.0004
43 | 0.0008
43 | 0.0006
43 | 0.0008
43 | 0.0009
43 | 0.0004
43 | 0.0007
43 | 0
43 | 0.0005
43 | 0.0004
43 | 0.0005
43 | 0.0007
43 | 0.0004
43 | 0.0007
43 | 0.0002
43 | 0.0002
43 | 0.0007
43 | 0
43 | 0.0011
43 | 0
43 | 0.0006
43 | 0.0002
43 | 0.0015
43 | 0.0014
43 | 0.0012
43 | 0.0003
43 | 0.0004
43 | 0.0012
43 | 0.001
43 | 0.0003
43 | 0.0001
43 | 0.0002
43 | 0.0002
43 | 0.0004
43 | 0.0005
43 | 0.0002
43 | 0.0001
43 | 0.0008
43 | 0.0018
43 | 0.0003
43 | 0.0012
43 | 0.0008
43 | 0.0005
43 | 0.0007
43 | 0.0005
43 | 0.0007
43 | 0.0003
43 | 0.0018
43 | 0.0014
43 | 0.0003
43 | 0.0004
43 | 0
43 | 0.001
43 | 0.0011
43 | 0.0004
43 | 0
43 | 0
43 | 0.0014
43 | 0.0005
43 | 0.0001
43 | 0.0019
43 | 0.0005
43 | 0.0006
43 | 0.0002
43 | 0.0001
43 | 0.0003
43 | 0.0001
43 | 0
43 | 0.0003
43 | 0
43 | 0.0003
43 | 0.0005
43 | 0.0005
43 | 0.0004
43 | 0.0008
43 | 0
43 | 0.0001
43 | 0.0003
43 | 0.0007
43 | 0
43 | 0.0002
43 | 0.0001
43 | 0.0006
43 | 0.0004
43 | 0.0012
43 | 0
43 | 0.0006
43 | 0.0001
43 | 0.0007
43 | 0
43 | 0
43 | 0.0001
43 | 0
43 | 0
43 | 0.0002
43 | 0.0001
43 | 0.0004
43 | 0.001
43 | 0.0002
43 | 0.001
43 | 0
43 | 0.0001
43 | 0.0012
43 | 0.0003
43 | 0.0001
43 | 0.0006
43 | 0.0001
43 | 0.0004
43 | 0.0009
43 | 0.0006
43 | 0
43 | 0.0002
43 | 0.0007
43 | 0.0004
43 | 0.0001
43 | 0.0008
43 | 0
43 | 0
43 | 0.0006
43 | 0.0004
43 | 0.0009
43 | 0
43 | 0
43 | 0.0002
43 | 0.0002
43 | 0.0002
43 | 0.0007
43 | 0.0006
43 | 0.0007
43 | 0.0004
43 | 0.0001
43 | 0.0002
43 | 0.0003
43 | 0.0006
43 | 0
43 | 0.0003
43 | 0.0003
43 | 0.0001
43 | 0.0002
43 | 0.0002
43 | 0.0003
43 | 0.0003
43 | 0.0003
43 | 0.0003
43 | 0.0003
43 | 0
43 | 0.0002
44 | 0.0019
44 | 0.0001
46 | 0.0003
46 | 0.0012
46 | 0.0008
46 | 0.0006
46 | 0
46 | 0.0004
46 | 0
46 | 0.0005
46 | 0.0006
46 | 0.0008
46 | 0.0002
46 | 0
46 | 0
46 | 0.0006
46 | 0.0001
46 | 0
46 | 0.0004
46 | 0.0004
46 | 0.0005
46 | 0.0002
46 | 0.0006
46 | 0.0013
46 | 0.0016
46 | 0.0003
46 | 0
46 | 0.0001
46 | 0.0003
46 | 0.0003
46 | 0.0003
46 | 0.0003
46 | 0.0007
46 | 0.0005
46 | 0.0013
46 | 0.0006
46 | 0.0014
46 | 0.0014
46 | 0.0007
46 | 0.0004
46 | 0.0017
46 | 0.0003
46 | 0.0008
46 | 0.0008
46 | 0.0007
46 | 0.0002
46 | 0.0017
46 | 0.0011
46 | 0.0006
46 | 0.0011
46 | 0.0012
46 | 0.0004
46 | 0.0005
46 | 0.0015
46 | 0.0003
46 | 0.0015
46 | 0.0001
46 | 0.0002
46 | 0.0017
46 | 0.0004
46 | 0.0009
46 | 0.0004
46 | 0.0017
46 | 0.0006
46 | 0.0007
46 | 0.0003
46 | 0.0016
46 | 0.0006
46 | 0.0008
46 | 0.0005
46 | 0.0006
46 | 0.0007
46 | 0.0005
46 | 0.0017
46 | 0.0015
46 | 0.0007
46 | 0.0003
46 | 0.0014
46 | 0.0011
46 | 0.0008
46 | 0.0004
46 | 0.0002
46 | 0.0009
46 | 0.0019
46 | 0.0004
46 | 0.0019
46 | 0.0014
46 | 0.0014
46 | 0
46 | 0.0003
46 | 0
46 | 0.0005
46 | 0.0003
46 | 0.0001
46 | 0.0002
46 | 0
46 | 0.0002
46 | 0.0002
46 | 0.0016
46 | 0.0005
46 | 0.0003
46 | 0.0001
46 | 0.0002
46 | 0.0002
46 | 0
46 | 0.0002
46 | 0.001
46 | 0.0004
46 | 0
46 | 0.0003
46 | 0.0008
46 | 0.0005
46 | 0.0003
46 | 0.0003
46 | 0.0001
46 | 0.0012
46 | 0.0003
46 | 0.0017
46 | 0.0001
46 | 0.0003
46 | 0.0004
46 | 0.0001
46 | 0.0007
46 | 0.0013
46 | 0.0001
46 | 0.0001
46 | 0.0001
46 | 0
46 | 0.0005
46 | 0.0002
46 | 0.0001
46 | 0.0003
46 | 0.0012
46 | 0.0008
46 | 0.0006
46 | 0.0002
46 | 0.0002
46 | 0.0008
46 | 0.0001
46 | 0.0004
46 | 0.0002
46 | 0.0004
46 | 0
46 | 0
46 | 0.0001
46 | 0.0015
46 | 0
46 | 0.0001
46 | 0.0002
46 | 0.0014
46 | 0.0002
46 | 0.0003
46 | 0.0001
46 | 0.0002
46 | 0.0001
46 | 0
46 | 0.0001
46 | 0.0008
46 | 0
46 | 0
46 | 0.0008
46 | 0.0002
46 | 0.0001
46 | 0.0002
46 | 0.0001
46 | 0.0004
46 | 0.0004
46 | 0
46 | 0.0002
46 | 0.0001
46 | 0.0001
46 | 0.0001
46 | 0
46 | 0.0004
46 | 0.0003
46 | 0.0003
46 | 0.0001
46 | 0.0017
46 | 0
46 | 0
46 | 0.0001
46 | 0
46 | 0.0003
46 | 0
46 | 0.0003
46 | 0
46 | 0.0004
46 | 0.0002
46 | 0.0005
46 | 0
46 | 0
46 | 0.0001
46 | 0.0003
46 | 0.0004
46 | 0.0003
46 | 0
46 | 0.0002
46 | 0.0004
46 | 0
46 | 0.0002
46 | 0.0003
46 | 0.0003
46 | 0
46 | 0.0001
46 | 0.0007
46 | 0.0004
46 | 0
46 | 0.0004
46 | 0
46 | 0
46 | 0.0004
46 | 0.0003
46 | 0.0002
46 | 0.0003
46 | 0.0003
46 | 0.0002
46 | 0.0003
46 | 0.0002
46 | 0.0008
46 | 0.0007
46 | 0.0005
46 | 0.0006
46 | 0
46 | 0.0004
46 | 0.0003
46 | 0.0007
46 | 0
46 | 0
46 | 0.0006
46 | 0.0007
46 | 0.0012
46 | 0
46 | 0.0008
46 | 0.0004
46 | 0.0017
46 | 0.0007
46 | 0.0004
46 | 0.0005
46 | 0.0005
46 | 0.0005
46 | 0.0002
46 | 0.0001
46 | 0.0005
46 | 0.0015
46 | 0.0002
46 | 0.0008
46 | 0.0007
46 | 0
46 | 0.0007
46 | 0.0015
46 | 0.0004
46 | 0
46 | 0.0005
46 | 0.0011
46 | 0.0012
46 | 0.0007
46 | 0.0002
46 | 0.0009
46 | 0.0008
46 | 0.0006
46 | 0.0007
46 | 0.0009
46 | 0.0007
46 | 0.0003
46 | 0.0011
46 | 0.0013
46 | 0.0001
46 | 0.0003
46 | 0.0005
46 | 0.0005
46 | 0.0002
46 | 0
46 | 0.001
46 | 0.0003
46 | 0.0008
46 | 0.001
46 | 0.0002
46 | 0.0012
46 | 0.0013
46 | 0.0005
46 | 0.0003
46 | 0.0003
46 | 0.0003
46 | 0.0002
46 | 0.0002
46 | 0.0006
46 | 0.0009
46 | 0.0005
46 | 0.0002
46 | 0.0003
46 | 0.0001
46 | 0.001
46 | 0.0003
46 | 0.0004
46 | 0.0005
46 | 0.0005
46 | 0.0001
46 | 0.0011
46 | 0.0013
46 | 0.0009
46 | 0.0006
46 | 0.0003
46 | 0.0007
46 | 0.0002
46 | 0.0003
46 | 0
46 | 0.0001
46 | 0
46 | 0.0008
46 | 0
46 | 0.0004
46 | 0.0002
46 | 0.0006
46 | 0.0002
46 | 0.0017
46 | 0.0004
46 | 0.0004
46 | 0.0002
46 | 0.0003
46 | 0.0015
46 | 0.0001
46 | 0.0006
46 | 0.0006
46 | 0.0017
46 | 0
46 | 0.0002
46 | 0.0003
46 | 0.0008
46 | 0.0001
46 | 0.0005
46 | 0.0003
46 | 0.0009
46 | 0.0003
46 | 0.0004
46 | 0
46 | 0.0002
46 | 0
46 | 0.0003
46 | 0.0015
46 | 0.0004
46 | 0
46 | 0.0001
46 | 0
46 | 0.0008
46 | 0.0002
46 | 0.0016
46 | 0.0001
46 | 0.0003
46 | 0.0003
46 | 0
46 | 0.0012
46 | 0.0003
46 | 0
46 | 0.0014
46 | 0
46 | 0.0001
46 | 0.0012
46 | 0.0005
46 | 0
46 | 0.0004
46 | 0.0003
46 | 0.0001
46 | 0.0003
46 | 0.0001
46 | 0.0001
46 | 0.0015
46 | 0.0002
46 | 0
46 | 0
46 | 0.0012
46 | 0.0002
46 | 0
46 | 0.0003
46 | 0
46 | 0.0004
46 | 0.0002
46 | 0
46 | 0.0013
46 | 0.0002
46 | 0.0001
46 | 0
46 | 0.0001
46 | 0
46 | 0.0002
46 | 0.0003
46 | 0.0003
46 | 0
46 | 0
46 | 0
46 | 0.0005
46 | 0.0002
46 | 0
46 | 0.0008
46 | 0.0003
46 | 0.0002
46 | 0
46 | 0
46 | 0
46 | 0.0002
46 | 0
46 | 0.0005
46 | 0
46 | 0.0001
46 | 0
46 | 0
46 | 0.0007
46 | 0.0005
46 | 0.0001
46 | 0.0005
46 | 0.0011
46 | 0
46 | 0
46 | 0
46 | 0
46 | 0
46 | 0.0005
46 | 0.0005
46 | 0.0008
46 | 0.0008
46 | 0.0001
46 | 0
46 | 0.0001
46 | 0.0009
46 | 0.0003
46 | 0
46 | 0.0002
46 | 0.0007
46 | 0.0006
46 | 0.0001
46 | 0.0007
46 | 0
46 | 0
46 | 0.0004
46 | 0.0005
46 | 0.0008
46 | 0.0007
46 | 0.0004
46 | 0.0001
46 | 0.0011
46 | 0.0006
46 | 0
46 | 0
46 | 0.0003
46 | 0.0004
46 | 0.0002
46 | 0.0004
46 | 0
46 | 0
46 | 0.0003
46 | 0.0003
46 | 0.0005
46 | 0.0002
46 | 0
46 | 0
46 | 0.0013
46 | 0
46 | 0.0004
46 | 0.0005
46 | 0.0002
46 | 0.0002
46 | 0.0003
46 | 0.0002
46 | 0.0012
46 | 0.0012
46 | 0
46 | 0.0001
46 | 0
46 | 0.0002
46 | 0
46 | 0
46 | 0.0003
46 | 0.0016
46 | 0.0002
46 | 0.0002
46 | 0.0001
46 | 0.0019
46 | 0.0002
46 | 0.0003
46 | 0.0002
46 | 0.0006
46 | 0
46 | 0.0017
46 | 0.0001
46 | 0.0002
46 | 0.0005
46 | 0.0002
46 | 0.0016
46 | 0.0001
46 | 0.0007
46 | 0.0002
46 | 0.0003
46 | 0
46 | 0.0001
46 | 0.0006
46 | 0.0007
46 | 0
46 | 0.0002
46 | 0.0006
46 | 0.0002
46 | 0
46 | 0.0018
48 | 0.0011
48 | 0.0002
48 | 0.0013
48 | 0.0018
48 | 0.0008
48 | 0.0002
48 | 0.0003
48 | 0.0007
48 | 0.0001
48 | 0.0001
48 | 0.0003
48 | 0.0002
48 | 0
48 | 0.0012
48 | 0
48 | 0.0002
48 | 0.0003
48 | 0.0006
48 | 0.0002
48 | 0
48 | 0.0001
48 | 0.0016
48 | 0.0005
48 | 0.0006
48 | 0
48 | 0.0002
48 | 0.0004
48 | 0.0001
48 | 0.0002
48 | 0.0003
48 | 0.0002
48 | 0.0009
48 | 0.0003
48 | 0.0001
48 | 0
48 | 0.0001
48 | 0.0013
48 | 0.0011
48 | 0.0005
48 | 0.0003
48 | 0.0012
48 | 0.0007
48 | 0.0007
48 | 0.0013
48 | 0.0002
48 | 0.0005
48 | 0.0016
48 | 0.0009
48 | 0.0004
48 | 0.0012
48 | 0.0013
48 | 0.0012
48 | 0.0019
48 | 0.0003
48 | 0.0015
48 | 0.0014
48 | 0.0003
48 | 0.001
48 | 0.0018
48 | 0.0013
48 | 0.0015
48 | 0.0017
48 | 0
48 | 0.0014
48 | 0.0006
48 | 0.0012
48 | 0.0007
48 | 0
48 | 0.0016
48 | 0.0011
48 | 0.001
48 | 0.0001
48 | 0.0007
48 | 0.0003
48 | 0.0001
48 | 0.001
48 | 0.0007
48 | 0.0007
48 | 0
48 | 0.0003
48 | 0.0004
48 | 0.001
48 | 0.0009
48 | 0.0003
48 | 0.0005
48 | 0.0001
48 | 0
48 | 0.0006
48 | 0
48 | 0.0005
48 | 0.0003
48 | 0.0001
48 | 0.0007
48 | 0.0003
48 | 0.0003
48 | 0.0001
48 | 0.0003
48 | 0.0001
48 | 0.0003
48 | 0.0005
48 | 0
48 | 0.0016
48 | 0.0002
48 | 0.0003
48 | 0.0005
48 | 0.0003
48 | 0.0002
48 | 0
48 | 0.0002
48 | 0
48 | 0.0008
48 | 0.0002
48 | 0.0001
48 | 0.0003
48 | 0.0012
48 | 0.0002
48 | 0.0008
48 | 0.0002
48 | 0.0005
48 | 0.0004
48 | 0.0012
48 | 0
48 | 0.0017
48 | 0
48 | 0.0005
48 | 0.0004
48 | 0
48 | 0.001
48 | 0.0002
48 | 0.0003
48 | 0.0002
48 | 0.0016
48 | 0.0003
48 | 0.0004
48 | 0.0007
48 | 0.0008
48 | 0.0006
48 | 0.0003
48 | 0.0002
48 | 0.0002
48 | 0.0004
48 | 0.0001
48 | 0.0001
48 | 0
48 | 0.0003
48 | 0.0001
48 | 0.0002
48 | 0.0009
48 | 0.0002
48 | 0.0002
48 | 0.0005
48 | 0.0001
48 | 0.0003
48 | 0.0001
48 | 0
48 | 0.0001
48 | 0.0003
48 | 0.0001
48 | 0
48 | 0.0004
48 | 0.0002
48 | 0.0002
48 | 0
48 | 0.0008
48 | 0
48 | 0.0001
48 | 0.0002
48 | 0.0001
48 | 0
48 | 0.0001
48 | 0.0012
48 | 0.0003
48 | 0.0001
48 | 0.0005
48 | 0
48 | 0.0004
48 | 0.0006
48 | 0.0003
48 | 0.0002
48 | 0
48 | 0.0002
48 | 0
48 | 0
48 | 0.0004
48 | 0.0003
48 | 0.0003
48 | 0
48 | 0.0001
48 | 0.0001
48 | 0.0003
48 | 0.0002
48 | 0
48 | 0.0005
48 | 0.0006
48 | 0.0003
48 | 0
48 | 0.0004
48 | 0
48 | 0.0003
48 | 0
48 | 0.0011
48 | 0.0004
48 | 0.0003
48 | 0.0002
48 | 0
48 | 0.0002
48 | 0.0004
48 | 0
48 | 0.0015
48 | 0.0001
48 | 0.0003
48 | 0.0008
48 | 0
48 | 0.0008
48 | 0.001
48 | 0.0003
48 | 0.0001
48 | 0.0004
48 | 0.0002
48 | 0.0004
48 | 0.0007
48 | 0.0008
48 | 0.0007
48 | 0.0004
48 | 0
48 | 0.0008
48 | 0.0005
48 | 0.0003
48 | 0.0005
48 | 0.001
48 | 0.0008
48 | 0.0009
48 | 0.0006
48 | 0.0002
48 | 0.0005
48 | 0.0006
48 | 0.0004
48 | 0.0007
48 | 0.0016
48 | 0.0019
48 | 0.0007
48 | 0
48 | 0.0005
48 | 0.0005
48 | 0.0002
48 | 0
48 | 0.0006
48 | 0.0007
48 | 0
48 | 0.0005
48 | 0.0002
48 | 0.0003
48 | 0
48 | 0.0017
48 | 0.0017
48 | 0
48 | 0.0003
48 | 0.0011
48 | 0.0007
48 | 0.0002
48 | 0.0002
48 | 0
48 | 0.0003
48 | 0.0015
48 | 0.0001
48 | 0.0005
48 | 0.0005
48 | 0
48 | 0.0011
48 | 0.0003
48 | 0.0002
48 | 0.0009
48 | 0.0007
48 | 0.0006
48 | 0.0005
48 | 0.0001
48 | 0.0007
48 | 0.0008
48 | 0
48 | 0.0006
48 | 0.0008
48 | 0.0004
48 | 0.0013
48 | 0.0007
48 | 0.0011
48 | 0.0002
48 | 0
48 | 0.0019
48 | 0.0012
48 | 0.0003
48 | 0
48 | 0.0015
48 | 0.0006
48 | 0.0008
48 | 0.0007
48 | 0
48 | 0.0001
48 | 0.0007
48 | 0.0002
48 | 0
48 | 0.0015
48 | 0.0008
48 | 0.0001
48 | 0.0001
48 | 0
48 | 0.0016
48 | 0.0009
48 | 0.0003
48 | 0.0003
48 | 0
48 | 0.0013
48 | 0.0018
48 | 0.0011
48 | 0
48 | 0.0002
48 | 0.0006
48 | 0.0008
48 | 0.0006
48 | 0.0006
48 | 0.0004
48 | 0.0007
48 | 0.0003
48 | 0.0006
48 | 0.0002
48 | 0.0001
48 | 0.0007
48 | 0.0005
48 | 0
48 | 0
48 | 0
48 | 0.0006
48 | 0.0008
48 | 0.0002
48 | 0.0003
48 | 0.0012
48 | 0.0014
48 | 0.0003
48 | 0.0003
48 | 0.0004
48 | 0.0004
48 | 0.0002
48 | 0.0006
48 | 0
48 | 0.0002
48 | 0.0001
48 | 0.0005
48 | 0.0004
48 | 0.0002
48 | 0
48 | 0.0007
48 | 0.0002
48 | 0.0002
48 | 0
48 | 0
48 | 0
48 | 0.0001
48 | 0
48 | 0.0006
48 | 0.0001
48 | 0.0001
48 | 0.0001
48 | 0.0006
48 | 0.0001
48 | 0.0002
48 | 0.0013
48 | 0.0011
48 | 0
48 | 0.0014
48 | 0
48 | 0.0002
48 | 0.0001
48 | 0.0007
48 | 0
48 | 0.0003
48 | 0
48 | 0.0016
48 | 0
48 | 0
48 | 0
48 | 0.0012
48 | 0.0001
48 | 0
48 | 0.0006
48 | 0
48 | 0.0003
48 | 0
48 | 0.0004
48 | 0
48 | 0.0005
48 | 0.0006
48 | 0
48 | 0.001
48 | 0.0001
48 | 0
48 | 0.0002
48 | 0.0003
48 | 0
48 | 0
48 | 0.0002
48 | 0.0004
48 | 0
48 | 0.0002
48 | 0
48 | 0
48 | 0
48 | 0.0001
48 | 0.0005
48 | 0.0003
48 | 0
48 | 0.0009
48 | 0.0001
48 | 0.0005
48 | 0.0002
48 | 0
48 | 0.0013
48 | 0.0006
48 | 0
48 | 0.0002
48 | 0.0015
48 | 0.0007
48 | 0.0003
48 | 0.0013
48 | 0.0003
48 | 0.0007
48 | 0
48 | 0.0003
48 | 0
48 | 0
48 | 0.0001
48 | 0.0003
48 | 0.001
48 | 0.0004
48 | 0
48 | 0.0008
48 | 0.0005
48 | 0
49 | 0.0008
49 | 0.0009
49 | 0.0013
49 | 0.0006
49 | 0.0001
49 | 0.0008
49 | 0.0007
49 | 0.0004
49 | 0
49 | 0.001
49 | 0.0003
49 | 0.0003
49 | 0.0006
49 | 0.0005
49 | 0.0003
49 | 0.0005
49 | 0.0001
49 | 0.0003
49 | 0.0008
49 | 0
49 | 0.0012
49 | 0.0004
49 | 0.0008
49 | 0.0016
49 | 0.0013
49 | 0.0003
49 | 0.0015
49 | 0.0006
49 | 0.0004
49 | 0.0005
49 | 0.0014
49 | 0.0015
49 | 0.0007
49 | 0.0005
49 | 0
49 | 0.001
49 | 0.0007
49 | 0.0008
49 | 0.0013
49 | 0.0011
49 | 0
49 | 0.0009
49 | 0.0003
49 | 0.0016
49 | 0.0008
49 | 0.0009
49 | 0.0011
49 | 0
49 | 0.0003
49 | 0.0005
49 | 0.0002
49 | 0.0002
49 | 0.0013
49 | 0.0003
49 | 0
49 | 0.0006
49 | 0.0001
49 | 0.0003
49 | 0.0004
49 | 0.0011
49 | 0.0002
49 | 0.0015
49 | 0.0002
49 | 0.0012
49 | 0.0017
49 | 0.0001
49 | 0.0001
49 | 0.0002
49 | 0.0005
49 | 0.0009
49 | 0.0007
49 | 0.0002
49 | 0.0003
49 | 0.001
49 | 0.0005
49 | 0.0007
49 | 0.0009
49 | 0.0008
49 | 0.0019
49 | 0
49 | 0
49 | 0.0001
49 | 0.0014
49 | 0.0001
49 | 0.0005
49 | 0.0002
49 | 0.0001
49 | 0.0002
49 | 0.0006
49 | 0.0006
49 | 0.0002
49 | 0.0001
49 | 0.0005
49 | 0.0003
49 | 0
49 | 0
49 | 0.0006
49 | 0.0017
49 | 0.0002
49 | 0
49 | 0.0003
49 | 0.0002
49 | 0.0002
49 | 0.0002
49 | 0
49 | 0.0004
49 | 0.0004
49 | 0.0019
49 | 0.0003
49 | 0.0003
49 | 0
49 | 0.0003
49 | 0.0004
49 | 0.0003
49 | 0.0003
49 | 0.0003
49 | 0.0014
49 | 0.0008
49 | 0.0006
49 | 0.0005
49 | 0.0008
49 | 0.0001
49 | 0.0008
49 | 0.0007
49 | 0.0004
49 | 0.0005
49 | 0.001
49 | 0.0011
49 | 0.0014
49 | 0.0017
49 | 0.0004
49 | 0.0007
49 | 0
49 | 0.0006
49 | 0.0002
49 | 0.0008
49 | 0.001
49 | 0.0001
49 | 0.0004
49 | 0
49 | 0.0011
49 | 0.0008
49 | 0.001
49 | 0.0003
49 | 0.0001
49 | 0.0014
49 | 0.0004
49 | 0.0008
49 | 0.0006
49 | 0.0018
49 | 0.0016
49 | 0.0001
49 | 0.0008
49 | 0
49 | 0.0014
49 | 0.0008
49 | 0
49 | 0.0009
49 | 0.001
49 | 0.0015
49 | 0.0004
49 | 0.0004
49 | 0.0002
49 | 0.0001
49 | 0.0007
49 | 0.0002
49 | 0.0009
49 | 0.0004
49 | 0.0016
49 | 0
49 | 0.0002
49 | 0.0002
49 | 0.0002
49 | 0.0005
49 | 0.0005
49 | 0.0002
49 | 0.0005
49 | 0.0001
49 | 0.0017
49 | 0
49 | 0.0001
49 | 0.0001
49 | 0.0002
49 | 0.0013
49 | 0.0016
49 | 0
49 | 0.0002
49 | 0.0004
49 | 0.0002
49 | 0.0003
49 | 0.0004
49 | 0.0001
49 | 0.0001
49 | 0
49 | 0.0004
49 | 0.0004
49 | 0.0002
49 | 0.0009
49 | 0
49 | 0.0001
49 | 0.0005
49 | 0
49 | 0.0007
49 | 0.0007
49 | 0.0006
49 | 0.0007
49 | 0
49 | 0.001
49 | 0.0018
49 | 0.0005
49 | 0.0006
49 | 0.0003
50 | 0.0003
50 | 0.0003
50 | 0.0003
50 | 0.0009
50 | 0.0002
50 | 0.0001
50 | 0
50 | 0.0002
50 | 0.0004
50 | 0.0003
50 | 0.0003
50 | 0.0001
50 | 0.0004
50 | 0.0003
50 | 0
50 | 0.0003
50 | 0.0001
50 | 0.0015
50 | 0
50 | 0
50 | 0
50 | 0
50 | 0.0005
50 | 0.0003
50 | 0.0003
50 | 0
50 | 0
50 | 0.0004
50 | 0
50 | 0.0005
50 | 0.0002
50 | 0.0004
50 | 0.0015
50 | 0.0006
50 | 0.0002
50 | 0.0005
50 | 0.0018
50 | 0.0002
50 | 0.0006
50 | 0.0001
50 | 0.0013
50 | 0.0012
50 | 0.0002
50 | 0.0004
50 | 0.0002
50 | 0.0003
50 | 0.0006
50 | 0.0015
50 | 0.0003
50 | 0.0002
50 | 0
50 | 0.0003
50 | 0.0002
50 | 0
50 | 0.0001
50 | 0.0002
50 | 0.0009
50 | 0.0002
50 | 0
50 | 0
50 | 0.0006
50 | 0.0011
50 | 0
50 | 0.0003
50 | 0.0012
50 | 0.0001
50 | 0.0002
50 | 0.0002
50 | 0.0003
50 | 0.0002
50 | 0.0009
50 | 0.0002
50 | 0.0004
50 | 0.0001
50 | 0.0005
50 | 0.0002
50 | 0.0003
50 | 0
50 | 0.0003
50 | 0.0005
50 | 0.0003
50 | 0.0005
50 | 0.001
50 | 0.0001
50 | 0.0002
50 | 0.0003
50 | 0.0003
50 | 0.0004
50 | 0
50 | 0.0004
50 | 0.0004
50 | 0
50 | 0.0005
50 | 0