## FTN Venture Test Data in 2018
mode | start-time | end-time | Station | Operator | Category | part-no | serial-no | Test | Outcome | Jumper Configuration >> Jumper Configuration >> LSL=129 | Tray Cal Offset >> Tray Cal Offset >> LSL=-180 >> USL=180 | R Offset >> R Offset >> LSL=-3 >> USL=3 | Z Offset >> Z Offset >> LSL=-3 >> USL=3 | Needle Seal Position >> Needle Seal Position >> LSL=48 >> USL=68 | Needle Seal Force >> Needle Seal Force >> LSL=48 >> USL=68 | Zp Offset >> Zp Offset >> LSL=-3 >> USL=3 | 1st Move Error >> 1st Move Error >> LSL=-0.3 >> USL=0.3 | Zp Pierce >> Zp Pierce >> LSL=-20 >> USL=20 | Z Pierce >> Z Pierce >> LSL=-20 >> USL=20 | R Lost Steps >> R Lost Steps >> LSL=-20 >> USL=20 | Sample Cooler Test >> Sample Cooler Test >> USL=-0.6 | Test Pressure Transducer >> Test Pressure Transducer >> LSL=0.45 >> USL=0.55 | Sample Syringe Decay >> Sample Syringe Decay >> LSL=-5 >> USL=15 | Dynamic Leak Test >> Dynamic Leak Test >> USL=500 | Test Preblank Injections >> Test Preblank Injections >> USL=792.9 | Carryover >> Carryover >> USL=0.128 | Chemistry Precision Area >> Chemistry Precision Area >> LSL=0 >> USL=0.5 | Chemistry Precision Height >> Chemistry Precision Height >> LSL=0 >> USL=0.9 | Chemistry X-Intercept >> Chemistry X-Intercept >> LSL=-0.2 >> USL=0.2 | Chemistry Linearity (R^2) >> Chemistry Linearity (R^2) >> LSL=0.999 >> USL=1 | Test Preblank Injections >> Test Preblank Injections >> USL=776.9 | Test Preblank Injections >> Test Preblank Injections >> USL=808.3 | Test Preblank Injections >> Test Preblank Injections >> USL=736.7 | Test Preblank Injections >> Test Preblank Injections >> USL=761.9 | Jumper Configuration >> Jumper Configuration >> LSL=133 | Test Preblank Injections >> Test Preblank Injections >> USL=713.6 | Carryover >> Carryover >> USL=0.04 | Test Preblank Injections >> Test Preblank Injections >> USL=782.8 | Test Preblank Injections >> Test Preblank Injections >> USL=781.1 | Test Preblank Injections >> Test Preblank Injections >> USL=701.9 | Test Preblank Injections >> Test Preblank Injections >> USL=800.6 | Test Preblank Injections >> Test Preblank Injections >> USL=726.5 | Test Preblank Injections >> Test Preblank Injections >> USL=731.6 | Test Preblank Injections >> Test Preblank Injections >> USL=766.8 | Test Preblank Injections >> Test Preblank Injections >> USL=650.8 | Test Preblank Injections >> Test Preblank Injections >> USL=711.7 | Test Preblank Injections >> Test Preblank Injections >> USL=775.5 | Test Preblank Injections >> Test Preblank Injections >> USL=757.7 | Test Preblank Injections >> Test Preblank Injections >> USL=701 | Test Preblank Injections >> Test Preblank Injections >> USL=723.5 | Test Preblank Injections >> Test Preblank Injections >> USL=754.1 | Carryover >> Carryover >> USL=0.06 | Test Preblank Injections >> Test Preblank Injections >> USL=721.5 | Test Preblank Injections >> Test Preblank Injections >> USL=661 | Test Preblank Injections >> Test Preblank Injections >> USL=739.3 | Test Preblank Injections >> Test Preblank Injections >> USL=781.6 | Test Preblank Injections >> Test Preblank Injections >> USL=705.2 | ReStart Initialization >> ReStart Initialization | Test Preblank Injections >> Test Preblank Injections >> USL=721.9 | Test Preblank Injections >> Test Preblank Injections >> USL=712.9 | Test Preblank Injections >> Test Preblank Injections >> USL=741.4 | Test Preblank Injections >> Test Preblank Injections >> USL=749.9 | Test Preblank Injections >> Test Preblank Injections >> USL=479.1 | Test Preblank Injections >> Test Preblank Injections >> USL=778.5 | Test Preblank Injections >> Test Preblank Injections >> USL=0 | Test Preblank Injections >> Test Preblank Injections >> USL=702.7 | Test Preblank Injections >> Test Preblank Injections >> USL=769 | Test Preblank Injections >> Test Preblank Injections >> USL=705.7 | Test Preblank Injections >> Test Preblank Injections >> USL=778.8 | Test Preblank Injections >> Test Preblank Injections >> USL=691.6 | Test Preblank Injections >> Test Preblank Injections >> USL=830.6 | Test Preblank Injections >> Test Preblank Injections >> USL=723.2 | Test Preblank Injections >> Test Preblank Injections >> USL=648.4 | Test Preblank Injections >> Test Preblank Injections >> USL=771.8 | Test Preblank Injections >> Test Preblank Injections >> USL=722.1 | Test Preblank Injections >> Test Preblank Injections >> USL=752.3 | Test Preblank Injections >> Test Preblank Injections >> USL=770.4 | Test Preblank Injections >> Test Preblank Injections >> USL=715.3 | Test Preblank Injections >> Test Preblank Injections >> USL=678.6 | Test Preblank Injections >> Test Preblank Injections >> USL=572.3 | Test Preblank Injections >> Test Preblank Injections >> USL=728.1 | Test Preblank Injections >> Test Preblank Injections >> USL=732.4 | Test Preblank Injections >> Test Preblank Injections >> USL=743 | Test Preblank Injections >> Test Preblank Injections >> USL=775.3 | Test Preblank Injections >> Test Preblank Injections >> USL=610.6 | Test Preblank Injections >> Test Preblank Injections >> USL=15961 | Test Preblank Injections >> Test Preblank Injections >> USL=722.9 | Test Preblank Injections >> Test Preblank Injections >> USL=744.6 | Test Preblank Injections >> Test Preblank Injections >> USL=785.7 | Test Preblank Injections >> Test Preblank Injections >> USL=735.3 | Test Preblank Injections >> Test Preblank Injections >> USL=786.8 | Test Preblank Injections >> Test Preblank Injections >> USL=741 | Test Ecord Interface >> Test Ecord Interface | Test Preblank Injections >> Test Preblank Injections >> USL=635 | Test Preblank Injections >> Test Preblank Injections >> USL=800.3 | Test Preblank Injections >> Test Preblank Injections >> USL=775 | Power-on Initialization >> Power-on Initialization | Test Preblank Injections >> Test Preblank Injections >> USL=724 | Test Preblank Injections >> Test Preblank Injections >> USL=745.8 | Test Preblank Injections >> Test Preblank Injections >> USL=734.5 | Test Preblank Injections >> Test Preblank Injections >> USL=760.3 | Test Preblank Injections >> Test Preblank Injections >> USL=663.1 | Test Preblank Injections >> Test Preblank Injections >> USL=792.7 | Test Preblank Injections >> Test Preblank Injections >> USL=736.5 | Test Preblank Injections >> Test Preblank Injections >> USL=775.4 | Test Preblank Injections >> Test Preblank Injections >> USL=961.5 | Test Preblank Injections >> Test Preblank Injections >> USL=717.5 | Test Preblank Injections >> Test Preblank Injections >> USL=738.6 | Test Preblank Injections >> Test Preblank Injections >> USL=727.2 | Test Preblank Injections >> Test Preblank Injections >> USL=731.3 | Test Preblank Injections >> Test Preblank Injections >> USL=667.6 | Test Preblank Injections >> Test Preblank Injections >> USL=710.1 | Test Preblank Injections >> Test Preblank Injections >> USL=658.6 | Test Preblank Injections >> Test Preblank Injections >> USL=764.6 | Test Preblank Injections >> Test Preblank Injections >> USL=944.4 | Test Preblank Injections >> Test Preblank Injections >> USL=773.5 | Test Preblank Injections >> Test Preblank Injections >> USL=965.6 | Test Preblank Injections >> Test Preblank Injections >> USL=752.8 | Test Preblank Injections >> Test Preblank Injections >> USL=715.5 | Test Preblank Injections >> Test Preblank Injections >> USL=751.4 | Test Preblank Injections >> Test Preblank Injections >> USL=962.8 | Test Preblank Injections >> Test Preblank Injections >> USL=730.3 | Test Preblank Injections >> Test Preblank Injections >> USL=791.8 | Test Event Outputs >> Test Event Outputs | Test Preblank Injections >> Test Preblank Injections >> USL=681.4 | Test Preblank Injections >> Test Preblank Injections >> USL=727.7 | Test Preblank Injections >> Test Preblank Injections >> USL=758.6 | Test Preblank Injections >> Test Preblank Injections >> USL=731.2 | Test Preblank Injections >> Test Preblank Injections >> USL=909.5 | Test Preblank Injections >> Test Preblank Injections >> USL=252.5 | Test Preblank Injections >> Test Preblank Injections >> USL=967.1 | Test Preblank Injections >> Test Preblank Injections >> USL=601.3 | Test Preblank Injections >> Test Preblank Injections >> USL=635.1 | Test Preblank Injections >> Test Preblank Injections >> USL=621.4 | Test Preblank Injections >> Test Preblank Injections >> USL=763.1 | Test Preblank Injections >> Test Preblank Injections >> USL=757.3 | Test Preblank Injections >> Test Preblank Injections >> USL=740.6 | Test Preblank Injections >> Test Preblank Injections >> USL=790.6 | Test Preblank Injections >> Test Preblank Injections >> USL=625 | Test Preblank Injections >> Test Preblank Injections >> USL=943.3 | Test Preblank Injections >> Test Preblank Injections >> USL=768.6 | Test Preblank Injections >> Test Preblank Injections >> USL=707.5 | Test Preblank Injections >> Test Preblank Injections >> USL=934.5 | Test Preblank Injections >> Test Preblank Injections >> USL=733.6 | Test Preblank Injections >> Test Preblank Injections >> USL=965.9 | Test Preblank Injections >> Test Preblank Injections >> USL=712.6 | Test Preblank Injections >> Test Preblank Injections >> USL=802.1 | Test Preblank Injections >> Test Preblank Injections >> USL=730.9 | Test Preblank Injections >> Test Preblank Injections >> USL=784.2 | Test Preblank Injections >> Test Preblank Injections >> USL=763.3 | Test Preblank Injections >> Test Preblank Injections >> USL=718.9 | Test Preblank Injections >> Test Preblank Injections >> USL=990.7 | Test Preblank Injections >> Test Preblank Injections >> USL=771.3 | Test Preblank Injections >> Test Preblank Injections >> USL=772.1 | Test Preblank Injections >> Test Preblank Injections >> USL=741.9 | Test Preblank Injections >> Test Preblank Injections >> USL=735.9 | Test Preblank Injections >> Test Preblank Injections >> USL=976.4 | Test Preblank Injections >> Test Preblank Injections >> USL=740.3 | Test Preblank Injections >> Test Preblank Injections >> USL=777.1 | Test Preblank Injections >> Test Preblank Injections >> USL=738.9 | Test Preblank Injections >> Test Preblank Injections >> USL=791.7 | Test Preblank Injections >> Test Preblank Injections >> USL=761.3 | Test Preblank Injections >> Test Preblank Injections >> USL=725.9 | Test Preblank Injections >> Test Preblank Injections >> USL=774 | Test Preblank Injections >> Test Preblank Injections >> USL=788.5 | Test Preblank Injections >> Test Preblank Injections >> USL=595.5 | Test Preblank Injections >> Test Preblank Injections >> USL=971.1 | Test Preblank Injections >> Test Preblank Injections >> USL=755.4 | Test Preblank Injections >> Test Preblank Injections >> USL=710.3 | Test Preblank Injections >> Test Preblank Injections >> USL=781.5 | Test Preblank Injections >> Test Preblank Injections >> USL=768.8 | Test Preblank Injections >> Test Preblank Injections >> USL=784.4 | Test Preblank Injections >> Test Preblank Injections >> USL=796.7 | Test Preblank Injections >> Test Preblank Injections >> USL=731 | Test Preblank Injections >> Test Preblank Injections >> USL=779.6 | Test Preblank Injections >> Test Preblank Injections >> USL=733.1 | Test Preblank Injections >> Test Preblank Injections >> USL=795.2 | Test Preblank Injections >> Test Preblank Injections >> USL=751.2 | Test Preblank Injections >> Test Preblank Injections >> USL=758.7 | Test Preblank Injections >> Test Preblank Injections >> USL=760.7 | Test Preblank Injections >> Test Preblank Injections >> USL=914.2 | Test Preblank Injections >> Test Preblank Injections >> USL=755 | Test Preblank Injections >> Test Preblank Injections >> USL=771.6 | Test Preblank Injections >> Test Preblank Injections >> USL=653.4 | Test Preblank Injections >> Test Preblank Injections >> USL=746.5 | Test Preblank Injections >> Test Preblank Injections >> USL=766.9 | Test Preblank Injections >> Test Preblank Injections >> USL=753.5 | Test Preblank Injections >> Test Preblank Injections >> USL=764.1 | Test Preblank Injections >> Test Preblank Injections >> USL=739.9 | Test Preblank Injections >> Test Preblank Injections >> USL=745.6 | Test Preblank Injections >> Test Preblank Injections >> USL=793 | Test Preblank Injections >> Test Preblank Injections >> USL=794.2 | Test Preblank Injections >> Test Preblank Injections >> USL=866.2 | Test Preblank Injections >> Test Preblank Injections >> USL=714.5 | Test Preblank Injections >> Test Preblank Injections >> USL=749.3 | Test Preblank Injections >> Test Preblank Injections >> USL=739.4 | Test Preblank Injections >> Test Preblank Injections >> USL=841.4 | Test Preblank Injections >> Test Preblank Injections >> USL=1059.9 | Test Preblank Injections >> Test Preblank Injections >> USL=747.3 | Test Preblank Injections >> Test Preblank Injections >> USL=753.1 | Test Preblank Injections >> Test Preblank Injections >> USL=1398.9 | Test Preblank Injections >> Test Preblank Injections >> USL=757.2 | Test Preblank Injections >> Test Preblank Injections >> USL=749.2 | Test Preblank Injections >> Test Preblank Injections >> USL=759.4 | Test Event Inputs >> Test Event Inputs | Test Preblank Injections >> Test Preblank Injections >> USL=761.8 | Test Preblank Injections >> Test Preblank Injections >> USL=2116.3 | Test Preblank Injections >> Test Preblank Injections >> USL=711.1 | Test Preblank Injections >> Test Preblank Injections >> USL=745.5 | Test Preblank Injections >> Test Preblank Injections >> USL=713.9 | Test Leak Detectors >> Test Leak Detectors | Test Preblank Injections >> Test Preblank Injections >> USL=758.8 | Test Preblank Injections >> Test Preblank Injections >> USL=1475.6 | R/Theta Positioner Test >> R/Theta Positioner Test | Test Preblank Injections >> Test Preblank Injections >> USL=833 | Test Preblank Injections >> Test Preblank Injections >> USL=1543.8 | Test Preblank Injections >> Test Preblank Injections >> USL=828.3 | Test Preblank Injections >> Test Preblank Injections >> USL=756.3 | Test Preblank Injections >> Test Preblank Injections >> USL=791.3 | Test Preblank Injections >> Test Preblank Injections >> USL=795.7 | Test Preblank Injections >> Test Preblank Injections >> USL=777.8 | Test Preblank Injections >> Test Preblank Injections >> USL=744.5 | Test Preblank Injections >> Test Preblank Injections >> USL=919.7 | Test Preblank Injections >> Test Preblank Injections >> USL=655.4 | Test Preblank Injections >> Test Preblank Injections >> USL=732.1 | Test Preblank Injections >> Test Preblank Injections >> USL=711.5 | Test Preblank Injections >> Test Preblank Injections >> USL=771.2 | Test Preblank Injections >> Test Preblank Injections >> USL=717.7 | Test Preblank Injections >> Test Preblank Injections >> USL=784.1 | Test Preblank Injections >> Test Preblank Injections >> USL=695.1 | Test Preblank Injections >> Test Preblank Injections >> USL=783.2 | Test Preblank Injections >> Test Preblank Injections >> USL=781.8 | Test Preblank Injections >> Test Preblank Injections >> USL=662 | Test Preblank Injections >> Test Preblank Injections >> USL=771.7 | Test Preblank Injections >> Test Preblank Injections >> USL=774.2 | Test Preblank Injections >> Test Preblank Injections >> USL=767.3 | Test Preblank Injections >> Test Preblank Injections >> USL=770 | Test Preblank Injections >> Test Preblank Injections >> USL=766.7 | Test Preblank Injections >> Test Preblank Injections >> USL=713.3 | Test Preblank Injections >> Test Preblank Injections >> USL=749.8 | Test Preblank Injections >> Test Preblank Injections >> USL=767.4 | Test Preblank Injections >> Test Preblank Injections >> USL=795.5 | Test Preblank Injections >> Test Preblank Injections >> USL=770.5 | Test Preblank Injections >> Test Preblank Injections >> USL=793.6 | Test Preblank Injections >> Test Preblank Injections >> USL=674.1 | Test Preblank Injections >> Test Preblank Injections >> USL=747.1 | Test Preblank Injections >> Test Preblank Injections >> USL=690.3 | Test Preblank Injections >> Test Preblank Injections >> USL=805.3 | Test Preblank Injections >> Test Preblank Injections >> USL=776 | Test Preblank Injections >> Test Preblank Injections >> USL=769.4 | Test Preblank Injections >> Test Preblank Injections >> USL=763.5 | Test Preblank Injections >> Test Preblank Injections >> USL=796.9 | Test Preblank Injections >> Test Preblank Injections >> USL=807.9 | Test Preblank Injections >> Test Preblank Injections >> USL=792.8 | Test Preblank Injections >> Test Preblank Injections >> USL=639 | Test Preblank Injections >> Test Preblank Injections >> USL=768.3 | Test Preblank Injections >> Test Preblank Injections >> USL=731.1 | Test Preblank Injections >> Test Preblank Injections >> USL=782.6 | Test Preblank Injections >> Test Preblank Injections >> USL=782.4 | Test Preblank Injections >> Test Preblank Injections >> USL=753.2 | Test Preblank Injections >> Test Preblank Injections >> USL=776.6 | Test Preblank Injections >> Test Preblank Injections >> USL=779.1 | Test Preblank Injections >> Test Preblank Injections >> USL=646.4 | Test Preblank Injections >> Test Preblank Injections >> USL=754.5 | Test Preblank Injections >> Test Preblank Injections >> USL=785.9 | Test Preblank Injections >> Test Preblank Injections >> USL=463.9 | Test Preblank Injections >> Test Preblank Injections >> USL=777.9 | Test Preblank Injections >> Test Preblank Injections >> USL=715.8 | Test Preblank Injections >> Test Preblank Injections >> USL=737.5 | Test Preblank Injections >> Test Preblank Injections >> USL=497.5 | Test Preblank Injections >> Test Preblank Injections >> USL=816.5 | Test Preblank Injections >> Test Preblank Injections >> USL=779.3 | Test Preblank Injections >> Test Preblank Injections >> USL=677.9 | Test Preblank Injections >> Test Preblank Injections >> USL=662.8 | Test Preblank Injections >> Test Preblank Injections >> USL=716.9 | Test Preblank Injections >> Test Preblank Injections >> USL=642.1 | Test Preblank Injections >> Test Preblank Injections >> USL=700.9 | Test Preblank Injections >> Test Preblank Injections >> USL=722.5 | Test Preblank Injections >> Test Preblank Injections >> USL=697.5 | Test Preblank Injections >> Test Preblank Injections >> USL=750.9 | Test Preblank Injections >> Test Preblank Injections >> USL=741.7 | Test Preblank Injections >> Test Preblank Injections >> USL=735.7 | Test Preblank Injections >> Test Preblank Injections >> USL=758.2 | Test Preblank Injections >> Test Preblank Injections >> USL=607.3 | Test Preblank Injections >> Test Preblank Injections >> USL=627.6 | Test Preblank Injections >> Test Preblank Injections >> USL=790.1 | Test Preblank Injections >> Test Preblank Injections >> USL=744.8 | Test Preblank Injections >> Test Preblank Injections >> USL=538 | Test Preblank Injections >> Test Preblank Injections >> USL=770.7 | Test Preblank Injections >> Test Preblank Injections >> USL=458.4 | Test Preblank Injections >> Test Preblank Injections >> USL=751.8 | Test Preblank Injections >> Test Preblank Injections >> USL=779.2 | Test Preblank Injections >> Test Preblank Injections >> USL=663.2 | Test Preblank Injections >> Test Preblank Injections >> USL=706.6 | Test Preblank Injections >> Test Preblank Injections >> USL=717.8 | Test Preblank Injections >> Test Preblank Injections >> USL=635.7 | Test Preblank Injections >> Test Preblank Injections >> USL=718.4 | Test Preblank Injections >> Test Preblank Injections >> USL=786.9 | Test Preblank Injections >> Test Preblank Injections >> USL=536.4 | Test Preblank Injections >> Test Preblank Injections >> USL=703.8 | Test Preblank Injections >> Test Preblank Injections >> USL=738.5 | Test Preblank Injections >> Test Preblank Injections >> USL=761.6 | Test Preblank Injections >> Test Preblank Injections >> USL=752.1 | Test Preblank Injections >> Test Preblank Injections >> USL=622.8 | Test Preblank Injections >> Test Preblank Injections >> USL=836 | Test Preblank Injections >> Test Preblank Injections >> USL=568.5 | Test Preblank Injections >> Test Preblank Injections >> USL=699.7 | Test Preblank Injections >> Test Preblank Injections >> USL=760.9 | Test Preblank Injections >> Test Preblank Injections >> USL=663.7 | Test Preblank Injections >> Test Preblank Injections >> USL=739.2 | Test Preblank Injections >> Test Preblank Injections >> USL=767.9 | Test Preblank Injections >> Test Preblank Injections >> USL=757.6 | Test Preblank Injections >> Test Preblank Injections >> USL=762.1 | Test Preblank Injections >> Test Preblank Injections >> USL=652.4 | Test Preblank Injections >> Test Preblank Injections >> USL=686.5 | Test Preblank Injections >> Test Preblank Injections >> USL=801.5 | Test Preblank Injections >> Test Preblank Injections >> USL=776.4 | Test Preblank Injections >> Test Preblank Injections >> USL=799.3 | Test Preblank Injections >> Test Preblank Injections >> USL=721.2 | Test Preblank Injections >> Test Preblank Injections >> USL=785.1 | Test Preblank Injections >> Test Preblank Injections >> USL=700.3 | Test Preblank Injections >> Test Preblank Injections >> USL=802.7 | Test Preblank Injections >> Test Preblank Injections >> USL=727 | Test Preblank Injections >> Test Preblank Injections >> USL=602.4 | Test Preblank Injections >> Test Preblank Injections >> USL=538.9 | Test Preblank Injections >> Test Preblank Injections >> USL=752 | Test Preblank Injections >> Test Preblank Injections >> USL=765.9 | Test Preblank Injections >> Test Preblank Injections >> USL=668.6 | Test Preblank Injections >> Test Preblank Injections >> USL=652.8 | Test Preblank Injections >> Test Preblank Injections >> USL=768.2 | Test Preblank Injections >> Test Preblank Injections >> USL=855.9 | Test Preblank Injections >> Test Preblank Injections >> USL=727.5 | Test Preblank Injections >> Test Preblank Injections >> USL=782 | Test Preblank Injections >> Test Preblank Injections >> USL=460.3 | Test Preblank Injections >> Test Preblank Injections >> USL=617.4 | Test Preblank Injections >> Test Preblank Injections >> USL=581.7 | Test Preblank Injections >> Test Preblank Injections >> USL=794 | Test Preblank Injections >> Test Preblank Injections >> USL=761.5 | Test Preblank Injections >> Test Preblank Injections >> USL=766.3 | Test Preblank Injections >> Test Preblank Injections >> USL=713.5 | Test Preblank Injections >> Test Preblank Injections >> USL=723.6 | Test Preblank Injections >> Test Preblank Injections >> USL=728.7 | Test Preblank Injections >> Test Preblank Injections >> USL=740.5 | Test Preblank Injections >> Test Preblank Injections >> USL=790.4 | Test Preblank Injections >> Test Preblank Injections >> USL=780.4 | Test Preblank Injections >> Test Preblank Injections >> USL=699.1 | Test Preblank Injections >> Test Preblank Injections >> USL=851.2 | Test Preblank Injections >> Test Preblank Injections >> USL=815 | Test Preblank Injections >> Test Preblank Injections >> USL=770.8 | Test Preblank Injections >> Test Preblank Injections >> USL=729.2 | Test Preblank Injections >> Test Preblank Injections >> USL=794.1 | Test Preblank Injections >> Test Preblank Injections >> USL=717 | Test Preblank Injections >> Test Preblank Injections >> USL=612.3 | Test Preblank Injections >> Test Preblank Injections >> USL=687.3 | Test Preblank Injections >> Test Preblank Injections >> USL=735.2 | Test Preblank Injections >> Test Preblank Injections >> USL=849 | Test Preblank Injections >> Test Preblank Injections >> USL=719.3 | Check Theta Drive Belt >> Check Theta Drive Belt | Test Preblank Injections >> Test Preblank Injections >> USL=862 | Test Preblank Injections >> Test Preblank Injections >> USL=789.4 | Test Preblank Injections >> Test Preblank Injections >> USL=639.1 | Test Preblank Injections >> Test Preblank Injections >> USL=629.6 | Test Preblank Injections >> Test Preblank Injections >> USL=759.2 | Test Preblank Injections >> Test Preblank Injections >> USL=656.7 | Test Preblank Injections >> Test Preblank Injections >> USL=785.4 | Test Preblank Injections >> Test Preblank Injections >> USL=626.4 | Test Preblank Injections >> Test Preblank Injections >> USL=696 | Test Preblank Injections >> Test Preblank Injections >> USL=865.2 | Test Preblank Injections >> Test Preblank Injections >> USL=799.4 | Test Preblank Injections >> Test Preblank Injections >> USL=699.9 | Test Preblank Injections >> Test Preblank Injections >> USL=720.4 | Test Preblank Injections >> Test Preblank Injections >> USL=729.4 | Test Preblank Injections >> Test Preblank Injections >> USL=737.3 | Test Preblank Injections >> Test Preblank Injections >> USL=522.8 | Test Preblank Injections >> Test Preblank Injections >> USL=750.3 | Test Preblank Injections >> Test Preblank Injections >> USL=732.5 | Test Preblank Injections >> Test Preblank Injections >> USL=729.6 | Test Preblank Injections >> Test Preblank Injections >> USL=689 | Test Preblank Injections >> Test Preblank Injections >> USL=762.2 | Test Preblank Injections >> Test Preblank Injections >> USL=667.9 | Test Preblank Injections >> Test Preblank Injections >> USL=772.6 | Test Preblank Injections >> Test Preblank Injections >> USL=771 | Test Preblank Injections >> Test Preblank Injections >> USL=777.5 | Test Preblank Injections >> Test Preblank Injections >> USL=774.6 | Test Preblank Injections >> Test Preblank Injections >> USL=782.1 | Test Preblank Injections >> Test Preblank Injections >> USL=681.9 | Test Preblank Injections >> Test Preblank Injections >> USL=708 | Test Preblank Injections >> Test Preblank Injections >> USL=789.8 | Test Preblank Injections >> Test Preblank Injections >> USL=802.6 | Test Preblank Injections >> Test Preblank Injections >> USL=730 | Test Preblank Injections >> Test Preblank Injections >> USL=775.2 | Test Preblank Injections >> Test Preblank Injections >> USL=746.4 | Test Preblank Injections >> Test Preblank Injections >> USL=810.6 | Test Preblank Injections >> Test Preblank Injections >> USL=743.9 | Test Preblank Injections >> Test Preblank Injections >> USL=791.1 | Test Preblank Injections >> Test Preblank Injections >> USL=798 | Test Preblank Injections >> Test Preblank Injections >> USL=813.3 | Test Preblank Injections >> Test Preblank Injections >> USL=783.6 | Test Preblank Injections >> Test Preblank Injections >> USL=738.3 | Test Preblank Injections >> Test Preblank Injections >> USL=753.6 | Test Preblank Injections >> Test Preblank Injections >> USL=829.2 | Test Preblank Injections >> Test Preblank Injections >> USL=743.1 | Test Preblank Injections >> Test Preblank Injections >> USL=781.9 | Test Preblank Injections >> Test Preblank Injections >> USL=783.3 | Test Preblank Injections >> Test Preblank Injections >> USL=733.2 | Test Preblank Injections >> Test Preblank Injections >> USL=822.8 | Test Preblank Injections >> Test Preblank Injections >> USL=695 | Test Preblank Injections >> Test Preblank Injections >> USL=741.2 | Test Preblank Injections >> Test Preblank Injections >> USL=656.2 | Sample Compartment Door & Lights >> Sample Compartment Door & Lights | Test Preblank Injections >> Test Preblank Injections >> USL=786.5 | Test Preblank Injections >> Test Preblank Injections >> USL=787.4 | Test Preblank Injections >> Test Preblank Injections >> USL=742.6 | Test Preblank Injections >> Test Preblank Injections >> USL=595.3 | Test Preblank Injections >> Test Preblank Injections >> USL=784.5 | Test Preblank Injections >> Test Preblank Injections >> USL=728.9 | Test Preblank Injections >> Test Preblank Injections >> USL=607.5 | Test Preblank Injections >> Test Preblank Injections >> USL=774.3 | Test Preblank Injections >> Test Preblank Injections >> USL=748 | Test Preblank Injections >> Test Preblank Injections >> USL=787.1 | Test Preblank Injections >> Test Preblank Injections >> USL=740.4 | Test Preblank Injections >> Test Preblank Injections >> USL=734.1 | Test Preblank Injections >> Test Preblank Injections >> USL=797.2 | Test Preblank Injections >> Test Preblank Injections >> USL=745.4 | Test Preblank Injections >> Test Preblank Injections >> USL=772.2 | Test Preblank Injections >> Test Preblank Injections >> USL=631.9 | Test Preblank Injections >> Test Preblank Injections >> USL=759.8 | Test Preblank Injections >> Test Preblank Injections >> USL=798.6 | Test Preblank Injections >> Test Preblank Injections >> USL=603.6 | Test Preblank Injections >> Test Preblank Injections >> USL=439.7 | Test Preblank Injections >> Test Preblank Injections >> USL=652.7 | Test Preblank Injections >> Test Preblank Injections >> USL=678.2 | Test Preblank Injections >> Test Preblank Injections >> USL=772.4 | Test Preblank Injections >> Test Preblank Injections >> USL=740.8 | Test Preblank Injections >> Test Preblank Injections >> USL=762.9 | Test Preblank Injections >> Test Preblank Injections >> USL=534.5 | Test Preblank Injections >> Test Preblank Injections >> USL=698.3 | Test Preblank Injections >> Test Preblank Injections >> USL=650.6 | Test Preblank Injections >> Test Preblank Injections >> USL=555 | Test Preblank Injections >> Test Preblank Injections >> USL=754.3 | Test Preblank Injections >> Test Preblank Injections >> USL=789.2 | Test Preblank Injections >> Test Preblank Injections >> USL=582.9 | Test Preblank Injections >> Test Preblank Injections >> USL=712.3 | Test Preblank Injections >> Test Preblank Injections >> USL=791 | Test Preblank Injections >> Test Preblank Injections >> USL=767.6 | Test Preblank Injections >> Test Preblank Injections >> USL=726.9 | Test Preblank Injections >> Test Preblank Injections >> USL=756 | Test Preblank Injections >> Test Preblank Injections >> USL=758 | Test Preblank Injections >> Test Preblank Injections >> USL=686.8 | Test Preblank Injections >> Test Preblank Injections >> USL=406.8 | Test Preblank Injections >> Test Preblank Injections >> USL=660.4 | Test Preblank Injections >> Test Preblank Injections >> USL=719 | Test Preblank Injections >> Test Preblank Injections >> USL=711.3 | Test Preblank Injections >> Test Preblank Injections >> USL=726.1 | Test Preblank Injections >> Test Preblank Injections >> USL=704.1 | Test Preblank Injections >> Test Preblank Injections >> USL=29920.4 | Test Preblank Injections >> Test Preblank Injections >> USL=658.5 | Test Preblank Injections >> Test Preblank Injections >> USL=619.2 | Test Preblank Injections >> Test Preblank Injections >> USL=665.3 | Test Preblank Injections >> Test Preblank Injections >> USL=698.5 | Test Preblank Injections >> Test Preblank Injections >> USL=671.1 | Test Preblank Injections >> Test Preblank Injections >> USL=639.6 | Test Preblank Injections >> Test Preblank Injections >> USL=748.5 | Test Preblank Injections >> Test Preblank Injections >> USL=685 | Test Preblank Injections >> Test Preblank Injections >> USL=714.8 | Test Preblank Injections >> Test Preblank Injections >> USL=564.6 | Test Preblank Injections >> Test Preblank Injections >> USL=624.6 | Sample Tray Select Switch >> Sample Tray Select Switch | Test Preblank Injections >> Test Preblank Injections >> USL=566.6 | Test Preblank Injections >> Test Preblank Injections >> USL=718.2 | Test Preblank Injections >> Test Preblank Injections >> USL=714.4 | Test Preblank Injections >> Test Preblank Injections >> USL=586.4 | Test Preblank Injections >> Test Preblank Injections >> USL=669.9 | Test Preblank Injections >> Test Preblank Injections >> USL=762.6 | Test Preblank Injections >> Test Preblank Injections >> USL=645.1 | Test Preblank Injections >> Test Preblank Injections >> USL=711 | Test Preblank Injections >> Test Preblank Injections >> USL=769.5 | Test Preblank Injections >> Test Preblank Injections >> USL=730.8 | Test Preblank Injections >> Test Preblank Injections >> USL=647 | Test Preblank Injections >> Test Preblank Injections >> USL=683.5 | Test Preblank Injections >> Test Preblank Injections >> USL=603.9 | Test Preblank Injections >> Test Preblank Injections >> USL=744.9 | Test Preblank Injections >> Test Preblank Injections >> USL=700.2 | Test Preblank Injections >> Test Preblank Injections >> USL=706.3 | Test Preblank Injections >> Test Preblank Injections >> USL=589.8 | Test Preblank Injections >> Test Preblank Injections >> USL=736.2 | Test Preblank Injections >> Test Preblank Injections >> USL=766.6 | Test Preblank Injections >> Test Preblank Injections >> USL=524 | Test Preblank Injections >> Test Preblank Injections >> USL=698 | Test Preblank Injections >> Test Preblank Injections >> USL=615.1 | Test Preblank Injections >> Test Preblank Injections >> USL=707.3 | Test Preblank Injections >> Test Preblank Injections >> USL=545.3 | Test Preblank Injections >> Test Preblank Injections >> USL=562.7 | Test Preblank Injections >> Test Preblank Injections >> USL=668.7 | Test Preblank Injections >> Test Preblank Injections >> USL=674.3 | Test Preblank Injections >> Test Preblank Injections >> USL=659.8 | Test Preblank Injections >> Test Preblank Injections >> USL=609.5 | Test Preblank Injections >> Test Preblank Injections >> USL=742 | Test Preblank Injections >> Test Preblank Injections >> USL=614.8 | Test Preblank Injections >> Test Preblank Injections >> USL=517.1 | Test Preblank Injections >> Test Preblank Injections >> USL=750.4 | Test Preblank Injections >> Test Preblank Injections >> USL=676.5 | Test Preblank Injections >> Test Preblank Injections >> USL=694.5 | Test Preblank Injections >> Test Preblank Injections >> USL=719.2 | Test Preblank Injections >> Test Preblank Injections >> USL=572.6 | Test Preblank Injections >> Test Preblank Injections >> USL=724.4 | Test Preblank Injections >> Test Preblank Injections >> USL=776.1 | Test Preblank Injections >> Test Preblank Injections >> USL=770.9 | Test Preblank Injections >> Test Preblank Injections >> USL=569 | Test Preblank Injections >> Test Preblank Injections >> USL=692.8 | Test Preblank Injections >> Test Preblank Injections >> USL=705.5 | Test Preblank Injections >> Test Preblank Injections >> USL=617.2 | Test Preblank Injections >> Test Preblank Injections >> USL=513.3 | Test Preblank Injections >> Test Preblank Injections >> USL=482 | Test Preblank Injections >> Test Preblank Injections >> USL=681.8 | Test Preblank Injections >> Test Preblank Injections >> USL=665 | Test Preblank Injections >> Test Preblank Injections >> USL=476.2 | Test Preblank Injections >> Test Preblank Injections >> USL=710.5 | Test Preblank Injections >> Test Preblank Injections >> USL=661.1 | Test Preblank Injections >> Test Preblank Injections >> USL=711.6 | Test Preblank Injections >> Test Preblank Injections >> USL=812.2 | Test Preblank Injections >> Test Preblank Injections >> USL=679.2 | Test Preblank Injections >> Test Preblank Injections >> USL=589.9 | Test Preblank Injections >> Test Preblank Injections >> USL=521.9 | Test Preblank Injections >> Test Preblank Injections >> USL=673.2 | Test Preblank Injections >> Test Preblank Injections >> USL=683.7 | Test Preblank Injections >> Test Preblank Injections >> USL=645.6 | Test Preblank Injections >> Test Preblank Injections >> USL=715.7 | Test Preblank Injections >> Test Preblank Injections >> USL=679.6 | Test Preblank Injections >> Test Preblank Injections >> USL=462.7 | Test Preblank Injections >> Test Preblank Injections >> USL=699.6 | Test Preblank Injections >> Test Preblank Injections >> USL=668.9 | Test Preblank Injections >> Test Preblank Injections >> USL=707.7 | Test Preblank Injections >> Test Preblank Injections >> USL=636.7 | Test Preblank Injections >> Test Preblank Injections >> USL=376.9 | Test Preblank Injections >> Test Preblank Injections >> USL=269.4 | Test Preblank Injections >> Test Preblank Injections >> USL=619.3 | Test Preblank Injections >> Test Preblank Injections >> USL=739.8 | Test Preblank Injections >> Test Preblank Injections >> USL=778.6 | Test Preblank Injections >> Test Preblank Injections >> USL=725.1 | Test Preblank Injections >> Test Preblank Injections >> USL=735.5 | Test Preblank Injections >> Test Preblank Injections >> USL=773 | Test Preblank Injections >> Test Preblank Injections >> USL=693.5 | Test Preblank Injections >> Test Preblank Injections >> USL=748.2 | Test Preblank Injections >> Test Preblank Injections >> USL=757.8 | Test Preblank Injections >> Test Preblank Injections >> USL=795.6 | Test Preblank Injections >> Test Preblank Injections >> USL=716 | Test Preblank Injections >> Test Preblank Injections >> USL=729 | Test Preblank Injections >> Test Preblank Injections >> USL=688.3 | Test Preblank Injections >> Test Preblank Injections >> USL=760 | Test Preblank Injections >> Test Preblank Injections >> USL=588 | Test Preblank Injections >> Test Preblank Injections >> USL=792.4 | Test Preblank Injections >> Test Preblank Injections >> USL=757.1 | Test Preblank Injections >> Test Preblank Injections >> USL=681.6 | Test Preblank Injections >> Test Preblank Injections >> USL=664.2 | Test Preblank Injections >> Test Preblank Injections >> USL=933.5 | Test Preblank Injections >> Test Preblank Injections >> USL=668.1 | Test Preblank Injections >> Test Preblank Injections >> USL=689.2 | Test Preblank Injections >> Test Preblank Injections >> USL=761.1 | Test Preblank Injections >> Test Preblank Injections >> USL=737.9 | Test Preblank Injections >> Test Preblank Injections >> USL=667.2 | Test Preblank Injections >> Test Preblank Injections >> USL=719.9 | Test Preblank Injections >> Test Preblank Injections >> USL=786.3 | Test Preblank Injections >> Test Preblank Injections >> USL=710.6 | Test Preblank Injections >> Test Preblank Injections >> USL=762 | Test Preblank Injections >> Test Preblank Injections >> USL=644.2 | Test Preblank Injections >> Test Preblank Injections >> USL=612.4 | BOS Pump Test >> BOS Pump Test | Test Preblank Injections >> Test Preblank Injections >> USL=684.9 | Test Preblank Injections >> Test Preblank Injections >> USL=765.7 | Test Preblank Injections >> Test Preblank Injections >> USL=628.1 | Test Preblank Injections >> Test Preblank Injections >> USL=651.8 | Test Preblank Injections >> Test Preblank Injections >> USL=781.3 | Test Preblank Injections >> Test Preblank Injections >> USL=715 | Test Preblank Injections >> Test Preblank Injections >> USL=739.7 | Test Preblank Injections >> Test Preblank Injections >> USL=561.9 | Test Preblank Injections >> Test Preblank Injections >> USL=617.6 | Test Preblank Injections >> Test Preblank Injections >> USL=491.5 | Test Preblank Injections >> Test Preblank Injections >> USL=748.7 | Test Preblank Injections >> Test Preblank Injections >> USL=783.5 | Test Preblank Injections >> Test Preblank Injections >> USL=728.3 | Test Preblank Injections >> Test Preblank Injections >> USL=783 | Test Preblank Injections >> Test Preblank Injections >> USL=706.5 | Test Preblank Injections >> Test Preblank Injections >> USL=707 | Test Preblank Injections >> Test Preblank Injections >> USL=689.6 | Test Preblank Injections >> Test Preblank Injections >> USL=799.6 | Test Preblank Injections >> Test Preblank Injections >> USL=733.7 | Test Preblank Injections >> Test Preblank Injections >> USL=754.2 | Test Preblank Injections >> Test Preblank Injections >> USL=746.2 | Test Preblank Injections >> Test Preblank Injections >> USL=793.5 | Test Preblank Injections >> Test Preblank Injections >> USL=758.3 | Test Preblank Injections >> Test Preblank Injections >> USL=742.2 | Test Preblank Injections >> Test Preblank Injections >> USL=739.6 | Test Preblank Injections >> Test Preblank Injections >> USL=733.8 | Test Preblank Injections >> Test Preblank Injections >> USL=694.4 | Test Preblank Injections >> Test Preblank Injections >> USL=762.7 | Test Preblank Injections >> Test Preblank Injections >> USL=765.5 | Test Preblank Injections >> Test Preblank Injections >> USL=731.4 | Test Preblank Injections >> Test Preblank Injections >> USL=760.4 | Test Preblank Injections >> Test Preblank Injections >> USL=682.9 | Test Preblank Injections >> Test Preblank Injections >> USL=785.8 | Test Preblank Injections >> Test Preblank Injections >> USL=686.9 | Test Preblank Injections >> Test Preblank Injections >> USL=770.3 | Test Preblank Injections >> Test Preblank Injections >> USL=629.2 | Test Preblank Injections >> Test Preblank Injections >> USL=619.4 | Test Preblank Injections >> Test Preblank Injections >> USL=302.4 | Test Preblank Injections >> Test Preblank Injections >> USL=718 | Test Preblank Injections >> Test Preblank Injections >> USL=984.9 | Test Preblank Injections >> Test Preblank Injections >> USL=779.5 | Test Preblank Injections >> Test Preblank Injections >> USL=777 | Test Preblank Injections >> Test Preblank Injections >> USL=530.6 | Test Preblank Injections >> Test Preblank Injections >> USL=674.4 | Test Preblank Injections >> Test Preblank Injections >> USL=675.2 | Test Preblank Injections >> Test Preblank Injections >> USL=764.7 | Test Preblank Injections >> Test Preblank Injections >> USL=753.3 | Test Preblank Injections >> Test Preblank Injections >> USL=638.2 | Test Preblank Injections >> Test Preblank Injections >> USL=719.4 | Test Preblank Injections >> Test Preblank Injections >> USL=736.3 | Test Preblank Injections >> Test Preblank Injections >> USL=760.5 | Test Preblank Injections >> Test Preblank Injections >> USL=638.7 | Test Preblank Injections >> Test Preblank Injections >> USL=648.8 | Test Preblank Injections >> Test Preblank Injections >> USL=761 | Test Preblank Injections >> Test Preblank Injections >> USL=766.2 | Test Preblank Injections >> Test Preblank Injections >> USL=615.3 | Test Preblank Injections >> Test Preblank Injections >> USL=790 | Test Preblank Injections >> Test Preblank Injections >> USL=511.8 | Test Preblank Injections >> Test Preblank Injections >> USL=702.4 | Test Preblank Injections >> Test Preblank Injections >> USL=746.6 | Test Preblank Injections >> Test Preblank Injections >> USL=603.8 | Test Preblank Injections >> Test Preblank Injections >> USL=763.7 | Test Preblank Injections >> Test Preblank Injections >> USL=662.9 | Test Preblank Injections >> Test Preblank Injections >> USL=712.5 | Test Preblank Injections >> Test Preblank Injections >> USL=704.7 | Test Preblank Injections >> Test Preblank Injections >> USL=698.2 | Test Preblank Injections >> Test Preblank Injections >> USL=766.1 | Test Preblank Injections >> Test Preblank Injections >> USL=603.7 | Test Preblank Injections >> Test Preblank Injections >> USL=754.8 | Test Preblank Injections >> Test Preblank Injections >> USL=895.5 | Test Preblank Injections >> Test Preblank Injections >> USL=788.1 | Test Preblank Injections >> Test Preblank Injections >> USL=728.4 | Test Preblank Injections >> Test Preblank Injections >> USL=1071.5 | Test Preblank Injections >> Test Preblank Injections >> USL=543.4 | Test Preblank Injections >> Test Preblank Injections >> USL=751.3 | Test Preblank Injections >> Test Preblank Injections >> USL=644.6 | Test Preblank Injections >> Test Preblank Injections >> USL=697.2 | Test Preblank Injections >> Test Preblank Injections >> USL=720 | Test Preblank Injections >> Test Preblank Injections >> USL=756.7 | Test Preblank Injections >> Test Preblank Injections >> USL=748.9 | Test Preblank Injections >> Test Preblank Injections >> USL=738 | Test Preblank Injections >> Test Preblank Injections >> USL=755.3 | Test Preblank Injections >> Test Preblank Injections >> USL=709.3 | Test Preblank Injections >> Test Preblank Injections >> USL=592.3 | Test Preblank Injections >> Test Preblank Injections >> USL=690.6 | Test Preblank Injections >> Test Preblank Injections >> USL=736 | Test Preblank Injections >> Test Preblank Injections >> USL=769.6 | Test Preblank Injections >> Test Preblank Injections >> USL=755.7 | Test Preblank Injections >> Test Preblank Injections >> USL=655.9 | Test Preblank Injections >> Test Preblank Injections >> USL=706.4 | Test Preblank Injections >> Test Preblank Injections >> USL=679.7 | Test Preblank Injections >> Test Preblank Injections >> USL=742.5 | Test Preblank Injections >> Test Preblank Injections >> USL=616.1 | Test Preblank Injections >> Test Preblank Injections >> USL=635.6 | Test Preblank Injections >> Test Preblank Injections >> USL=607 | Test Preblank Injections >> Test Preblank Injections >> USL=678.5 | Test Preblank Injections >> Test Preblank Injections >> USL=611.3 | Test Preblank Injections >> Test Preblank Injections >> USL=638.1 | Test Preblank Injections >> Test Preblank Injections >> USL=700.4 | Test Preblank Injections >> Test Preblank Injections >> USL=670.4 | Test Preblank Injections >> Test Preblank Injections >> USL=724.6 | Test Preblank Injections >> Test Preblank Injections >> USL=701.8 | Test Preblank Injections >> Test Preblank Injections >> USL=693.7 | Test Preblank Injections >> Test Preblank Injections >> USL=678 | Test Preblank Injections >> Test Preblank Injections >> USL=620.7 | Test Preblank Injections >> Test Preblank Injections >> USL=698.6 | Test Preblank Injections >> Test Preblank Injections >> USL=543.9 | Test Preblank Injections >> Test Preblank Injections >> USL=678.9 | Test Preblank Injections >> Test Preblank Injections >> USL=653.1 | Test Preblank Injections >> Test Preblank Injections >> USL=591 | Test Preblank Injections >> Test Preblank Injections >> USL=692.3 | Test Preblank Injections >> Test Preblank Injections >> USL=594.1 | Test Preblank Injections >> Test Preblank Injections >> USL=692.9 | Test Preblank Injections >> Test Preblank Injections >> USL=619 | Test Preblank Injections >> Test Preblank Injections >> USL=550 | Test Preblank Injections >> Test Preblank Injections >> USL=720.1 | Test Preblank Injections >> Test Preblank Injections >> USL=647.1 | Test Preblank Injections >> Test Preblank Injections >> USL=535.5 | Test Preblank Injections >> Test Preblank Injections >> USL=586.3 | Test Preblank Injections >> Test Preblank Injections >> USL=637.7 | Test Preblank Injections >> Test Preblank Injections >> USL=574.4 | Test Preblank Injections >> Test Preblank Injections >> USL=765.6 | Test Preblank Injections >> Test Preblank Injections >> USL=583.8 | Test Preblank Injections >> Test Preblank Injections >> USL=732.7 | Test Preblank Injections >> Test Preblank Injections >> USL=735.6 | Test Preblank Injections >> Test Preblank Injections >> USL=691 | Test Preblank Injections >> Test Preblank Injections >> USL=718.1 | Test Preblank Injections >> Test Preblank Injections >> USL=594.2 | Test Preblank Injections >> Test Preblank Injections >> USL=536.2 | Test Preblank Injections >> Test Preblank Injections >> USL=669.8 | Test Preblank Injections >> Test Preblank Injections >> USL=686.2 | Test Preblank Injections >> Test Preblank Injections >> USL=734.9 | Test Preblank Injections >> Test Preblank Injections >> USL=676.2 | Test Preblank Injections >> Test Preblank Injections >> USL=467 | Test Preblank Injections >> Test Preblank Injections >> USL=438.4 | Test Preblank Injections >> Test Preblank Injections >> USL=758.4 | Test Preblank Injections >> Test Preblank Injections >> USL=772.7 | Test Preblank Injections >> Test Preblank Injections >> USL=686.1 | Test Preblank Injections >> Test Preblank Injections >> USL=623 | Test Preblank Injections >> Test Preblank Injections >> USL=713.8 | Test Preblank Injections >> Test Preblank Injections >> USL=732.6 | Test Preblank Injections >> Test Preblank Injections >> USL=725.6 | Test Preblank Injections >> Test Preblank Injections >> USL=743.2 | Test Preblank Injections >> Test Preblank Injections >> USL=764.3 | Test Preblank Injections >> Test Preblank Injections >> USL=811.4 | Test Preblank Injections >> Test Preblank Injections >> USL=743.7 | Test Preblank Injections >> Test Preblank Injections >> USL=796.8 | Test Preblank Injections >> Test Preblank Injections >> USL=758.1 | Test Preblank Injections >> Test Preblank Injections >> USL=786.6 | Test Preblank Injections >> Test Preblank Injections >> USL=725.7 | Test Preblank Injections >> Test Preblank Injections >> USL=462.6 | Test Preblank Injections >> Test Preblank Injections >> USL=641.2 | Test Preblank Injections >> Test Preblank Injections >> USL=657.1 | Test Preblank Injections >> Test Preblank Injections >> USL=646.3 | Test Preblank Injections >> Test Preblank Injections >> USL=531.1 | Test Preblank Injections >> Test Preblank Injections >> USL=664.3 | Test Preblank Injections >> Test Preblank Injections >> USL=690.5 | Test Preblank Injections >> Test Preblank Injections >> USL=731.8 | Test Preblank Injections >> Test Preblank Injections >> USL=806.4 | Test Preblank Injections >> Test Preblank Injections >> USL=710 | Test Preblank Injections >> Test Preblank Injections >> USL=767.8 | Test Preblank Injections >> Test Preblank Injections >> USL=748.4 | Test Preblank Injections >> Test Preblank Injections >> USL=796 | Test Preblank Injections >> Test Preblank Injections >> USL=751 | Test Preblank Injections >> Test Preblank Injections >> USL=771.1 | Test Preblank Injections >> Test Preblank Injections >> USL=722 | Test Preblank Injections >> Test Preblank Injections >> USL=760.6 | Test Preblank Injections >> Test Preblank Injections >> USL=697.4 | Test Preblank Injections >> Test Preblank Injections >> USL=757 | Test Preblank Injections >> Test Preblank Injections >> USL=737.4 | Test Preblank Injections >> Test Preblank Injections >> USL=864.3 | Test Preblank Injections >> Test Preblank Injections >> USL=724.3 | Test Preblank Injections >> Test Preblank Injections >> USL=780 | Test Preblank Injections >> Test Preblank Injections >> USL=783.1 | Test Preblank Injections >> Test Preblank Injections >> USL=714.2 | Test Preblank Injections >> Test Preblank Injections >> USL=716.8 | Test Preblank Injections >> Test Preblank Injections >> USL=775.7 | Test Preblank Injections >> Test Preblank Injections >> USL=950.9 | Test Preblank Injections >> Test Preblank Injections >> USL=724.1 | Test Preblank Injections >> Test Preblank Injections >> USL=667.7 | Test Preblank Injections >> Test Preblank Injections >> USL=774.5 | Test Preblank Injections >> Test Preblank Injections >> USL=745 | Test Preblank Injections >> Test Preblank Injections >> USL=721.6 | Test Preblank Injections >> Test Preblank Injections >> USL=704.4 | Test Preblank Injections >> Test Preblank Injections >> USL=946.1 | Test Preblank Injections >> Test Preblank Injections >> USL=748.1 | Test Preblank Injections >> Test Preblank Injections >> USL=722.3 | Test Preblank Injections >> Test Preblank Injections >> USL=969.1 | Test Preblank Injections >> Test Preblank Injections >> USL=1366.5 | Test Preblank Injections >> Test Preblank Injections >> USL=911.2 | Test Preblank Injections >> Test Preblank Injections >> USL=778.7 | Test Preblank Injections >> Test Preblank Injections >> USL=690.7 | Test Preblank Injections >> Test Preblank Injections >> USL=493.7 | Test Preblank Injections >> Test Preblank Injections >> USL=727.3 | Test Preblank Injections >> Test Preblank Injections >> USL=617 | Test Preblank Injections >> Test Preblank Injections >> USL=907.8 | Test Preblank Injections >> Test Preblank Injections >> USL=1373.7 | Test Preblank Injections >> Test Preblank Injections >> USL=1219.9 | Test Preblank Injections >> Test Preblank Injections >> USL=793.4 | Test Preblank Injections >> Test Preblank Injections >> USL=750.5 | Test Preblank Injections >> Test Preblank Injections >> USL=795.4 | Test Preblank Injections >> Test Preblank Injections >> USL=650.4 | Test Preblank Injections >> Test Preblank Injections >> USL=792.6 | Test Preblank Injections >> Test Preblank Injections >> USL=712.4 | Test Preblank Injections >> Test Preblank Injections >> USL=1100.1 | Test Preblank Injections >> Test Preblank Injections >> USL=632.1 | Test Preblank Injections >> Test Preblank Injections >> USL=781 | Test Preblank Injections >> Test Preblank Injections >> USL=1311.9 | Test Preblank Injections >> Test Preblank Injections >> USL=753.9 | Test Preblank Injections >> Test Preblank Injections >> USL=544.5 | Test Preblank Injections >> Test Preblank Injections >> USL=745.1 | Test Preblank Injections >> Test Preblank Injections >> USL=787.6 | Test Preblank Injections >> Test Preblank Injections >> USL=753.8 | Test Preblank Injections >> Test Preblank Injections >> USL=652.9 | Test Preblank Injections >> Test Preblank Injections >> USL=716.1 | Test Preblank Injections >> Test Preblank Injections >> USL=726.4 | Test Preblank Injections >> Test Preblank Injections >> USL=754.7 | Test Preblank Injections >> Test Preblank Injections >> USL=702.8 | Test Preblank Injections >> Test Preblank Injections >> USL=714.7 | Test Preblank Injections >> Test Preblank Injections >> USL=682.7 | Test Preblank Injections >> Test Preblank Injections >> USL=584.2 | Test Preblank Injections >> Test Preblank Injections >> USL=664 | Test Preblank Injections >> Test Preblank Injections >> USL=742.3 | Test Preblank Injections >> Test Preblank Injections >> USL=745.2 | Test Preblank Injections >> Test Preblank Injections >> USL=744.7 | Test Preblank Injections >> Test Preblank Injections >> USL=716.2 | Test Preblank Injections >> Test Preblank Injections >> USL=738.4 | Test Preblank Injections >> Test Preblank Injections >> USL=708.2 | Test Preblank Injections >> Test Preblank Injections >> USL=665.2 | Test Preblank Injections >> Test Preblank Injections >> USL=670.2 | Test Preblank Injections >> Test Preblank Injections >> USL=597.6 | Test Preblank Injections >> Test Preblank Injections >> USL=772.3 | Test Preblank Injections >> Test Preblank Injections >> USL=688.4 | Test Preblank Injections >> Test Preblank Injections >> USL=743.5 | Test Preblank Injections >> Test Preblank Injections >> USL=732.8 | Test Preblank Injections >> Test Preblank Injections >> USL=756.1 | Test Preblank Injections >> Test Preblank Injections >> USL=751.9 | Test Preblank Injections >> Test Preblank Injections >> USL=763.6 | Test Preblank Injections >> Test Preblank Injections >> USL=605.2 | Test Preblank Injections >> Test Preblank Injections >> USL=699.2 | Test Preblank Injections >> Test Preblank Injections >> USL=775.9 | Test Preblank Injections >> Test Preblank Injections >> USL=734.4 | Test Preblank Injections >> Test Preblank Injections >> USL=745.7 | Test Preblank Injections >> Test Preblank Injections >> USL=761.2 | Test Preblank Injections >> Test Preblank Injections >> USL=781.7 | Test Preblank Injections >> Test Preblank Injections >> USL=635.8 | Test Preblank Injections >> Test Preblank Injections >> USL=697.7 | Test Preblank Injections >> Test Preblank Injections >> USL=786.4 | Test Preblank Injections >> Test Preblank Injections >> USL=688 | Test Preblank Injections >> Test Preblank Injections >> USL=801.3 | Test Preblank Injections >> Test Preblank Injections >> USL=608.5 | Test Preblank Injections >> Test Preblank Injections >> USL=710.7 | Test Preblank Injections >> Test Preblank Injections >> USL=792.1 | Test Preblank Injections >> Test Preblank Injections >> USL=721.1 | Test Preblank Injections >> Test Preblank Injections >> USL=695.6 | Test Preblank Injections >> Test Preblank Injections >> USL=757.9 | Test Preblank Injections >> Test Preblank Injections >> USL=720.5 | Test Preblank Injections >> Test Preblank Injections >> USL=670.5 | Test Preblank Injections >> Test Preblank Injections >> USL=787.2 | Test Preblank Injections >> Test Preblank Injections >> USL=730.5 | Test Preblank Injections >> Test Preblank Injections >> USL=754.9 | Test Preblank Injections >> Test Preblank Injections >> USL=730.7 | Test Preblank Injections >> Test Preblank Injections >> USL=774.4 | Test Preblank Injections >> Test Preblank Injections >> USL=784.6 | Test Preblank Injections >> Test Preblank Injections >> USL=777.6 | Test Preblank Injections >> Test Preblank Injections >> USL=797.9 | Test Preblank Injections >> Test Preblank Injections >> USL=788.9 | Test Preblank Injections >> Test Preblank Injections >> USL=788.2 | Test Preblank Injections >> Test Preblank Injections >> USL=785.6 | Test Preblank Injections >> Test Preblank Injections >> USL=783.9 | Test Preblank Injections >> Test Preblank Injections >> USL=783.4 | Test Preblank Injections >> Test Preblank Injections >> USL=773.7 | Test Preblank Injections >> Test Preblank Injections >> USL=755.1 | Test Preblank Injections >> Test Preblank Injections >> USL=764.9 | Test Preblank Injections >> Test Preblank Injections >> USL=756.4 | Test Preblank Injections >> Test Preblank Injections >> USL=819.7 | Test Preblank Injections >> Test Preblank Injections >> USL=768.7 | Test Preblank Injections >> Test Preblank Injections >> USL=784.7 | Test Preblank Injections >> Test Preblank Injections >> USL=791.5 | Test Preblank Injections >> Test Preblank Injections >> USL=790.8 | Test Preblank Injections >> Test Preblank Injections >> USL=778.1 | Test Preblank Injections >> Test Preblank Injections >> USL=768.4 | Firmware Download >> Firmware Download | Test Preblank Injections >> Test Preblank Injections >> USL=788.7 | Test Preblank Injections >> Test Preblank Injections >> USL=755.9 | Test Preblank Injections >> Test Preblank Injections >> USL=767.5 | Test Preblank Injections >> Test Preblank Injections >> USL=795.8 | Test Preblank Injections >> Test Preblank Injections >> USL=779 | Test Preblank Injections >> Test Preblank Injections >> USL=772.9 | Test Preblank Injections >> Test Preblank Injections >> USL=773.4 | Test Preblank Injections >> Test Preblank Injections >> USL=779.4 | Test Preblank Injections >> Test Preblank Injections >> USL=790.9 | Test Preblank Injections >> Test Preblank Injections >> USL=798.5 | Test Preblank Injections >> Test Preblank Injections >> USL=764 | Test Preblank Injections >> Test Preblank Injections >> USL=734.7 | Test Preblank Injections >> Test Preblank Injections >> USL=785.3 | Test Preblank Injections >> Test Preblank Injections >> USL=753.4 | Test Preblank Injections >> Test Preblank Injections >> USL=772 | Test Preblank Injections >> Test Preblank Injections >> USL=786.1 | Test Preblank Injections >> Test Preblank Injections >> USL=758.5 | Test Preblank Injections >> Test Preblank Injections >> USL=793.2 | Test Preblank Injections >> Test Preblank Injections >> USL=777.4 | Test Preblank Injections >> Test Preblank Injections >> USL=787.3 | Test Preblank Injections >> Test Preblank Injections >> USL=789.9 | Test Preblank Injections >> Test Preblank Injections >> USL=763.4 | Test Preblank Injections >> Test Preblank Injections >> USL=806.8 | Test Preblank Injections >> Test Preblank Injections >> USL=774.7 | Test Preblank Injections >> Test Preblank Injections >> USL=801.8 | Test Preblank Injections >> Test Preblank Injections >> USL=788.4 | Test Preblank Injections >> Test Preblank Injections >> USL=783.8 | Test Preblank Injections >> Test Preblank Injections >> USL=771.9 | Test Preblank Injections >> Test Preblank Injections >> USL=780.3 | Test Preblank Injections >> Test Preblank Injections >> USL=788.8 | Test Preblank Injections >> Test Preblank Injections >> USL=780.5 | Test Preblank Injections >> Test Preblank Injections >> USL=480.2 | Test Preblank Injections >> Test Preblank Injections >> USL=783.7 | Test Preblank Injections >> Test Preblank Injections >> USL=776.5 | Test Preblank Injections >> Test Preblank Injections >> USL=784.8 | Test Preblank Injections >> Test Preblank Injections >> USL=771.4 | Test Preblank Injections >> Test Preblank Injections >> USL=772.5 | Test Preblank Injections >> Test Preblank Injections >> USL=793.1 | Test Preblank Injections >> Test Preblank Injections >> USL=752.4 | Test Preblank Injections >> Test Preblank Injections >> USL=780.2 | Test Preblank Injections >> Test Preblank Injections >> USL=751.7 | Test Preblank Injections >> Test Preblank Injections >> USL=743.4 | Test Preblank Injections >> Test Preblank Injections >> USL=778.3 | Test Preblank Injections >> Test Preblank Injections >> USL=789.5 | Test Preblank Injections >> Test Preblank Injections >> USL=771.5 | Test Preblank Injections >> Test Preblank Injections >> USL=756.6 | Test Preblank Injections >> Test Preblank Injections >> USL=786.2 | Test Preblank Injections >> Test Preblank Injections >> USL=745.3 | Test Preblank Injections >> Test Preblank Injections >> USL=762.3 | Test Preblank Injections >> Test Preblank Injections >> USL=791.6 | Test Preblank Injections >> Test Preblank Injections >> USL=804.4 | Test Preblank Injections >> Test Preblank Injections >> USL=735.4 | Test Preblank Injections >> Test Preblank Injections >> USL=738.1 | Test Preblank Injections >> Test Preblank Injections >> USL=797.7 | Test Preblank Injections >> Test Preblank Injections >> USL=1052.6 | Test Preblank Injections >> Test Preblank Injections >> USL=750 | Test Preblank Injections >> Test Preblank Injections >> USL=784.3 | Test Preblank Injections >> Test Preblank Injections >> USL=818.1 | Test Preblank Injections >> Test Preblank Injections >> USL=789 | Test Preblank Injections >> Test Preblank Injections >> USL=769.3 | Test Preblank Injections >> Test Preblank Injections >> USL=789.6 | Test Preblank Injections >> Test Preblank Injections >> USL=829.5 | Test Preblank Injections >> Test Preblank Injections >> USL=789.1 | Test Preblank Injections >> Test Preblank Injections >> USL=543.5 | Test Preblank Injections >> Test Preblank Injections >> USL=767 | Test Preblank Injections >> Test Preblank Injections >> USL=747.7 | Test Preblank Injections >> Test Preblank Injections >> USL=800.9 | Test Preblank Injections >> Test Preblank Injections >> USL=658.8 | Test Preblank Injections >> Test Preblank Injections >> USL=727.8 | Test Preblank Injections >> Test Preblank Injections >> USL=772.8 | Test Preblank Injections >> Test Preblank Injections >> USL=742.1 | Test Preblank Injections >> Test Preblank Injections >> USL=684.2 | Test Preblank Injections >> Test Preblank Injections >> USL=767.1 | Test Preblank Injections >> Test Preblank Injections >> USL=515.4 | Test Preblank Injections >> Test Preblank Injections >> USL=733.9 | Test Preblank Injections >> Test Preblank Injections >> USL=686.7 | Test Preblank Injections >> Test Preblank Injections >> USL=815.8 | Test Preblank Injections >> Test Preblank Injections >> USL=683.8 | Test Preblank Injections >> Test Preblank Injections >> USL=726.6 | Test Preblank Injections >> Test Preblank Injections >> USL=1013.4 | Test Preblank Injections >> Test Preblank Injections >> USL=697.1 | Test Preblank Injections >> Test Preblank Injections >> USL=595.4 | Test Preblank Injections >> Test Preblank Injections >> USL=671.3 | Test Preblank Injections >> Test Preblank Injections >> USL=605.9 | Test SPI Interface >> Test SPI Interface | Test Preblank Injections >> Test Preblank Injections >> USL=707.4 | Test Preblank Injections >> Test Preblank Injections >> USL=751.1 | Test Preblank Injections >> Test Preblank Injections >> USL=689.7 | Test Preblank Injections >> Test Preblank Injections >> USL=674 | Test Preblank Injections >> Test Preblank Injections >> USL=692.4 | Test Preblank Injections >> Test Preblank Injections >> USL=540.8 | Test Preblank Injections >> Test Preblank Injections >> USL=749 | Test Preblank Injections >> Test Preblank Injections >> USL=774.8 | Test Preblank Injections >> Test Preblank Injections >> USL=678.8 | Test Preblank Injections >> Test Preblank Injections >> USL=646.2 | Test Preblank Injections >> Test Preblank Injections >> USL=776.3 | Test Preblank Injections >> Test Preblank Injections >> USL=798.3 | Test Preblank Injections >> Test Preblank Injections >> USL=752.7 | Test Preblank Injections >> Test Preblank Injections >> USL=707.9 | Test Preblank Injections >> Test Preblank Injections >> USL=705.4 | Test Preblank Injections >> Test Preblank Injections >> USL=548.3 | Test Preblank Injections >> Test Preblank Injections >> USL=635.5 | Test Preblank Injections >> Test Preblank Injections >> USL=622.3 | Test Preblank Injections >> Test Preblank Injections >> USL=791.2 | Test Preblank Injections >> Test Preblank Injections >> USL=623.3 | Test Preblank Injections >> Test Preblank Injections >> USL=764.2 | Test Preblank Injections >> Test Preblank Injections >> USL=509.7 | Test Preblank Injections >> Test Preblank Injections >> USL=738.7 | Test Preblank Injections >> Test Preblank Injections >> USL=680.5 | Test Preblank Injections >> Test Preblank Injections >> USL=696.6 | Test Preblank Injections >> Test Preblank Injections >> USL=736.6 | Test Preblank Injections >> Test Preblank Injections >> USL=671.8 | Test Preblank Injections >> Test Preblank Injections >> USL=675.1 | Test Preblank Injections >> Test Preblank Injections >> USL=674.9 | Test Preblank Injections >> Test Preblank Injections >> USL=587.4 | Test Preblank Injections >> Test Preblank Injections >> USL=587.1 | Test Preblank Injections >> Test Preblank Injections >> USL=654.4 | Test Preblank Injections >> Test Preblank Injections >> USL=692.5 | Test Preblank Injections >> Test Preblank Injections >> USL=703.2 | Test Preblank Injections >> Test Preblank Injections >> USL=644.8 | Test Preblank Injections >> Test Preblank Injections >> USL=794.8 | Test Preblank Injections >> Test Preblank Injections >> USL=768.5 | Test Preblank Injections >> Test Preblank Injections >> USL=763.2 | Test Preblank Injections >> Test Preblank Injections >> USL=759 | Test Preblank Injections >> Test Preblank Injections >> USL=728.2 | Test Preblank Injections >> Test Preblank Injections >> USL=808.1 | Test Preblank Injections >> Test Preblank Injections >> USL=703.5 | Test Preblank Injections >> Test Preblank Injections >> USL=784.9 | Test Preblank Injections >> Test Preblank Injections >> USL=804.6 | Test Preblank Injections >> Test Preblank Injections >> USL=798.4 | Test Preblank Injections >> Test Preblank Injections >> USL=785 | Test Preblank Injections >> Test Preblank Injections >> USL=797.6 | Test Preblank Injections >> Test Preblank Injections >> USL=576.6 | Test Preblank Injections >> Test Preblank Injections >> USL=630.4 | Test Preblank Injections >> Test Preblank Injections >> USL=790.5 | Test Preblank Injections >> Test Preblank Injections >> USL=765.2 | Test Preblank Injections >> Test Preblank Injections >> USL=787 | Test Preblank Injections >> Test Preblank Injections >> USL=807.8 | Test Preblank Injections >> Test Preblank Injections >> USL=778 | Test Preblank Injections >> Test Preblank Injections >> USL=759.9 | Test Preblank Injections >> Test Preblank Injections >> USL=563.6 | Test Preblank Injections >> Test Preblank Injections >> USL=680.7 | Test Preblank Injections >> Test Preblank Injections >> USL=610 | Test Preblank Injections >> Test Preblank Injections >> USL=774.9 | Test Preblank Injections >> Test Preblank Injections >> USL=413.9 | Test Preblank Injections >> Test Preblank Injections >> USL=568.9 | Test Preblank Injections >> Test Preblank Injections >> USL=643.3 | Test Preblank Injections >> Test Preblank Injections >> USL=729.1 | Test Preblank Injections >> Test Preblank Injections >> USL=813.2 | Test Preblank Injections >> Test Preblank Injections >> USL=685.3 | Test Preblank Injections >> Test Preblank Injections >> USL=667.4 | Test Preblank Injections >> Test Preblank Injections >> USL=708.9 | Test Preblank Injections >> Test Preblank Injections >> USL=660.9 | Test Preblank Injections >> Test Preblank Injections >> USL=613.1 | Test Preblank Injections >> Test Preblank Injections >> USL=577.9 | Test Preblank Injections >> Test Preblank Injections >> USL=487.9 | Test Preblank Injections >> Test Preblank Injections >> USL=653.6 | Test Preblank Injections >> Test Preblank Injections >> USL=687 | Test Preblank Injections >> Test Preblank Injections >> USL=682.3 | Test Preblank Injections >> Test Preblank Injections >> USL=658.7 | Test Preblank Injections >> Test Preblank Injections >> USL=696.3 | Test Preblank Injections >> Test Preblank Injections >> USL=614.9 | Test Preblank Injections >> Test Preblank Injections >> USL=709.9 | Test Preblank Injections >> Test Preblank Injections >> USL=694.8 | Test Preblank Injections >> Test Preblank Injections >> USL=474.3 | Test Preblank Injections >> Test Preblank Injections >> USL=464 | Test Preblank Injections >> Test Preblank Injections >> USL=687.8 | Test Preblank Injections >> Test Preblank Injections >> USL=737.7 | Test Preblank Injections >> Test Preblank Injections >> USL=496 | Test Preblank Injections >> Test Preblank Injections >> USL=672.1 | Test Preblank Injections >> Test Preblank Injections >> USL=582 | Test Preblank Injections >> Test Preblank Injections >> USL=662.1 | Test Preblank Injections >> Test Preblank Injections >> USL=747.2 | Test Preblank Injections >> Test Preblank Injections >> USL=803.8 | Test Preblank Injections >> Test Preblank Injections >> USL=725.4 | Test Preblank Injections >> Test Preblank Injections >> USL=756.8 | Test Preblank Injections >> Test Preblank Injections >> USL=734.3 | Test Preblank Injections >> Test Preblank Injections >> USL=627.5 | Test Preblank Injections >> Test Preblank Injections >> USL=718.5 | Test Preblank Injections >> Test Preblank Injections >> USL=744.2 | Test Preblank Injections >> Test Preblank Injections >> USL=538.6 | Test Preblank Injections >> Test Preblank Injections >> USL=539.3 | Test Preblank Injections >> Test Preblank Injections >> USL=578.6 | Test Preblank Injections >> Test Preblank Injections >> USL=655.2 | Test Preblank Injections >> Test Preblank Injections >> USL=739.5 | Test Preblank Injections >> Test Preblank Injections >> USL=649.4 | Test Preblank Injections >> Test Preblank Injections >> USL=572.9 | Test Preblank Injections >> Test Preblank Injections >> USL=769.7 | Test Preblank Injections >> Test Preblank Injections >> USL=695.4 | Test Preblank Injections >> Test Preblank Injections >> USL=626.1 | Test Preblank Injections >> Test Preblank Injections >> USL=748.3 | Test Preblank Injections >> Test Preblank Injections >> USL=508.5 | Test Preblank Injections >> Test Preblank Injections >> USL=805.4 | Test Preblank Injections >> Test Preblank Injections >> USL=757.4 | Test Preblank Injections >> Test Preblank Injections >> USL=788.3 | Test Preblank Injections >> Test Preblank Injections >> USL=631.2 | Test Preblank Injections >> Test Preblank Injections >> USL=562.8 | Test Preblank Injections >> Test Preblank Injections >> USL=660.1 | Test Preblank Injections >> Test Preblank Injections >> USL=512.5 | Test Preblank Injections >> Test Preblank Injections >> USL=667.8 | Test Preblank Injections >> Test Preblank Injections >> USL=2426.8 | Test Preblank Injections >> Test Preblank Injections >> USL=417.3 | Test Preblank Injections >> Test Preblank Injections >> USL=634.6 | Test Preblank Injections >> Test Preblank Injections >> USL=704.5 | Test Preblank Injections >> Test Preblank Injections >> USL=498.2 | Test Preblank Injections >> Test Preblank Injections >> USL=563 | Test Preblank Injections >> Test Preblank Injections >> USL=716.6 | Test Preblank Injections >> Test Preblank Injections >> USL=715.1 | Test Preblank Injections >> Test Preblank Injections >> USL=747.9 | Test Preblank Injections >> Test Preblank Injections >> USL=557.8 | Test Preblank Injections >> Test Preblank Injections >> USL=657.9 | Test Preblank Injections >> Test Preblank Injections >> USL=660 | Test Preblank Injections >> Test Preblank Injections >> USL=711.4 | Test Preblank Injections >> Test Preblank Injections >> USL=621.1 | Test Preblank Injections >> Test Preblank Injections >> USL=713.2 | Test Preblank Injections >> Test Preblank Injections >> USL=721.8 | Test Preblank Injections >> Test Preblank Injections >> USL=780.8 | Test Preblank Injections >> Test Preblank Injections >> USL=679.1 | Test Preblank Injections >> Test Preblank Injections >> USL=625.5 | Test Preblank Injections >> Test Preblank Injections >> USL=561.6 | Test Preblank Injections >> Test Preblank Injections >> USL=539.5 | Test Preblank Injections >> Test Preblank Injections >> USL=545.6 | Test Preblank Injections >> Test Preblank Injections >> USL=535.2 | Test Preblank Injections >> Test Preblank Injections >> USL=700.1 | Test Preblank Injections >> Test Preblank Injections >> USL=624.7 | Test Preblank Injections >> Test Preblank Injections >> USL=582.3 | Test Preblank Injections >> Test Preblank Injections >> USL=591.3 | Test Preblank Injections >> Test Preblank Injections >> USL=773.1 | Test Preblank Injections >> Test Preblank Injections >> USL=376 | Test Preblank Injections >> Test Preblank Injections >> USL=626 | Test Preblank Injections >> Test Preblank Injections >> USL=694.3 | Test Preblank Injections >> Test Preblank Injections >> USL=677.3 | Test Preblank Injections >> Test Preblank Injections >> USL=385.3 | Test Preblank Injections >> Test Preblank Injections >> USL=724.2 | Test Preblank Injections >> Test Preblank Injections >> USL=297.9 | Test Preblank Injections >> Test Preblank Injections >> USL=741.6 | Test Preblank Injections >> Test Preblank Injections >> USL=754.6 | Test Preblank Injections >> Test Preblank Injections >> USL=510.1 | Test Preblank Injections >> Test Preblank Injections >> USL=551.6 | Test Preblank Injections >> Test Preblank Injections >> USL=732.3 | Test Preblank Injections >> Test Preblank Injections >> USL=501.1 | Test Preblank Injections >> Test Preblank Injections >> USL=681.2 | Test Preblank Injections >> Test Preblank Injections >> USL=665.8 | Test Preblank Injections >> Test Preblank Injections >> USL=693.1 | Test Preblank Injections >> Test Preblank Injections >> USL=739.1 | Test Preblank Injections >> Test Preblank Injections >> USL=695.9 | Test Preblank Injections >> Test Preblank Injections >> USL=639.4 | Test Preblank Injections >> Test Preblank Injections >> USL=759.3 | Test Preblank Injections >> Test Preblank Injections >> USL=338.7 | Test Preblank Injections >> Test Preblank Injections >> USL=674.8 | Test Preblank Injections >> Test Preblank Injections >> USL=625.9 | Test Preblank Injections >> Test Preblank Injections >> USL=477 | Test Preblank Injections >> Test Preblank Injections >> USL=357 | Test Preblank Injections >> Test Preblank Injections >> USL=880.6 | Test Preblank Injections >> Test Preblank Injections >> USL=1261.1 | Test Preblank Injections >> Test Preblank Injections >> USL=634.2 | Test Preblank Injections >> Test Preblank Injections >> USL=804.8 | Test Preblank Injections >> Test Preblank Injections >> USL=495.3 | Test Preblank Injections >> Test Preblank Injections >> USL=661.3 | Test Preblank Injections >> Test Preblank Injections >> USL=1157.9 | Test Preblank Injections >> Test Preblank Injections >> USL=1272.1 | Test Preblank Injections >> Test Preblank Injections >> USL=571.6 | Test Preblank Injections >> Test Preblank Injections >> USL=611.7 | Test Preblank Injections >> Test Preblank Injections >> USL=595 | Test Preblank Injections >> Test Preblank Injections >> USL=730.1 | Test Preblank Injections >> Test Preblank Injections >> USL=510.8 | Test Preblank Injections >> Test Preblank Injections >> USL=728.6 | Test Preblank Injections >> Test Preblank Injections >> USL=816 | Test Preblank Injections >> Test Preblank Injections >> USL=655 | Test Preblank Injections >> Test Preblank Injections >> USL=608.4 | Test Preblank Injections >> Test Preblank Injections >> USL=679.8 | Test Preblank Injections >> Test Preblank Injections >> USL=777.2 | Test Preblank Injections >> Test Preblank Injections >> USL=605.8 | Test Preblank Injections >> Test Preblank Injections >> USL=654.7 | Test Preblank Injections >> Test Preblank Injections >> USL=740.7 | Test Preblank Injections >> Test Preblank Injections >> USL=696.5 | Test Preblank Injections >> Test Preblank Injections >> USL=558.1 | Test Preblank Injections >> Test Preblank Injections >> USL=588.3 | Test Preblank Injections >> Test Preblank Injections >> USL=590 | Test Preblank Injections >> Test Preblank Injections >> USL=644.7 | Test Preblank Injections >> Test Preblank Injections >> USL=682.1 | Test Preblank Injections >> Test Preblank Injections >> USL=899.8 | Test Preblank Injections >> Test Preblank Injections >> USL=736.4 | Test Preblank Injections >> Test Preblank Injections >> USL=591.6 | Test Preblank Injections >> Test Preblank Injections >> USL=660.6 | Test Preblank Injections >> Test Preblank Injections >> USL=449.4 | Test Preblank Injections >> Test Preblank Injections >> USL=609.6 | Test Preblank Injections >> Test Preblank Injections >> USL=522.7 | Test Preblank Injections >> Test Preblank Injections >> USL=737.2 | Test Preblank Injections >> Test Preblank Injections >> USL=1559.3 | Test Preblank Injections >> Test Preblank Injections >> USL=669.6 | Test Preblank Injections >> Test Preblank Injections >> USL=759.5 | Test Preblank Injections >> Test Preblank Injections >> USL=658.1 | Test Preblank Injections >> Test Preblank Injections >> USL=679.9 | Test Preblank Injections >> Test Preblank Injections >> USL=675 | Test Preblank Injections >> Test Preblank Injections >> USL=696.2 | Test Preblank Injections >> Test Preblank Injections >> USL=780.7 | Test Preblank Injections >> Test Preblank Injections >> USL=488.1 | Test Preblank Injections >> Test Preblank Injections >> USL=765 | Test Preblank Injections >> Test Preblank Injections >> USL=618.1 | Test Preblank Injections >> Test Preblank Injections >> USL=874.4 | Test Preblank Injections >> Test Preblank Injections >> USL=593 | Test Preblank Injections >> Test Preblank Injections >> USL=656.9 | Test Preblank Injections >> Test Preblank Injections >> USL=599.1 | Test Preblank Injections >> Test Preblank Injections >> USL=685.9 | Test Preblank Injections >> Test Preblank Injections >> USL=688.5 | Test Preblank Injections >> Test Preblank Injections >> USL=592.8 | Test Preblank Injections >> Test Preblank Injections >> USL=576.4 | Test Preblank Injections >> Test Preblank Injections >> USL=725 | Test Preblank Injections >> Test Preblank Injections >> USL=613.3 | Test Preblank Injections >> Test Preblank Injections >> USL=742.4 | Test Preblank Injections >> Test Preblank Injections >> USL=688.9 | Test Preblank Injections >> Test Preblank Injections >> USL=724.8 | Test Preblank Injections >> Test Preblank Injections >> USL=664.6 | Test Preblank Injections >> Test Preblank Injections >> USL=596.5 | Test Preblank Injections >> Test Preblank Injections >> USL=532.7 | Test Preblank Injections >> Test Preblank Injections >> USL=693.4 | Test Preblank Injections >> Test Preblank Injections >> USL=750.8 | Test Preblank Injections >> Test Preblank Injections >> USL=813.1 | Test Preblank Injections >> Test Preblank Injections >> USL=789.7 | Test Preblank Injections >> Test Preblank Injections >> USL=796.3 | Test Preblank Injections >> Test Preblank Injections >> USL=795.9 | Test Preblank Injections >> Test Preblank Injections >> USL=797.1 | Test Preblank Injections >> Test Preblank Injections >> USL=780.9 | Test Preblank Injections >> Test Preblank Injections >> USL=768.9 | Test Preblank Injections >> Test Preblank Injections >> USL=763.8 | Test Preblank Injections >> Test Preblank Injections >> USL=725.3 | Test Preblank Injections >> Test Preblank Injections >> USL=737.1 | Test Preblank Injections >> Test Preblank Injections >> USL=789.3 | Test Preblank Injections >> Test Preblank Injections >> USL=1739.1 | Test Preblank Injections >> Test Preblank Injections >> USL=797.8 | Test Preblank Injections >> Test Preblank Injections >> USL=745.9 | Test Preblank Injections >> Test Preblank Injections >> USL=781.2 | Test Preblank Injections >> Test Preblank Injections >> USL=800.1 | Test Preblank Injections >> Test Preblank Injections >> USL=755.2 | Test Preblank Injections >> Test Preblank Injections >> USL=798.1 | Test Preblank Injections >> Test Preblank Injections >> USL=790.2 | Test Preblank Injections >> Test Preblank Injections >> USL=790.3 | Test Preblank Injections >> Test Preblank Injections >> USL=822.7 | Test Preblank Injections >> Test Preblank Injections >> USL=751.5 | Test Preblank Injections >> Test Preblank Injections >> USL=779.7 | Test Preblank Injections >> Test Preblank Injections >> USL=759.7 | Test Preblank Injections >> Test Preblank Injections >> USL=746.1 | Test Preblank Injections >> Test Preblank Injections >> USL=765.8 | Test Preblank Injections >> Test Preblank Injections >> USL=796.2 | Test Preblank Injections >> Test Preblank Injections >> USL=775.8 | Test Preblank Injections >> Test Preblank Injections >> USL=811.9 | Test Preblank Injections >> Test Preblank Injections >> USL=792.3 | Test Preblank Injections >> Test Preblank Injections >> USL=736.9 | Test Preblank Injections >> Test Preblank Injections >> USL=761.4 | Test Preblank Injections >> Test Preblank Injections >> USL=779.9 | Test Preblank Injections >> Test Preblank Injections >> USL=744 | Test Preblank Injections >> Test Preblank Injections >> USL=797.5 | Test Preblank Injections >> Test Preblank Injections >> USL=794.4 | Test Preblank Injections >> Test Preblank Injections >> USL=777.3 | Test Preblank Injections >> Test Preblank Injections >> USL=788 | Test Preblank Injections >> Test Preblank Injections >> USL=749.1 | Test Preblank Injections >> Test Preblank Injections >> USL=781.4 | Test Preblank Injections >> Test Preblank Injections >> USL=775.6 | Test Preblank Injections >> Test Preblank Injections >> USL=700.5 | Test Preblank Injections >> Test Preblank Injections >> USL=769.1 | Test Preblank Injections >> Test Preblank Injections >> USL=739 | Test Preblank Injections >> Test Preblank Injections >> USL=780.6 | Test Preblank Injections >> Test Preblank Injections >> USL=782.5 | Test Preblank Injections >> Test Preblank Injections >> USL=2597.3 | Test Preblank Injections >> Test Preblank Injections >> USL=741.5 | Test Preblank Injections >> Test Preblank Injections >> USL=738.2 | Test Preblank Injections >> Test Preblank Injections >> USL=765.3 | Test Preblank Injections >> Test Preblank Injections >> USL=784 | Test Preblank Injections >> Test Preblank Injections >> USL=761.7 | Test Preblank Injections >> Test Preblank Injections >> USL=741.3 | Test Preblank Injections >> Test Preblank Injections >> USL=723.3 | Test Preblank Injections >> Test Preblank Injections >> USL=742.8 | Test Preblank Injections >> Test Preblank Injections >> USL=796.4 | Test Preblank Injections >> Test Preblank Injections >> USL=749.5 | Test Preblank Injections >> Test Preblank Injections >> USL=773.8 | Test Preblank Injections >> Test Preblank Injections >> USL=722.4 | Test Preblank Injections >> Test Preblank Injections >> USL=770.2 | Test Preblank Injections >> Test Preblank Injections >> USL=731.9 | Test Preblank Injections >> Test Preblank Injections >> USL=768.1 | Test Preblank Injections >> Test Preblank Injections >> USL=798.9 | Test Preblank Injections >> Test Preblank Injections >> USL=794.3 | Test Preblank Injections >> Test Preblank Injections >> USL=802.3 | Test Preblank Injections >> Test Preblank Injections >> USL=899.2 | Test Preblank Injections >> Test Preblank Injections >> USL=801.1 | Test Preblank Injections >> Test Preblank Injections >> USL=756.2 | Test Preblank Injections >> Test Preblank Injections >> USL=809.9 | Test Preblank Injections >> Test Preblank Injections >> USL=744.3 | Test Preblank Injections >> Test Preblank Injections >> USL=798.2 | Test Preblank Injections >> Test Preblank Injections >> USL=810.7 | Test Preblank Injections >> Test Preblank Injections >> USL=782.9 | Test Preblank Injections >> Test Preblank Injections >> USL=801.2 | Test Preblank Injections >> Test Preblank Injections >> USL=796.6 | Test Preblank Injections >> Test Preblank Injections >> USL=812 | Test Preblank Injections >> Test Preblank Injections >> USL=769.9 | Test Preblank Injections >> Test Preblank Injections >> USL=775.1 | Test Preblank Injections >> Test Preblank Injections >> USL=749.4 | Test Preblank Injections >> Test Preblank Injections >> USL=778.9 | Test Preblank Injections >> Test Preblank Injections >> USL=806.6 | Test Preblank Injections >> Test Preblank Injections >> USL=729.5 | Test Preblank Injections >> Test Preblank Injections >> USL=792 | Test Preblank Injections >> Test Preblank Injections >> USL=753.7 | Test Preblank Injections >> Test Preblank Injections >> USL=801.9 | Test Preblank Injections >> Test Preblank Injections >> USL=815.9 | Test Preblank Injections >> Test Preblank Injections >> USL=786 | Test Preblank Injections >> Test Preblank Injections >> USL=810.1 | Test Preblank Injections >> Test Preblank Injections >> USL=776.2 | Test Preblank Injections >> Test Preblank Injections >> USL=830.5 | Test Preblank Injections >> Test Preblank Injections >> USL=778.4 | Test Preblank Injections >> Test Preblank Injections >> USL=7239.5 | Test Preblank Injections >> Test Preblank Injections >> USL=797.3 | Test Preblank Injections >> Test Preblank Injections >> USL=814.6 | Test Preblank Injections >> Test Preblank Injections >> USL=814.7 | Test Preblank Injections >> Test Preblank Injections >> USL=817.9 | Test Preblank Injections >> Test Preblank Injections >> USL=821.8 | Test Preblank Injections >> Test Preblank Injections >> USL=791.9 | Test Preblank Injections >> Test Preblank Injections >> USL=740.2 | Test Preblank Injections >> Test Preblank Injections >> USL=812.7 | Test Preblank Injections >> Test Preblank Injections >> USL=7366.9 | Test Preblank Injections >> Test Preblank Injections >> USL=812.3 | Test Preblank Injections >> Test Preblank Injections >> USL=589.2 | Test Preblank Injections >> Test Preblank Injections >> USL=770.6 | Test Preblank Injections >> Test Preblank Injections >> USL=7653.4 | Test Preblank Injections >> Test Preblank Injections >> USL=782.2 | Test Preblank Injections >> Test Preblank Injections >> USL=759.1 | Test Preblank Injections >> Test Preblank Injections >> USL=7610.2 | Test Preblank Injections >> Test Preblank Injections >> USL=1571 | Test Preblank Injections >> Test Preblank Injections >> USL=801.6 | Test Preblank Injections >> Test Preblank Injections >> USL=7707.8 | Test Preblank Injections >> Test Preblank Injections >> USL=1581.4 | Test Preblank Injections >> Test Preblank Injections >> USL=766.5 | Test Preblank Injections >> Test Preblank Injections >> USL=1680.8 | Test Preblank Injections >> Test Preblank Injections >> USL=7781 | Test Preblank Injections >> Test Preblank Injections >> USL=794.7 | Test Preblank Injections >> Test Preblank Injections >> USL=7716.8 | Test Preblank Injections >> Test Preblank Injections >> USL=1659.7 | Test Preblank Injections >> Test Preblank Injections >> USL=796.5 | Test Preblank Injections >> Test Preblank Injections >> USL=818.9 | Test Preblank Injections >> Test Preblank Injections >> USL=676 | Test Preblank Injections >> Test Preblank Injections >> USL=7818.5 | Test Preblank Injections >> Test Preblank Injections >> USL=587.7 | Test Preblank Injections >> Test Preblank Injections >> USL=1774 | Test Preblank Injections >> Test Preblank Injections >> USL=795.1 | Test Preblank Injections >> Test Preblank Injections >> USL=652.5 | Test Preblank Injections >> Test Preblank Injections >> USL=1722.1 | Test Preblank Injections >> Test Preblank Injections >> USL=7778.7 | Test Preblank Injections >> Test Preblank Injections >> USL=565.5 | Test Preblank Injections >> Test Preblank Injections >> USL=700.8 | Test Preblank Injections >> Test Preblank Injections >> USL=687.4 | Test Preblank Injections >> Test Preblank Injections >> USL=1729.8 | Test Preblank Injections >> Test Preblank Injections >> USL=7719 | Test Preblank Injections >> Test Preblank Injections >> USL=7721.6 | Test Preblank Injections >> Test Preblank Injections >> USL=828.4 | Test Preblank Injections >> Test Preblank Injections >> USL=825.1 | Test Preblank Injections >> Test Preblank Injections >> USL=824.3 | Test Preblank Injections >> Test Preblank Injections >> USL=773.2 | Test Preblank Injections >> Test Preblank Injections >> USL=773.9 | Test Preblank Injections >> Test Preblank Injections >> USL=1761.4 | Test Preblank Injections >> Test Preblank Injections >> USL=7944.2 | Test Preblank Injections >> Test Preblank Injections >> USL=816.8 | Test Preblank Injections >> Test Preblank Injections >> USL=1787 | Test Preblank Injections >> Test Preblank Injections >> USL=1111.2 | Test Preblank Injections >> Test Preblank Injections >> USL=831.7 | Test Preblank Injections >> Test Preblank Injections >> USL=829.8 | Test Preblank Injections >> Test Preblank Injections >> USL=1736.5 | Test Preblank Injections >> Test Preblank Injections >> USL=808.8 | Test Preblank Injections >> Test Preblank Injections >> USL=7915.6 | Test Preblank Injections >> Test Preblank Injections >> USL=790.7 | Test Preblank Injections >> Test Preblank Injections >> USL=1780.7 | Test Preblank Injections >> Test Preblank Injections >> USL=694 | Test Preblank Injections >> Test Preblank Injections >> USL=818.4 | Test Preblank Injections >> Test Preblank Injections >> USL=724.9 | Test Preblank Injections >> Test Preblank Injections >> USL=8068.9 | Test Preblank Injections >> Test Preblank Injections >> USL=787.9 | Test Preblank Injections >> Test Preblank Injections >> USL=795 | Test Preblank Injections >> Test Preblank Injections >> USL=816.9 | Test Preblank Injections >> Test Preblank Injections >> USL=1766.3 | Test Preblank Injections >> Test Preblank Injections >> USL=816.4 | Test Preblank Injections >> Test Preblank Injections >> USL=7801.5 | Test Preblank Injections >> Test Preblank Injections >> USL=825.5 | Test Preblank Injections >> Test Preblank Injections >> USL=7503.8 | Test Preblank Injections >> Test Preblank Injections >> USL=1797.7 | Test Preblank Injections >> Test Preblank Injections >> USL=800.8 | Test Preblank Injections >> Test Preblank Injections >> USL=814.2 | Test Preblank Injections >> Test Preblank Injections >> USL=760.8 | Test Preblank Injections >> Test Preblank Injections >> USL=7906.1 | Test Preblank Injections >> Test Preblank Injections >> USL=1783.7 | Test Preblank Injections >> Test Preblank Injections >> USL=820.9 | Test Preblank Injections >> Test Preblank Injections >> USL=7844.6 | Test Preblank Injections >> Test Preblank Injections >> USL=1782.4 | Test Preblank Injections >> Test Preblank Injections >> USL=746.3 | Test Preblank Injections >> Test Preblank Injections >> USL=795.3 | Test Preblank Injections >> Test Preblank Injections >> USL=786.7 | Test Preblank Injections >> Test Preblank Injections >> USL=682.6 | Test Preblank Injections >> Test Preblank Injections >> USL=801 | Test Preblank Injections >> Test Preblank Injections >> USL=755.6 | Test Preblank Injections >> Test Preblank Injections >> USL=747.5 | Test Preblank Injections >> Test Preblank Injections >> USL=794.6 | Test Preblank Injections >> Test Preblank Injections >> USL=794.5 | Test Preblank Injections >> Test Preblank Injections >> USL=780.1 | Test Preblank Injections >> Test Preblank Injections >> USL=817.1 | Test Preblank Injections >> Test Preblank Injections >> USL=759.6 | Test Preblank Injections >> Test Preblank Injections >> USL=797.4 | Test Preblank Injections >> Test Preblank Injections >> USL=766.4 | Test Preblank Injections >> Test Preblank Injections >> USL=2583.6 | Test Preblank Injections >> Test Preblank Injections >> USL=796.1 | Test Preblank Injections >> Test Preblank Injections >> USL=934.7 | Test Preblank Injections >> Test Preblank Injections >> USL=746 | Test Preblank Injections >> Test Preblank Injections >> USL=765.4 | Test Preblank Injections >> Test Preblank Injections >> USL=593.5
LSL | [9 empty] | 129 | -180 | -3 | -3 | 48 | 48 | -3 | -0.3 | -20 | -20 | -20 |  | 0.45 | -5 |  |  |  | 0 | 0 | -0.2 | 0.999 |  |  |  |  | 133
USL | [10 empty] | 180 | 3 | 3 | 68 | 68 | 3 | 0.3 | 20 | 20 | 20 | -0.6 | 0.55 | 15 | 500 | 792.9 | 0.128 | 0.5 | 0.9 | 0.2 | 1 | 776.9 | 808.3 | 736.7 | 761.9 |  | 713.6 | 0.04 | 782.8 | 781.1 | 701.9 | 800.6 | 726.5 | 731.6 | 766.8 | 650.8 | 711.7 | 775.5 | 757.7 | 701 | 723.5 | 754.1 | 0.06 | 721.5 | 661 | 739.3 | 781.6 | 705.2 |  | 721.9 | 712.9 | 741.4 | 749.9 | 479.1 | 778.5 | 0 | 702.7 | 769 | 705.7 | 778.8 | 691.6 | 830.6 | 723.2 | 648.4 | 771.8 | 722.1 | 752.3 | 770.4 | 715.3 | 678.6 | 572.3 | 728.1 | 732.4 | 743 | 775.3 | 610.6 | 15961 | 722.9 | 744.6 | 785.7 | 735.3 | 786.8 | 741 |  | 635 | 800.3 | 775 |  | 724 | 745.8 | 734.5 | 760.3 | 663.1 | 792.7 | 736.5 | 775.4 | 961.5 | 717.5 | 738.6 | 727.2 | 731.3 | 667.6 | 710.1 | 658.6 | 764.6 | 944.4 | 773.5 | 965.6 | 752.8 | 715.5 | 751.4 | 962.8 | 730.3 | 791.8 |  | 681.4 | 727.7 | 758.6 | 731.2 | 909.5 | 252.5 | 967.1 | 601.3 | 635.1 | 621.4 | 763.1 | 757.3 | 740.6 | 790.6 | 625 | 943.3 | 768.6 | 707.5 | 934.5 | 733.6 | 965.9 | 712.6 | 802.1 | 730.9 | 784.2 | 763.3 | 718.9 | 990.7 | 771.3 | 772.1 | 741.9 | 735.9 | 976.4 | 740.3 | 777.1 | 738.9 | 791.7 | 761.3 | 725.9 | 774 | 788.5 | 595.5 | 971.1 | 755.4 | 710.3 | 781.5 | 768.8 | 784.4 | 796.7 | 731 | 779.6 | 733.1 | 795.2 | 751.2 | 758.7 | 760.7 | 914.2 | 755 | 771.6 | 653.4 | 746.5 | 766.9 | 753.5 | 764.1 | 739.9 | 745.6 | 793 | 794.2 | 866.2 | 714.5 | 749.3 | 739.4 | 841.4 | 1059.9 | 747.3 | 753.1 | 1398.9 | 757.2 | 749.2 | 759.4 |  | 761.8 | 2116.3 | 711.1 | 745.5 | 713.9 |  | 758.8 | 1475.6 |  | 833 | 1543.8 | 828.3 | 756.3 | 791.3 | 795.7 | 777.8 | 744.5 | 919.7 | 655.4 | 732.1 | 711.5 | 771.2 | 717.7 | 784.1 | 695.1 | 783.2 | 781.8 | 662 | 771.7 | 774.2 | 767.3 | 770 | 766.7 | 713.3 | 749.8 | 767.4 | 795.5 | 770.5 | 793.6 | 674.1 | 747.1 | 690.3 | 805.3 | 776 | 769.4 | 763.5 | 796.9 | 807.9 | 792.8 | 639 | 768.3 | 731.1 | 782.6 | 782.4 | 753.2 | 776.6 | 779.1 | 646.4 | 754.5 | 785.9 | 463.9 | 777.9 | 715.8 | 737.5 | 497.5 | 816.5 | 779.3 | 677.9 | 662.8 | 716.9 | 642.1 | 700.9 | 722.5 | 697.5 | 750.9 | 741.7 | 735.7 | 758.2 | 607.3 | 627.6 | 790.1 | 744.8 | 538 | 770.7 | 458.4 | 751.8 | 779.2 | 663.2 | 706.6 | 717.8 | 635.7 | 718.4 | 786.9 | 536.4 | 703.8 | 738.5 | 761.6 | 752.1 | 622.8 | 836 | 568.5 | 699.7 | 760.9 | 663.7 | 739.2 | 767.9 | 757.6 | 762.1 | 652.4 | 686.5 | 801.5 | 776.4 | 799.3 | 721.2 | 785.1 | 700.3 | 802.7 | 727 | 602.4 | 538.9 | 752 | 765.9 | 668.6 | 652.8 | 768.2 | 855.9 | 727.5 | 782 | 460.3 | 617.4 | 581.7 | 794 | 761.5 | 766.3 | 713.5 | 723.6 | 728.7 | 740.5 | 790.4 | 780.4 | 699.1 | 851.2 | 815 | 770.8 | 729.2 | 794.1 | 717 | 612.3 | 687.3 | 735.2 | 849 | 719.3 |  | 862 | 789.4 | 639.1 | 629.6 | 759.2 | 656.7 | 785.4 | 626.4 | 696 | 865.2 | 799.4 | 699.9 | 720.4 | 729.4 | 737.3 | 522.8 | 750.3 | 732.5 | 729.6 | 689 | 762.2 | 667.9 | 772.6 | 771 | 777.5 | 774.6 | 782.1 | 681.9 | 708 | 789.8 | 802.6 | 730 | 775.2 | 746.4 | 810.6 | 743.9 | 791.1 | 798 | 813.3 | 783.6 | 738.3 | 753.6 | 829.2 | 743.1 | 781.9 | 783.3 | 733.2 | 822.8 | 695 | 741.2 | 656.2 |  | 786.5 | 787.4 | 742.6 | 595.3 | 784.5 | 728.9 | 607.5 | 774.3 | 748 | 787.1 | 740.4 | 734.1 | 797.2 | 745.4 | 772.2 | 631.9 | 759.8 | 798.6 | 603.6 | 439.7 | 652.7 | 678.2 | 772.4 | 740.8 | 762.9 | 534.5 | 698.3 | 650.6 | 555 | 754.3 | 789.2 | 582.9 | 712.3 | 791 | 767.6 | 726.9 | 756 | 758 | 686.8 | 406.8 | 660.4 | 719 | 711.3 | 726.1 | 704.1 | 29920.4 | 658.5 | 619.2 | 665.3 | 698.5 | 671.1 | 639.6 | 748.5 | 685 | 714.8 | 564.6 | 624.6 |  | 566.6 | 718.2 | 714.4 | 586.4 | 669.9 | 762.6 | 645.1 | 711 | 769.5 | 730.8 | 647 | 683.5 | 603.9 | 744.9 | 700.2 | 706.3 | 589.8 | 736.2 | 766.6 | 524 | 698 | 615.1 | 707.3 | 545.3 | 562.7 | 668.7 | 674.3 | 659.8 | 609.5 | 742 | 614.8 | 517.1 | 750.4 | 676.5 | 694.5 | 719.2 | 572.6 | 724.4 | 776.1 | 770.9 | 569 | 692.8 | 705.5 | 617.2 | 513.3 | 482 | 681.8 | 665 | 476.2 | 710.5 | 661.1 | 711.6 | 812.2 | 679.2 | 589.9 | 521.9 | 673.2 | 683.7 | 645.6 | 715.7 | 679.6 | 462.7 | 699.6 | 668.9 | 707.7 | 636.7 | 376.9 | 269.4 | 619.3 | 739.8 | 778.6 | 725.1 | 735.5 | 773 | 693.5 | 748.2 | 757.8 | 795.6 | 716 | 729 | 688.3 | 760 | 588 | 792.4 | 757.1 | 681.6 | 664.2 | 933.5 | 668.1 | 689.2 | 761.1 | 737.9 | 667.2 | 719.9 | 786.3 | 710.6 | 762 | 644.2 | 612.4 |  | 684.9 | 765.7 | 628.1 | 651.8 | 781.3 | 715 | 739.7 | 561.9 | 617.6 | 491.5 | 748.7 | 783.5 | 728.3 | 783 | 706.5 | 707 | 689.6 | 799.6 | 733.7 | 754.2 | 746.2 | 793.5 | 758.3 | 742.2 | 739.6 | 733.8 | 694.4 | 762.7 | 765.5 | 731.4 | 760.4 | 682.9 | 785.8 | 686.9 | 770.3 | 629.2 | 619.4 | 302.4 | 718 | 984.9 | 779.5 | 777 | 530.6 | 674.4 | 675.2 | 764.7 | 753.3 | 638.2 | 719.4 | 736.3 | 760.5 | 638.7 | 648.8 | 761 | 766.2 | 615.3 | 790 | 511.8 | 702.4 | 746.6 | 603.8 | 763.7 | 662.9 | 712.5 | 704.7 | 698.2 | 766.1 | 603.7 | 754.8 | 895.5 | 788.1 | 728.4 | 1071.5 | 543.4 | 751.3 | 644.6 | 697.2 | 720 | 756.7 | 748.9 | 738 | 755.3 | 709.3 | 592.3 | 690.6 | 736 | 769.6 | 755.7 | 655.9 | 706.4 | 679.7 | 742.5 | 616.1 | 635.6 | 607 | 678.5 | 611.3 | 638.1 | 700.4 | 670.4 | 724.6 | 701.8 | 693.7 | 678 | 620.7 | 698.6 | 543.9 | 678.9 | 653.1 | 591 | 692.3 | 594.1 | 692.9 | 619 | 550 | 720.1 | 647.1 | 535.5 | 586.3 | 637.7 | 574.4 | 765.6 | 583.8 | 732.7 | 735.6 | 691 | 718.1 | 594.2 | 536.2 | 669.8 | 686.2 | 734.9 | 676.2 | 467 | 438.4 | 758.4 | 772.7 | 686.1 | 623 | 713.8 | 732.6 | 725.6 | 743.2 | 764.3 | 811.4 | 743.7 | 796.8 | 758.1 | 786.6 | 725.7 | 462.6 | 641.2 | 657.1 | 646.3 | 531.1 | 664.3 | 690.5 | 731.8 | 806.4 | 710 | 767.8 | 748.4 | 796 | 751 | 771.1 | 722 | 760.6 | 697.4 | 757 | 737.4 | 864.3 | 724.3 | 780 | 783.1 | 714.2 | 716.8 | 775.7 | 950.9 | 724.1 | 667.7 | 774.5 | 745 | 721.6 | 704.4 | 946.1 | 748.1 | 722.3 | 969.1 | 1366.5 | 911.2 | 778.7 | 690.7 | 493.7 | 727.3 | 617 | 907.8 | 1373.7 | 1219.9 | 793.4 | 750.5 | 795.4 | 650.4 | 792.6 | 712.4 | 1100.1 | 632.1 | 781 | 1311.9 | 753.9 | 544.5 | 745.1 | 787.6 | 753.8 | 652.9 | 716.1 | 726.4 | 754.7 | 702.8 | 714.7 | 682.7 | 584.2 | 664 | 742.3 | 745.2 | 744.7 | 716.2 | 738.4 | 708.2 | 665.2 | 670.2 | 597.6 | 772.3 | 688.4 | 743.5 | 732.8 | 756.1 | 751.9 | 763.6 | 605.2 | 699.2 | 775.9 | 734.4 | 745.7 | 761.2 | 781.7 | 635.8 | 697.7 | 786.4 | 688 | 801.3 | 608.5 | 710.7 | 792.1 | 721.1 | 695.6 | 757.9 | 720.5 | 670.5 | 787.2 | 730.5 | 754.9 | 730.7 | 774.4 | 784.6 | 777.6 | 797.9 | 788.9 | 788.2 | 785.6 | 783.9 | 783.4 | 773.7 | 755.1 | 764.9 | 756.4 | 819.7 | 768.7 | 784.7 | 791.5 | 790.8 | 778.1 | 768.4 |  | 788.7 | 755.9 | 767.5 | 795.8 | 779 | 772.9 | 773.4 | 779.4 | 790.9 | 798.5 | 764 | 734.7 | 785.3 | 753.4 | 772 | 786.1 | 758.5 | 793.2 | 777.4 | 787.3 | 789.9 | 763.4 | 806.8 | 774.7 | 801.8 | 788.4 | 783.8 | 771.9 | 780.3 | 788.8 | 780.5 | 480.2 | 783.7 | 776.5 | 784.8 | 771.4 | 772.5 | 793.1 | 752.4 | 780.2 | 751.7 | 743.4 | 778.3 | 789.5 | 771.5 | 756.6 | 786.2 | 745.3 | 762.3 | 791.6 | 804.4 | 735.4 | 738.1 | 797.7 | 1052.6 | 750 | 784.3 | 818.1 | 789 | 769.3 | 789.6 | 829.5 | 789.1 | 543.5 | 767 | 747.7 | 800.9 | 658.8 | 727.8 | 772.8 | 742.1 | 684.2 | 767.1 | 515.4 | 733.9 | 686.7 | 815.8 | 683.8 | 726.6 | 1013.4 | 697.1 | 595.4 | 671.3 | 605.9 |  | 707.4 | 751.1 | 689.7 | 674 | 692.4 | 540.8 | 749 | 774.8 | 678.8 | 646.2 | 776.3 | 798.3 | 752.7 | 707.9 | 705.4 | 548.3 | 635.5 | 622.3 | 791.2 | 623.3 | 764.2 | 509.7 | 738.7 | 680.5 | 696.6 | 736.6 | 671.8 | 675.1 | 674.9 | 587.4 | 587.1 | 654.4 | 692.5 | 703.2 | 644.8 | 794.8 | 768.5 | 763.2 | 759 | 728.2 | 808.1 | 703.5 | 784.9 | 804.6 | 798.4 | 785 | 797.6 | 576.6 | 630.4 | 790.5 | 765.2 | 787 | 807.8 | 778 | 759.9 | 563.6 | 680.7 | 610 | 774.9 | 413.9 | 568.9 | 643.3 | 729.1 | 813.2 | 685.3 | 667.4 | 708.9 | 660.9 | 613.1 | 577.9 | 487.9 | 653.6 | 687 | 682.3 | 658.7 | 696.3 | 614.9 | 709.9 | 694.8 | 474.3 | 464 | 687.8 | 737.7 | 496 | 672.1 | 582 | 662.1 | 747.2 | 803.8 | 725.4 | 756.8 | 734.3 | 627.5 | 718.5 | 744.2 | 538.6 | 539.3 | 578.6 | 655.2 | 739.5 | 649.4 | 572.9 | 769.7 | 695.4 | 626.1 | 748.3 | 508.5 | 805.4 | 757.4 | 788.3 | 631.2 | 562.8 | 660.1 | 512.5 | 667.8 | 2426.8 | 417.3 | 634.6 | 704.5 | 498.2 | 563 | 716.6 | 715.1 | 747.9 | 557.8 | 657.9 | 660 | 711.4 | 621.1 | 713.2 | 721.8 | 780.8 | 679.1 | 625.5 | 561.6 | 539.5 | 545.6 | 535.2 | 700.1 | 624.7 | 582.3 | 591.3 | 773.1 | 376 | 626 | 694.3 | 677.3 | 385.3 | 724.2 | 297.9 | 741.6 | 754.6 | 510.1 | 551.6 | 732.3 | 501.1 | 681.2 | 665.8 | 693.1 | 739.1 | 695.9 | 639.4 | 759.3 | 338.7 | 674.8 | 625.9 | 477 | 357 | 880.6 | 1261.1 | 634.2 | 804.8 | 495.3 | 661.3 | 1157.9 | 1272.1 | 571.6 | 611.7 | 595 | 730.1 | 510.8 | 728.6 | 816 | 655 | 608.4 | 679.8 | 777.2 | 605.8 | 654.7 | 740.7 | 696.5 | 558.1 | 588.3 | 590 | 644.7 | 682.1 | 899.8 | 736.4 | 591.6 | 660.6 | 449.4 | 609.6 | 522.7 | 737.2 | 1559.3 | 669.6 | 759.5 | 658.1 | 679.9 | 675 | 696.2 | 780.7 | 488.1 | 765 | 618.1 | 874.4 | 593 | 656.9 | 599.1 | 685.9 | 688.5 | 592.8 | 576.4 | 725 | 613.3 | 742.4 | 688.9 | 724.8 | 664.6 | 596.5 | 532.7 | 693.4 | 750.8 | 813.1 | 789.7 | 796.3 | 795.9 | 797.1 | 780.9 | 768.9 | 763.8 | 725.3 | 737.1 | 789.3 | 1739.1 | 797.8 | 745.9 | 781.2 | 800.1 | 755.2 | 798.1 | 790.2 | 790.3 | 822.7 | 751.5 | 779.7 | 759.7 | 746.1 | 765.8 | 796.2 | 775.8 | 811.9 | 792.3 | 736.9 | 761.4 | 779.9 | 744 | 797.5 | 794.4 | 777.3 | 788 | 749.1 | 781.4 | 775.6 | 700.5 | 769.1 | 739 | 780.6 | 782.5 | 2597.3 | 741.5 | 738.2 | 765.3 | 784 | 761.7 | 741.3 | 723.3 | 742.8 | 796.4 | 749.5 | 773.8 | 722.4 | 770.2 | 731.9 | 768.1 | 798.9 | 794.3 | 802.3 | 899.2 | 801.1 | 756.2 | 809.9 | 744.3 | 798.2 | 810.7 | 782.9 | 801.2 | 796.6 | 812 | 769.9 | 775.1 | 749.4 | 778.9 | 806.6 | 729.5 | 792 | 753.7 | 801.9 | 815.9 | 786 | 810.1 | 776.2 | 830.5 | 778.4 | 7239.5 | 797.3 | 814.6 | 814.7 | 817.9 | 821.8 | 791.9 | 740.2 | 812.7 | 7366.9 | 812.3 | 589.2 | 770.6 | 7653.4 | 782.2 | 759.1 | 7610.2 | 1571 | 801.6 | 7707.8 | 1581.4 | 766.5 | 1680.8 | 7781 | 794.7 | 7716.8 | 1659.7 | 796.5 | 818.9 | 676 | 7818.5 | 587.7 | 1774 | 795.1 | 652.5 | 1722.1 | 7778.7 | 565.5 | 700.8 | 687.4 | 1729.8 | 7719 | 7721.6 | 828.4 | 825.1 | 824.3 | 773.2 | 773.9 | 1761.4 | 7944.2 | 816.8 | 1787 | 1111.2 | 831.7 | 829.8 | 1736.5 | 808.8 | 7915.6 | 790.7 | 1780.7 | 694 | 818.4 | 724.9 | 8068.9 | 787.9 | 795 | 816.9 | 1766.3 | 816.4 | 7801.5 | 825.5 | 7503.8 | 1797.7 | 800.8 | 814.2 | 760.8 | 7906.1 | 1783.7 | 820.9 | 7844.6 | 1782.4 | 746.3 | 795.3 | 786.7 | 682.6 | 801 | 755.6 | 747.5 | 794.6 | 794.5 | 780.1 | 817.1 | 759.6 | 797.4 | 766.4 | 2583.6 | 796.1 | 934.7 | 746 | 765.4 | 593.5
UOM | [10 empty] | Deg | mm | mm | mm | lb/inch | mm |  | Lost Steps | Lost Steps | Lost Steps | Deg C/Min | PSI | PSI | nL/min | area | ppm | %RSD | %RSD | ul |  | area | area | area | area |  | area | ppm | area | area | area | area | area | area | area | area | area | area | area | area | area | area | ppm | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area |  | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area |  | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area | area
P | 2018-07-29T12:31:22+08:00 | 2018-07-29T16:07:58+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | G18SDI112G | CPSM Final Test | PASS | 129 | 165.92 | 0.08 | -1.07 | 56.41 | 52.07 | 1.21 | 0.13 | 1 | -2 | 1 | -0.8759 | 0.524 | -2.6 | 4 | 0 | 0 | 0.0315 | 0.2693 | 0.0066 | 1
P | 2018-07-29T15:49:07+08:00 | 2018-07-29T16:22:59+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI122G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-07-29T13:34:14+08:00 | 2018-07-29T16:27:23+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI118G | CPSM Final Test | PASS | 129 | 162.32 | 0.27 | -0.69 | 56.72 | 53.85 | 1.74 | 0.13 | -1 | 1 | 0 | -0.8201 | 0.514 | 12.1 | 144 |  | 0.01732 | 0.0336 | 0.259 | 0.0123 | 1 | 0
P | 2018-07-29T14:38:05+08:00 | 2018-07-29T16:47:48+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI123G | CPSM Final Test | FAIL | [13 empty] | 20.3
P | 2018-07-29T14:33:11+08:00 | 2018-07-29T17:38:55+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI122G | CPSM Final Test | FAIL | [14 empty] | 979
P | 2018-07-29T15:38:09+08:00 | 2018-07-29T18:18:28+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI120G | CPSM Final Test | PASS | 129 | 165.83 | -0.02 | -0.77 | 56.52 | 49.95 | 1.82 | 0.13 | 0 | -1 | 0 | -0.8161 | 0.4942 | 8.5 | 14 |  | 0 | 0.2957 | 0.3403 | -0.0098 | 1 |  | 0
P | 2018-07-29T15:42:56+08:00 | 2018-07-29T18:48:06+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI123G | CPSM Final Test | PASS | 129 | 160.79 | 0.17 | -1.38 | 56.14 | 49.28 | 1.51 | 0.13 | -1 | -2 | 0 | -0.7982 | 0.51 | 4.3 | 23 |  | 0 | 0.1049 | 0.2721 | -0.0234 | 0.99996 |  |  | 0
P | 2018-07-29T16:06:03+08:00 | 2018-07-29T19:23:35+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI122G | CPSM Final Test | PASS | 129 | 159.08 | -0.02 | -0.46 | 57.14 | 49.36 | 1.97 | 0.13 | 0 | -1 | 0 | -0.8002 | 0.5362 | 2.4 | -7 | 0 | 0.02371 | 0.0361 | 0.1861 | -0.028 | 0.99999
P | 2018-07-30T10:54:04+08:00 | 2018-07-30T12:33:38+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | G18SDI141G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-07-30T11:43:51+08:00 | 2018-07-30T15:33:52+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | G18SDI141G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -0.84 | 55.91 | 51.22 | 1.36 | 0.13 | 1 | -2 | 0 | -0.852 | 0.5148 | 9.3 | 35 |  | 0 | 0.1942 | 0.1595 | 0.0344 | 0.99995 |  |  |  | 0
P | 2018-07-30T13:24:55+08:00 | 2018-07-30T16:14:47+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP560G | CPSM Final Test | PASS |  | 163.58 | 0.17 | -0.38 | 56.49 | 50.63 | 1.06 | 0.13 | 0 | 1 | 0 | -0.9518 | 0.541 | 4.8 | -2 |  |  | 0.0999 | 0.3408 | -0.0236 | 1 |  |  |  |  | 133 | 0 | 0
P | 2018-07-30T13:23:47+08:00 | 2018-07-30T16:17:22+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP562G | CPSM Final Test | PASS |  | 161.51 | -0.02 | -0.84 | 56.1 | 49.95 | 1.36 | 0.13 | 1 | -2 | 0 | -0.9878 | 0.508 | 5.3 | 17 |  |  | 0.1118 | 0.0859 | 0.0226 | 1 |  |  |  |  | 133 |  | 0 | 0
P | 2018-07-30T13:21:45+08:00 | 2018-07-30T16:19:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP561G | CPSM Final Test | PASS |  | 163.13 | 0.27 | -0.77 | 56.19 | 49.11 | 1.13 | 0.13 | 0 | -1 | 0 | -0.9318 | 0.5258 | 5.5 | 36 |  |  | 0.0857 | 0.2792 | 0.0082 | 0.99999 |  |  |  |  | 133 |  | 0 |  | 0
P | 2018-07-30T13:22:49+08:00 | 2018-07-30T16:21:27+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | G18FNP558G | CPSM Final Test | PASS |  | 163.58 | 0.46 | -0.46 | 56.32 | 52.92 | 1.21 | 0.13 | -1 | -3 | 0 | -0.9817 | 0.4804 | 9.7 | 18 |  |  | 0.153 | 0.1697 | 0.0156 | 0.99998 |  |  |  |  | 133 |  | 0 |  |  | 0
P | 2018-07-30T13:26:03+08:00 | 2018-07-30T16:21:58+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP559G | CPSM Final Test | PASS |  | 161.33 | 0.17 | -0.69 | 56.45 | 50.97 | 0.83 | 0.13 | 0 | 2 | 0 | -0.9139 | 0.505 | 5.8 | 13 |  |  | 0.0354 | 0.2416 | -0.0214 | 0.99999 |  |  |  |  | 133 |  | 0 |  |  |  | 0
P | 2018-07-30T13:32:09+08:00 | 2018-07-30T16:29:53+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | G18FNP557G | CPSM Final Test | PASS |  | 159.35 | 0.27 | -1.07 | 56.24 | 48.43 | 1.36 | 0.13 | 1 | -2 | 0 | -1.0197 | 0.488 | 4.3 | -27 |  |  | 0.0718 | 0.423 | -0.0542 | 1 |  |  |  |  | 133 |  | 0 |  |  |  |  | 0
P | 2018-07-30T13:27:51+08:00 | 2018-07-30T16:25:34+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | G18FNP555G | CPSM Final Test | PASS |  | 163.4 | -0.02 | -1.07 | 56.27 | 49.95 | 0.98 | 0.13 | 1 | -3 | 0 | -1.0176 | 0.482 | 6.1 | -8 |  |  | 0.2145 | 0.2414 | -0.0511 | 0.99999 |  |  |  |  | 133 |  | 0 | [5 empty] | 0
P | 2018-07-30T15:05:31+08:00 | 2018-07-30T19:06:56+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP571G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-07-30T14:41:26+08:00 | 2018-07-30T19:38:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP563G | CPSM Final Test | FAIL
P | 2018-07-30T14:32:25+08:00 | 2018-07-30T19:55:42+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP565G | CPSM Final Test | FAIL
P | 2018-07-30T14:44:30+08:00 | 2018-07-30T20:07:28+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | G18FNP569G | CPSM Final Test | FAIL | [13 empty] | 21.4
P | 2018-07-30T13:40:14+08:00 | 2018-07-30T20:24:30+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP572G | CPSM Final Test | FAIL
P | 2018-07-30T17:21:34+08:00 | 2018-07-30T20:37:32+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | G18FNP570G | CPSM Final Test | PASS |  | 161.42 | -0.21 | -0.61 | 56.52 | 48.51 | 1.29 | 0.13 | -2 | -2 | 0 | -0.9598 | 0.5188 | 5.5 | 7 |  |  | 0.0705 | 0.1212 | -0.046 | 0.99999 |  |  |  |  | 133 |  | 0 | [6 empty] | 0
P | 2018-07-30T17:23:03+08:00 | 2018-07-30T20:45:23+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | G18FNP566G | CPSM Final Test | PASS |  | 161.24 | 0.37 | 0.15 | 56.32 | 51.48 | 1.36 | 0.13 | -1 | -2 | 0 | -0.9477 | 0.5132 | 8.9 | 6 |  |  | 0.1274 | 0.412 | 0.0129 | 0.99996 |  |  |  |  | 133 |  | 0 | [7 empty] | 0
P | 2018-07-30T16:58:59+08:00 | 2018-07-30T20:48:00+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | G18FNP564G | CPSM Final Test | PASS |  | 165.11 | 0.27 | -0.54 | 56.47 | 51.48 | 1.29 | 0.13 | -1 | -4 | 0 | -0.9318 | 0.5218 | 4.4 | -42 |  |  | 0.0746 | 0.1297 | -0.025 | 0.99999 |  |  |  |  | 133 |  | 0 | [8 empty] | 0
P | 2018-07-30T17:14:04+08:00 | 2018-07-30T21:01:06+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP565G | CPSM Final Test | PASS |  | 162.68 | 0.27 | -0.38 | 56.62 | 49.7 | 2.05 | 0.13 | 0 | -2 | 0 | -0.9359 | 0.4952 | 2.3 | 23 |  |  | 0.0857 | 0.2264 | -0.0527 | 0.99999 |  |  |  |  | 133 |  | 0 | [9 empty] | 0
P | 2018-07-30T17:06:14+08:00 | 2018-07-30T21:43:38+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP571G | CPSM Final Test | PASS |  | 162.86 | 0.27 | -0.61 | 56.5 | 48.09 | 1.21 | 0.13 | -1 | -1 | 0 | -1.0094 | 0.5016 | 5.5 | 93 |  |  | 0.0521 | 0.3443 | -0.0432 | 0.99999 |  |  |  |  | 133 |  | 0 | [10 empty] | 0
P | 2018-07-30T13:26:12+08:00 | 2018-07-30T21:25:46+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | G18FNP569G | CPSM Final Test | FAIL
P | 2018-07-30T01:51:42+08:00 | 2018-07-31T08:28:37+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP563G | CPSM Final Test | FAIL
P | 2018-07-30T17:10:09+08:00 | 2018-07-31T08:29:13+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP563G | CPSM Final Test | PASS |  | 161.96 | 0.37 | -0.31 | 56.54 | 48.94 | 1.82 | 0.13 | 0 | -1 | 0 | -0.9358 | 0.5192 | 2 | 12 |  |  | 0.062 | 0.2901 | -0.0899 | 1 |  |  |  |  | 133 |  | 0 | [11 empty] | 0
P | 2018-07-30T00:18:08+08:00 | 2018-07-31T09:46:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP572G | CPSM Final Test | FAIL
P | 2018-07-30T00:02:24+08:00 | 2018-07-31T10:49:34+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | G18FNP569G | CPSM Final Test | FAIL
P | 2018-07-30T17:02:22+08:00 | 2018-07-31T10:52:10+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP572G | CPSM Final Test | PASS |  | 163.94 | -0.02 | -0.99 | 56.41 | 54.02 | 1.51 | 0.13 | 0 | 1 | 0 | -0.8939 | 0.4952 | 9.3 | 39 |  |  | 0.251 | 0.0931 | -0.0104 | 0.99998 |  |  |  |  | 133 |  | 0 | [12 empty] | 0
P | 2018-07-31T09:35:43+08:00 | 2018-07-31T11:35:35+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP799G | CPSM Final Test | FAIL | [5 empty] | 45.97
P | 2018-07-30T17:25:59+08:00 | 2018-07-31T11:52:22+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | G18FNP569G | CPSM Final Test | PASS |  | 160.61 | 0.27 | -0.84 | 56.5 | 51.22 | 1.82 | 0.13 | 0 | -4 | 1 | -0.852 | 0.5228 | 8.3 | 12 |  |  | 0.1363 | 0.2537 | 0.0875 | 0.99991 |  |  |  |  | 133 |  | 0.00324 | [13 empty] | 0
P | 2018-07-31T09:03:43+08:00 | 2018-07-31T12:07:35+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP796G | CPSM Final Test | FAIL | [5 empty] | 45.47
P | 2018-07-31T08:53:03+08:00 | 2018-07-31T12:18:15+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP799G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-07-31T08:32:52+08:00 | 2018-07-31T12:18:04+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP797G | CPSM Final Test | FAIL
P | 2018-07-31T08:57:31+08:00 | 2018-07-31T13:42:11+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP556G | CPSM Final Test | FAIL
P | 2018-07-31T11:03:08+08:00 | 2018-07-31T13:53:44+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP800G | CPSM Final Test | PASS | 129 | 160.25 | -0.02 | -0.31 | 56.62 | 50.04 | 1.67 | 0.13 | 1 | 2 | 0 | -0.8978 | 0.5214 | 2.3 | 28 |  |  | 0.0992 | 0.2801 | -0.0363 | 0.99999 | [19 empty] | 0 |  | 0
P | 2018-07-31T11:06:47+08:00 | 2018-07-31T13:59:10+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP567G | CPSM Final Test | PASS |  | 162.5 | -0.02 | -0.54 | 56.39 | 52.32 | 1.13 | 0.13 | 1 | -2 | -1 | -0.9442 | 0.5166 | 4.6 | -6 |  |  | 0.1098 | 0.3687 | -0.0061 | 1 |  |  |  |  | 133 |  | 0 | [15 empty] | 0
P | 2018-07-31T11:05:37+08:00 | 2018-07-31T14:01:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | G18FNP568G | CPSM Final Test | PASS |  | 159.71 | 0.27 | -0.69 | 56.41 | 48.34 | 1.67 | 0.13 | 1 | -3 | 1 | -0.9717 | 0.5206 | 11.7 | -3 |  |  | 0.1205 | 0.203 | -0.0281 | 0.99996 |  |  |  |  | 133 |  | 0 | [16 empty] | 0
P | 2018-07-31T10:25:28+08:00 | 2018-07-31T14:04:02+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP797G | CPSM Final Test | PASS | 129 | 159.53 | 0.27 | -0.69 | 56.61 | 50.8 | 1.59 | 0.13 | 0 | -4 | 0 | -0.9019 | 0.4924 | 5 | 5 |  |  | 0.1656 | 0.2839 | -0.0827 | 1 | [16 empty] | 0 |  |  |  |  | 0.05639
P | 2018-07-31T11:19:51+08:00 | 2018-07-31T14:47:47+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP556G | CPSM Final Test | PASS |  | 153.41 | -0.02 | -0.46 | 56.37 | 50.38 | 1.51 | 0.13 | 0 | -5 | 0 | -0.8521 | 0.492 | -0.5 | 151 |  |  | 0.1772 | 0.1374 | -0.0113 | 0.99998 |  |  |  |  | 133 |  | 0 | [17 empty] | 0
P | 2018-07-31T10:35:39+08:00 | 2018-07-31T15:32:37+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP796G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.16 | 56.81 | 48.6 | 1.74 | 0.13 | 0 | 1 | 0 | -0.898 | 0.5232 | 3.3 | 30 |  |  | 0.0402 | 0.0816 | -0.0137 | 1 | [21 empty] | 0 |  |  |  | 0
P | 2018-07-31T10:35:39+08:00 | 2018-07-31T15:38:05+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP799G | CPSM Final Test | PASS | 129 | 161.87 | 0.17 | -0.46 | 56.09 | 49.36 | 1.51 | 0.13 | -1 | -2 | 0 | -0.8839 | 0.5122 | -1.5 | 0 |  |  | 0.1017 | 0.2511 | -0.0104 | 1 | [21 empty] | 0 |  |  |  |  | 0
P | 2018-07-31T15:45:45+08:00 | 2018-07-31T17:20:03+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP798G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-07-31T15:40:59+08:00 | 2018-07-31T17:22:58+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP801G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-07-31T15:28:17+08:00 | 2018-07-31T17:35:40+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP801G | CPSM Final Test | FAIL
P | 2018-07-31T13:41:40+08:00 | 2018-07-31T18:50:00+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP804G | CPSM Final Test | FAIL | [13 empty] | 26.9
P | 2018-07-31T13:11:11+08:00 | 2018-07-31T19:37:16+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP795G | CPSM Final Test | FAIL
P | 2018-07-31T12:44:29+08:00 | 2018-07-31T19:38:52+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP802G | CPSM Final Test | FAIL
P | 2018-07-31T12:24:02+08:00 | 2018-07-31T20:07:38+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP804G | CPSM Final Test | FAIL
P | 2018-07-31T16:07:36+08:00 | 2018-07-31T20:21:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP805G | CPSM Final Test | PASS | 129 | 169.07 | 0.37 | -0.84 | 56.73 | 53.68 | 1.51 | 0.13 | 1 | -2 | 0 | -0.9178 | 0.521 | 5.2 | -29 |  |  | 0.06 | 0.1002 | -0.001 | 0.99999 | [21 empty] | 0.00944 | [6 empty] | 0
P | 2018-07-31T12:17:37+08:00 | 2018-07-31T20:30:50+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP795G | CPSM Final Test | FAIL
P | 2018-07-31T16:32:54+08:00 | 2018-07-31T20:37:05+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP798G | CPSM Final Test | PASS | 129 | 161.51 | 0.37 | -0.08 | 56.63 | 50.04 | 1.67 | 0.13 | 0 | 1 | 0 | -0.9159 | 0.518 | 3.5 | 46 |  |  | 0.1174 | 0.2923 | 0.0148 | 0.99999 | [21 empty] | 0 | [7 empty] | 0
P | 2018-07-31T16:31:58+08:00 | 2018-07-31T20:32:50+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP801G | CPSM Final Test | PASS | 129 | 157.73 | -0.02 | -2.29 | 56.43 | 51.31 | 1.44 | 0.13 | -1 | -2 | 1 | -0.9678 | 0.496 | -1.4 | 15 |  |  | 0.1388 | 0.2523 | 0.014 | 1 | [21 empty] | 0 | [8 empty] | 0
P | 2018-07-31T16:15:50+08:00 | 2018-07-31T21:22:12+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP804G | CPSM Final Test | PASS | 129 | 158.45 | 0.37 | -0.84 | 56.9 | 49.95 | 1.9 | 0.13 | 0 | -2 | 0 | -0.886 | 0.525 | 14 | 6 |  |  | 0.0671 | 0.2184 | -0.0215 | 0.99999 | [21 empty] | 0 | [9 empty] | 0
P | 2018-07-31T16:11:40+08:00 | 2018-07-31T21:23:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP802G | CPSM Final Test | PASS | 129 | 158 | 0.37 | -1.3 | 56.71 | 53.34 | 1.44 | 0.13 | -1 | 1 | 0 | -0.8978 | 0.519 | 3.8 | 1 |  |  | 0.2399 | 0.1958 | 0.0443 | 0.99998 | [21 empty] | 0 | [10 empty] | 0
P | 2018-07-31T16:24:13+08:00 | 2018-07-31T22:15:00+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP795G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.92 | 56.51 | 48.18 | 1.51 | 0.13 | 1 | -3 | 0 | -0.9978 | 0.527 | 4.4 | 1 |  |  | 0.0763 | 0.2913 | 0.0132 | 0.99997 | [21 empty] | 0.02906 | [11 empty] | 0
P | 2018-08-01T09:44:40+08:00 | 2018-08-01T12:00:44+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP806G | CPSM Final Test | FAIL | [5 empty] | 45.55
P | 2018-07-31T03:47:46+08:00 | 2018-08-01T12:31:12+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP803G | CPSM Final Test | FAIL | [55 empty] | 0
P | 2018-08-01T10:15:16+08:00 | 2018-08-01T13:21:38+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP810G | CPSM Final Test | PASS | 129 | 164.57 | -0.02 | -0.84 | 56.49 | 50.12 | 1.21 | 0.13 | -1 | 1 | 0 | -0.8421 | 0.5386 | 6.4 | 9 |  |  | 0.1796 | 0.4105 | 0.0226 | 0.99994 | [20 empty] | 0 | 0
P | 2018-07-31T20:09:29+08:00 | 2018-08-01T13:37:46+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP803G | CPSM Final Test | PASS | 129 | 159.53 | 0.08 | 0.07 | 57.07 | 50.55 | 1.29 | 0.13 | 1 | 3 | -1 | -0.9458 | 0.4912 | -0.3 | 11 |  |  | 0.1709 | 0.1895 | -0.0174 | 1 | [21 empty] | 0 |  |  | 0
P | 2018-08-01T11:02:30+08:00 | 2018-08-01T13:56:35+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP808G | CPSM Final Test | PASS | 129 | 159.17 | 0.27 | -0.54 | 56.2 | 49.53 | 1.36 | 0.13 | -1 | 1 | -1 | -0.89 | 0.541 | -0.4 | 4 |  |  | 0.0565 | 0.2434 | 0.0335 | 1 | [21 empty] | 0 | [13 empty] | 0
P | 2018-08-01T10:49:05+08:00 | 2018-08-01T13:55:06+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP809G | CPSM Final Test | PASS | 129 | 160.43 | 0.17 | -0.46 | 56.4 | 53.34 | 1.44 | 0.13 | 0 | -3 | 0 | -0.85 | 0.476 | 1.9 | 42 |  |  | 0.1014 | 0.7733 | 0.029 | 0.99998 | [21 empty] | 0 | [14 empty] | 0
P | 2018-08-01T10:45:29+08:00 | 2018-08-01T13:57:00+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP811G | CPSM Final Test | PASS | 129 | 161.24 | 0.27 | -0.61 | 56.76 | 50.12 | 0.91 | 0.13 | -1 | 1 | 0 | -0.8859 | 0.5278 | 4.3 | -19 |  |  | 0.1013 | 0.1591 | 0.0096 | 0.99999 | [21 empty] | 0.03652 | [15 empty] | 0
P | 2018-08-01T11:30:44+08:00 | 2018-08-01T15:18:06+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP813G | CPSM Final Test | PASS | 129 | 166.55 | -0.02 | 0.3 | 56.48 | 50.55 | 1.82 | 0.13 | 1 | 2 | -1 | -0.9198 | 0.5218 | 4.4 | 8 |  |  | 0.0295 | 0.1167 | -0.0069 | 0.99999 | [21 empty] | 0 | [16 empty] | 0
P | 2018-08-01T10:52:42+08:00 | 2018-08-01T15:20:16+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP806G | CPSM Final Test | PASS | 129 | 162.05 | 0.56 | 0.61 | 56.7 | 50.12 | 1.59 | 0.13 | 0 | 3 | -1 | -0.902 | 0.5186 | 9.6 | 21 |  |  | 0.1169 | 0.2793 | -0.0019 | 0.99999 | [21 empty] | 0 | [17 empty] | 0
P | 2018-08-01T11:27:08+08:00 | 2018-08-01T15:20:26+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP807G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | 0.3 | 56.77 | 49.11 | 1.36 | 0.13 | 1 | 2 | 0 | -0.9099 | 0.4936 | 8 | 1 |  |  | 0.0824 | 0.4041 | -0.0219 | 0.99999 | [21 empty] | 0.02916 | [18 empty] | 0
P | 2018-08-01T15:21:52+08:00 | 2018-08-01T16:33:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP817G | CPSM Final Test | FAIL
P | 2018-08-01T19:24:25+08:00 | 2018-08-01T19:46:26+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP820G | CPSM Final Test | FAIL
P | 2018-08-01T11:48:43+08:00 | 2018-08-01T20:15:02+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP812G | CPSM Final Test | FAIL | [42 empty] | 0.17296
P | 2018-08-01T16:16:01+08:00 | 2018-08-01T20:18:50+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP814G | CPSM Final Test | PASS | 129 | 158.99 | -0.02 | -0.61 | 56.44 | 52.83 | 1.9 | 0.13 | -1 | -4 | 0 | -0.908 | 0.5118 | 6.8 | -7 |  |  | 0.0486 | 0.249 | -0.05 | 1 | [21 empty] | 0 |  |  | 0
P | 2018-08-01T16:11:01+08:00 | 2018-08-01T20:28:01+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP815G | CPSM Final Test | PASS | 129 | 162.77 | -0.02 | 0.15 | 56.73 | 53.17 | 1.74 | 0.13 | 1 | 2 | 0 | -0.9079 | 0.4908 | 11.2 | 2 |  |  | 0.03 | 0.4155 | 0.0074 | 0.99998 | [21 empty] | 0 | [19 empty] | 0
P | 2018-08-01T15:54:04+08:00 | 2018-08-01T20:25:42+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP816G | CPSM Final Test | PASS | 129 | 156.65 | -0.02 | -0.77 | 56.44 | 54.44 | 1.36 | 0.13 | 0 | -3 | 0 | -0.7782 | 0.5318 | 4.2 | 51 |  |  | 0.3215 | 0.5283 | 0.0169 | 0.99993 | [21 empty] | 0 | [20 empty] | 0
P | 2018-08-01T15:57:31+08:00 | 2018-08-01T20:29:16+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP817G | CPSM Final Test | PASS | 129 | 165.92 | 0.27 | -0.77 | 56.51 | 49.87 | 1.59 | 0.22 | 0 | -1 | 0 | -0.9059 | 0.521 | 2.7 | 12 |  |  | 0.0577 | 0.1376 | -0.0331 | 1 | [21 empty] | 0 | [21 empty] | 0
P | 2018-08-01T16:01:52+08:00 | 2018-08-01T21:18:04+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP812G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -0.69 | 56.49 | 50.21 | 1.67 | 0.13 | -1 | 1 | 0 | -0.9178 | 0.495 | 7.1 | 5 |  |  | 0.1105 | 0.3637 | -0.0849 | 1 | [21 empty] | 0 | [22 empty] | 0
P | 2018-08-02T09:43:41+08:00 | 2018-08-02T11:41:32+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP824G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-08-02T08:43:57+08:00 | 2018-08-02T12:45:17+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP822G | CPSM Final Test | FAIL | [66 empty] | 0
P | 2018-08-02T10:29:26+08:00 | 2018-08-02T13:34:04+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP821G | CPSM Final Test | PASS | 129 | 162.5 | 0.68 | -0.54 | 56.63 | 48.85 | 1.51 | 0.13 | 0 | -3 | 0 | -0.8819 | 0.5102 | 6.7 | -1 |  |  | 0.133 | 0.2743 | -0.0706 | 1 | [21 empty] | 0 | [24 empty] | 0
P | 2018-08-02T10:33:42+08:00 | 2018-08-02T13:35:44+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP823G | CPSM Final Test | PASS | 129 | 161.96 | -0.33 | -0.77 | 56.61 | 49.95 | 1.82 | 0.13 | 1 | 3 | 0 | -0.9178 | 0.5114 | 3.7 | 1 |  |  | 0.1097 | 0.409 | -0.057 | 1 | [21 empty] | 0 | [25 empty] | 0
P | 2018-08-02T10:44:37+08:00 | 2018-08-02T14:19:09+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP822G | CPSM Final Test | PASS | 129 | 159.98 | 0.37 | -0.77 | 56.46 | 49.11 | 1.21 | 0.13 | -1 | 1 | 0 | -0.83 | 0.502 | -0.3 | 21 |  |  | 0.1154 | 0.6076 | 0.0043 | 0.99999 | [21 empty] | 0 | [23 empty] | 0
P | 2018-08-01T00:21:40+08:00 | 2018-08-02T14:19:02+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP819G | CPSM Final Test | FAIL
P | 2018-08-01T19:20:21+08:00 | 2018-08-02T14:19:42+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP819G | CPSM Final Test | PASS | 129 | 163.85 | -0.11 | -0.61 | 56.72 | 50.29 | 1.44 | 0.13 | 0 | -3 | 0 | -0.858 | 0.5374 | 2.7 | -32 |  |  | 0.0835 | 0.0805 | -0.0232 | 0.99999 | [21 empty] | 0 | [26 empty] | 0
P | 2018-08-02T10:42:36+08:00 | 2018-08-02T14:24:28+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP824G | CPSM Final Test | PASS | 129 | 157.37 | 0.37 | -0.54 | 56.43 | 48.6 | 1.21 | 0.13 | 1 | 3 | 0 | -0.9239 | 0.523 | 7.3 | -26 |  |  | 0.0899 | 0.171 | -0.0231 | 1 | [19 empty] | 0 |  | 0
P | 2018-08-02T09:56:33+08:00 | 2018-08-02T14:42:50+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP825G | CPSM Final Test | FAIL
P | 2018-08-02T08:55:09+08:00 | 2018-08-02T15:44:14+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP825G | CPSM Final Test | FAIL | [70 empty] | 0
P | 2018-08-02T12:19:41+08:00 | 2018-08-02T16:49:36+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP825G | CPSM Final Test | PASS | 129 | 161.06 | -0.02 | -0.69 | 56.47 | 48.43 | 1.67 | 0.13 | -1 | 1 | 0 | -0.7702 | 0.5358 | 1.6 | 41 |  |  | 0.106 | 0.1895 | 0.0302 | 0.99992 | [21 empty] | 0 | [27 empty] | 0
P | 2018-08-02T14:54:54+08:00 | 2018-08-02T17:54:48+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP826G | CPSM Final Test | PASS | 129 | 152.96 | 0.17 | -0.77 | 56.51 | 50.55 | 1.21 | 0.13 | 0 | -1 | -1 | -0.8779 | 0.5152 | 2.4 | 20 |  |  | 0.0569 | 0.1484 | -0.0087 | 0.99999 | [21 empty] | 0 | [28 empty] | 0
P | 2018-08-02T14:56:26+08:00 | 2018-08-02T18:01:49+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP828G | CPSM Final Test | PASS | 129 | 167 | -0.02 | -0.38 | 56.51 | 50.04 | 2.05 | 0.13 | 0 | -2 | 0 | -0.9318 | 0.528 | 2.1 | 1 |  |  | 0.1998 | 0.2348 | 0.0193 | 1 | [21 empty] | 0 | [29 empty] | 0
P | 2018-08-02T15:16:47+08:00 | 2018-08-02T17:57:50+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP829G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.54 | 56.53 | 49.87 | 1.9 | 0.13 | -1 | 1 | 1 | -0.9658 | 0.51 | 5.3 | 3 |  |  | 0.0614 | 0.2346 | -0.0003 | 1 | [21 empty] | 0 | [30 empty] | 0
P | 2018-08-02T14:58:23+08:00 | 2018-08-02T17:59:42+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP827G | CPSM Final Test | PASS | 129 | 162.77 | 0.37 | -0.69 | 56.32 | 51.05 | 1.29 | 0.13 | -1 | 1 | 0 | -0.89 | 0.4896 | 5.7 | 3 |  |  | 0.1464 | 0.0568 | -0.0411 | 1 | [21 empty] | 0 | [31 empty] | 0
P | 2018-08-02T01:35:54+08:00 | 2018-08-02T19:34:32+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP820G | CPSM Final Test | FAIL | [12 empty] | 0.5312
P | 2018-08-02T17:55:57+08:00 | 2018-08-02T19:51:06+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP835G | CPSM Final Test | FAIL
P | 2018-08-02T00:39:02+08:00 | 2018-08-02T20:31:24+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP820G | CPSM Final Test | FAIL
P | 2018-08-02T19:11:51+08:00 | 2018-08-02T22:21:50+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP818G | CPSM Final Test | PASS | 129 | 160.79 | -0.02 | -0.54 | 56.42 | 52.32 | 1.29 | 0.13 | 0 | 2 | 0 | -0.8839 | 0.535 | 5.5 | 9 |  |  | 0.1231 | 0.2775 | -0.0651 | 1 | [21 empty] | 0 | [32 empty] | 0
P | 2018-08-02T18:35:32+08:00 | 2018-08-02T22:23:18+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP833G | CPSM Final Test | PASS | 129 | 165.2 | 0.27 | -0.77 | 56.5 | 48.01 | 2.12 | 0.13 | 0 | 3 | 0 | -0.9278 | 0.51 | 9.6 | -21 |  |  | 0.0996 | 0.0912 | -0.0163 | 0.99999 | [21 empty] | 0 | [31 empty] | 0
P | 2018-08-02T18:52:34+08:00 | 2018-08-02T22:45:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP830G | CPSM Final Test | PASS | 129 | 163.31 | -0.02 | -0.99 | 56.23 | 49.78 | 1.13 | 0.13 | 0 | -3 | 1 | -0.9557 | 0.519 | 8.1 | 13 |  |  | 0.3915 | 0.7033 | 0.0227 | 0.99986 | [21 empty] | 0 | [33 empty] | 0
P | 2018-08-02T04:15:44+08:00 | 2018-08-03T09:03:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP831G | CPSM Final Test | FAIL | [5 empty] | 45.97
P | 2018-08-02T06:01:03+08:00 | 2018-08-03T09:17:57+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP834G | CPSM Final Test | FAIL | [14 empty] | 616
P | 2018-08-02T03:57:00+08:00 | 2018-08-03T09:28:30+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP832G | CPSM Final Test | FAIL
P | 2018-08-02T19:39:30+08:00 | 2018-08-03T11:02:23+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP834G | CPSM Final Test | PASS | 129 | 162.23 | 0.37 | -0.23 | 56.41 | 50.8 | 1.67 | 0.13 | 1 | 1 | 0 | -0.9198 | 0.5148 | 0.4 | 48 |  |  | 0.0665 | 0.0969 | -0.0051 | 1 | [21 empty] | 0 | [34 empty] | 0
P | 2018-08-02T02:02:54+08:00 | 2018-08-03T11:22:36+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP832G | CPSM Final Test | FAIL | [78 empty] | 0
P | 2018-08-02T01:35:14+08:00 | 2018-08-03T11:43:56+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP831G | CPSM Final Test | FAIL | [79 empty] | 0
P | 2018-08-02T01:20:06+08:00 | 2018-08-03T12:05:24+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP832G | CPSM Final Test | FAIL | [78 empty] | 0
P | 2018-08-02T18:39:35+08:00 | 2018-08-03T12:49:24+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP831G | CPSM Final Test | PASS | 129 | 166.28 | -0.02 | -0.61 | 56.32 | 48.09 | 1.74 | 0.21 | 1 | 0 | 0 | -1.0052 | 0.4824 | -3 | 21 |  |  | 0.1214 | 0.4972 | -0.0181 | 1 | [21 empty] | 0 | [36 empty] | 0
P | 2018-08-02T18:42:45+08:00 | 2018-08-03T13:11:06+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP832G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.38 | 56.11 | 50.29 | 1.59 | 0.15 | 0 | 1 | 1 | -0.9498 | 0.499 | 9.2 | -6 |  |  | 0.0551 | 0.203 | -0.0135 | 1 | [21 empty] | 0.04233 | [35 empty] | 0
P | 2018-08-03T10:44:00+08:00 | 2018-08-03T13:40:20+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP838G | CPSM Final Test | PASS | 129 | 162.86 | 0.27 | -0.61 | 56.46 | 48.34 | 1.36 | 0.13 | 0 | -1 | 0 | -0.8919 | 0.5246 | 9.7 | -17 |  |  | 0.0829 | 0.1511 | 0.0229 | 0.99999 | [21 empty] | 0 | [37 empty] | 0
P | 2018-08-03T09:59:41+08:00 | 2018-08-03T13:46:58+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP820G | CPSM Final Test | PASS | 129 | 160.16 | -0.02 | -0.61 | 56.25 | 51.48 | 1.59 | 0.13 | 1 | -2 | 0 | -1.0138 | 0.5052 | 6.4 | -10 |  |  | 0.1595 | 0.1462 | -0.0026 | 0.99998 | [21 empty] | 0 | [38 empty] | 0
P | 2018-08-01T23:44:11+08:00 | 2018-08-03T14:02:52+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP835G | CPSM Final Test | FAIL
P | 2018-08-03T07:25:56+08:00 | 2018-08-03T14:09:32+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP837G | CPSM Final Test | FAIL
P | 2018-08-03T10:47:44+08:00 | 2018-08-03T14:10:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP837G | CPSM Final Test | PASS | 129 | 159.08 | 0.17 | -0.69 | 56.49 | 49.87 | 1.21 | 0.13 | 0 | 1 | 0 | -0.9499 | 0.5088 | 2.5 | -3 |  |  | 0.1864 | 0.3348 | -0.0299 | 0.99999 | [21 empty] | 0 | [39 empty] | 0
P | 2018-08-03T09:50:20+08:00 | 2018-08-03T14:27:16+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP836G | CPSM Final Test | FAIL | [83 empty] | FALSE
P | 2018-08-03T12:11:32+08:00 | 2018-08-03T14:57:56+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP840G | CPSM Final Test | FAIL | [13 empty] | 15.8
P | 2018-08-03T06:28:21+08:00 | 2018-08-03T16:17:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP839G | CPSM Final Test | FAIL | [84 empty] | 0
P | 2018-08-03T11:01:44+08:00 | 2018-08-03T16:07:44+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP840G | CPSM Final Test | FAIL | [85 empty] | 0
P | 2018-08-01T21:03:23+08:00 | 2018-08-03T16:43:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP835G | CPSM Final Test | FAIL
P | 2018-08-03T10:24:38+08:00 | 2018-08-03T16:44:50+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP840G | CPSM Final Test | FAIL | [85 empty] | 0
P | 2018-08-03T11:23:01+08:00 | 2018-08-03T17:23:05+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP839G | CPSM Final Test | PASS | 129 | 156.83 | -0.02 | -0.92 | 56.33 | 48.18 | 1.51 | 0.13 | 1 | -2 | 0 | -0.9379 | 0.511 | 6 | -4 |  |  | 0.2644 | 0.5296 | -0.0468 | 1 | [21 empty] | 0 | [41 empty] | 0
P | 2018-08-03T12:08:48+08:00 | 2018-08-03T17:23:04+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP836G | CPSM Final Test | PASS | 129 | 157.37 | 0.27 | -0.84 | 56.34 | 51.82 | 1.36 | 0.13 | 0 | -2 | 0 | -0.9518 | 0.5052 | 0.3 | 76 |  |  | 0.0971 | 0.1412 | -0.0079 | 0.99998 | [21 empty] | 0 | [43 empty] | 0
P | 2018-08-03T13:34:44+08:00 | 2018-08-03T17:50:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP840G | CPSM Final Test | PASS | 129 | 157.01 | 0.27 | -0.92 | 56.44 | 52.58 | 1.29 | 0.13 | 0 | -1 | 0 | -0.88 | 0.5184 | 10.6 | 7 |  |  | 0.0528 | 0.2417 | -0.0557 | 1 | [21 empty] | 0 | [42 empty] | 0
P | 2018-08-02T18:53:32+08:00 | 2018-08-03T18:28:06+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP835G | CPSM Final Test | PASS | 129 | 162.05 | 0.46 | -1.07 | 56.44 | 51.73 | 1.29 | 0.13 | -1 | 1 | 0 | -0.8659 | 0.4912 | 5.8 | 14 |  |  | 0.1054 | 0.2241 | 0.0371 | 0.99995 | [21 empty] | 0 | [27 empty] | 0
P | 2018-08-04T08:00:51+08:00 | 2018-08-04T10:09:34+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | G18USM064G | CPSM Final Test | FAIL | [13 empty] | -13.2
P | 2018-08-04T06:53:37+08:00 | 2018-08-04T10:31:30+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | G18USM065G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-08-04T11:00:13+08:00 | 2018-08-04T11:06:47+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | G18USM061G | CPSM Final Test | FAIL | [87 empty] | FALSE
P | 2018-08-04T08:56:36+08:00 | 2018-08-04T11:45:54+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | G18USM063G | CPSM Final Test | PASS |  | 159.44 | 0.46 | -0.61 | 56.92 | 49.7 | 1.36 | 0.13 | 0 | 4 | 0 | -0.8759 | 0.5148 | 8.3 | 2 |  |  | 0.023 | 0.2165 | 0.0051 | 0.99999 |  |  |  |  | 133 |  | 0 | [57 empty] | 0
P | 2018-08-04T09:25:29+08:00 | 2018-08-04T12:13:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | G18USM066G | CPSM Final Test | PASS |  | 160.52 | 0.27 | -0.61 | 56.6 | 49.02 | 1.51 | 0.13 | 0 | -2 | 0 | -0.9039 | 0.5016 | 2 | -20 |  |  | 0.3577 | 0.3398 | 0.0571 | 0.99988 |  |  |  |  | 133 |  | 0 | [56 empty] | 0
P | 2018-08-04T09:05:12+08:00 | 2018-08-04T12:13:46+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | G18USM064G | CPSM Final Test | PASS |  | 149.9 | -0.02 | -0.84 | 56.74 | 55.03 | 1.51 | 0.13 | 1 | -2 | 0 | -0.844 | 0.519 | -2 | 10 |  |  | 0.0448 | 0.3345 | -0.0387 | 0.99999 |  |  |  |  | 133 |  | 0.03519 | [50 empty] | 0
P | 2018-08-04T09:38:26+08:00 | 2018-08-04T12:31:38+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | G18USM062G | CPSM Final Test | PASS |  | 166.73 | 0.65 | -0.69 | 56.76 | 51.48 | 1.82 | 0.13 | 0 | 1 | 0 | -0.8899 | 0.49 | -4.3 | -10 |  |  | 0.0753 | 0.0657 | -0.0509 | 0.99999 |  |  |  |  | 133 |  | 0 | [60 empty] | 0
P | 2018-08-04T10:14:09+08:00 | 2018-08-04T12:06:44+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | G18USM068G | CPSM Final Test | FAIL | [13 empty] | -5.3
P | 2018-08-04T08:42:34+08:00 | 2018-08-04T13:22:26+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | G18USM065G | CPSM Final Test | PASS |  | 160.07 | -0.02 | 0 | 56.46 | 49.19 | 1.21 | 0.13 | 1 | 2 | 0 | -0.858 | 0.517 | -0.2 | 45 |  |  | 0.3619 | 0.2252 | -0.0265 | 0.99999 |  |  |  |  | 133 |  | 0 | [61 empty] | 0
P | 2018-08-04T11:03:30+08:00 | 2018-08-04T13:51:05+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | G18USM061G | CPSM Final Test | PASS |  | 166.91 | 0.27 | -0.84 | 56.62 | 48.6 | 1.51 | 0.13 | 1 | 3 | -1 | -0.8859 | 0.5112 | -5 | -37 |  |  | 0.0991 | 0.0817 | -0.0031 | 1 |  |  |  |  | 133 |  | 0 | [62 empty] | 0
P | 2018-08-04T07:47:01+08:00 | 2018-08-04T14:33:52+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | G18USM068G | CPSM Final Test | FAIL | [71 empty] | 0
P | 2018-08-04T11:10:27+08:00 | 2018-08-04T15:39:46+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | G18USM068G | CPSM Final Test | PASS |  | 164.57 | -0.02 | -0.92 | 56.61 | 50.29 | 1.51 | 0.13 | 0 | -1 | 1 | -0.7323 | 0.528 | 5.9 | -3 |  |  | 0.1819 | 0.1902 | 0.0805 | 0.99995 |  |  |  |  | 133 |  | 0 | [43 empty] | 0
P | 2018-08-04T11:42:06+08:00 | 2018-08-04T15:41:17+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI143G | CPSM Final Test | FAIL | [78 empty] | 1380
P | 2018-08-04T10:56:12+08:00 | 2018-08-04T15:42:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI144G | CPSM Final Test | FAIL | [16 empty] | 0.25818
P | 2018-08-04T12:58:02+08:00 | 2018-08-04T15:46:58+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | G18USM067G | CPSM Final Test | PASS |  | 169.61 | 0.27 | -1.07 | 56.47 | 50.04 | 1.13 | 0.21 | 0 | -3 | 0 | -0.85 | 0.5032 | 2.1 | 2 |  |  | 0.1011 | 0.1177 | -0.0677 | 0.99999 |  |  |  |  | 133 |  | 0 | [63 empty] | 0
P | 2018-08-04T13:20:13+08:00 | 2018-08-04T16:07:53+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI146G | CPSM Final Test | PASS | 129 | 165.02 | 0.17 | -0.61 | 56.51 | 51.22 | 1.44 | 0.13 | 1 | 3 | 1 | -0.8401 | 0.499 | 6.7 | 2 |  | 0.02945 | 0.2466 | 0.5964 | 0.0525 | 1 | [71 empty] | 0
P | 2018-08-04T13:41:41+08:00 | 2018-08-04T16:47:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI143G | CPSM Final Test | PASS | 129 | 160.07 | 0.17 | -1.07 | 56.34 | 49.61 | 1.29 | 0.13 | 0 | -1 | 0 | -0.8959 | 0.506 | -1.3 | -5 |  | 0.12302 | 0.062 | 0.3063 | 0.0076 | 1 | [57 empty] | 0
P | 2018-08-06T07:51:40+08:00 | 2018-08-06T11:24:26+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI148G | CPSM Final Test | FAIL | [13 empty] | -7.3
P | 2018-08-06T08:25:52+08:00 | 2018-08-06T11:25:46+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI147G | CPSM Final Test | FAIL | [13 empty] | 18
P | 2018-08-06T09:11:34+08:00 | 2018-08-06T12:07:10+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI144G | CPSM Final Test | PASS | 129 | 163.4 | 0.17 | -1.07 | 56.54 | 51.31 | 1.36 | 0.13 | 0 | -3 | 1 | -0.9697 | 0.489 | 1.3 | -13 |  | 0.01747 | 0.3052 | 0.3746 | -0.0141 | 0.99999 | [72 empty] | 0
P | 2018-08-06T08:18:57+08:00 | 2018-08-06T12:20:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI150G | CPSM Final Test | FAIL | [13 empty] | -15.8
P | 2018-08-06T08:14:19+08:00 | 2018-08-06T12:33:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI145G | CPSM Final Test | FAIL | [13 empty] | 16.7
P | 2018-08-06T06:31:04+08:00 | 2018-08-06T13:20:34+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI147G | CPSM Final Test | FAIL | [94 empty] | 0
P | 2018-08-06T09:38:03+08:00 | 2018-08-06T13:39:36+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI148G | CPSM Final Test | PASS | 129 | 156.74 | 0.08 | -0.54 | 56.78 | 49.11 | 1.29 | 0.13 | -1 | 1 | 0 | -0.9419 | 0.464 | -2.2 | 7 |  | 0.015 | 0.2342 | 0.3053 | -0.033 | 1 | [31 empty] | 0
P | 2018-08-06T06:36:51+08:00 | 2018-08-06T14:02:29+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI150G | CPSM Final Test | FAIL | [86 empty] | 5659
P | 2018-08-06T06:50:45+08:00 | 2018-08-06T13:57:12+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI145G | CPSM Final Test | FAIL | [78 empty] | 1918
P | 2018-08-06T10:35:51+08:00 | 2018-08-06T14:26:05+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI152G | CPSM Final Test | PASS | 129 | 162.59 | 0.68 | -0.92 | 56.41 | 52.66 | 1.36 | 0.13 | 1 | -2 | 0 | -0.8681 | 0.549 | -3.8 | 24 |  | 0.00315 | 0.146 | 0.3284 | 0.0683 | 0.99991 | [50 empty] | 0
P | 2018-08-06T09:55:49+08:00 | 2018-08-06T14:26:06+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI147G | CPSM Final Test | PASS | 129 | 165.83 | 0.49 | -0.84 | 56.34 | 48.94 | 1.06 | 0.13 | -1 | 1 | 0 | -0.862 | 0.4976 | 10.3 | 5 |  | 0 | 0.1248 | 0.0745 | -0.0767 | 0.99999 | [73 empty] | 0
P | 2018-08-06T10:19:40+08:00 | 2018-08-06T15:08:03+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI150G | CPSM Final Test | PASS | 129 | 159.8 | 0.08 | -0.31 | 56.6 | 49.78 | 1.59 | 0.13 | -1 | -3 | 0 | -0.8759 | 0.5056 | 3 | 26 |  | 0.02756 | 0.1096 | 0.2115 | 0.007 | 0.99998 | [65 empty] | 0
P | 2018-08-06T10:23:58+08:00 | 2018-08-06T15:02:54+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI145G | CPSM Final Test | PASS | 129 | 157.55 | -0.02 | -0.84 | 56.72 | 54.69 | 1.29 | 0.13 | 0 | -2 | 0 | -0.9278 | 0.5098 | 8.2 | 60 |  | 0 | 0.0575 | 0.265 | -0.006 | 0.99999 | [57 empty] | 0
P | 2018-08-06T15:02:08+08:00 | 2018-08-06T18:51:00+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI157G | CPSM Final Test | FAIL | [13 empty] | -8.3
P | 2018-08-06T03:16:02+08:00 | 2018-08-06T19:33:32+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI151G | CPSM Final Test | FAIL
P | 2018-08-06T13:38:41+08:00 | 2018-08-06T19:56:57+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI158G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-08-06T17:30:29+08:00 | 2018-08-06T20:13:29+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI142G | CPSM Final Test | FAIL | [13 empty] | 21.7
P | 2018-08-06T17:10:29+08:00 | 2018-08-06T20:28:55+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI156G | CPSM Final Test | PASS | 129 | 161.69 | 0.08 | -1.22 | 56.34 | 48.18 | 0.83 | 0.13 | -1 | -2 | 0 | -0.9816 | 0.511 | 0 | -24 |  | 0.03197 | 0.2015 | 0.3208 | 0.0362 | 0.99999 | [65 empty] | 0
P | 2018-08-06T11:24:47+08:00 | 2018-08-06T20:37:34+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI151G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.54 | 56.43 | 53.42 | 1.51 | 0.13 | 1 | 1 | 0 | -0.9338 | 0.5032 | 1.9 | -2 |  | 0 | 0.1276 | 0.1998 | -0.0396 | 1 | [74 empty] | 0
P | 2018-08-06T16:56:34+08:00 | 2018-08-06T20:51:06+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI157G | CPSM Final Test | PASS | 129 | 162.95 | 0.08 | -0.84 | 56.71 | 55.8 | 1.13 | 0.13 | 1 | -3 | 0 | -0.9955 | 0.5194 | 0.3 | 12 |  | 0.01205 | 0.0657 | 0.1206 | -0.0364 | 1 | [75 empty] | 0
P | 2018-08-06T17:20:40+08:00 | 2018-08-06T21:01:50+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI159G | CPSM Final Test | PASS | 129 | 161.24 | 0.27 | -1.22 | 56.49 | 48.77 | 0.83 | 0.13 | 1 | 4 | 1 | -0.9079 | 0.5292 | 5 | 15 |  | 0 | 0.0543 | 0.2157 | -0.004 | 0.99999 | [73 empty] | 0
P | 2018-08-06T16:47:49+08:00 | 2018-08-06T21:57:29+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI158G | CPSM Final Test | PASS | 129 | 160.25 | 0.46 | -1.38 | 56.47 | 51.99 | 0.98 | 0.13 | -1 | -2 | 0 | -0.7942 | 0.5462 | -3.5 | -1 |  | 0 | 0.0733 | 0.3049 | 0.0613 | 0.99988 | [76 empty] | 0
P | 2018-08-06T13:48:15+08:00 | 2018-08-06T20:37:08+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI155G | CPSM Final Test | FAIL
P | 2018-08-06T17:12:41+08:00 | 2018-08-07T08:33:40+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI155G | CPSM Final Test | PASS | 129 | 157.46 | 0.37 | -0.84 | 56.79 | 50.29 | 0.91 | 0.13 | 1 | 10 | 0 | -1.0017 | 0.4968 | 3.5 | 10 |  | 0.01209 | 0.05 | 0.3188 | -0.0384 | 0.99999 | [57 empty] | 0
P | 2018-08-06T02:21:43+08:00 | 2018-08-07T08:34:52+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI149G | CPSM Final Test | FAIL | [98 empty] | 67588
P | 2018-08-06T05:07:43+08:00 | 2018-08-07T08:36:15+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI142G | CPSM Final Test | FAIL | [13 empty] | 28.7
P | 2018-08-06T01:52:59+08:00 | 2018-08-07T09:03:36+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI149G | CPSM Final Test | FAIL | [98 empty] | 7502
P | 2018-08-06T18:51:59+08:00 | 2018-08-07T10:36:45+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI142G | CPSM Final Test | PASS | 129 | 167.27 | -0.21 | -1.07 | 56.53 | 51.14 | 1.29 | 0.13 | 1 | -2 | -1 | -0.9578 | 0.5114 | -3.1 | -3 |  | 0.00606 | 0.0438 | 0.2231 | 0.0114 | 0.99999 | [78 empty] | 0
P | 2018-08-06T00:01:35+08:00 | 2018-08-07T10:55:00+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI149G | CPSM Final Test | FAIL | [16 empty] | 0.18481
P | 2018-08-05T22:55:05+08:00 | 2018-08-07T12:01:30+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI149G | CPSM Final Test | FAIL
P | 2018-08-06T17:28:18+08:00 | 2018-08-07T12:02:00+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI149G | CPSM Final Test | PASS | 129 | 159.62 | -0.02 | -0.77 | 56.47 | 49.7 | 1.51 | 0.13 | -1 | -4 | 0 | -0.88 | 0.5014 | 12.3 | -17 |  | 0.0538 | 0.0835 | 0.2409 | -0.0087 | 0.99999 | [77 empty] | 0
P | 2018-08-07T12:03:01+08:00 | 2018-08-07T12:18:14+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI166G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-08-07T13:16:41+08:00 | 2018-08-07T14:06:12+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI163G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-08-07T09:44:08+08:00 | 2018-08-07T14:19:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI167G | CPSM Final Test | FAIL | [13 empty] | -10
P | 2018-08-07T10:12:26+08:00 | 2018-08-07T14:27:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI168G | CPSM Final Test | FAIL | [13 empty] | -5.2
P | 2018-08-07T10:32:05+08:00 | 2018-08-07T14:29:23+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI164G | CPSM Final Test | FAIL | [13 empty] | -8.8
P | 2018-08-07T08:48:26+08:00 | 2018-08-07T14:46:39+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI169G | CPSM Final Test | FAIL
P | 2018-08-07T08:54:26+08:00 | 2018-08-07T15:11:43+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI165G | CPSM Final Test | FAIL
P | 2018-08-07T08:00:22+08:00 | 2018-08-07T15:34:43+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI169G | CPSM Final Test | FAIL
P | 2018-08-07T08:15:18+08:00 | 2018-08-07T15:48:31+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI167G | CPSM Final Test | FAIL | [100 empty] | 1368
P | 2018-08-07T12:10:37+08:00 | 2018-08-07T15:59:16+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI166G | CPSM Final Test | PASS | 129 | 160.88 | 0.37 | -1.07 | 56.61 | 48.18 | 1.21 | 0.13 | 1 | 3 | 0 | -0.856 | 0.5482 | 7.1 | 9 |  | 0.02103 | 0.0852 | 0.2314 | -0.0168 | 1 | [80 empty] | 0
P | 2018-08-07T09:01:51+08:00 | 2018-08-07T15:59:37+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI164G | CPSM Final Test | FAIL | [102 empty] | 1855
P | 2018-08-07T12:03:05+08:00 | 2018-08-07T16:17:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI165G | CPSM Final Test | PASS | 129 | 162.05 | 0.37 | -0.99 | 56.48 | 52.15 | 1.29 | 0.13 | 0 | 1 | 0 | -0.846 | 0.5078 | -0.7 | 2 |  | 0.01424 | 0.174 | 0.1148 | 0.0121 | 0.99999 | [82 empty] | 0
P | 2018-08-07T11:04:13+08:00 | 2018-08-07T16:18:40+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI163G | CPSM Final Test | FAIL | [98 empty] | 2283
P | 2018-08-07T12:19:47+08:00 | 2018-08-07T16:27:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI168G | CPSM Final Test | PASS | 129 | 157.28 | -0.02 | -0.77 | 56.49 | 49.36 | 1.44 | 0.13 | -1 | 2 | 0 | -0.818 | 0.493 | -2.1 | -18 |  | 0 | 0.0732 | 0.1311 | 0.0123 | 1 | [83 empty] | 0
P | 2018-08-07T12:01:54+08:00 | 2018-08-07T16:54:03+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI167G | CPSM Final Test | PASS | 129 | 166.64 | -0.11 | -0.46 | 56.63 | 48.6 | 1.59 | 0.13 | 0 | -4 | 0 | -0.9098 | 0.502 | 6.5 | 15 |  | 0 | 0.1122 | 0.0535 | 0.0545 | 0.99999 | [79 empty] | 0
P | 2018-08-07T12:30:44+08:00 | 2018-08-07T17:05:13+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI164G | CPSM Final Test | PASS | 129 | 163.76 | -0.21 | -1.3 | 56.41 | 56.39 | 1.51 | 0.13 | -1 | 2 | 1 | -0.8599 | 0.495 | 0.8 | 12 |  | 0.01572 | 0.1423 | 0.3129 | 0.0035 | 0.99999 | [81 empty] | 0
P | 2018-08-07T13:41:26+08:00 | 2018-08-07T17:24:28+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI163G | CPSM Final Test | PASS | 129 | 159.44 | -0.11 | -0.77 | 56.71 | 48.77 | 1.36 | 0.13 | 1 | 1 | -1 | -0.9119 | 0.505 | -3.8 | 41 |  | 0.01072 | 0.1319 | 0.3042 | -0.0148 | 0.99999 | [77 empty] | 0
P | 2018-08-07T04:33:02+08:00 | 2018-08-07T19:02:03+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI169G | CPSM Final Test | FAIL
P | 2018-08-07T03:48:03+08:00 | 2018-08-07T19:29:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI161G | CPSM Final Test | FAIL | [14 empty] | 871
P | 2018-08-07T11:47:33+08:00 | 2018-08-07T20:47:59+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI169G | CPSM Final Test | PASS | 129 | 156.56 | 1.06 | -0.99 | 56.54 | 48.6 | 0.6 | 0.13 | 0 | 2 | 0 | -0.85 | 0.5332 | 2.8 | -8 |  | 0.01508 | 0.0351 | 0.0713 | 0.0022 | 1 | [84 empty] | 0
P | 2018-08-07T11:38:31+08:00 | 2018-08-07T21:13:36+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI161G | CPSM Final Test | PASS | 129 | 158.9 | 0.08 | -0.99 | 56.63 | 49.19 | 1.29 | 0.13 | -2 | -2 | 0 | -0.834 | 0.496 | -0.9 | 80 |  | 0 | 0.1144 | 0.1588 | 0.0756 | 0.99984 | [85 empty] | 0
P | 2018-08-07T03:48:14+08:00 | 2018-08-08T10:51:31+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI154G | CPSM Final Test | FAIL | [13 empty] | -10.1
P | 2018-08-08T07:59:14+08:00 | 2018-08-08T12:09:46+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI179G | CPSM Final Test | FAIL | [14 empty] | 1427
P | 2018-08-07T20:05:30+08:00 | 2018-08-08T12:28:51+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI172G | CPSM Final Test | PASS | 129 | 162.14 | 0.17 | 0.45 | 56.8 | 51.65 | 1.13 | 0.13 | 1 | 2 | 1 | -0.9817 | 0.541 | -1.2 | -6 |  | 0 | 0.1022 | 0.1085 | 0.0073 | 0.99999 | [79 empty] | 0
P | 2018-08-07T19:42:53+08:00 | 2018-08-08T12:27:44+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI174G | CPSM Final Test | PASS | 129 | 153.23 | 0.17 | -0.54 | 56.49 | 49.36 | 0.91 | 0.13 | 1 | -2 | 0 | -1.0116 | 0.5238 | 0.3 | 11 |  | 0 | 0.1004 | 0.1782 | -0.0104 | 1 | [80 empty] | 0
P | 2018-08-08T06:13:18+08:00 | 2018-08-08T12:37:30+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI153G | CPSM Final Test | FAIL | [98 empty] | 0
P | 2018-08-08T06:42:08+08:00 | 2018-08-08T13:08:21+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI171G | CPSM Final Test | FAIL
P | 2018-08-08T09:55:15+08:00 | 2018-08-08T13:08:49+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI171G | CPSM Final Test | PASS | 129 | 160.34 | 0.08 | -0.38 | 56.7 | 52.32 | 1.51 | 0.13 | 0 | 1 | 0 | -0.9238 | 0.5198 | -1.3 | -47 |  | 0.01828 | 0.1155 | 0.1611 | -0.0386 | 1 | [86 empty] | 0
P | 2018-08-08T09:25:24+08:00 | 2018-08-08T13:43:06+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI153G | CPSM Final Test | PASS | 129 | 159.35 | 0.08 | -0.69 | 56.98 | 51.82 | 1.59 | 0.13 | 1 | -3 | 0 | -0.9138 | 0.5296 | 7.9 | 9 |  | 0 | 0.0611 | 0.2153 | -0.0454 | 0.99999 | [77 empty] | 0
P | 2018-08-08T06:07:34+08:00 | 2018-08-08T14:01:26+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI179G | CPSM Final Test | FAIL
P | 2018-08-08T10:04:30+08:00 | 2018-08-08T14:08:08+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI179G | CPSM Final Test | PASS | 129 | 156.83 | 0.27 | -0.61 | 56.63 | 49.28 | 1.36 | 0.13 | -1 | -4 | 0 | -0.848 | 0.521 | -4.3 | -2 |  | 0.00649 | 0.1381 | 0.1455 | 0.0556 | 0.99992 | [87 empty] | 0
P | 2018-08-07T01:29:56+08:00 | 2018-08-08T14:10:47+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI177G | CPSM Final Test | FAIL | [103 empty] | 0
P | 2018-08-06T22:29:56+08:00 | 2018-08-08T16:09:49+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI154G | CPSM Final Test | FAIL
P | 2018-08-06T22:34:12+08:00 | 2018-08-08T16:25:37+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI173G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-08-08T13:08:03+08:00 | 2018-08-08T16:35:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI184G | CPSM Final Test | FAIL
P | 2018-08-08T12:21:00+08:00 | 2018-08-08T16:39:08+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI186G | CPSM Final Test | FAIL | [13 empty] | 16.8
P | 2018-08-08T12:20:19+08:00 | 2018-08-08T16:53:49+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI185G | CPSM Final Test | FAIL | [14 empty] | -126
P | 2018-08-06T22:29:16+08:00 | 2018-08-08T17:11:27+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI177G | CPSM Final Test | FAIL | [16 empty] | 0
P | 2018-08-06T22:27:36+08:00 | 2018-08-08T17:13:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI177G | CPSM Final Test | FAIL | [16 empty] | 0
P | 2018-08-08T13:07:03+08:00 | 2018-08-08T17:21:24+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI180G | CPSM Final Test | FAIL | [98 empty] | 3485
P | 2018-08-07T19:19:53+08:00 | 2018-08-08T17:54:25+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI154G | CPSM Final Test | PASS | 129 | 163.49 | 0.17 | -0.54 | 56.52 | 50.46 | 1.51 | 0.13 | 0 | -2 | 1 | -0.9478 | 0.5056 | -3.2 | -6 |  | 0.01619 | 0.0688 | 0.4545 | 0.0205 | 1 | [81 empty] | 0
P | 2018-08-08T14:49:39+08:00 | 2018-08-08T18:05:29+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI178G | CPSM Final Test | PASS | 129 | 163.49 | -0.11 | -0.99 | 56.49 | 49.45 | 1.21 | 0.13 | 1 | -4 | 0 | -0.9277 | 0.5346 | -0.1 | -9 |  | 0 | 0.0749 | 0.1707 | 0.0395 | 0.99999 | [88 empty] | 0
P | 2018-08-08T14:51:57+08:00 | 2018-08-08T18:20:34+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI184G | CPSM Final Test | PASS | 129 | 166.19 | 0.49 | -1.07 | 56.47 | 48.01 | 0.6 | 0.13 | 1 | 2 | 0 | -0.9178 | 0.502 | -3.3 | 12 |  | 0 | 0.116 | 0.1614 | -0.026 | 1 | [89 empty] | 0
P | 2018-08-08T15:14:14+08:00 | 2018-08-08T18:27:04+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI180G | CPSM Final Test | PASS | 129 | 155.39 | 0.46 | -0.84 | 56.47 | 53 | 0.75 | 0.13 | 0 | 1 | 0 | -0.9438 | 0.544 | -0.8 | -4 |  | 0.00731 | 0.0904 | 0.1843 | -0.0217 | 0.99997 | [77 empty] | 0
P | 2018-08-08T14:37:04+08:00 | 2018-08-08T18:38:45+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI185G | CPSM Final Test | PASS | 129 | 157.46 | 0.27 | -0.61 | 56.62 | 52.07 | 1.21 | 0.13 | 0 | 2 | 0 | -0.9438 | 0.5026 | -1.4 | -11 |  | 0.01104 | 0.0428 | 0.0419 | -0.019 | 1 | [90 empty] | 0
P | 2018-08-08T14:30:04+08:00 | 2018-08-08T18:43:18+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI186G | CPSM Final Test | PASS | 129 | 160.25 | 0.46 | -1.3 | 56.61 | 50.21 | 1.44 | 0.13 | 0 | 1 | 0 | -0.8819 | 0.535 | 10.3 | 34 |  | 0 | 0.2152 | 0.351 | 0.0803 | 0.99983 | [91 empty] | 0
P | 2018-08-07T19:29:55+08:00 | 2018-08-08T19:17:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI173G | CPSM Final Test | PASS | 129 | 172.4 | 0.08 | -0.69 | 56.63 | 49.02 | 0.91 | 0.13 | -1 | 1 | 0 | -0.9776 | 0.4922 | 2 | -25 |  | 0.02582 | 0.0651 | 0.293 | -0.039 | 1 | [92 empty] | 0
P | 2018-08-06T20:22:38+08:00 | 2018-08-08T19:18:05+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI177G | CPSM Final Test | FAIL | [17 empty] | 0
P | 2018-08-08T17:48:37+08:00 | 2018-08-08T19:26:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI183G | CPSM Final Test | FAIL | [5 empty] | 45.72
P | 2018-08-08T19:13:36+08:00 | 2018-08-08T20:05:33+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI191G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-08-08T17:43:10+08:00 | 2018-08-08T20:10:41+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI190G | CPSM Final Test | FAIL | [13 empty] | -14.5
P | 2018-08-07T19:50:22+08:00 | 2018-08-08T20:21:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI177G | CPSM Final Test | PASS | 129 | 162.95 | 0.68 | -0.08 | 56.61 | 50.8 | 1.51 | 0.13 | 1 | 3 | 0 | -0.914 | 0.4736 | -4.4 | 8 |  | 0.04277 | 0.077 | 0.1946 | -0.004 | 1 | [82 empty] | 0
P | 2018-08-08T16:31:45+08:00 | 2018-08-08T20:43:44+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI183G | CPSM Final Test | FAIL | [14 empty] | 1774
P | 2018-08-08T18:34:58+08:00 | 2018-08-08T21:30:51+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI182G | CPSM Final Test | PASS | 129 | 156.92 | 0.46 | -0.92 | 56.7 | 53.76 | 1.06 | 0.13 | -1 | -3 | 0 | -0.9777 | 0.519 | 2.6 | -42 |  | 0 | 0.1342 | 0.1515 | 0.0248 | 0.99997 | [94 empty] | 0
P | 2018-08-08T19:05:19+08:00 | 2018-08-08T21:53:52+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI181G | CPSM Final Test | PASS | 129 | 154.94 | 0.97 | -0.54 | 56.77 | 51.39 | 0.91 | 0.13 | 0 | 1 | 0 | -0.8679 | 0.5268 | 0.7 | -10 |  | 0.00843 | 0.1206 | 0.5043 | 0.0035 | 0.99999 | [77 empty] | 0
P | 2018-08-08T18:56:56+08:00 | 2018-08-08T22:10:57+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI190G | CPSM Final Test | PASS | 129 | 159.53 | 0.56 | -0.31 | 56.73 | 53.85 | 1.13 | 0.13 | 1 | 3 | 0 | -0.8859 | 0.5072 | -1.5 | -4 |  | 0.11422 | 0.2787 | 0.407 | 0.0334 | 0.99998 | [95 empty] | 0
P | 2018-08-07T04:51:33+08:00 | 2018-08-10T08:23:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI183G | CPSM Final Test | FAIL
P | 2018-08-08T18:37:44+08:00 | 2018-08-10T08:24:22+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI183G | CPSM Final Test | PASS | 129 | 163.49 | 0.56 | -0.23 | 57.04 | 51.73 | 1.44 | 0.13 | 1 | -2 | -1 | -0.9758 | 0.53 | 3 | -4 |  | 0 | 0.1814 | 0.3725 | 0.0125 | 0.99999 | [96 empty] | 0
P | 2018-08-07T06:00:08+08:00 | 2018-08-10T09:52:30+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI187G | CPSM Final Test | FAIL
P | 2018-08-08T19:56:19+08:00 | 2018-08-10T11:37:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI187G | CPSM Final Test | PASS | 129 | 168.89 | 0.17 | -0.46 | 56.8 | 49.28 | 0.91 | 0.13 | 0 | -1 | 0 | -0.8041 | 0.51 | -2.9 | -2 |  | 0.00136 | 0.1609 | 0.191 | 0.0674 | 0.9998 | [97 empty] | 0
P | 2018-08-08T20:08:16+08:00 | 2018-08-10T11:38:59+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI191G | CPSM Final Test | PASS | 129 | 159.71 | 0.46 | -0.84 | 56.79 | 49.36 | 0.91 | 0.13 | -1 | -4 | 1 | -0.836 | 0.4994 | -2.6 | -20 |  | 0.01702 | 0.0477 | 0.0921 | 0.016 | 0.99998 | [98 empty] | 0
P | 2018-08-10T11:55:32+08:00 | 2018-08-10T14:44:59+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI162G | CPSM Final Test | FAIL | [5 empty] | 46.23
P | 2018-08-10T12:53:59+08:00 | 2018-08-10T15:47:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI189G | CPSM Final Test | FAIL | [13 empty] | -10.9
P | 2018-08-10T11:12:25+08:00 | 2018-08-10T15:55:31+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI170G | CPSM Final Test | FAIL | [13 empty] | -6.4
P | 2018-08-10T13:09:10+08:00 | 2018-08-10T15:52:21+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI175G | CPSM Final Test | FAIL | [5 empty] | 46.57
P | 2018-08-10T13:55:25+08:00 | 2018-08-10T16:06:02+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI188G | CPSM Final Test | FAIL | [5 empty] | 46.06
P | 2018-08-10T10:54:03+08:00 | 2018-08-10T16:13:53+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI170G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-08-10T13:56:29+08:00 | 2018-08-10T16:36:37+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI176G | CPSM Final Test | FAIL | [87 empty] | FALSE
P | 2018-08-10T10:04:01+08:00 | 2018-08-10T17:03:55+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI170G | CPSM Final Test | FAIL
P | 2018-08-10T14:13:03+08:00 | 2018-08-10T17:27:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI160G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.99 | 56.73 | 54.44 | 1.59 | 0.13 | 1 | 3 | 0 | -0.9537 | 0.5044 | 3.7 | 21 |  | 0.01712 | 0.0649 | 0.1926 | -0.0216 | 0.99999 | [99 empty] | 0
P | 2018-08-10T13:20:16+08:00 | 2018-08-10T17:33:15+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI162G | CPSM Final Test | PASS | 129 | 162.68 | 0.17 | -1.15 | 56.92 | 48.34 | 1.06 | 0.13 | 0 | 1 | 0 | -0.8659 | 0.4624 | 1.1 | 136 |  | 0.01645 | 0.0527 | 0.0973 | -0.0429 | 1 | [100 empty] | 0
P | 2018-08-10T12:47:07+08:00 | 2018-08-10T17:45:59+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI176G | CPSM Final Test | FAIL | [13 empty] | -7.7
P | 2018-08-10T14:20:57+08:00 | 2018-08-10T17:48:42+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI189G | CPSM Final Test | PASS | 129 | 163.31 | 0.27 | -0.61 | 56.47 | 48.01 | 0.98 | 0.13 | 0 | -1 | 1 | -0.9219 | 0.5148 | 3.9 | -21 |  | 0 | 0.0358 | 0.194 | -0.0414 | 1 | [101 empty] | 0
P | 2018-08-10T10:46:06+08:00 | 2018-08-10T18:15:25+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI175G | CPSM Final Test | FAIL
P | 2018-08-10T13:33:58+08:00 | 2018-08-10T18:48:27+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI170G | CPSM Final Test | PASS | 129 | 162.23 | 0.37 | -0.38 | 56.72 | 51.99 | 1.13 | 0.13 | 1 | -2 | 0 | -0.9756 | 0.4912 | -2.3 | -11 |  | 0.01078 | 0.2256 | 0.1162 | 0.034 | 0.99994 | [102 empty] | 0
P | 2018-08-10T15:00:43+08:00 | 2018-08-10T18:47:42+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI188G | CPSM Final Test | PASS | 129 | 155.93 | 0.37 | -0.61 | 56.56 | 49.45 | 0.91 | 0.13 | -2 | -1 | 0 | -0.9139 | 0.471 | 0.8 | -19 |  | 0.02482 | 0.1066 | 0.1225 | -0.0138 | 1 | [103 empty] | 0
P | 2018-08-10T09:24:06+08:00 | 2018-08-10T19:37:25+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI175G | CPSM Final Test | FAIL | [125 empty] | 1454
P | 2018-08-10T10:43:21+08:00 | 2018-08-10T19:49:45+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI176G | CPSM Final Test | FAIL | [126 empty] | 0
P | 2018-08-10T14:30:46+08:00 | 2018-08-10T20:42:59+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI175G | CPSM Final Test | PASS | 129 | 164.93 | 0.08 | -0.46 | 56.54 | 48.43 | 0.91 | 0.13 | -2 | -3 | -1 | -0.854 | 0.4956 | -0.3 | 67 |  | 0 | 0.1568 | 0.2373 | -0.0157 | 1 | [104 empty] | 0
P | 2018-08-10T15:16:33+08:00 | 2018-08-10T20:55:23+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI176G | CPSM Final Test | PASS | 129 | 157.46 | 0.17 | -0.99 | 56.44 | 50.29 | 1.06 | 0.16 | 0 | -2 | 0 | -0.8939 | 0.5098 | -0.7 | 42 |  | 0.00845 | 0.1541 | 0.2699 | 0.015 | 0.99999 | [105 empty] | 0
P | 2018-08-11T13:19:46+08:00 | 2018-08-11T15:32:45+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP580G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-08-11T12:42:18+08:00 | 2018-08-11T16:10:13+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP580G | CPSM Final Test | FAIL
P | 2018-08-11T11:55:02+08:00 | 2018-08-11T16:59:30+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | H18FNP581G | CPSM Final Test | FAIL | [127 empty] | 2161
P | 2018-08-11T11:48:49+08:00 | 2018-08-11T17:10:30+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP578G | CPSM Final Test | FAIL | [128 empty] | 0
P | 2018-08-11T11:45:42+08:00 | 2018-08-11T17:12:19+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP579G | CPSM Final Test | FAIL
P | 2018-08-11T14:30:33+08:00 | 2018-08-11T17:41:51+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP585G | CPSM Final Test | PASS |  | 164.3 | 0.17 | -0.54 | 56.61 | 50.97 | 0.98 | 0.13 | -1 | 3 | 0 | -0.9159 | 0.4944 | 2.1 | 5 |  |  | 0.0662 | 0.1284 | 0.0622 | 0.99997 |  |  |  |  | 133 |  | 0.01157 | [101 empty] | 0
P | 2018-08-11T14:32:02+08:00 | 2018-08-11T17:38:37+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | H18FNP574G | CPSM Final Test | PASS |  | 157.55 | -0.02 | -0.23 | 56.67 | 50.38 | 1.29 | 0.13 | 0 | -4 | -1 | -0.9638 | 0.533 | -0.5 | 24 |  |  | 0.112 | 0.14 | -0.0459 | 0.99999 |  |  |  |  | 133 |  | 0.01883 | [102 empty] | 0
P | 2018-08-11T14:41:33+08:00 | 2018-08-11T17:44:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP582G | CPSM Final Test | PASS |  | 162.41 | 0.17 | -0.61 | 56.3 | 50.12 | 0.83 | 0.13 | 1 | -4 | -1 | -0.9476 | 0.464 | 4.3 | 13 |  |  | 0.0537 | 0.0912 | -0.0418 | 0.99999 |  |  |  |  | 133 |  | 0 | [97 empty] | 0
P | 2018-08-11T14:27:16+08:00 | 2018-08-11T18:05:02+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | H18FNP581G | CPSM Final Test | PASS |  | 162.95 | -0.02 | -0.54 | 56.79 | 53 | 1.21 | 0.13 | -1 | 1 | 0 | -0.894 | 0.5062 | 2.2 | -15 |  |  | 0.06 | 0.208 | -0.0456 | 0.99999 |  |  |  |  | 133 |  | 0 | [99 empty] | 0
P | 2018-08-11T14:26:15+08:00 | 2018-08-11T18:14:31+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP580G | CPSM Final Test | PASS |  | 164.3 | 0.37 | -0.38 | 56.32 | 49.7 | 1.06 | 0.13 | -2 | -3 | -1 | -0.9378 | 0.4802 | 14.3 | 11 |  |  | 0.1346 | 0.3243 | -0.0417 | 0.99999 |  |  |  |  | 133 |  | 0.01992 | [103 empty] | 0
P | 2018-08-11T14:29:00+08:00 | 2018-08-11T18:15:25+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP579G | CPSM Final Test | PASS |  | 166.37 | 0.27 | -0.77 | 56.81 | 50.29 | 1.44 | 0.13 | -1 | -2 | 0 | -0.8759 | 0.5238 | 1.5 | 15 |  |  | 0.3175 | 0.4633 | -0.0123 | 0.99999 |  |  |  |  | 133 |  | 0 | [98 empty] | 0
P | 2018-08-11T14:29:39+08:00 | 2018-08-11T18:15:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP578G | CPSM Final Test | PASS |  | 164.21 | 0.27 | -0.61 | 56.95 | 50.63 | 0.83 | 0.13 | 0 | -1 | 0 | -0.9318 | 0.4896 | 7.9 | -2 |  |  | 0.0813 | 0.0954 | -0.0355 | 1 |  |  |  |  | 133 |  | 0 | [100 empty] | 0
P | 2018-08-11T15:59:11+08:00 | 2018-08-11T18:44:28+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | H18FNP573G | CPSM Final Test | PASS |  | 167 | -0.02 | -0.54 | 56.6 | 50.88 | 1.13 | 0.13 | 0 | 3 | 0 | -0.8879 | 0.53 | 1.3 | -5 |  |  | 0.2411 | 0.2095 | 0.1101 | 0.9998 |  |  |  |  | 133 |  | 0.00142 | [104 empty] | 0
P | 2018-08-13T08:17:27+08:00 | 2018-08-13T12:08:03+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP576G | CPSM Final Test | FAIL
P | 2018-08-13T10:16:29+08:00 | 2018-08-13T13:21:20+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | H18FNP583G | CPSM Final Test | PASS |  | 158.36 | 0.27 | -0.46 | 56.63 | 49.19 | 1.13 | 0.13 | -2 | -2 | 0 | -0.9339 | 0.521 | 8.8 | 14 |  |  | 0.0786 | 0.1953 | 0.004 | 1 |  |  |  |  | 133 |  | 0 | [99 empty] | 0
P | 2018-08-13T10:05:47+08:00 | 2018-08-13T13:28:26+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP575G | CPSM Final Test | PASS |  | 171.59 | 0.27 | -0.92 | 56.52 | 50.38 | 0.6 | 0.13 | -1 | -3 | 0 | -0.9478 | 0.5024 | 1.5 | 39 |  |  | 0.1495 | 0.0481 | 0.0132 | 1 |  |  |  |  | 133 |  | 0 | [100 empty] | 0
P | 2018-08-13T10:37:42+08:00 | 2018-08-13T13:37:40+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | H18FNP586G | CPSM Final Test | PASS |  | 159.26 | 0.27 | -0.92 | 56.53 | 50.21 | 1.13 | 0.13 | 0 | 2 | 1 | -1.0077 | 0.5472 | 8 | 45 |  |  | 0.1816 | 0.2403 | -0.0142 | 0.99999 |  |  |  |  | 133 |  | 0 | [105 empty] | 0
P | 2018-08-13T10:09:03+08:00 | 2018-08-13T13:41:55+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP584G | CPSM Final Test | PASS |  | 164.57 | -0.02 | -0.92 | 56.46 | 50.29 | 0.98 | 0.13 | -1 | -1 | 1 | -0.9318 | 0.516 | 11.3 | 14 |  |  | 0.0944 | 0.2502 | 0.0044 | 0.99996 |  |  |  |  | 133 |  | 0 | [106 empty] | 0
P | 2018-08-13T07:29:09+08:00 | 2018-08-13T12:56:21+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP576G | CPSM Final Test | FAIL | [91 empty] | 0
P | 2018-08-13T07:04:31+08:00 | 2018-08-13T14:04:42+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP847G | CPSM Final Test | FAIL
P | 2018-08-13T10:12:45+08:00 | 2018-08-13T14:55:19+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP576G | CPSM Final Test | PASS |  | 159.89 | 0.27 | -0.84 | 56.64 | 48.77 | 0.75 | 0.13 | 0 | 1 | 0 | -0.9099 | 0.532 | 2.5 | 3 |  |  | 0.1517 | 0.045 | 0.0228 | 0.99996 |  |  |  |  | 133 |  | 0 | [63 empty] | 0
P | 2018-08-13T14:14:44+08:00 | 2018-08-13T17:30:13+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP854G | CPSM Final Test | FAIL | [13 empty] | 16.1
P | 2018-08-13T15:05:21+08:00 | 2018-08-13T17:45:07+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP588G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-08-13T12:52:53+08:00 | 2018-08-13T18:43:29+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP852G | CPSM Final Test | FAIL
P | 2018-08-13T15:56:06+08:00 | 2018-08-13T19:00:26+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP851G | CPSM Final Test | PASS | 129 | 164.12 | 0.27 | -1.3 | 56.31 | 48.68 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9478 | 0.5048 | 3.5 | 22 |  |  | 0.0642 | 0.0668 | -0.0275 | 1 | [21 empty] | 0 | [92 empty] | 0
P | 2018-08-13T15:52:29+08:00 | 2018-08-13T19:30:39+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP854G | CPSM Final Test | PASS | 129 | 160.88 | 0.56 | -1.3 | 56.34 | 50.46 | 0.91 | 0.13 | -1 | 1 | 0 | -0.9518 | 0.4604 | 4.1 | 11 |  |  | 0.1115 | 0.2358 | 0.0303 | 0.99999 | [21 empty] | 0 | [93 empty] | 0
P | 2018-08-13T13:19:01+08:00 | 2018-08-13T19:31:27+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP588G | CPSM Final Test | FAIL
P | 2018-08-13T15:48:11+08:00 | 2018-08-13T20:28:01+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP852G | CPSM Final Test | PASS | 129 | 156.47 | 0.78 | -0.61 | 56.72 | 54.1 | 1.36 | 0.13 | -1 | -2 | 1 | -0.8999 | 0.4944 | 3.5 | -7 |  |  | 0.0204 | 0.0833 | 0.0216 | 0.99997 | [21 empty] | 0 | [48 empty] | 0
P | 2018-08-13T17:08:24+08:00 | 2018-08-13T20:31:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP855G | CPSM Final Test | PASS | 129 | 157.19 | -0.02 | -0.77 | 56.47 | 50.72 | 1.36 | 0.13 | 1 | -2 | -1 | -0.9638 | 0.518 | 8 | -17 |  |  | 0.0424 | 0.0881 | -0.0489 | 1 | [21 empty] | 0 | [94 empty] | 0
P | 2018-08-12T22:30:17+08:00 | 2018-08-14T10:20:11+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP588G | CPSM Final Test | FAIL
P | 2018-08-14T09:46:13+08:00 | 2018-08-14T12:00:29+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP856G | CPSM Final Test | FAIL | [13 empty] | 17.9
P | 2018-08-13T16:25:14+08:00 | 2018-08-14T12:08:41+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP588G | CPSM Final Test | PASS |  | 164.21 | 0.46 | -0.61 | 56.56 | 49.87 | 1.13 | 0.13 | 1 | -1 | 0 | -0.9697 | 0.4996 | 9.6 | 13 |  |  | 0.0909 | 0.2684 | -0.0212 | 1 |  |  |  |  | 133 |  | 0 |  | 0
P | 2018-08-14T10:21:25+08:00 | 2018-08-14T13:09:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | H18FNP587G | CPSM Final Test | PASS |  | 161.15 | 1.06 | -1.07 | 56.86 | 51.22 | 1.36 | 0.13 | 1 | -2 | 0 | -0.9219 | 0.5148 | 3.9 | 10 |  |  | 0.0706 | 0.1133 | -0.0158 | 1 |  |  |  |  | 133 |  | 0 | [110 empty] | 0
P | 2018-08-14T10:40:32+08:00 | 2018-08-14T13:27:57+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP858G | CPSM Final Test | PASS | 129 | 157.01 | 0.37 | -0.38 | 56.79 | 51.39 | 1.36 | 0.13 | -1 | 1 | 0 | -0.9139 | 0.5346 | 3 | 3 |  |  | 0.1518 | 0.0807 | -0.0003 | 0.99999 | [21 empty] | 0 | [96 empty] | 0
P | 2018-08-14T10:53:42+08:00 | 2018-08-14T13:29:59+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP857G | CPSM Final Test | PASS | 129 | 164.57 | -0.02 | -0.08 | 57.14 | 50.38 | 0.6 | 0.13 | -3 | 3 | 0 | -0.87 | 0.5174 | 7.6 | 2 |  |  | 0.0743 | 0.1156 | 0.018 | 0.99998 | [21 empty] | 0 | [97 empty] | 0
P | 2018-08-14T10:59:57+08:00 | 2018-08-14T13:37:12+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP860G | CPSM Final Test | PASS | 129 | 158.18 | 0.17 | -0.77 | 56.78 | 51.22 | 0.98 | 0.13 | 0 | -1 | 0 | -0.9198 | 0.5364 | 2.4 | 1 |  |  | 0.1001 | 0.2408 | 0.0427 | 0.99998 | [21 empty] | 0.01694 | [98 empty] | 0
P | 2018-08-14T10:55:01+08:00 | 2018-08-14T13:33:22+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP850G | CPSM Final Test | PASS | 129 | 162.59 | 0.37 | -0.99 | 56.53 | 51.22 | 1.29 | 0.13 | -2 | 2 | 0 | -0.9219 | 0.4916 | 4.3 | 17 |  |  | 0.1031 | 0.1062 | -0.0269 | 1 | [21 empty] | 0 | [99 empty] | 0
P | 2018-08-14T10:53:21+08:00 | 2018-08-14T14:00:51+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP856G | CPSM Final Test | PASS | 129 | 158.54 | 0.17 | -0.46 | 56.76 | 48.51 | 1.51 | 0.13 | 1 | -2 | 1 | -0.87 | 0.55 | 2.8 | 18 |  |  | 0.1519 | 0.3143 | -0.0051 | 0.99997 | [21 empty] | 0 | [100 empty] | 0
P | 2018-08-14T15:02:47+08:00 | 2018-08-14T17:23:46+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP577G | CPSM Final Test | FAIL | [5 empty] | 47.16
P | 2018-08-12T19:27:42+08:00 | 2018-08-14T18:02:41+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP847G | CPSM Final Test | FAIL | [144 empty] | 77897
P | 2018-08-14T15:37:52+08:00 | 2018-08-14T18:52:24+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP865G | CPSM Final Test | PASS | 129 | 153.77 | 0.27 | -0.84 | 56.53 | 50.55 | 1.51 | 0.13 | 1 | -2 | 0 | -0.9457 | 0.4946 | 10.1 | 37 |  |  | 0.2979 | 0.545 | 0.0383 | 0.99998 | [21 empty] | 0 | [102 empty] | 0
P | 2018-08-14T15:44:12+08:00 | 2018-08-14T18:52:47+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP853G | CPSM Final Test | PASS | 129 | 142.7 | -0.02 | -0.38 | 56.77 | 48.26 | 1.13 | 0.13 | 0 | -2 | 0 | -0.9597 | 0.5458 | 13.6 | 10 |  |  | 0.0527 | 0.4435 | 0.0171 | 0.99999 | [21 empty] | 0 | [21 empty] | 0
P | 2018-08-14T15:47:22+08:00 | 2018-08-14T18:55:21+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP859G | CPSM Final Test | PASS | 129 | 156.83 | -0.02 | -0.84 | 56.32 | 48.09 | 1.06 | 0.13 | 1 | -2 | 1 | -0.8799 | 0.5384 | 7.6 | 24 |  |  | 0.3969 | 0.6827 | -0.015 | 0.99999 | [21 empty] | 0 | [103 empty] | 0
P | 2018-08-14T16:06:34+08:00 | 2018-08-14T19:03:32+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP849G | CPSM Final Test | PASS | 129 | 166.37 | 0.27 | -0.54 | 56.78 | 51.48 | 1.51 | 0.13 | 0 | 1 | 0 | -0.9479 | 0.4984 | 0.7 | 15 |  |  | 0.1355 | 0.1708 | -0.0157 | 0.99999 | [21 empty] | 0 | [104 empty] | 0
P | 2018-08-14T11:38:59+08:00 | 2018-08-14T19:43:21+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | H18FTP862G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-08-14T11:51:30+08:00 | 2018-08-14T19:50:51+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP861G | CPSM Final Test | FAIL | [13 empty] | 16.7
P | 2018-08-14T11:11:49+08:00 | 2018-08-14T20:10:31+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | H18FTP862G | CPSM Final Test | FAIL | [5 empty] | 45.13
P | 2018-08-14T16:13:17+08:00 | 2018-08-14T20:43:08+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP577G | CPSM Final Test | PASS |  | 165.2 | -0.02 | -0.84 | 56.86 | 49.02 | 1.29 | 0.13 | -1 | 1 | -1 | -0.9557 | 0.4958 | 5.1 | 5 |  |  | 0.261 | 0.4016 | 0.0176 | 0.99998 |  |  |  |  | 133 |  | 0 | [120 empty] | 0
P | 2018-08-14T19:57:28+08:00 | 2018-08-14T22:35:57+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP869G | CPSM Final Test | PASS | 129 | 161.96 | -0.21 | -0.31 | 56.71 | 54.95 | 1.67 | 0.13 | 0 | -1 | 0 | -0.9238 | 0.494 | 9 | 11 |  |  | 0.105 | 0.3121 | -0.0092 | 1 | [21 empty] | 0 | [106 empty] | 0
P | 2018-08-14T19:55:20+08:00 | 2018-08-14T22:37:01+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP870G | CPSM Final Test | PASS | 129 | 160.07 | 0.17 | -0.31 | 56.71 | 53.76 | 1.13 | 0.13 | 0 | -2 | 0 | -0.9478 | 0.5044 | 8.2 | 69 |  |  | 0.0634 | 0.3677 | -0.0222 | 1 | [21 empty] | 0 | [107 empty] | 0
P | 2018-08-13T21:59:39+08:00 | 2018-08-15T09:22:41+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | H18FTP862G | CPSM Final Test | FAIL | [14 empty] | 1754
P | 2018-08-14T15:51:11+08:00 | 2018-08-15T09:48:35+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP861G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.46 | 56.53 | 49.78 | 1.29 | 0.13 | 1 | -3 | 1 | -0.9219 | 0.5144 | 1.8 | -9 |  |  | 0.1297 | 0.0288 | 0.0095 | 0.99999 | [21 empty] | 0.03937 | [108 empty] | 0
P | 2018-08-15T09:25:54+08:00 | 2018-08-15T10:39:56+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP868G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-08-15T08:58:24+08:00 | 2018-08-15T11:07:26+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP868G | CPSM Final Test | FAIL | [5 empty] | 46.91
P | 2018-08-13T19:59:31+08:00 | 2018-08-15T11:22:49+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | H18FTP862G | CPSM Final Test | FAIL | [42 empty] | 0.25748
P | 2018-08-15T08:11:21+08:00 | 2018-08-15T12:22:37+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP863G | CPSM Final Test | FAIL
P | 2018-08-12T01:06:12+08:00 | 2018-08-15T12:24:11+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP847G | CPSM Final Test | FAIL | [17 empty] | 0
P | 2018-08-14T15:41:10+08:00 | 2018-08-15T12:25:45+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | H18FTP862G | CPSM Final Test | PASS | 129 | 162.14 | 0.17 | -0.69 | 56.55 | 48.6 | 1.82 | 0.13 | 1 | -2 | 0 | -0.9557 | 0.5038 | 3.2 | -20 |  |  | 0.025 | 0.1951 | -0.0067 | 1 | [21 empty] | 0 | [109 empty] | 0
P | 2018-08-15T10:07:57+08:00 | 2018-08-15T13:21:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP864G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -0.31 | 56.72 | 50.88 | 1.13 | 0.13 | 1 | -2 | 0 | -0.8939 | 0.5036 | 3 | 10 |  |  | 0.1836 | 0.1226 | -0.0013 | 0.99999 | [21 empty] | 0 | [110 empty] | 0
P | 2018-08-13T18:45:11+08:00 | 2018-08-15T13:27:09+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP847G | CPSM Final Test | PASS | 129 | 159.53 | -0.02 | -0.69 | 57.08 | 50.46 | 0.91 | 0.13 | 1 | -2 | 0 | -0.7702 | 0.472 | 8.1 | 12 |  |  | 0.1315 | 0.1152 | 0.1219 | 0.99979 | [21 empty] | 0 | [101 empty] | 0
P | 2018-08-15T10:04:17+08:00 | 2018-08-15T13:28:22+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP867G | CPSM Final Test | PASS | 129 | 157.55 | 0.27 | -0.77 | 56.54 | 48.43 | 2.05 | 0.13 | -1 | 3 | -1 | -0.9038 | 0.489 | 7.4 | 45 |  |  | 0.1652 | 0.3853 | 0.0062 | 1 | [21 empty] | 0 | [111 empty] | 0
P | 2018-08-15T10:44:36+08:00 | 2018-08-15T13:31:53+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP866G | CPSM Final Test | PASS | 129 | 163.04 | 0.17 | -0.46 | 56.63 | 48.6 | 1.36 | 0.13 | 1 | -4 | 0 | -0.852 | 0.4978 | 4.2 | 44 |  |  | 0.0982 | 0.2827 | -0.0586 | 1 | [21 empty] | 0.00923 | [112 empty] | 0
P | 2018-08-15T10:02:55+08:00 | 2018-08-15T13:39:20+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP868G | CPSM Final Test | PASS | 129 | 159.71 | 0.46 | -0.77 | 56.5 | 49.7 | 0.83 | 0.13 | -1 | 1 | -1 | -0.9398 | 0.5114 | 0 | -6 |  |  | 0.1323 | 0.1275 | 0.0406 | 0.99998 | [21 empty] | 0.02469 | [113 empty] | 0
P | 2018-08-15T10:16:59+08:00 | 2018-08-15T13:35:43+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP863G | CPSM Final Test | PASS | 129 | 164.21 | 0.37 | -0.54 | 56.49 | 49.53 | 1.44 | 0.13 | -1 | 1 | 1 | -0.8919 | 0.497 | -1.3 | -25 |  |  | 0.0978 | 0.0693 | -0.019 | 1 | [21 empty] | 0 | [114 empty] | 0
P | 2018-08-15T12:44:08+08:00 | 2018-08-15T16:07:31+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP841G | CPSM Final Test | FAIL | [13 empty] | -5.6
P | 2018-08-15T14:40:10+08:00 | 2018-08-15T17:50:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | H18FTP848G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -0.31 | 56.82 | 51.56 | 1.06 | 0.13 | 1 | -2 | 0 | -0.846 | 0.4626 | 6.8 | 11 |  |  | 0.0415 | 0.2278 | -0.0406 | 0.99999 | [21 empty] | 0 | [115 empty] | 0
P | 2018-08-15T14:38:14+08:00 | 2018-08-15T17:55:52+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP844G | CPSM Final Test | PASS | 129 | 158.45 | -0.02 | -0.92 | 56.5 | 49.36 | 0.83 | 0.13 | -1 | 1 | 0 | -0.862 | 0.547 | 3.8 | 8 |  |  | 0.4216 | 0.2745 | 0.0153 | 0.99999 | [21 empty] | 0.01367 | [116 empty] | 0
P | 2018-08-15T14:29:02+08:00 | 2018-08-15T17:50:45+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP843G | CPSM Final Test | PASS | 129 | 157.82 | -0.02 | 0.23 | 56.64 | 49.11 | 1.21 | 0.13 | 1 | 1 | 0 | -0.9737 | 0.5196 | 0.5 | 18 |  |  | 0.0959 | 0.1357 | 0.0074 | 1 | [21 empty] | 0 | [117 empty] | 0
P | 2018-08-15T14:54:01+08:00 | 2018-08-15T17:51:37+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP845G | CPSM Final Test | PASS | 129 | 157.55 | 0.37 | -0.84 | 56.62 | 49.78 | 0.91 | 0.13 | 1 | -3 | 0 | -0.9059 | 0.499 | 1.3 | 5 |  |  | 0.019 | 0.234 | -0.0393 | 1 | [21 empty] | 0 | [118 empty] | 0
P | 2018-08-15T14:44:03+08:00 | 2018-08-15T17:53:31+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP842G | CPSM Final Test | PASS | 129 | 159.53 | -0.02 | -0.31 | 56.42 | 55.03 | 1.51 | 0.13 | 1 | 1 | 0 | -0.9237 | 0.515 | 3.8 | 60 |  |  | 0.0846 | 0.3724 | 0.0176 | 0.99999 | [21 empty] | 0 | [100 empty] | 0
P | 2018-08-15T14:47:23+08:00 | 2018-08-15T17:59:31+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP871G | CPSM Final Test | PASS | 129 | 159.62 | 0.27 | -0.77 | 56.51 | 49.45 | 0.91 | 0.13 | 1 | 2 | 0 | -0.9338 | 0.525 | 5.7 | 3 |  |  | 0.1534 | 0.203 | -0.0023 | 0.99999 | [21 empty] | 0.01421 | [119 empty] | 0
P | 2018-08-15T14:25:50+08:00 | 2018-08-15T18:07:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP841G | CPSM Final Test | PASS | 129 | 161.06 | 0.37 | -0.92 | 56.61 | 49.53 | 1.44 | 0.13 | -1 | 1 | 0 | -0.858 | 0.5318 | -2.3 | 21 |  |  | 0.1552 | 0.3334 | 0.1084 | 0.99984 | [21 empty] | 0 | [101 empty] | 0
P | 2018-08-15T18:20:33+08:00 | 2018-08-15T19:59:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP594G | CPSM Final Test | FAIL | [5 empty] | 47.16
P | 2018-08-15T16:38:59+08:00 | 2018-08-15T20:43:37+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP846G | CPSM Final Test | FAIL | [163 empty] | 8759
P | 2018-08-15T18:36:41+08:00 | 2018-08-15T21:33:11+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP872G | CPSM Final Test | PASS | 129 | 160.52 | 0.17 | -0.84 | 56.52 | 50.46 | 1.36 | 0.13 | -1 | -2 | 0 | -0.9537 | 0.53 | 8.7 | -9 |  |  | 0.1849 | 0.4641 | 0.0461 | 0.99998 | [21 empty] | 0 | [121 empty] | 0
P | 2018-08-15T18:41:18+08:00 | 2018-08-15T21:49:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP846G | CPSM Final Test | PASS | 129 | 159.98 | 0.78 | -0.38 | 56.82 | 48.51 | 1.06 | 0.13 | 1 | 2 | 0 | -0.9338 | 0.5034 | 7.8 | 0 |  |  | 0.1085 | 0.2055 | -0.0149 | 1 | [21 empty] | 0 | [120 empty] | 0
P | 2018-08-15T03:41:09+08:00 | 2018-08-16T10:38:29+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP594G | CPSM Final Test | FAIL
P | 2018-08-15T04:04:51+08:00 | 2018-08-16T10:48:31+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | H18FNP593G | CPSM Final Test | FAIL | [14 empty] | 1111
P | 2018-08-15T19:22:41+08:00 | 2018-08-16T11:40:43+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP590G | CPSM Final Test | PASS |  | 159.71 | 0.17 | -0.99 | 56.23 | 50.55 | 2.12 | 0.13 | 0 | -4 | -1 | -0.9458 | 0.4866 | -0.1 | 10 |  |  | 0.0779 | 0.8648 | -0.0583 | 0.99999 |  |  |  |  | 133 |  | 0 | [137 empty] | 0
P | 2018-08-15T19:29:45+08:00 | 2018-08-16T11:49:50+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | H18FNP589G | CPSM Final Test | PASS |  | 160.79 | -0.02 | -1.07 | 56.23 | 51.14 | 0.91 | 0.13 | 0 | -2 | 0 | -0.9637 | 0.5416 | 1.6 | 71 |  |  | 0.3955 | 0.6251 | 0.0072 | 0.99988 |  |  |  |  | 133 |  | 0 | [138 empty] | 0
P | 2018-08-15T03:02:27+08:00 | 2018-08-16T11:47:37+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP592G | CPSM Final Test | FAIL
P | 2018-08-15T19:25:02+08:00 | 2018-08-16T11:48:03+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP592G | CPSM Final Test | PASS |  | -26.59 | -0.02 | -0.92 | 56.5 | 50.55 | 1.9 | 0.13 | 1 | 1 | 0 | -0.9617 | 0.5336 | 2.9 | 4 |  |  | 0.0534 | 0.0155 | -0.0222 | 0.99999 |  |  |  |  | 133 |  | 0 | [139 empty] | 0
P | 2018-08-15T19:09:49+08:00 | 2018-08-16T12:26:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP594G | CPSM Final Test | PASS |  | 164.3 | -0.02 | -0.92 | 56.73 | 49.95 | 1.44 | 0.13 | -1 | 2 | 0 | -0.9558 | 0.5032 | 2.1 | 4 |  |  | 0.0519 | 0.1823 | -0.0281 | 0.99998 |  |  |  |  | 133 |  | 0 | [140 empty] | 0
P | 2018-08-15T01:32:24+08:00 | 2018-08-16T13:11:19+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | H18FNP591G | CPSM Final Test | FAIL | [169 empty] | 0
P | 2018-08-15T19:21:52+08:00 | 2018-08-16T14:16:57+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | H18FNP591G | CPSM Final Test | PASS |  | 160.43 | 0.27 | -0.31 | 56.32 | 49.36 | 1.59 | 0.13 | 1 | 2 | 0 | -0.9458 | 0.539 | 2.4 | 15 |  |  | 0.0807 | 0.374 | -0.0802 | 1 |  |  |  |  | 133 |  | 0 | [141 empty] | 0
P | 2018-08-16T07:38:29+08:00 | 2018-08-16T15:16:49+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | H18USM072G | CPSM Final Test | FAIL | [170 empty] | 0
P | 2018-08-16T11:27:39+08:00 | 2018-08-16T16:22:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | H18USM072G | CPSM Final Test | PASS |  | 161.51 | 0.37 | -0.92 | 56.54 | 49.19 | 1.67 | 0.13 | -1 | -1 | 0 | -0.8799 | 0.5188 | 3.3 | 110 |  |  | 0.1952 | 0.1597 | -0.0106 | 0.99999 |  |  |  |  | 133 |  | 0 | [142 empty] | 0
P | 2018-08-16T16:09:49+08:00 | 2018-08-16T17:13:20+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | H18USM071G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-08-14T21:07:07+08:00 | 2018-08-16T17:46:15+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | H18FNP593G | CPSM Final Test | FAIL | [171 empty] | 0
P | 2018-08-16T10:49:33+08:00 | 2018-08-16T18:32:29+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | H18USM070G | CPSM Final Test | FAIL | [172 empty] | 0
P | 2018-08-16T16:33:05+08:00 | 2018-08-16T18:34:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | H18FNP596G | CPSM Final Test | FAIL | [5 empty] | 46.14
P | 2018-08-16T15:33:09+08:00 | 2018-08-16T18:41:55+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI192G | CPSM Final Test | PASS | 129 | 157.91 | 0.08 | -0.61 | 56.53 | 48.26 | 1.74 | 0.13 | 1 | -2 | 0 | -0.9438 | 0.5224 | -0.8 | -12 |  | 0 | 0.078 | 0.0505 | -0.0678 | 1 | [152 empty] | 0
P | 2018-08-15T19:26:41+08:00 | 2018-08-16T18:51:47+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | H18FNP593G | CPSM Final Test | PASS |  | 160.97 | 0.08 | -1.22 | 56.2 | 49.36 | 1.59 | 0.13 | -1 | -3 | 0 | -0.852 | 0.5106 | -4.2 | -6 |  |  | 0.127 | 0.52 | -0.0273 | 0.9999 |  |  |  |  | 133 |  | 0 | [143 empty] | 0
P | 2018-08-16T13:48:07+08:00 | 2018-08-16T19:20:46+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | H18FNP595G | CPSM Final Test | FAIL | [14 empty] | 610
P | 2018-08-16T15:29:05+08:00 | 2018-08-16T19:38:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | H18FNP596G | CPSM Final Test | FAIL | [13 empty] | -6.9
P | 2018-08-16T14:41:01+08:00 | 2018-08-16T19:38:05+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | H18USM070G | CPSM Final Test | PASS |  | 160.34 | -0.02 | -0.77 | 56.94 | 49.45 | 1.36 | 0.13 | -1 | -2 | 0 | -0.9238 | 0.516 | 7.6 | -7 |  |  | 0.1236 | 0.1491 | -0.0433 | 1 |  |  |  |  | 133 |  | 0 | [144 empty] | 0
P | 2018-08-16T12:58:29+08:00 | 2018-08-16T20:10:24+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | H18FNP595G | CPSM Final Test | FAIL
P | 2018-08-16T17:33:39+08:00 | 2018-08-16T21:42:19+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | H18FNP596G | CPSM Final Test | PASS |  | 164.03 | 0.17 | -0.77 | 56.31 | 48.34 | 1.67 | 0.13 | -1 | -3 | 0 | -0.9119 | 0.548 | 3.4 | -16 |  |  | 0.066 | 0.0736 | -0.0177 | 0.99999 |  |  |  |  | 133 |  | 0 | [142 empty] | 0
P | 2018-08-15T23:24:41+08:00 | 2018-08-17T09:58:28+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | H18USM071G | CPSM Final Test | FAIL | [174 empty] | 0
P | 2018-08-15T23:15:45+08:00 | 2018-08-17T09:53:08+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | H18FNP595G | CPSM Final Test | FAIL | [175 empty] | 0
P | 2018-08-16T16:41:35+08:00 | 2018-08-17T11:04:00+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | H18USM071G | CPSM Final Test | PASS |  | 164.12 | 0.27 | -0.99 | 56.76 | 48.68 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9258 | 0.5092 | 12.8 | -4 |  |  | 0.0932 | 0.2382 | 0.0175 | 0.99998 |  |  |  |  | 133 |  | 0 | [146 empty] | 0
P | 2018-08-15T21:32:43+08:00 | 2018-08-17T11:36:10+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | H18FNP595G | CPSM Final Test | FAIL
P | 2018-08-16T16:34:26+08:00 | 2018-08-17T11:36:40+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | H18FNP595G | CPSM Final Test | PASS |  | 160.61 | -0.11 | -0.61 | 56.73 | 52.83 | 1.67 | 0.13 | 0 | 2 | 0 | -0.8619 | 0.539 | -2.8 | 21 |  |  | 0.1429 | 0.1992 | -0.0206 | 0.99999 |  |  |  |  | 133 |  | 0 | [147 empty] | 0
P | 2018-08-17T08:02:48+08:00 | 2018-08-17T16:15:23+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI203G | CPSM Final Test | FAIL | [176 empty] | 0
P | 2018-08-17T12:15:02+08:00 | 2018-08-17T16:20:08+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI204G | CPSM Final Test | PASS | 129 | 165.92 | -0.3 | -0.54 | 56.53 | 49.87 | 1.44 | 0.13 | -1 | -3 | 0 | -0.842 | 0.5312 | 5.1 | 31 |  | 0.02886 | 0.1326 | 0.1822 | 0.0067 | 0.99999 | [154 empty] | 0
P | 2018-08-17T12:37:18+08:00 | 2018-08-17T16:33:26+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | H18USM074G | CPSM Final Test | PASS |  | 160.97 | 0.27 | -0.16 | 56.52 | 48.01 | 1.51 | 0.13 | -1 | 3 | 0 | -0.8739 | 0.5156 | 2.8 | 15 |  |  | 0.1014 | 0.1336 | -0.006 | 0.99998 |  |  |  |  | 133 |  | 0 | [146 empty] | 0
P | 2018-08-17T12:09:05+08:00 | 2018-08-17T17:21:03+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI203G | CPSM Final Test | PASS | 129 | 153.77 | -0.02 | -1.22 | 56.2 | 50.72 | 0.91 | 0.13 | 1 | -2 | 0 | -0.9059 | 0.504 | -1.3 | -14 |  | 0 | 0.0665 | 0.0663 | 0.0221 | 0.99999 | [155 empty] | 0
P | 2018-08-17T14:44:28+08:00 | 2018-08-17T17:30:51+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI200G | CPSM Final Test | FAIL | [13 empty] | -6.3
P | 2018-08-17T06:00:16+08:00 | 2018-08-17T19:16:29+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI205G | CPSM Final Test | FAIL | [13 empty] | 18.8
P | 2018-08-17T09:47:38+08:00 | 2018-08-17T18:17:25+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI207G | CPSM Final Test | FAIL
P | 2018-08-17T14:02:31+08:00 | 2018-08-17T18:17:55+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI207G | CPSM Final Test | PASS | 129 | 162.95 | 0.37 | -0.54 | 56.22 | 50.8 | 0.75 | 0.13 | 1 | 3 | 0 | -0.854 | 0.496 | -2.4 | -5 |  | 0.00486 | 0.1056 | 0.1521 | 0.0119 | 0.99996 | [150 empty] | 0
P | 2018-08-17T11:47:24+08:00 | 2018-08-17T19:29:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | H18USM069G | CPSM Final Test | FAIL | [27 empty] | 0.05628
P | 2018-08-17T11:45:22+08:00 | 2018-08-17T19:31:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | H18USM069G | CPSM Final Test | FAIL
P | 2018-08-17T16:07:39+08:00 | 2018-08-17T19:30:55+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI200G | CPSM Final Test | PASS | 129 | -6.07 | 0.08 | -0.84 | 56.49 | 51.73 | 1.29 | 0.13 | -1 | 1 | 0 | -0.9218 | 0.526 | -4.3 | 12 |  | 0.05353 | 0.1119 | 0.1638 | -0.0264 | 0.99997 | [156 empty] | 0
P | 2018-08-17T19:52:30+08:00 | 2018-08-17T20:28:06+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI212G | CPSM Final Test | FAIL | [5 empty] | 44.7
P | 2018-08-16T22:43:24+08:00 | 2018-08-18T08:33:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | H18USM069G | CPSM Final Test | FAIL
P | 2018-08-17T15:38:33+08:00 | 2018-08-18T08:34:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | H18USM069G | CPSM Final Test | PASS |  | 161.78 | 1.03 | -0.31 | 56.69 | 50.12 | 1.51 | 0.13 | 0 | -1 | 1 | -0.7842 | 0.523 | 4.3 | -2 |  |  | 0.0401 | 0.0619 | -0.0318 | 1 |  |  |  |  | 133 |  | 0.02348 | [150 empty] | 0
P | 2018-08-17T06:56:01+08:00 | 2018-08-18T09:01:23+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI210G | CPSM Final Test | FAIL | [13 empty] | -6.5
P | 2018-08-16T16:06:12+08:00 | 2018-08-18T09:10:33+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI205G | CPSM Final Test | FAIL | [16 empty] | 0.17758
P | 2018-08-17T12:38:22+08:00 | 2018-08-18T10:13:31+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI205G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.99 | 56.47 | 50.72 | 1.36 | 0.13 | 0 | -2 | 0 | -0.9018 | 0.5154 | 14 | 1 |  | 0.01367 | 0.0796 | 0.2807 | -0.001 | 0.99999 | [151 empty] | 0
P | 2018-08-18T08:02:12+08:00 | 2018-08-18T10:37:07+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI211G | CPSM Final Test | FAIL | [13 empty] | -6.9
P | 2018-08-18T08:20:22+08:00 | 2018-08-18T10:36:58+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI208G | CPSM Final Test | FAIL | [13 empty] | -12.1
P | 2018-08-17T20:10:18+08:00 | 2018-08-18T10:44:02+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI212G | CPSM Final Test | PASS | 129 | 162.14 | 0.17 | -0.77 | 56.31 | 49.02 | 1.36 | 0.13 | 0 | -2 | 1 | -0.9578 | 0.5372 | -2.3 | 0 |  | 0.02746 | 0.04 | 0.3783 | 0.0031 | 0.99999 | [158 empty] | 0
P | 2018-08-17T04:58:04+08:00 | 2018-08-18T10:47:15+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI209G | CPSM Final Test | FAIL | [16 empty] | 5.0161
P | 2018-08-17T04:43:17+08:00 | 2018-08-18T11:14:07+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI210G | CPSM Final Test | FAIL | [28 empty] | 2102
P | 2018-08-18T10:33:40+08:00 | 2018-08-18T11:20:32+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI213G | CPSM Final Test | FAIL | [5 empty] | 46.57
P | 2018-08-18T09:28:11+08:00 | 2018-08-18T11:41:51+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI198G | CPSM Final Test | FAIL | [13 empty] | -10.2
P | 2018-08-17T19:52:39+08:00 | 2018-08-18T11:50:09+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI209G | CPSM Final Test | PASS | 129 | 160.34 | 0.27 | -0.77 | 56.34 | 49.7 | 1.51 | 0.13 | 1 | -2 | 1 | -0.9019 | 0.5176 | 3 | 0 |  | 0 | 0.1217 | 0.2557 | 0.018 | 0.99987 | [159 empty] | 0
P | 2018-08-17T04:02:57+08:00 | 2018-08-18T11:54:27+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI210G | CPSM Final Test | FAIL | [28 empty] | 1367
P | 2018-08-18T09:36:14+08:00 | 2018-08-18T12:21:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | H18SDI214G | CPSM Final Test | PASS | 129 | 163.4 | 0.08 | -0.61 | 56.52 | 49.02 | 1.36 | 0.13 | -1 | -1 | 0 | -0.8281 | 0.5062 | 6.2 | -3 |  | 0.02287 | 0.1338 | 0.2223 | 0.0028 | 0.99999 | [160 empty] | 0
P | 2018-08-18T09:19:39+08:00 | 2018-08-18T12:38:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI211G | CPSM Final Test | PASS | 129 | 161.78 | 0.08 | -0.99 | 56.18 | 49.78 | 1.06 | 0.13 | 0 | 1 | 0 | -0.8559 | 0.517 | 0.8 | 9 |  | 0 | 0.1099 | 0.185 | -0.017 | 1 | [155 empty] | 0
P | 2018-08-18T05:56:36+08:00 | 2018-08-18T13:00:44+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI208G | CPSM Final Test | FAIL | [182 empty] | 115933
P | 2018-08-17T19:58:42+08:00 | 2018-08-18T13:14:07+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI210G | CPSM Final Test | PASS | 129 | 160.52 | 0.17 | -0.61 | 56.46 | 50.97 | 0.98 | 0.13 | -1 | -2 | 1 | -0.9579 | 0.5212 | -3.5 | -18 |  | 0.0126 | 0.1184 | 0.0566 | 0.0431 | 1 | [7 empty] | 0
P | 2018-08-18T07:27:37+08:00 | 2018-08-18T13:42:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI198G | CPSM Final Test | FAIL
P | 2018-08-18T10:35:01+08:00 | 2018-08-18T13:42:55+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI198G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -0.77 | 56.61 | 49.7 | 1.44 | 0.13 | -1 | -3 | 0 | -0.834 | 0.485 | 4.8 | 32 |  | 0.03868 | 0.093 | 0.2212 | -0.0178 | 0.99999 | [162 empty] | 0
P | 2018-08-18T10:57:06+08:00 | 2018-08-18T13:51:12+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI213G | CPSM Final Test | PASS | 129 | 165.56 | -0.02 | -0.46 | 56.46 | 48.09 | 1.44 | 0.13 | -1 | -2 | 0 | -0.7661 | 0.512 | 4 | 24 |  | 0 | 0.0883 | 0.1062 | 0.0268 | 0.99996 | [163 empty] | 0
P | 2018-08-18T09:28:40+08:00 | 2018-08-18T14:06:24+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI208G | CPSM Final Test | PASS | 129 | 161.78 | 0.17 | -0.77 | 56.33 | 49.53 | 0.98 | 0.13 | 1 | -2 | 1 | -0.86 | 0.5254 | -4 | -10 |  | 0.01957 | 0.0581 | 0.1997 | 0.0269 | 0.99999 | [161 empty] | 0
P | 2018-08-18T15:15:18+08:00 | 2018-08-18T16:11:36+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI193G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-08-18T15:41:16+08:00 | 2018-08-18T18:55:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI194G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.23 | 56.61 | 48.6 | 1.51 | 0.13 | 1 | -2 | 0 | -0.8839 | 0.5292 | 6 | -3 |  | 0 | 0.1085 | 0.4125 | 0.0222 | 0.99999 | [164 empty] | 0
P | 2018-08-18T15:59:44+08:00 | 2018-08-18T19:05:58+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI206G | CPSM Final Test | PASS | 129 | 164.39 | 0.37 | -0.77 | 56.44 | 50.8 | 1.36 | 0.13 | 1 | -2 | 0 | -0.8559 | 0.519 | -1 | 7 |  | 0 | 0.0738 | 0.162 | 0.0541 | 0.99999 | [165 empty] | 0
P | 2018-08-18T15:46:37+08:00 | 2018-08-18T19:01:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | H18USM073G | CPSM Final Test | PASS |  | 158.36 | -0.02 | -0.99 | 56.4 | 50.12 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9638 | 0.4654 | 4.3 | -6 |  |  | 0.1605 | 0.2851 | -0.0022 | 0.99999 |  |  |  |  | 133 |  | 0.01614 | [159 empty] | 0
P | 2018-08-18T15:57:44+08:00 | 2018-08-18T19:01:33+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI197G | CPSM Final Test | PASS | 129 | 162.68 | 0.08 | -0.46 | 56.64 | 50.04 | 1.21 | 0.13 | 0 | 1 | 0 | -0.9518 | 0.539 | -3.3 | 40 |  | 0.01556 | 0.0549 | 0.1391 | 0.005 | 0.99999 | [7 empty] | 0
P | 2018-08-18T15:43:27+08:00 | 2018-08-18T19:04:56+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI193G | CPSM Final Test | PASS | 129 | 154.4 | -0.02 | -0.84 | 56.76 | 48.01 | 1.59 | 0.13 | 1 | 1 | 0 | -0.8639 | 0.5306 | 7.8 | 71 |  | 0.00747 | 0.0503 | 0.2708 | 0.0097 | 0.99999 | [161 empty] | 0
P | 2018-08-16T23:46:06+08:00 | 2018-08-20T08:14:51+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI196G | CPSM Final Test | FAIL | [14 empty] | 761
P | 2018-08-16T22:16:52+08:00 | 2018-08-20T09:26:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI202G | CPSM Final Test | FAIL
P | 2018-08-16T22:32:22+08:00 | 2018-08-20T09:28:35+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI196G | CPSM Final Test | FAIL | [14 empty] | 886
P | 2018-08-16T20:55:32+08:00 | 2018-08-20T10:47:21+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI202G | CPSM Final Test | FAIL | [188 empty] | 2114
P | 2018-08-18T16:00:28+08:00 | 2018-08-20T11:22:57+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI196G | CPSM Final Test | PASS | 129 | 155.66 | -0.02 | -0.84 | 56.6 | 50.12 | 1.67 | 0.13 | -1 | -3 | -1 | -0.9178 | 0.5446 | 11.3 | -3 |  | 0 | 0.1593 | 0.08 | 0.0093 | 0.99991 | [168 empty] | 0
P | 2018-08-18T15:51:26+08:00 | 2018-08-20T11:53:05+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI202G | CPSM Final Test | PASS | 129 | 161.06 | 0.87 | -1.38 | 56.03 | 48.09 | 1.13 | 0.13 | 1 | -2 | 0 | -0.86 | 0.53 | -1.7 | 9 |  | 0.01013 | 0.4724 | 0.5247 | -0.0362 | 0.99999 | [167 empty] | 0
P | 2018-08-20T11:02:45+08:00 | 2018-08-20T13:54:40+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI220G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.92 | 56.17 | 48.77 | 1.36 | 0.13 | -2 | -2 | 0 | -0.9059 | 0.507 | 4.8 | -5 |  | 0.00654 | 0.1165 | 0.2915 | 0.034 | 1 | [169 empty] | 0
P | 2018-08-20T10:58:00+08:00 | 2018-08-20T13:50:05+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI216G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.92 | 56.43 | 48.85 | 1.21 | 0.13 | -2 | -2 | 0 | -0.9537 | 0.5114 | 0.8 | -1 |  | 0.0085 | 0.167 | 0.0477 | 0.0052 | 1 | [170 empty] | 0
P | 2018-08-20T06:06:20+08:00 | 2018-08-20T15:09:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI222G | CPSM Final Test | FAIL
P | 2018-08-20T11:36:37+08:00 | 2018-08-20T15:27:45+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI215G | CPSM Final Test | FAIL
P | 2018-08-20T11:02:24+08:00 | 2018-08-20T16:04:23+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI217G | CPSM Final Test | FAIL | [188 empty] | 1903
P | 2018-08-20T15:12:14+08:00 | 2018-08-20T16:17:34+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI224G | CPSM Final Test | FAIL | [5 empty] | 45.8
P | 2018-08-20T10:37:40+08:00 | 2018-08-20T16:12:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI222G | CPSM Final Test | PASS | 129 | 159.08 | -0.02 | -0.23 | 56.6 | 49.87 | 1.44 | 0.13 | 1 | 2 | 0 | -0.8918 | 0.5048 | 1.9 | 39 |  | 0 | 0.2145 | 0.3067 | 0.0141 | 0.99998 | 0
P | 2018-08-20T14:46:54+08:00 | 2018-08-20T16:56:55+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | H18SDI195G | CPSM Final Test | FAIL | [13 empty] | -7.3
P | 2018-08-16T14:39:11+08:00 | 2018-08-20T16:58:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | H18SDI201G | CPSM Final Test | FAIL | [5 empty] | 45.38
P | 2018-08-20T13:33:23+08:00 | 2018-08-20T17:10:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI217G | CPSM Final Test | PASS | 129 | 160.52 | -0.02 | -0.77 | 56.3 | 51.48 | 1.21 | 0.13 | -1 | -2 | 0 | -0.9199 | 0.5 | -1.1 | 0 |  | 0 | 0.0317 | 0.2534 | -0.011 | 1 | [167 empty] | 0
P | 2018-08-20T13:32:11+08:00 | 2018-08-20T17:15:21+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI215G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -1.07 | 56.25 | 50.55 | 1.13 | 0.13 | -1 | -3 | 0 | -0.8859 | 0.497 | -2.4 | -1 |  | 0 | 0.1641 | 0.2781 | 0.0438 | 0.99991 | [171 empty] | 0
P | 2018-08-20T03:08:28+08:00 | 2018-08-20T18:05:59+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI221G | CPSM Final Test | FAIL | [16 empty] | 0.29083
P | 2018-08-20T15:44:54+08:00 | 2018-08-20T18:53:14+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI224G | CPSM Final Test | PASS | 129 | 156.11 | 0.27 | -0.54 | 56.53 | 51.05 | 1.29 | 0.13 | -1 | 2 | -1 | -0.9797 | 0.4838 | 7.4 | -1 |  | 0.01714 | 0.1501 | 0.143 | 0.0259 | 0.99999 | [172 empty] | 0
P | 2018-08-20T15:51:55+08:00 | 2018-08-20T18:58:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | H18SDI195G | CPSM Final Test | PASS | 129 | 163.58 | 0.17 | -0.54 | 56.78 | 50.38 | 1.59 | 0.13 | 1 | 1 | 0 | -0.9178 | 0.5134 | -0.1 | 8 |  | 0.02318 | 0.1374 | 0.135 | -0.0215 | 1 | [173 empty] | 0
P | 2018-08-20T15:43:46+08:00 | 2018-08-20T18:59:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI225G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-08-20T17:46:18+08:00 | 2018-08-20T19:07:21+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI226G | CPSM Final Test | FAIL | [83 empty] | FALSE
P | 2018-08-20T17:44:04+08:00 | 2018-08-20T19:09:35+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI226G | CPSM Final Test | FAIL | [195 empty] | FALSE
P | 2018-08-20T01:55:26+08:00 | 2018-08-20T19:19:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI221G | CPSM Final Test | FAIL
P | 2018-08-20T10:37:14+08:00 | 2018-08-20T19:19:33+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI221G | CPSM Final Test | PASS | 129 | 160.52 | -0.02 | 0.07 | 56.48 | 53.59 | 0.91 | 0.13 | -1 | 1 | 1 | -0.9578 | 0.491 | 3.1 | 18 |  | 0 | 0.0905 | 0.0213 | -0.0257 | 0.99998 | [175 empty] | 0
P | 2018-08-16T12:10:33+08:00 | 2018-08-20T19:27:16+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | H18SDI201G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-08-20T14:59:51+08:00 | 2018-08-20T20:06:29+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI227G | CPSM Final Test | FAIL
P | 2018-08-20T16:44:32+08:00 | 2018-08-20T20:09:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI226G | CPSM Final Test | FAIL | [13 empty] | -6.3
P | 2018-08-20T17:21:31+08:00 | 2018-08-20T20:58:54+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI225G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -1.68 | 56.21 | 50.12 | 1.36 | 0.13 | -1 | -1 | 0 | -0.9956 | 0.4922 | -3.8 | 7 |  | 0.0117 | 0.095 | 0.2894 | -0.0198 | 1 | [176 empty] | 0
P | 2018-08-20T17:44:36+08:00 | 2018-08-20T21:04:51+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | H18SDI228G | CPSM Final Test | PASS | 129 | 159.35 | 0.37 | -0.08 | 56.73 | 49.28 | 1.67 | 0.13 | -1 | 1 | 1 | -0.9158 | 0.497 | -4.2 | 22 |  | 0.00202 | 0.1448 | 0.0525 | 0.0982 | 0.99989 | [177 empty] | 0
P | 2018-08-20T17:33:10+08:00 | 2018-08-20T21:12:03+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI227G | CPSM Final Test | PASS | 129 | 156.92 | -0.02 | -0.54 | 56.41 | 49.87 | 1.51 | 0.13 | -1 | -2 | 0 | -0.9258 | 0.4892 | 3.2 | 36 |  | 0.01274 | 0.1858 | 0.0969 | -0.0371 | 1 | [167 empty] | 0
P | 2018-08-20T19:18:28+08:00 | 2018-08-20T21:58:37+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | H18SDI218G | CPSM Final Test | PASS | 129 | 158.54 | -0.02 | -0.61 | 56.16 | 50.8 | 1.29 | 0.13 | 1 | 3 | 0 | -0.9277 | 0.504 | -4.8 | 4 |  | 0.02238 | 0.1216 | 0.114 | -0.0295 | 1 | [173 empty] | 0
P | 2018-08-20T19:18:49+08:00 | 2018-08-20T22:06:32+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI199G | CPSM Final Test | PASS | 129 | 164.75 | 0.27 | -0.69 | 56.31 | 49.45 | 1.29 | 0.13 | 0 | 1 | 0 | -0.9617 | 0.5268 | 3.8 | 89 |  | 0.00619 | 0.0953 | 0.1588 | 0.0316 | 0.99999 | [178 empty] | 0
P | 2018-08-20T06:29:19+08:00 | 2018-08-21T08:22:23+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI219G | CPSM Final Test | FAIL
P | 2018-08-20T03:38:18+08:00 | 2018-08-21T09:15:21+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI226G | CPSM Final Test | FAIL | [21 empty] | 733137
P | 2018-08-20T19:25:51+08:00 | 2018-08-21T10:06:55+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | H18SDI219G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.69 | 56.33 | 49.28 | 1.36 | 0.13 | 1 | 3 | 1 | -0.9277 | 0.519 | 7.2 | -57 |  | 0 | 0.0788 | 0.101 | -0.0057 | 1 | [179 empty] | 0
P | 2018-08-20T18:26:49+08:00 | 2018-08-21T10:21:21+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | H18SDI226G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -0.69 | 56.38 | 51.05 | 1.21 | 0.13 | 0 | -3 | 0 | -0.9638 | 0.5374 | 4.6 | -13 |  | 0 | 0.1042 | 0.1702 | 0.0554 | 0.99999 | 0
P | 2018-08-21T10:30:53+08:00 | 2018-08-21T10:51:43+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI223G | CPSM Final Test | FAIL | [201 empty] | FALSE
P | 2018-08-21T08:09:45+08:00 | 2018-08-21T12:01:39+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI201G | CPSM Final Test | FAIL
P | 2018-08-21T09:39:56+08:00 | 2018-08-21T12:24:10+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | H18SDI229G | CPSM Final Test | PASS | 129 | 155.93 | -0.02 | -0.31 | 56.31 | 50.12 | 1.06 | 0.13 | 1 | 4 | 1 | -0.9278 | 0.5324 | 6.5 | 80 |  | 0.01539 | 0.1253 | 0.1489 | 0.0295 | 0.99999 | [181 empty] | 0
P | 2018-08-21T10:05:42+08:00 | 2018-08-21T13:07:31+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | H18SDI201G | CPSM Final Test | PASS | 129 | 155.57 | 0.46 | -0.46 | 56.3 | 49.02 | 1.51 | 0.13 | -1 | -1 | 0 | -0.9738 | 0.5158 | 3.8 | 17 |  | 0.01562 | 0.042 | 0.3161 | -0.0207 | 0.99999 | [167 empty] | 0
P | 2018-08-21T10:41:18+08:00 | 2018-08-21T13:22:33+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | H18SDI223G | CPSM Final Test | PASS | 129 | 154.22 | 0.27 | -0.61 | 56.43 | 50.8 | 1.29 | 0.13 | -1 | -1 | 1 | -0.9357 | 0.5326 | -4.8 | -10 |  | 0.01524 | 0.0839 | 0.1309 | -0.0281 | 0.99998 | [182 empty] | 0
P | 2018-08-23T10:34:25+08:00 | 2018-08-23T11:23:33+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP597G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-08-23T10:18:03+08:00 | 2018-08-23T11:39:55+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP597G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-08-23T09:03:27+08:00 | 2018-08-23T12:12:06+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | H18FNP600G | CPSM Final Test | PASS |  | 159.89 | -0.02 | -0.92 | 56.25 | 50.8 | 1.13 | 0.13 | -2 | -2 | 0 | -0.836 | 0.5068 | 5.3 | 19 |  |  | 0.1795 | 0.355 | 0.1127 | 0.99993 |  |  |  |  | 133 |  | 0.00436 | [177 empty] | 0
P | 2018-08-23T09:26:29+08:00 | 2018-08-23T12:14:24+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | H18FNP598G | CPSM Final Test | PASS |  | 158.9 | 0.37 | -0.54 | 56.78 | 56.13 | 0.91 | 0.13 | 1 | -4 | 0 | -0.9737 | 0.4974 | 3.7 | 1 |  |  | 0.1284 | 0.1209 | -0.0547 | 0.99999 |  |  |  |  | 133 |  | 0 | [178 empty] | 0
P | 2018-08-23T07:02:46+08:00 | 2018-08-23T12:33:17+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP599G | CPSM Final Test | FAIL | [188 empty] | 0
P | 2018-08-23T10:13:47+08:00 | 2018-08-23T13:15:22+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP601G | CPSM Final Test | PASS |  | 168.08 | -0.02 | -0.54 | 56.07 | 49.78 | 1.13 | 0.13 | -2 | 3 | -1 | -0.9317 | 0.5074 | 5.2 | 11 |  |  | 0.2763 | 0.5353 | 0.0074 | 0.99998 |  |  |  |  | 133 |  | 0 | [179 empty] | 0
P | 2018-08-23T09:48:02+08:00 | 2018-08-23T13:38:49+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP599G | CPSM Final Test | PASS |  | 163.58 | -0.02 | -0.61 | 56.53 | 50.55 | 1.29 | 0.13 | 0 | -1 | 0 | -0.9738 | 0.515 | 7.3 | 8 |  |  | 0.0556 | 0.2182 | -0.0058 | 0.99999 |  |  |  |  | 133 |  | 0 | [160 empty] | 0
P | 2018-08-23T10:58:59+08:00 | 2018-08-23T14:09:07+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP597G | CPSM Final Test | PASS |  | 160.16 | -0.02 | -1.07 | 56.27 | 49.36 | 0.91 | 0.13 | 0 | 0 | -1 | -0.9577 | 0.528 | 2.9 | 3 |  |  | 0.207 | 0.1275 | 0.0131 | 0.99999 |  |  |  |  | 133 |  | 0 | [180 empty] | 0
P | 2018-08-23T13:20:20+08:00 | 2018-08-23T16:24:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP603G | CPSM Final Test | PASS |  | 161.69 | -0.02 | -0.69 | 56.42 | 49.02 | 1.82 | 0.13 | 1 | -2 | 0 | -0.9238 | 0.511 | 2.1 | -3 |  |  | 0.1359 | 0.1535 | 0.0135 | 0.99999 |  |  |  |  | 133 |  | 0 | [181 empty] | 0
P | 2018-08-23T10:17:31+08:00 | 2018-08-23T16:23:49+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP602G | CPSM Final Test | FAIL | [21 empty] | 0
P | 2018-08-23T11:37:46+08:00 | 2018-08-23T16:37:39+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP605G | CPSM Final Test | FAIL
P | 2018-08-23T14:07:45+08:00 | 2018-08-23T17:21:58+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP604G | CPSM Final Test | PASS |  | 165.29 | -0.02 | -0.84 | 56.14 | 48.18 | 1.36 | 0.13 | 1 | 3 | 0 | -0.9339 | 0.5354 | 8.2 | 8 |  |  | 0.1796 | 0.8418 | -0.0641 | 0.99999 |  |  |  |  | 133 |  | 0 | [182 empty] | 0
P | 2018-08-23T09:00:17+08:00 | 2018-08-23T17:41:03+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP602G | CPSM Final Test | FAIL | [21 empty] | 0
P | 2018-08-23T14:07:42+08:00 | 2018-08-23T18:27:37+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP605G | CPSM Final Test | PASS |  | 163.67 | -0.02 | -0.31 | 56.32 | 49.45 | 0.98 | 0.13 | 0 | -1 | 0 | -0.9218 | 0.54 | 5.8 | 5 |  |  | 0.0626 | 0.2003 | -0.0347 | 0.99999 |  |  |  |  | 133 |  | 0 | [160 empty] | 0
P | 2018-08-23T13:20:40+08:00 | 2018-08-23T18:46:45+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP602G | CPSM Final Test | PASS |  | 165.02 | 0.27 | -0.69 | 56.42 | 50.46 | 0.91 | 0.13 | -1 | -2 | 0 | -0.8879 | 0.5174 | 8.4 | 44 |  |  | 0.1543 | 0.1302 | 0.0092 | 0.99998 | 0 |  |  |  | 133 |  | 0
P | 2018-08-24T09:10:12+08:00 | 2018-08-24T10:36:50+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP611G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-08-23T20:17:58+08:00 | 2018-08-24T11:08:25+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP607G | CPSM Final Test | PASS |  | 161.24 | -0.02 | -0.46 | 56.71 | 56.73 | 1.21 | 0.13 | 0 | -3 | 1 | -0.918 | 0.502 | 3.8 | 37 |  |  | 0.1183 | 0.2501 | 0.0058 | 0.99996 | 0 |  |  |  | 133 |  | 0
P | 2018-08-24T08:34:52+08:00 | 2018-08-24T11:07:39+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP610G | CPSM Final Test | FAIL
P | 2018-08-24T06:31:04+08:00 | 2018-08-24T11:28:07+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP606G | CPSM Final Test | FAIL | [14 empty] | 803
P | 2018-08-24T07:35:48+08:00 | 2018-08-24T12:06:43+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP610G | CPSM Final Test | FAIL
P | 2018-08-24T05:16:08+08:00 | 2018-08-24T12:33:37+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP608G | CPSM Final Test | FAIL | [104 empty] | 0
P | 2018-08-24T05:33:42+08:00 | 2018-08-24T12:25:29+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP606G | CPSM Final Test | FAIL | [208 empty] | 132964
P | 2018-08-24T08:44:11+08:00 | 2018-08-24T13:25:04+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | H18FNP609G | CPSM Final Test | FAIL
P | 2018-08-24T08:54:52+08:00 | 2018-08-24T13:39:21+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP608G | CPSM Final Test | PASS |  | 161.06 | 0.17 | -0.84 | 56.79 | 50.8 | 1.9 | 0.13 | -1 | -3 | 0 | -0.9618 | 0.5096 | 3 | -5 |  |  | 0.0514 | 0.2029 | -0.0034 | 0.99999 |  |  |  |  | 133 |  | 0 | [76 empty] | 0
P | 2018-08-24T08:59:36+08:00 | 2018-08-24T13:47:33+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP606G | CPSM Final Test | PASS |  | 159.53 | 0.17 | -0.99 | 56.26 | 49.45 | 1.59 | 0.13 | 1 | -4 | 0 | -0.9278 | 0.515 | -0.1 | -4 |  |  | 0.067 | 0.1486 | 0.0192 | 0.99997 |  |  |  |  | 133 |  | 0 | [180 empty] | 0
P | 2018-08-24T09:51:16+08:00 | 2018-08-24T13:54:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP610G | CPSM Final Test | PASS |  | 163.58 | -0.02 | -0.46 | 56.48 | 50.55 | 2.2 | 0.13 | 1 | 3 | -1 | -0.874 | 0.5106 | 4 | 1 |  |  | 0.16 | 0.072 | 0.0169 | 0.99999 |  |  |  |  | 133 |  | 0 | [183 empty] | 0
P | 2018-08-24T07:20:38+08:00 | 2018-08-24T14:22:18+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP611G | CPSM Final Test | FAIL | [55 empty] | 0
P | 2018-08-24T13:33:38+08:00 | 2018-08-24T14:54:35+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP614G | CPSM Final Test | FAIL | [5 empty] | 44.2
P | 2018-08-24T07:15:49+08:00 | 2018-08-24T14:53:26+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | H18FNP609G | CPSM Final Test | FAIL
P | 2018-08-24T11:04:37+08:00 | 2018-08-24T15:00:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | H18FNP609G | CPSM Final Test | PASS |  | 165.02 | -0.02 | -0.54 | 56.47 | 50.12 | 1.21 | 0.13 | 1 | -3 | 0 | -1.0076 | 0.537 | 5.9 | -2 |  |  | 0.1138 | 0.0902 | -0.0415 | 1 |  |  |  |  | 133 |  | 0 | [184 empty] | 0
P | 2018-08-24T10:51:28+08:00 | 2018-08-24T15:27:52+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP611G | CPSM Final Test | PASS |  | 162.5 | -0.02 | 0.61 | 56.42 | 53.59 | 2.05 | 0.13 | 0 | 3 | 0 | -0.9198 | 0.51 | 4.5 | 56 |  |  | 0.1089 | 0.2381 | 0.006 | 1 |  |  |  |  | 133 |  | 0 | [185 empty] | 0
P | 2018-08-24T12:32:26+08:00 | 2018-08-24T15:52:54+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | H18FNP612G | CPSM Final Test | PASS |  | 166.82 | 0.17 | -0.54 | 56.56 | 51.56 | 0.91 | 0.13 | 1 | -4 | -1 | -0.8699 | 0.528 | 8 | 49 |  |  | 0.1466 | 0.4093 | 0.039 | 0.99975 |  |  |  |  | 133 |  | 0 | [186 empty] | 0
P | 2018-08-24T14:11:27+08:00 | 2018-08-24T16:58:47+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | H18FNP613G | CPSM Final Test | PASS |  | 164.21 | 0.65 | -0.84 | 56.4 | 50.21 | 1.44 | 0.13 | 1 | 3 | 0 | -0.9297 | 0.511 | 2.4 | 3 |  |  | 0.1474 | 0.0749 | -0.0487 | 0.99999 |  |  |  |  | 133 |  | 0 | [76 empty] | 0
P | 2018-08-24T14:14:50+08:00 | 2018-08-24T17:21:49+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | H18FNP615G | CPSM Final Test | PASS |  | 164.3 | -0.02 | -0.77 | 56.5 | 49.45 | 1.36 | 0.13 | -1 | -3 | 0 | -0.852 | 0.497 | 5.9 | 33 |  |  | 0.1862 | 0.2095 | 0.0334 | 1 |  |  |  |  | 133 |  | 0.01208 | [187 empty] | 0
P | 2018-08-24T14:14:06+08:00 | 2018-08-24T17:38:59+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | H18FNP614G | CPSM Final Test | PASS |  | 158.9 | -0.3 | -0.84 | 55.76 | 48.18 | 1.9 | 0.13 | -2 | -5 | 0 | -0.9099 | 0.5028 | 6.9 | -14 |  |  | 0.1902 | 0.2124 | 0.0049 | 0.99998 |  |  |  |  | 133 |  | 0 | [180 empty] | 0
P | 2018-08-24T14:51:01+08:00 | 2018-08-24T20:17:11+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP875G | CPSM Final Test | FAIL
P | 2018-08-24T17:13:39+08:00 | 2018-08-24T21:18:09+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | H18FNP617G | CPSM Final Test | PASS |  | 157.37 | -0.02 | -0.92 | 56.43 | 54.78 | 1.44 | 0.13 | 0 | 4 | 0 | -0.9378 | 0.5296 | 6.1 | -7 |  |  | 0.0445 | 0.1219 | 0.0163 | 0.99998 |  |  |  |  | 133 |  | 0 | [183 empty] | 0
P | 2018-08-24T18:36:44+08:00 | 2018-08-24T21:27:07+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP873G | CPSM Final Test | PASS | 129 | 163.49 | 0.27 | -0.46 | 56.61 | 49.28 | 1.51 | 0.13 | -1 | -1 | 0 | -0.9358 | 0.5094 | 8.6 | 11 |  |  | 0.107 | 0.3884 | 0.0081 | 0.99995 | [21 empty] | 0 | [173 empty] | 0
P | 2018-08-24T17:34:06+08:00 | 2018-08-24T21:22:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP875G | CPSM Final Test | PASS | 129 | 169.16 | -0.02 | -0.84 | 56.51 | 51.56 | 1.74 | 0.13 | 0 | -2 | 0 | -0.9398 | 0.509 | 8.9 | 7 |  |  | 0.0932 | 0.2186 | 0.0269 | 0.99999 | [21 empty] | 0 | [174 empty] | 0
P | 2018-08-24T18:32:40+08:00 | 2018-08-24T21:22:32+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP874G | CPSM Final Test | PASS | 129 | 161.33 | 0.17 | -0.46 | 56.7 | 53.85 | 1.9 | 0.13 | 0 | -2 | -1 | -0.9798 | 0.5438 | 2.4 | -7 |  |  | 0.2367 | 0.3188 | -0.0126 | 1 | [21 empty] | 0 | [117 empty] | 0
P | 2018-08-24T02:34:08+08:00 | 2018-08-25T08:27:44+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP616G | CPSM Final Test | FAIL
P | 2018-08-24T17:30:56+08:00 | 2018-08-25T08:28:26+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | H18FNP616G | CPSM Final Test | PASS |  | 165.83 | 0.27 | 0.38 | 56.52 | 50.63 | 1.67 | 0.13 | 0 | 3 | 0 | -0.9398 | 0.509 | 3.7 | -2 |  |  | 0.0411 | 0.1922 | -0.0155 | 0.99999 |  |  |  |  | 133 |  | 0 | [185 empty] | 0
P | 2018-08-25T10:30:51+08:00 | 2018-08-25T11:04:35+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP876G | CPSM Final Test | FAIL | [201 empty] | FALSE
P | 2018-08-25T10:13:39+08:00 | 2018-08-25T11:21:47+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP876G | CPSM Final Test | FAIL | [5 empty] | 43.86
P | 2018-08-25T10:17:48+08:00 | 2018-08-25T13:05:03+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP880G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-08-25T12:42:35+08:00 | 2018-08-25T13:56:12+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP880G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-08-25T09:21:21+08:00 | 2018-08-25T14:10:17+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP879G | CPSM Final Test | FAIL | [218 empty] | 130399
P | 2018-08-25T13:49:06+08:00 | 2018-08-25T14:33:30+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP882G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-08-25T10:53:49+08:00 | 2018-08-25T14:43:48+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP878G | CPSM Final Test | PASS | 129 | 157.1 | -0.02 | -0.61 | 56.7 | 54.86 | 1.36 | 0.13 | 0 | -3 | 0 | -0.8459 | 0.5402 | 4.1 | 53 |  |  | 0.0534 | 0.1967 | -0.0614 | 0.99999 | [21 empty] | 0 | [176 empty] | 0
P | 2018-08-25T10:50:58+08:00 | 2018-08-25T14:46:13+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP877G | CPSM Final Test | PASS | 129 | 159.35 | -0.02 | -0.31 | 56.73 | 52.41 | 2.05 | 0.13 | 0 | -2 | 0 | -0.844 | 0.535 | 9.2 | 10 |  |  | 0.1277 | 0.2197 | 0.081 | 0.99992 | [21 empty] | 0 | [177 empty] | 0
P | 2018-08-25T10:47:43+08:00 | 2018-08-25T14:47:21+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP876G | CPSM Final Test | PASS | 129 | 166.73 | -0.02 | -0.54 | 56.4 | 49.11 | 1.74 | 0.13 | 0 | 2 | -1 | -0.8859 | 0.534 | 4.7 | 24 |  |  | 0.087 | 0.127 | 0.0064 | 0.99999 | [21 empty] | 0 | [178 empty] | 0
P | 2018-08-25T12:23:27+08:00 | 2018-08-25T16:21:51+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP881G | CPSM Final Test | FAIL
P | 2018-08-25T13:19:24+08:00 | 2018-08-25T16:26:44+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP880G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -1.15 | 56.48 | 51.82 | 1.36 | 0.13 | 0 | -2 | 0 | -0.9139 | 0.486 | 6.4 | 44 |  |  | 0.1131 | 0.2108 | -0.0075 | 0.99999 | [21 empty] | 0 | [179 empty] | 0
P | 2018-08-25T16:15:03+08:00 | 2018-08-25T17:12:33+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP883G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-08-25T14:11:18+08:00 | 2018-08-25T17:27:14+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP882G | CPSM Final Test | PASS | 129 | 156.74 | -0.02 | -0.69 | 56.56 | 48.6 | 1.74 | 0.13 | -1 | 1 | 1 | -0.9598 | 0.513 | 3.9 | 50 |  |  | 0.2126 | 0.401 | 0.0424 | 0.99995 | [21 empty] | 0 | [180 empty] | 0
P | 2018-08-25T14:22:39+08:00 | 2018-08-25T18:06:29+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP881G | CPSM Final Test | PASS | 129 | 159.71 | 0.27 | -0.84 | 56.34 | 52.24 | 1.59 | 0.13 | 1 | -2 | 0 | -0.9378 | 0.545 | 6.8 | 0 |  |  | 0.0503 | 0.0555 | 0.0085 | 0.99999 | [21 empty] | 0 | [181 empty] | 0
P | 2018-08-25T14:27:21+08:00 | 2018-08-25T19:00:15+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP883G | CPSM Final Test | FAIL
P | 2018-08-25T14:13:57+08:00 | 2018-08-25T19:50:56+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP884G | CPSM Final Test | FAIL
P | 2018-08-25T16:50:45+08:00 | 2018-08-25T20:17:52+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP886G | CPSM Final Test | PASS | 129 | 160.7 | -0.02 | 0.07 | 56.8 | 50.63 | 1.97 | 0.13 | -1 | -4 | 0 | -0.9378 | 0.5054 | 4 | 77 |  |  | 0.0749 | 0.1783 | 0.0046 | 1 | [21 empty] | 0 | [182 empty] | 0
P | 2018-08-25T17:00:09+08:00 | 2018-08-25T20:25:58+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP885G | CPSM Final Test | PASS | 129 | 164.39 | -0.02 | -0.92 | 56.32 | 48.94 | 1.44 | 0.13 | -1 | 1 | 1 | -0.9459 | 0.5228 | 6.3 | 10 |  |  | 0.1025 | 0.0459 | 0.0096 | 0.99999 | [21 empty] | 0 | [183 empty] | 0
P | 2018-08-25T16:43:48+08:00 | 2018-08-25T20:44:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP883G | CPSM Final Test | PASS | 129 | 163.31 | -0.02 | -0.54 | 56.65 | 52.07 | 1.9 | 0.13 | 1 | 1 | 0 | -0.9398 | 0.5358 | 2.9 | 1 |  |  | 0.203 | 0.1734 | -0.0336 | 0.99999 | [21 empty] | 0 | [184 empty] | 0
P | 2018-08-25T17:02:27+08:00 | 2018-08-25T20:56:32+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP884G | CPSM Final Test | PASS | 129 | 161.06 | -0.02 | 0 | 56.71 | 54.53 | 1.51 | 0.13 | 0 | -1 | 0 | -0.9516 | 0.522 | 3.4 | 37 |  |  | 0.1337 | 0.1386 | 0.0414 | 0.99998 | [21 empty] | 0 | [185 empty] | 0
P | 2018-08-25T07:35:08+08:00 | 2018-08-26T08:52:07+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP890G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-08-26T09:04:40+08:00 | 2018-08-26T10:05:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP891G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-08-26T09:13:58+08:00 | 2018-08-26T10:11:40+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP887G | CPSM Final Test | FAIL | [5 empty] | 46.06
P | 2018-08-26T08:17:22+08:00 | 2018-08-26T11:00:52+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP888G | CPSM Final Test | FAIL
P | 2018-08-25T20:13:38+08:00 | 2018-08-26T11:44:43+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP890G | CPSM Final Test | PASS | 129 | 159.89 | 0.17 | -0.38 | 56.62 | 48.85 | 1.51 | 0.13 | -1 | 2 | -1 | -0.9258 | 0.538 | 8.5 | 22 |  |  | 0.0803 | 0.103 | -0.0118 | 1 | [21 empty] | 0.00673 | [139 empty] | 0
P | 2018-08-25T20:10:21+08:00 | 2018-08-26T11:51:16+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP889G | CPSM Final Test | PASS | 129 | 160.43 | -0.02 | -0.46 | 56.82 | 49.45 | 1.9 | 0.13 | 1 | 3 | 0 | -0.868 | 0.5268 | 5 | 10 |  |  | 0.1537 | 0.1873 | 0.0875 | 0.99994 | [21 empty] | 0.0031 | [186 empty] | 0
P | 2018-08-26T09:31:47+08:00 | 2018-08-26T12:35:17+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP892G | CPSM Final Test | PASS | 129 | 162.68 | 0.17 | -0.61 | 56.41 | 52.92 | 1.44 | 0.13 | 1 | -1 | 0 | -0.856 | 0.5118 | 5.5 | 59 |  |  | 0.1246 | 0.186 | -0.0201 | 1 | [21 empty] | 0 | [187 empty] | 0
P | 2018-08-26T09:34:52+08:00 | 2018-08-26T12:44:30+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP891G | CPSM Final Test | PASS | 129 | 162.68 | 0.08 | 0.3 | 56.55 | 48.34 | 1.29 | 0.13 | 0 | 3 | 0 | -0.8261 | 0.542 | 1.3 | 58 |  |  | 0.136 | 0.2394 | 0.034 | 0.99997 | [21 empty] | 0 | [185 empty] | 0
P | 2018-08-26T09:39:07+08:00 | 2018-08-26T12:45:32+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP888G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.31 | 56.54 | 48.09 | 1.44 | 0.13 | 1 | -4 | -1 | -0.82 | 0.5098 | 7.3 | -13 |  |  | 0.0969 | 0.264 | -0.0105 | 1 | [21 empty] | 0.03914 | [188 empty] | 0
P | 2018-08-26T09:42:49+08:00 | 2018-08-26T12:46:22+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP887G | CPSM Final Test | PASS | 129 | 167.81 | 0.37 | -0.77 | 56.23 | 48.51 | 1.29 | 0.13 | -2 | -3 | 0 | -0.8799 | 0.5182 | -1.4 | -7 |  |  | 0.0996 | 0.1089 | 0.0068 | 1 | [21 empty] | 0 | [189 empty] | 0
P | 2018-08-26T07:55:01+08:00 | 2018-08-26T16:39:44+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP894G | CPSM Final Test | FAIL
P | 2018-08-26T14:10:05+08:00 | 2018-08-26T16:41:12+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP896G | CPSM Final Test | FAIL | [13 empty] | 19.5
P | 2018-08-26T14:00:22+08:00 | 2018-08-26T16:45:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP897G | CPSM Final Test | FAIL
P | 2018-08-26T07:47:41+08:00 | 2018-08-26T16:47:04+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP894G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-08-26T15:29:19+08:00 | 2018-08-26T18:11:20+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP893G | CPSM Final Test | PASS | 129 | 161.6 | -0.02 | -0.23 | 56.44 | 53.76 | 1.51 | 0.13 | 1 | 3 | 0 | -0.9339 | 0.532 | 3.1 | 9 |  |  | 0.1117 | 0.1352 | 0.0373 | 1 | [21 empty] | 0 | [190 empty] | 0
P | 2018-08-26T15:19:43+08:00 | 2018-08-26T18:13:38+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP895G | CPSM Final Test | PASS | 129 | 168.89 | 0.27 | -0.61 | 56.55 | 49.11 | 2.28 | 0.13 | 2 | -3 | 0 | -1.0137 | 0.539 | 3.8 | 4 |  |  | 0.1329 | 0.1091 | -0.027 | 0.99999 | [21 empty] | 0 | [191 empty] | 0
P | 2018-08-26T15:22:47+08:00 | 2018-08-26T18:29:55+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP897G | CPSM Final Test | PASS | 129 | 158.9 | -0.02 | -0.99 | 56.18 | 49.78 | 1.51 | 0.13 | -1 | 2 | 0 | -0.8899 | 0.5064 | 2.8 | -6 |  |  | 0.0995 | 0.2062 | 0.0093 | 0.99998 | [21 empty] | 0 | [192 empty] | 0
P | 2018-08-26T15:25:39+08:00 | 2018-08-26T18:41:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP896G | CPSM Final Test | PASS | 129 | 165.56 | 0.08 | -0.61 | 56.31 | 48.85 | 1.74 | 0.13 | 1 | -2 | 0 | -0.886 | 0.539 | 6.3 | -9 |  |  | 0.1117 | 0.3535 | 0.0028 | 0.99993 | [21 empty] | 0 | [193 empty] | 0
P | 2018-08-27T09:46:43+08:00 | 2018-08-27T10:51:02+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP898G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-08-27T09:36:00+08:00 | 2018-08-27T10:55:40+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP900G | CPSM Final Test | FAIL | [5 empty] | 46.23
P | 2018-08-25T12:58:53+08:00 | 2018-08-27T11:35:52+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP894G | CPSM Final Test | FAIL
P | 2018-08-26T12:17:23+08:00 | 2018-08-27T12:38:54+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP894G | CPSM Final Test | PASS | 129 | 165.29 | -0.02 | 0.15 | 56.43 | 49.19 | 1.9 | 0.13 | 0 | 1 | 1 | -0.9278 | 0.5182 | 3.6 | 41 |  |  | 0.3796 | 0.3086 | 0.0768 | 0.99987 | [21 empty] | 0.00472 | [194 empty] | 0
P | 2018-08-27T10:12:07+08:00 | 2018-08-27T13:04:17+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP899G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.54 | 56.71 | 49.11 | 0.91 | 0.13 | 0 | 1 | 1 | -0.9039 | 0.5236 | 5.3 | 1 |  |  | 0.0966 | 0.0636 | 0.0088 | 0.99996 | [21 empty] | 0 | [195 empty] | 0
P | 2018-08-27T10:18:52+08:00 | 2018-08-27T13:33:06+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP898G | CPSM Final Test | PASS | 129 | 160.16 | -0.21 | -0.38 | 56.71 | 48.94 | 1.97 | 0.13 | 1 | 3 | 0 | -0.9258 | 0.5092 | 3.3 | 0 |  |  | 0.1347 | 0.0972 | -0.0009 | 1 | [21 empty] | 0.02206 | [85 empty] | 0
P | 2018-08-27T10:54:42+08:00 | 2018-08-27T13:42:48+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP901G | CPSM Final Test | PASS | 129 | 152.51 | 0.46 | -0.61 | 56.53 | 49.95 | 1.36 | 0.13 | 1 | -1 | 0 | -0.9359 | 0.506 | 5.2 | 2 |  |  | 0.2027 | 0.0683 | -0.0286 | 1 | [21 empty] | 0 | [196 empty] | 0
P | 2018-08-27T11:12:55+08:00 | 2018-08-27T13:54:40+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP900G | CPSM Final Test | PASS | 129 | 159.44 | 0.17 | 0.68 | 56.24 | 48.51 | 1.9 | 0.13 | 1 | 3 | 0 | -0.832 | 0.484 | 4.7 | 49 |  |  | 0.0668 | 0.3177 | 0.0339 | 0.99996 | [21 empty] | 0.01137 | [197 empty] | 0
P | 2018-08-23T06:22:29+08:00 | 2018-08-27T17:09:09+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP879G | CPSM Final Test | FAIL | [42 empty] | 0.13005
P | 2018-08-25T11:45:49+08:00 | 2018-08-27T18:12:23+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP879G | CPSM Final Test | PASS | 129 | 162.68 | -0.02 | -0.54 | 56.63 | 48.77 | 2.12 | 0.13 | -1 | 2 | 1 | -0.8919 | 0.5414 | 3.2 | 36 |  |  | 0.0954 | 0.132 | 0.0013 | 0.99998 | [21 empty] | 0 | [175 empty] | 0
P | 2018-08-27T15:36:18+08:00 | 2018-08-27T18:21:40+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP904G | CPSM Final Test | PASS | 129 | 159.35 | -0.02 | -0.77 | 56.21 | 48.6 | 1.97 | 0.13 | -1 | -1 | 0 | -0.9159 | 0.5032 | 4.1 | 9 |  |  | 0.0346 | 0.2574 | -0.0072 | 1 | [21 empty] | 0 | [198 empty] | 0
P | 2018-08-27T15:39:12+08:00 | 2018-08-27T18:24:30+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP903G | CPSM Final Test | PASS | 129 | 162.14 | -0.21 | -0.69 | 56.31 | 48.34 | 1.51 | 0.13 | 1 | 4 | 0 | -0.9158 | 0.5088 | 7.3 | -15 |  |  | 0.1339 | 0.2287 | 0.0421 | 0.99997 | [21 empty] | 0.01409 | [97 empty] | 0
P | 2018-08-27T15:42:51+08:00 | 2018-08-27T18:25:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP902G | CPSM Final Test | PASS | 129 | 157.55 | -0.02 | -0.61 | 56.32 | 50.29 | 1.67 | 0.13 | 0 | 3 | 0 | -0.9678 | 0.5382 | 6.3 | -2 |  |  | 0.0498 | 0.0346 | 0.0204 | 1 | [21 empty] | 0 | [199 empty] | 0
P | 2018-08-27T17:27:51+08:00 | 2018-08-27T19:43:04+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP907G | CPSM Final Test | FAIL | [13 empty] | 20.5
P | 2018-08-27T18:27:40+08:00 | 2018-08-27T21:24:50+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP905G | CPSM Final Test | PASS | 129 | 158.27 | 0.17 | -0.61 | 56.72 | 53.26 | 1.67 | 0.13 | 0 | -1 | 0 | -1.0117 | 0.5272 | 2.8 | -7 |  |  | 0.2009 | 0.2265 | -0.0022 | 1 | [21 empty] | 0 | [200 empty] | 0
P | 2018-08-27T18:31:25+08:00 | 2018-08-27T21:27:32+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP906G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.46 | 56.53 | 49.7 | 1.59 | 0.13 | 0 | -1 | 0 | -0.9837 | 0.5248 | 10.3 | 17 |  |  | 0.0824 | 0.126 | -0.0065 | 0.99999 | [21 empty] | 0 | [201 empty] | 0
P | 2018-08-27T18:35:28+08:00 | 2018-08-27T21:43:56+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | H18FTP907G | CPSM Final Test | PASS | 129 | 156.83 | -0.21 | -0.77 | 56.32 | 48.51 | 1.9 | 0.13 | 0 | -1 | 1 | -0.9398 | 0.539 | 6.6 | -37 |  |  | 0.1097 | 0.2158 | 0.0926 | 0.9999 | [21 empty] | 0 | [194 empty] | 0
P | 2018-08-28T09:38:48+08:00 | 2018-08-28T12:37:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP909G | CPSM Final Test | PASS | 129 | 163.31 | -0.02 | -0.38 | 56.41 | 52.32 | 1.29 | 0.13 | -1 | 2 | 0 | -0.8141 | 0.507 | 3 | 59 |  |  | 0.0683 | 0.199 | 0.012 | 1 | [21 empty] | 0 | [202 empty] | 0
P | 2018-08-28T09:49:02+08:00 | 2018-08-28T12:40:00+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP908G | CPSM Final Test | PASS | 129 | 162.5 | 0.17 | 0.45 | 56.63 | 49.19 | 1.44 | 0.13 | -1 | -1 | 1 | -0.8759 | 0.5152 | 6 | 28 |  |  | 0.1259 | 0.2147 | 0.0265 | 0.99999 | [21 empty] | 0 | [203 empty] | 0
P | 2018-08-28T09:42:18+08:00 | 2018-08-28T12:40:32+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP910G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -0.77 | 56.33 | 50.72 | 1.13 | 0.13 | 1 | 1 | 1 | -0.858 | 0.5128 | 0.2 | 30 |  |  | 0.1751 | 0.0799 | 0.0576 | 0.99998 | [21 empty] | 0 | [204 empty] | 0
P | 2018-08-28T13:58:02+08:00 | 2018-08-28T15:48:24+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP912G | CPSM Final Test | FAIL
P | 2018-08-28T13:11:47+08:00 | 2018-08-28T16:39:05+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP911G | CPSM Final Test | FAIL
P | 2018-08-28T12:45:56+08:00 | 2018-08-28T17:00:30+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP912G | CPSM Final Test | FAIL
P | 2018-08-28T14:46:07+08:00 | 2018-08-28T17:56:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP913G | CPSM Final Test | PASS | 129 | 159.44 | -0.02 | 0.15 | 56.49 | 49.61 | 1.36 | 0.13 | 1 | 3 | 0 | -0.9438 | 0.5038 | 3.3 | 14 |  |  | 0.0685 | 0.1511 | -0.0288 | 0.99999 | [21 empty] | 0 | [205 empty] | 0
P | 2018-08-28T14:43:10+08:00 | 2018-08-28T17:58:22+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP914G | CPSM Final Test | PASS | 129 | 157.1 | -0.3 | -0.77 | 56.34 | 48.85 | 1.67 | 0.13 | -1 | 1 | -1 | -0.872 | 0.5042 | 1.6 | 70 |  |  | 0.0516 | 0.2006 | 0.0054 | 0.99999 | [21 empty] | 0 | [206 empty] | 0
P | 2018-08-28T14:55:26+08:00 | 2018-08-28T18:23:13+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | H18FTP911G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.38 | 56.32 | 50.72 | 1.44 | 0.13 | 1 | -2 | 0 | -0.866 | 0.5222 | 1.2 | 4 |  |  | 0.1354 | 0.3298 | 0.068 | 0.99996 | [21 empty] | 0 | [175 empty] | 0
P | 2018-08-28T18:02:38+08:00 | 2018-08-28T18:44:05+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP915G | CPSM Final Test | FAIL | [5 empty] | 44.79
P | 2018-08-28T14:53:13+08:00 | 2018-08-28T18:50:38+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | H18FTP912G | CPSM Final Test | PASS | 129 | 161.69 | 0.17 | -0.54 | 56.4 | 49.7 | 1.51 | 0.13 | -2 | -4 | 0 | -0.9319 | 0.537 | 3.8 | 16 |  |  | 0.0796 | 0.1021 | -0.0128 | 1 | [21 empty] | 0 | [207 empty] | 0
P | 2018-08-28T17:29:40+08:00 | 2018-08-28T19:17:03+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP915G | CPSM Final Test | FAIL | [5 empty] | 45.04
P | 2018-08-28T17:43:09+08:00 | 2018-08-28T20:18:20+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP918G | CPSM Final Test | FAIL
P | 2018-08-28T18:30:02+08:00 | 2018-08-28T21:22:10+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | H18FTP917G | CPSM Final Test | PASS | 129 | 165.38 | 0.17 | -0.92 | 56.19 | 48.85 | 1.44 | 0.13 | -1 | -2 | 0 | -0.9158 | 0.492 | 2.1 | 15 |  |  | 0.1224 | 0.1112 | 0.0171 | 1 | [21 empty] | 0 | [208 empty] | 0
P | 2018-08-28T04:50:54+08:00 | 2018-08-29T09:10:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP918G | CPSM Final Test | FAIL
P | 2018-08-28T05:49:23+08:00 | 2018-08-29T09:49:21+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP915G | CPSM Final Test | FAIL | [55 empty] | 0
P | 2018-08-28T02:42:20+08:00 | 2018-08-29T10:09:38+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP916G | CPSM Final Test | FAIL | [19 empty] | -0.8527
P | 2018-08-28T19:49:22+08:00 | 2018-08-29T10:54:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | H18FTP915G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -0.46 | 56.42 | 49.36 | 1.59 | 0.13 | -2 | -2 | 0 | -0.9058 | 0.511 | 12.7 | 111 |  |  | 0.0953 | 0.1465 | -0.0392 | 0.99998 | [21 empty] | 0 | [209 empty] | 0
P | 2018-08-28T19:00:44+08:00 | 2018-08-29T10:54:53+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | H18FTP918G | CPSM Final Test | PASS | 129 | 160.88 | 0.17 | -0.84 | 56.43 | 52.83 | 1.21 | 0.13 | 1 | -3 | 0 | -0.92 | 0.4966 | 7.9 | -6 |  |  | 0.0898 | 0.1194 | 0.0091 | 1 | [21 empty] | 0 | [210 empty] | 0
P | 2018-08-28T18:25:59+08:00 | 2018-08-29T11:12:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | H18FTP916G | CPSM Final Test | PASS | 129 | 163.31 | -0.02 | -1.07 | 56.12 | 48.09 | 1.51 | 0.13 | 0 | -1 | 0 | -0.9159 | 0.5138 | 4.3 | 15 |  |  | 0.0609 | 0.1669 | -0.0226 | 0.99999 | [21 empty] | 0 | [211 empty] | 0
P | 2018-08-30T08:58:30+08:00 | 2018-08-30T11:54:32+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | H18FNP619G | CPSM Final Test | PASS |  | 164.12 | 0.08 | -1.68 | 56.05 | 51.48 | 1.44 | 0.13 | 1 | -3 | 0 | -0.884 | 0.49 | 1.6 | -2 |  |  | 0.1227 | 0.3014 | 0.017 | 0.99999 |  |  |  |  | 133 |  | 0 | [224 empty] | 0
P | 2018-08-30T09:23:34+08:00 | 2018-08-30T12:18:24+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP620G | CPSM Final Test | PASS |  | 167.99 | -0.02 | -0.23 | 56.23 | 50.63 | 1.59 | 0.13 | -2 | 2 | 0 | -0.9318 | 0.4862 | 3.3 | -3 |  |  | 0.0823 | 0.2935 | -0.0252 | 0.99999 |  |  |  |  | 133 |  | 0.00808 | [227 empty] | 0
P | 2018-08-30T09:46:07+08:00 | 2018-08-30T12:49:25+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | H18FNP622G | CPSM Final Test | PASS |  | 161.51 | 0.17 | -1.22 | 56.4 | 48.68 | 1.21 | 0.13 | 1 | -2 | 0 | -0.8899 | 0.5384 | 4.6 | 76 |  |  | 0.155 | 0.2194 | -0.0068 | 0.99999 |  |  |  |  | 133 |  | 0 | [228 empty] | 0
P | 2018-08-30T07:03:53+08:00 | 2018-08-30T13:33:15+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP623G | CPSM Final Test | FAIL | [257 empty] | 0
P | 2018-08-30T10:55:07+08:00 | 2018-08-30T13:46:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | H18FNP624G | CPSM Final Test | PASS |  | 161.87 | -0.3 | -0.46 | 56.17 | 48.94 | 1.59 | 0.13 | -2 | 2 | 1 | -0.9018 | 0.548 | -0.9 | -3 |  |  | 0.219 | 0.4475 | -0.0344 | 0.99999 |  |  |  |  | 133 |  | 0 | [230 empty] | 0
P | 2018-08-30T11:14:46+08:00 | 2018-08-30T14:05:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP621G | CPSM Final Test | PASS |  | 157.73 | 0.27 | -1.3 | 56.13 | 49.45 | 1.44 | 0.13 | -1 | -3 | 0 | -0.8979 | 0.5258 | -4.1 | 6 |  |  | 0.3194 | 0.2398 | -0.0113 | 0.99997 |  |  |  |  | 133 |  | 0 | [53 empty] | 0
P | 2018-08-30T11:33:26+08:00 | 2018-08-30T14:16:19+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | H18FNP625G | CPSM Final Test | PASS |  | 173.48 | -0.11 | -1.07 | 56.41 | 55.8 | 1.67 | 0.13 | 1 | 1 | 1 | -0.7563 | 0.5364 | 6.6 | 32 |  |  | 0.206 | 0.3091 | 0.0864 | 0.99976 |  |  |  |  | 133 |  | 0.00249 | [231 empty] | 0
P | 2018-08-30T10:18:34+08:00 | 2018-08-30T14:38:47+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP623G | CPSM Final Test | PASS |  | 162.41 | -0.02 | -0.31 | 56.4 | 51.14 | 1.59 | 0.13 | 1 | -3 | 0 | -0.9358 | 0.5234 | 3.9 | 7 |  |  | 0.1452 | 0.293 | -0.0318 | 0.99999 |  |  |  |  | 133 |  | 0.01733 | [229 empty] | 0
P | 2018-08-30T10:50:58+08:00 | 2018-08-30T16:24:36+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | H18FNP626G | CPSM Final Test | FAIL | [260 empty] | 3396
P | 2018-08-30T11:22:40+08:00 | 2018-08-30T16:33:46+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP627G | CPSM Final Test | FAIL | [261 empty] | 0
P | 2018-08-30T13:37:47+08:00 | 2018-08-30T17:30:08+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | H18FNP626G | CPSM Final Test | PASS |  | 161.15 | -0.02 | -1.07 | 56.03 | 51.82 | 1.44 | 0.13 | -1 | -2 | 0 | -0.9359 | 0.5352 | -2.2 | -32 |  |  | 0.1168 | 0.1739 | -0.0164 | 1 |  |  |  |  | 133 |  | 0.00925 | [232 empty] | 0
P | 2018-08-30T13:58:13+08:00 | 2018-08-30T17:53:34+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP627G | CPSM Final Test | PASS |  | 163.04 | -0.02 | -0.92 | 56.23 | 51.39 | 1.82 | 0.13 | 1 | -3 | -1 | -0.9239 | 0.5348 | -0.6 | 0 |  |  | 0.1385 | 0.1031 | -0.005 | 1 |  |  |  |  | 133 |  | 0.02594 | [233 empty] | 0
P | 2018-08-30T15:32:14+08:00 | 2018-08-30T19:11:06+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP629G | CPSM Final Test | FAIL
P | 2018-08-30T17:47:56+08:00 | 2018-08-30T19:17:35+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | H18FNP618G | CPSM Final Test | FAIL | [5 empty] | 47.16
P | 2018-08-30T15:13:32+08:00 | 2018-08-30T19:29:48+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP629G | CPSM Final Test | FAIL | [13 empty] | -5.4
P | 2018-08-30T14:49:40+08:00 | 2018-08-30T19:53:40+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP629G | CPSM Final Test | FAIL | [13 empty] | -6.2
P | 2018-08-30T17:20:48+08:00 | 2018-08-30T21:04:01+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | H18FNP628G | CPSM Final Test | PASS |  | 165.11 | -0.11 | -1.15 | 56.33 | 48.85 | 1.74 | 0.13 | -1 | -2 | 0 | -0.842 | 0.4608 | 10 | 1 |  |  | 0.1618 | 0.4819 | 0.0451 | 0.99989 |  |  |  |  | 133 |  | 0 | [234 empty] | 0
P | 2018-08-30T18:32:46+08:00 | 2018-08-30T22:01:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | H18FNP618G | CPSM Final Test | PASS |  | 160.34 | 0.08 | -0.99 | 56.2 | 49.95 | 1.36 | 0.13 | -2 | -2 | -1 | -0.9836 | 0.5086 | -0.6 | 11 |  |  | 0.2541 | 0.195 | -0.0223 | 0.99996 |  |  |  |  | 133 |  | 0 | [235 empty] | 0
P | 2018-08-30T17:21:40+08:00 | 2018-08-30T21:57:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP629G | CPSM Final Test | PASS |  | 164.75 | 0.08 | -0.99 | 56.23 | 52.58 | 1.59 | 0.13 | 1 | 3 | 1 | -0.9298 | 0.5006 | 1.8 | 14 |  |  | 0.1389 | 0.0837 | -0.0115 | 1 |  |  |  |  | 133 |  | 0 | [100 empty] | 0
P | 2018-08-30T03:18:44+08:00 | 2018-08-31T09:34:57+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP630G | CPSM Final Test | FAIL
P | 2018-08-30T01:55:26+08:00 | 2018-08-31T10:58:15+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP630G | CPSM Final Test | FAIL | [257 empty] | 0
P | 2018-08-31T09:19:41+08:00 | 2018-08-31T11:57:52+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP633G | CPSM Final Test | FAIL | [13 empty] | -7.6
P | 2018-08-30T18:26:50+08:00 | 2018-08-31T12:03:51+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP630G | CPSM Final Test | PASS |  | 160.7 | 0.08 | -1.22 | 56.1 | 50.72 | 0.91 | 0.13 | -2 | -2 | 0 | -0.9638 | 0.549 | 1 | 24 |  |  | 0.1769 | 0.4531 | -0.0458 | 1 |  |  |  |  | 133 |  | 0.00928 | [229 empty] | 0
P | 2018-08-31T08:14:58+08:00 | 2018-08-31T12:09:20+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP631G | CPSM Final Test | FAIL
P | 2018-08-31T10:37:42+08:00 | 2018-08-31T13:35:33+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | H18FNP632G | CPSM Final Test | PASS |  | 155.21 | 0.08 | -1.07 | 56.46 | 50.29 | 1.74 | 0.13 | -1 | 1 | 0 | -0.8241 | 0.4996 | 5.2 | -8 |  |  | 0.16 | 0.3147 | 0.0642 | 0.99981 |  |  |  |  | 133 |  | 0 | [236 empty] | 0
P | 2018-08-31T10:12:09+08:00 | 2018-08-31T13:57:22+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP631G | CPSM Final Test | PASS |  | 165.92 | -0.02 | -0.92 | 56.16 | 48.43 | 1.36 | 0.13 | -1 | 3 | -1 | -0.8181 | 0.5054 | 4.1 | 10 |  |  | 0.1463 | 0.4382 | -0.0479 | 0.99999 |  |  |  |  | 133 |  | 0.03924 | [233 empty] | 0
P | 2018-08-31T10:38:46+08:00 | 2018-08-31T14:01:20+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | H18FNP633G | CPSM Final Test | PASS |  | 160.25 | 0.08 | -0.99 | 56.71 | 48.34 | 1.59 | 0.13 | -1 | -2 | 0 | -0.858 | 0.4936 | 7.9 | -5 |  |  | 0.1195 | 0.0726 | -0.0354 | 1 |  |  |  |  | 133 |  | 0.012 | [237 empty] | 0
P | 2018-08-31T11:18:43+08:00 | 2018-08-31T17:01:08+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP637G | CPSM Final Test | FAIL | [82 empty] | 1360
P | 2018-08-31T11:21:47+08:00 | 2018-08-31T17:15:21+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP636G | CPSM Final Test | FAIL | [14 empty] | 770
P | 2018-08-31T10:48:29+08:00 | 2018-08-31T17:31:22+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP637G | CPSM Final Test | FAIL | [82 empty] | 1521
P | 2018-08-31T15:02:06+08:00 | 2018-08-31T18:07:21+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | H18FNP638G | CPSM Final Test | PASS |  | 161.87 | 0.17 | -0.31 | 56.13 | 52.41 | 1.44 | 0.13 | 1 | 2 | 0 | -0.9637 | 0.524 | -3.3 | 61 |  |  | 0.1267 | 0.2032 | 0.0073 | 0.99997 |  |  |  |  | 133 |  | 0 | [238 empty] | 0
P | 2018-08-31T15:25:54+08:00 | 2018-08-31T18:20:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | H18FNP634G | CPSM Final Test | PASS |  | 167.99 | 0.46 | -1.07 | 56.4 | 49.53 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9698 | 0.4932 | 2.9 | -27 |  |  | 0.074 | 0.2359 | -0.0731 | 1 |  |  |  |  | 133 |  | 0 | [141 empty] | 0
P | 2018-08-31T14:09:56+08:00 | 2018-08-31T18:37:00+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | H18FNP637G | CPSM Final Test | PASS |  | 163.58 | -0.02 | -1.3 | 56.01 | 49.87 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9717 | 0.523 | 7.3 | 1 |  |  | 0.0508 | 0.1525 | -0.0347 | 0.99999 |  |  |  |  | 133 |  | 0 | [54 empty] | 0
P | 2018-08-31T14:18:34+08:00 | 2018-08-31T19:03:27+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | H18FNP636G | CPSM Final Test | PASS |  | 165.38 | -0.81 | -0.69 | 56.5 | 48.68 | 1.59 | 0.13 | -2 | 3 | 0 | -0.9579 | 0.4714 | 10.8 | 19 |  |  | 0.0229 | 0.2408 | 0.002 | 0.99999 |  |  |  |  | 133 |  | 0.0081 | [229 empty] | 0
P | 2018-08-31T18:34:35+08:00 | 2018-08-31T19:13:01+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP923G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-08-31T13:36:58+08:00 | 2018-08-31T19:19:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP919G | CPSM Final Test | FAIL | [91 empty] | 0
P | 2018-08-31T16:29:30+08:00 | 2018-08-31T19:24:26+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | H18FNP635G | CPSM Final Test | PASS |  | 167.09 | -0.02 | -0.08 | 56.57 | 48.09 | 1.59 | 0.13 | 1 | 3 | 1 | -0.9178 | 0.5412 | 1.3 | -3 |  |  | 0.0962 | 0.1182 | 0.0063 | 0.99999 |  |  |  |  | 133 |  | 0 | [232 empty] | 0
P | 2018-08-31T16:28:03+08:00 | 2018-08-31T20:24:45+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP919G | CPSM Final Test | PASS | 129 | 166.01 | -0.49 | -0.99 | 56.56 | 49.78 | 1.67 | 0.13 | 1 | 3 | 1 | -0.9558 | 0.537 | -0.7 | 1 |  |  | 0.1464 | 0.1669 | 0.0327 | 0.99997 | [21 empty] | 0 | [48 empty] | 0
P | 2018-08-31T17:41:42+08:00 | 2018-08-31T21:08:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP922G | CPSM Final Test | PASS | 129 | 157.1 | -0.21 | -1.38 | 56.2 | 50.97 | 0.83 | 0.13 | 1 | -2 | 0 | -0.8479 | 0.5154 | -3.7 | 62 |  |  | 0.1862 | 0.2526 | 0.0617 | 0.99994 | [21 empty] | 0 | [224 empty] | 0
P | 2018-08-31T18:46:18+08:00 | 2018-08-31T21:58:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP921G | CPSM Final Test | PASS | 129 | 160.16 | 0.37 | -1.22 | 56.31 | 50.72 | 0.98 | 0.13 | -1 | -1 | 0 | -0.9278 | 0.518 | 7.2 | 61 |  |  | 0.1049 | 0.2864 | 0.0579 | 0.99998 | [21 empty] | 0 | [225 empty] | 0
P | 2018-08-31T18:53:48+08:00 | 2018-08-31T21:54:47+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP923G | CPSM Final Test | PASS | 129 | 161.06 | 0.46 | -1.07 | 56.3 | 50.04 | 1.36 | 0.13 | 1 | -2 | 1 | -0.8062 | 0.532 | 6.8 | 20 |  |  | 0.0386 | 0.1094 | 0.0014 | 1 | [21 empty] | 0 | [226 empty] | 0
P | 2018-08-30T23:51:29+08:00 | 2018-09-01T10:59:44+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | H18FTP920G | CPSM Final Test | FAIL | [42 empty] | 0.14692
P | 2018-08-31T17:25:36+08:00 | 2018-09-01T12:02:32+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | H18FTP920G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -1.3 | 56.18 | 52.07 | 1.21 | 0.13 | -1 | -3 | 0 | -0.8859 | 0.538 | -1.9 | -11 |  |  | 0.0242 | 0.093 | -0.0173 | 0.99999 | [21 empty] | 0 | [31 empty] | 0
P | 2018-09-01T11:22:21+08:00 | 2018-09-01T14:43:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP926G | CPSM Final Test | PASS | 129 | 163.76 | -0.02 | -1.22 | 56.17 | 51.05 | 1.13 | 0.13 | 1 | 6 | 1 | -0.7762 | 0.538 | 1.5 | -3 |  |  | 0.0532 | 0.2629 | 0.0157 | 1 | [21 empty] | 0 | [196 empty] | 0
P | 2018-09-01T11:31:47+08:00 | 2018-09-01T14:59:30+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | H18FTP924G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.69 | 56.42 | 56.22 | 1.9 | 0.13 | 0 | -4 | 0 | -0.8321 | 0.52 | 2.3 | 18 |  |  | 0.0829 | 0.1773 | 0.0019 | 1 | [21 empty] | 0 | [39 empty] | 0
P | 2018-09-01T12:58:23+08:00 | 2018-09-01T15:47:33+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP927G | CPSM Final Test | PASS | 129 | 162.41 | 0.17 | -1.38 | 56.25 | 51.39 | 1.13 | 0.13 | -2 | -2 | 0 | -0.9138 | 0.5178 | -2.3 | -10 |  |  | 0.082 | 0.1917 | 0.0112 | 0.99997 | [21 empty] | 0 | [48 empty] | 0
P | 2018-09-01T13:27:10+08:00 | 2018-09-01T16:26:14+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | H18FTP928G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -1.38 | 56.17 | 49.45 | 0.98 | 0.13 | 1 | 2 | 0 | -0.842 | 0.5388 | 0.2 | -5 |  |  | 0.3193 | 0.373 | -0.0358 | 1 | [21 empty] | 0 | [121 empty] | 0
P | 2018-09-01T13:35:19+08:00 | 2018-09-01T16:28:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP930G | CPSM Final Test | PASS | 129 | 160.07 | -0.02 | -0.84 | 56.46 | 50.63 | 0.68 | 0.13 | 0 | 1 | 0 | -0.9517 | 0.5394 | -1 | 18 |  |  | 0.1624 | 0.4326 | -0.0046 | 0.99998 | [21 empty] | 0.05276 | [84 empty] | 0
P | 2018-09-01T16:00:37+08:00 | 2018-09-01T18:51:20+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | H18FTP929G | CPSM Final Test | PASS | 129 | 156.74 | 0.08 | -1.15 | 56.46 | 49.28 | 1.29 | 0.13 | 0 | 1 | 1 | -0.9539 | 0.5368 | -1 | 3 |  |  | 0.1314 | 0.1421 | 0.0122 | 1 | [21 empty] | 0 | [227 empty] | 0
P | 2018-09-01T16:13:38+08:00 | 2018-09-01T19:05:27+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP931G | CPSM Final Test | PASS | 129 | 163.67 | 0.17 | -0.31 | 56.41 | 54.78 | 0.6 | 0.13 | 1 | -2 | -1 | -0.9717 | 0.5178 | 3.1 | 18 |  |  | 0.2552 | 0.0656 | 0.0944 | 0.99983 | [21 empty] | 0 | [153 empty] | 0
P | 2018-09-01T16:05:57+08:00 | 2018-09-01T19:30:47+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP935G | CPSM Final Test | FAIL
P | 2018-09-01T16:22:26+08:00 | 2018-09-01T19:31:21+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP934G | CPSM Final Test | FAIL | [14 empty] | 734
P | 2018-09-01T16:52:33+08:00 | 2018-09-01T20:11:36+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP932G | CPSM Final Test | PASS | 129 | 160.43 | 0.08 | -1.15 | 56.32 | 51.05 | 1.29 | 0.13 | 0 | -2 | 0 | -0.8759 | 0.5228 | 2.9 | -1 |  |  | 0.205 | 0.1691 | 0.0278 | 0.99998 | [21 empty] | 0 | [228 empty] | 0
P | 2018-09-01T16:56:01+08:00 | 2018-09-01T20:11:10+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | H18FTP933G | CPSM Final Test | PASS | 129 | 163.04 | 0.27 | -1.22 | 56.18 | 51.14 | 0.75 | 0.13 | 0 | 3 | 0 | -0.9079 | 0.5304 | 3.7 | -1 |  |  | 0.0679 | 0.2863 | -0.0421 | 0.99999 | [21 empty] | 0 | [229 empty] | 0
P | 2018-09-01T17:48:22+08:00 | 2018-09-01T21:14:47+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP935G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -0.69 | 56.52 | 51.56 | 1.29 | 0.13 | 1 | -4 | -1 | -0.904 | 0.522 | 6 | 18 |  |  | 0.1328 | 0.0893 | 0.0104 | 0.99997 | [21 empty] | 0.01893 | [230 empty] | 0
P | 2018-09-01T17:56:53+08:00 | 2018-09-01T21:16:03+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP934G | CPSM Final Test | PASS | 129 | 157.73 | 0.17 | -1.15 | 56.5 | 50.97 | 0.75 | 0.13 | 0 | -2 | 0 | -0.9159 | 0.5294 | 3.3 | -10 |  |  | 0.0597 | 0.1585 | 0.0104 | 1 | [21 empty] | 0 | [231 empty] | 0
P | 2018-09-02T09:22:25+08:00 | 2018-09-02T10:31:35+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP946G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-09-02T10:09:18+08:00 | 2018-09-02T11:10:17+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP950G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-09-02T07:28:44+08:00 | 2018-09-02T11:34:05+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP937G | CPSM Final Test | FAIL
P | 2018-09-02T06:26:50+08:00 | 2018-09-02T12:35:59+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP937G | CPSM Final Test | FAIL
P | 2018-09-02T09:50:42+08:00 | 2018-09-02T12:57:27+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP947G | CPSM Final Test | PASS | 129 | 167.45 | -0.02 | -0.99 | 56.33 | 51.22 | 1.44 | 0.13 | 0 | 2 | 1 | -0.8321 | 0.5456 | 4.6 | -13 |  |  | 0.0535 | 0.1218 | -0.0185 | 1 | [21 empty] | 0 | [232 empty] | 0
P | 2018-09-02T09:38:30+08:00 | 2018-09-02T12:58:36+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | H18FTP938G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | 0.07 | 56.49 | 49.78 | 1.51 | 0.13 | 1 | 3 | 0 | -0.8241 | 0.5124 | -0.9 | 63 |  |  | 0.0827 | 0.1808 | 0.0301 | 0.99998 | [21 empty] | 0 | [233 empty] | 0
P | 2018-09-02T09:57:00+08:00 | 2018-09-02T13:06:19+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP946G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.99 | 56.54 | 49.53 | 0.75 | 0.13 | 1 | -4 | 1 | -0.8899 | 0.546 | 3 | 16 |  |  | 0.2044 | 0.1288 | 0.044 | 0.99997 | [21 empty] | 0.01488 | [234 empty] | 0
P | 2018-09-02T10:10:40+08:00 | 2018-09-02T13:14:30+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP949G | CPSM Final Test | PASS | 129 | 157.37 | -0.02 | -0.99 | 56.23 | 50.55 | 1.44 | 0.13 | 0 | -2 | 0 | -0.8501 | 0.514 | 6.8 | 11 |  |  | 0.0738 | 0.24 | 0.0223 | 0.99997 | [21 empty] | 0 | [235 empty] | 0
P | 2018-09-02T10:12:33+08:00 | 2018-09-02T13:14:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | H18FTP948G | CPSM Final Test | PASS | 129 | 162.41 | 0.27 | -0.92 | 56.43 | 53.85 | 1.29 | 0.13 | 0 | 3 | 0 | -0.858 | 0.5048 | 1 | 29 |  |  | 0.0827 | 0.0191 | -0.0305 | 1 | [21 empty] | 0 | [236 empty] | 0
P | 2018-09-02T09:24:41+08:00 | 2018-09-02T13:17:10+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | H18FTP936G | CPSM Final Test | PASS | 129 | 157.73 | 0.17 | -1.15 | 56.43 | 56.05 | 1.06 | 0.13 | -1 | 3 | 0 | -0.7882 | 0.502 | 12.6 | 20 |  |  | 0.1221 | 0.5414 | -0.029 | 0.99999 | [21 empty] | 0 | [237 empty] | 0
P | 2018-09-02T10:39:47+08:00 | 2018-09-02T13:39:19+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP950G | CPSM Final Test | PASS | 129 | 164.75 | -0.11 | -1.07 | 56.31 | 49.7 | 0.91 | 0.13 | 1 | -4 | -1 | -0.86 | 0.5002 | -1.5 | 8 |  |  | 0.1813 | 0.3347 | 0.0573 | 0.99989 | [21 empty] | 0 | [238 empty] | 0
P | 2018-09-02T09:31:25+08:00 | 2018-09-02T13:41:41+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP937G | CPSM Final Test | PASS | 129 | 165.65 | -0.21 | -1.15 | 56.49 | 52.07 | 1.51 | 0.13 | 1 | 2 | 0 | -0.834 | 0.465 | -1.3 | 3 |  |  | 0.0979 | 0.2242 | 0.0392 | 0.99997 | [21 empty] | 0 | [239 empty] | 0
P | 2018-09-02T13:13:29+08:00 | 2018-09-02T16:08:08+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | H18FTP956G | CPSM Final Test | FAIL | [5 empty] | 46.91
P | 2018-09-02T14:53:31+08:00 | 2018-09-02T17:03:50+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP953G | CPSM Final Test | FAIL | [13 empty] | -13.3
P | 2018-09-02T14:01:54+08:00 | 2018-09-02T17:17:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | H18FTP954G | CPSM Final Test | FAIL
P | 2018-09-02T14:55:52+08:00 | 2018-09-02T18:30:24+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | H18FTP955G | CPSM Final Test | PASS | 129 | 166.19 | -0.3 | -0.77 | 56.42 | 54.36 | 1.21 | 0.13 | -1 | -1 | 0 | -0.864 | 0.5136 | 7.9 | 51 |  |  | 0.0587 | 0.3379 | -0.0514 | 1 | [21 empty] | 0 | [240 empty] | 0
P | 2018-09-02T15:13:21+08:00 | 2018-09-02T18:32:09+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP939G | CPSM Final Test | PASS | 129 | 169.7 | -0.21 | -1.07 | 56.42 | 56.13 | 1.13 | 0.13 | 0 | -2 | 1 | -0.876 | 0.5428 | -3.4 | 3 |  |  | 0.0864 | 0.3964 | -0.0138 | 1 | [21 empty] | 0 | [241 empty] | 0
P | 2018-09-02T14:48:28+08:00 | 2018-09-02T18:33:51+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP951G | CPSM Final Test | PASS | 129 | 160.34 | -0.02 | -0.77 | 56.33 | 50.21 | 1.29 | 0.13 | 1 | 2 | 0 | -0.892 | 0.4906 | 2 | 20 |  |  | 0.1166 | 0.0887 | -0.0209 | 0.99997 | [21 empty] | 0 | [239 empty] | 0
P | 2018-09-02T15:20:59+08:00 | 2018-09-02T18:41:33+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP952G | CPSM Final Test | PASS | 129 | 164.66 | 0.37 | -0.92 | 56.46 | 51.56 | 1.29 | 0.13 | 1 | 3 | -1 | -0.9039 | 0.5262 | 6.4 | 13 |  |  | 0.1554 | 0.1907 | 0.0262 | 0.99998 | [21 empty] | 0.01959 | [242 empty] | 0
P | 2018-09-02T14:40:48+08:00 | 2018-09-02T18:47:14+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | H18FTP956G | CPSM Final Test | PASS | 129 | 165.11 | 0.08 | -1.22 | 56.08 | 48.51 | 1.29 | 0.13 | -1 | -1 | 0 | -0.8441 | 0.5272 | 0.8 | -7 |  |  | 0.0953 | 0.2437 | -0.0133 | 0.99999 | [21 empty] | 0 | [243 empty] | 0
P | 2018-09-02T15:58:41+08:00 | 2018-09-02T19:03:46+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP953G | CPSM Final Test | PASS | 129 | 165.47 | -0.3 | -1.15 | 56.49 | 49.28 | 1.44 | 0.13 | -1 | 1 | 0 | -0.9358 | 0.531 | 12 | 3 |  |  | 0.122 | 0.1232 | 0.0428 | 0.99997 | [21 empty] | 0 | [244 empty] | 0
P | 2018-09-02T15:39:29+08:00 | 2018-09-02T19:01:22+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | H18FTP954G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -1.22 | 56.2 | 51.31 | 1.29 | 0.13 | 1 | 2 | 1 | -0.9319 | 0.512 | 5.3 | -20 |  |  | 0.0485 | 0.1599 | -0.0941 | 0.99995 | [21 empty] | 0 | [245 empty] | 0
P | 2018-09-01T22:57:31+08:00 | 2018-09-03T08:45:09+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP940G | CPSM Final Test | FAIL | [289 empty] | 2211
P | 2018-09-02T15:51:20+08:00 | 2018-09-03T09:50:33+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP940G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -1.38 | 56.4 | 54.61 | 1.21 | 0.13 | 0 | -1 | -1 | -0.8998 | 0.5264 | -1.6 | 37 |  |  | 0.1447 | 0.3744 | 0.0782 | 0.99984 | [21 empty] | 0.01372 | [246 empty] | 0
P | 2018-09-03T11:04:55+08:00 | 2018-09-03T14:01:06+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | H18FTP941G | CPSM Final Test | FAIL | [5 empty] | 45.38
P | 2018-09-03T12:33:07+08:00 | 2018-09-03T14:08:09+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T12:28:21+08:00 | 2018-09-03T14:12:55+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T12:25:11+08:00 | 2018-09-03T14:16:05+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T12:10:53+08:00 | 2018-09-03T14:30:23+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T11:52:09+08:00 | 2018-09-03T14:49:07+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T11:47:35+08:00 | 2018-09-03T14:53:41+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T11:47:13+08:00 | 2018-09-03T14:54:03+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T10:55:39+08:00 | 2018-09-03T15:45:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T10:47:53+08:00 | 2018-09-03T15:53:23+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-09-03T12:17:14+08:00 | 2018-09-03T16:20:58+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | H18FTP925G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -1.15 | 56.44 | 52.49 | 1.06 | 0.13 | 1 | 2 | -1 | -0.9359 | 0.529 | 0.8 | -28 |  |  | 0.1639 | 0.2471 | 0.02 | 0.99997 | [21 empty] | 0 | [247 empty] | 0
P | 2018-09-03T15:41:16+08:00 | 2018-09-03T16:21:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP958G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-09-03T12:25:30+08:00 | 2018-09-03T16:22:35+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | H18FTP944G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -1.45 | 56.01 | 51.39 | 1.59 | 0.13 | 1 | -2 | 0 | -0.84 | 0.51 | 3.2 | 10 |  |  | 0.0855 | 0.5487 | 0.0057 | 0.99999 | [21 empty] | 0 | [248 empty] | 0
P | 2018-09-03T13:28:25+08:00 | 2018-09-03T16:22:43+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP942G | CPSM Final Test | PASS | 129 | 166.64 | -0.02 | -1.07 | 56.2 | 48.85 | 1.59 | 0.13 | 1 | -2 | 0 | -0.9199 | 0.4862 | 8.4 | -7 |  |  | 0.1071 | 0.1184 | -0.0299 | 0.99997 | [21 empty] | 0.02446 | [249 empty] | 0
P | 2018-09-03T13:42:21+08:00 | 2018-09-03T16:30:28+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP945G | CPSM Final Test | PASS | 129 | 167.45 | 0.08 | -0.92 | 56.42 | 52.92 | 1.44 | 0.13 | 0 | -1 | 0 | -0.9677 | 0.535 | 2.3 | 31 |  |  | 0.0945 | 0.2176 | 0.0533 | 0.99999 | [21 empty] | 0 | [250 empty] | 0
P | 2018-09-03T12:33:00+08:00 | 2018-09-03T16:33:40+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | H18FTP941G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -1.3 | 55.95 | 50.63 | 1.36 | 0.13 | -2 | 1 | -1 | -0.9178 | 0.5278 | 3.8 | 25 |  |  | 0.092 | 0.0643 | 0.0065 | 0.99999 | [21 empty] | 0 | [251 empty] | 0
P | 2018-09-03T14:15:28+08:00 | 2018-09-03T16:55:25+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP959G | CPSM Final Test | FAIL | [13 empty] | -9.4
P | 2018-09-03T16:00:31+08:00 | 2018-09-03T18:41:38+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | H18FTP943G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -1.83 | 56.21 | 52.49 | 0.98 | 0.13 | 0 | -3 | 0 | -0.8979 | 0.5232 | 3.3 | 11 |  |  | 0.0876 | 0.1397 | -0.0494 | 1 | [21 empty] | 0 | [252 empty] | 0
P | 2018-09-03T15:35:26+08:00 | 2018-09-03T18:55:27+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | H18FTP959G | CPSM Final Test | PASS | 129 | 162.05 | -0.21 | -1.22 | 56.33 | 50.04 | 1.67 | 0.13 | 1 | -4 | 0 | -0.856 | 0.498 | 10.7 | 14 |  |  | 0.1656 | 0.3876 | 0.0587 | 0.99981 | [21 empty] | 0.01066 | [246 empty] | 0
P | 2018-09-03T16:01:27+08:00 | 2018-09-03T18:56:30+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP958G | CPSM Final Test | PASS | 129 | 156.29 | 0.17 | -0.92 | 56.38 | 49.36 | 1.06 | 0.13 | -1 | 1 | 0 | -0.9459 | 0.4952 | 2.8 | 6 |  |  | 0.1108 | 0.2465 | -0.021 | 1 | [21 empty] | 0 | [253 empty] | 0
P | 2018-09-03T17:59:11+08:00 | 2018-09-03T19:11:10+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP962G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-09-03T18:46:35+08:00 | 2018-09-03T21:29:20+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | H18FTP960G | CPSM Final Test | PASS | 129 | 158.81 | -0.02 | -0.99 | 56.18 | 51.05 | 1.51 | 0.13 | 1 | -4 | -1 | -0.9577 | 0.5436 | -4.3 | 0 |  |  | 0.4156 | 0.5507 | -0.0306 | 0.99999 | [21 empty] | 0 | [254 empty] | 0
P | 2018-09-03T18:55:31+08:00 | 2018-09-03T21:39:48+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | H18FTP957G | CPSM Final Test | PASS | 129 | 157.19 | -0.02 | -0.77 | 56.42 | 51.22 | 1.51 | 0.13 | -1 | 1 | 0 | -0.9039 | 0.476 | -2.2 | 14 |  |  | 0.1108 | 0.2591 | -0.0075 | 0.99998 | [21 empty] | 0 | [255 empty] | 0
P | 2018-09-03T04:24:49+08:00 | 2018-09-04T08:45:32+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP962G | CPSM Final Test | FAIL | [14 empty] | 872
P | 2018-09-03T04:12:42+08:00 | 2018-09-04T09:02:55+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP961G | CPSM Final Test | FAIL
P | 2018-09-03T18:35:11+08:00 | 2018-09-04T10:29:54+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | H18FTP962G | CPSM Final Test | PASS | 129 | 167.18 | -0.11 | -0.08 | 56.38 | 49.36 | 1.29 | 0.13 | 0 | 2 | 0 | -0.9637 | 0.5152 | 0.6 | 11 |  |  | 0.1173 | 0.2421 | 0.0265 | 0.99995 | [21 empty] | 0 | [256 empty] | 0
P | 2018-09-03T18:37:48+08:00 | 2018-09-04T10:47:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP961G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | -1.15 | 56.03 | 48.01 | 1.59 | 0.13 | -1 | -2 | 0 | -0.9398 | 0.5392 | 4.8 | -1 |  |  | 0.0973 | 0.1769 | -0.0228 | 1 | [21 empty] | 0.01906 | [257 empty] | 0
P | 2018-09-04T08:33:02+08:00 | 2018-09-04T14:31:40+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP963G | CPSM Final Test | FAIL
P | 2018-09-04T09:57:03+08:00 | 2018-09-04T14:32:49+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP964G | CPSM Final Test | FAIL
P | 2018-09-04T11:32:21+08:00 | 2018-09-04T15:34:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | H18FTP963G | CPSM Final Test | PASS | 129 | 160.43 | -0.02 | -0.77 | 56.6 | 49.95 | 1.82 | 0.13 | -1 | -2 | 0 | -0.9098 | 0.5074 | 8.6 | 35 |  |  | 0.039 | 0.2369 | -0.055 | 1 | [21 empty] | 0 | [258 empty] | 0
P | 2018-09-04T12:14:56+08:00 | 2018-09-04T15:38:19+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | H18FTP964G | CPSM Final Test | PASS | 129 | 166.19 | -0.02 | -0.92 | 56.17 | 51.14 | 1.59 | 0.13 | 1 | -2 | 0 | -0.8979 | 0.529 | 5.8 | 26 |  |  | 0.0837 | 0.1596 | 0.0151 | 0.99998 | [21 empty] | 0.01948 | [259 empty] | 0
P | 2018-09-05T13:06:26+08:00 | 2018-09-05T14:29:51+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP639G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-09-05T11:19:43+08:00 | 2018-09-05T15:45:22+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP642G | CPSM Final Test | FAIL
P | 2018-09-05T11:30:42+08:00 | 2018-09-05T16:05:35+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP639G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-09-05T11:21:14+08:00 | 2018-09-05T16:15:03+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP639G | CPSM Final Test | FAIL | [87 empty] | FALSE
P | 2018-09-05T10:41:52+08:00 | 2018-09-05T16:57:40+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP640G | CPSM Final Test | FAIL
P | 2018-09-05T13:32:33+08:00 | 2018-09-05T17:51:00+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP642G | CPSM Final Test | PASS |  | 165.11 | -0.02 | -0.77 | 56.43 | 51.31 | 1.59 | 0.13 | 1 | -2 | 0 | -0.9479 | 0.5428 | 6.9 | -14 |  |  | 0.1747 | 0.2148 | -0.0455 | 1 |  |  |  |  | 133 |  | 0.03053 | [275 empty] | 0
P | 2018-09-05T15:20:06+08:00 | 2018-09-05T18:14:51+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP643G | CPSM Final Test | PASS |  | 165.56 | -0.02 | 0 | 56.36 | 50.04 | 1.36 | 0.13 | 1 | 2 | 0 | -0.9718 | 0.5268 | 4.8 | -6 |  |  | 0.0518 | 0.1342 | 0.0132 | 0.99998 |  |  |  |  | 133 |  | 0 | [274 empty] | 0
P | 2018-09-05T15:27:16+08:00 | 2018-09-05T18:21:52+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | J18FNP641G | CPSM Final Test | PASS |  | 165.02 | -0.02 | -0.54 | 56.77 | 49.78 | 1.44 | 0.13 | -2 | -2 | 0 | -0.9718 | 0.5144 | 9.2 | 4 |  |  | 0.1131 | 0.1271 | -0.0065 | 0.99993 |  |  |  |  | 133 |  | 0 | [276 empty] | 0
P | 2018-09-05T13:49:46+08:00 | 2018-09-05T18:50:26+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP640G | CPSM Final Test | PASS |  | 165.29 | -0.02 | -0.84 | 56.23 | 51.31 | 1.36 | 0.13 | 0 | 1 | 0 | -0.8999 | 0.509 | 5.8 | 41 |  |  | 0.2082 | 0.3378 | -0.0095 | 0.99998 |  |  |  |  | 133 |  | 0 | [277 empty] | 0
P | 2018-09-05T16:26:16+08:00 | 2018-09-05T19:23:01+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP639G | CPSM Final Test | PASS |  | 165.02 | -0.3 | -0.84 | 56.43 | 49.53 | 1.36 | 0.13 | 1 | -3 | -1 | -0.882 | 0.492 | 4.1 | 3 |  |  | 0.2324 | 0.2923 | -0.0375 | 1 |  |  |  |  | 133 |  | 0 | [278 empty] | 0
P | 2018-09-05T17:15:55+08:00 | 2018-09-05T20:53:46+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP646G | CPSM Final Test | PASS |  | 165.02 | -0.02 | -1.07 | 56.4 | 48.09 | 1.74 | 0.13 | 0 | -1 | 1 | -0.856 | 0.5132 | 3.3 | -3 |  |  | 0.1587 | 0.1252 | 0.0694 | 0.99989 |  |  |  |  | 133 |  | 0.00386 | [279 empty] | 0
P | 2018-09-05T01:50:38+08:00 | 2018-09-06T08:39:38+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | J18FNP644G | CPSM Final Test | FAIL
P | 2018-09-05T17:15:08+08:00 | 2018-09-06T08:40:14+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | J18FNP644G | CPSM Final Test | PASS |  | 162.05 | -0.02 | -0.92 | 56.43 | 48.77 | 1.44 | 0.13 | 1 | -4 | 0 | -0.8999 | 0.5016 | 1.3 | 19 |  |  | 0.0661 | 0.1651 | 0.0037 | 0.99997 |  |  |  |  | 133 |  | 0 | [280 empty] | 0
P | 2018-09-06T08:34:07+08:00 | 2018-09-06T11:07:09+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP649G | CPSM Final Test | FAIL
P | 2018-09-06T09:47:02+08:00 | 2018-09-06T12:51:42+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP648G | CPSM Final Test | PASS |  | 165.29 | -0.02 | -0.92 | 56.23 | 51.22 | 0.6 | 0.13 | 1 | -2 | -1 | -0.9458 | 0.5254 | 4.2 | 60 |  |  | 0.114 | 0.2161 | 0.0401 | 0.99996 |  |  |  |  | 133 |  | 0 | [281 empty] | 0
P | 2018-09-06T09:50:38+08:00 | 2018-09-06T13:11:07+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP649G | CPSM Final Test | PASS |  | 166.19 | 0.27 | -0.69 | 56.22 | 50.12 | 1.21 | 0.13 | 1 | -2 | 0 | -0.93 | 0.491 | 8.2 | 33 |  |  | 0.1092 | 0.0845 | -0.0158 | 1 |  |  |  |  | 133 |  | 0.00886 | [282 empty] | 0
P | 2018-09-06T09:20:06+08:00 | 2018-09-06T10:06:02+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP647G | CPSM Final Test | FAIL
P | 2018-09-06T09:42:45+08:00 | 2018-09-06T15:31:59+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP650G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-09-06T15:32:58+08:00 | 2018-09-06T16:43:40+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP654G | CPSM Final Test | FAIL | [5 empty] | 45.8
P | 2018-09-06T12:40:47+08:00 | 2018-09-06T18:26:02+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | J18FNP651G | CPSM Final Test | PASS |  | 163.49 | -0.02 | -0.61 | 56.33 | 51.73 | 1.9 | 0.13 | 0 | -1 | 1 | -0.9498 | 0.4846 | 6.5 | 29 |  |  | 0.1662 | 0.2386 | 0.0071 | 0.99997 |  |  |  |  | 133 |  | 0 | [283 empty] | 0
P | 2018-09-06T12:37:22+08:00 | 2018-09-06T18:24:29+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP650G | CPSM Final Test | PASS |  | 163.4 | -0.02 | -1.15 | 55.89 | 52.32 | 1.29 | 0.13 | 1 | -3 | 0 | -0.9298 | 0.535 | 1.2 | 21 |  |  | 0.1604 | 0.1284 | 0.0174 | 1 |  |  |  |  | 133 |  | 0 | [284 empty] | 0
P | 2018-09-06T09:43:04+08:00 | 2018-09-06T18:28:06+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP647G | CPSM Final Test | PASS |  | 164.39 | -0.02 | -0.99 | 56.34 | 52.49 | 1.29 | 0.13 | 0 | 2 | 1 | -0.9159 | 0.4956 | 7.8 | 30 |  |  | 0.2508 | 0.2537 | 0.0113 | 0.99999 |  |  |  |  | 133 |  | 0 | [285 empty] | 0
P | 2018-09-06T15:37:43+08:00 | 2018-09-06T18:45:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | J18FNP652G | CPSM Final Test | PASS |  | 167.54 | 0.37 | -1.15 | 56.11 | 49.45 | 0.98 | 0.13 | 0 | 1 | 0 | -0.9218 | 0.5206 | 4 | -3 |  |  | 0.2003 | 0.1315 | -0.008 | 0.99996 |  |  |  |  | 133 |  | 0 | [218 empty] | 0
P | 2018-09-06T16:11:46+08:00 | 2018-09-06T18:56:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP653G | CPSM Final Test | PASS |  | 162.41 | -0.02 | -0.99 | 56.17 | 49.19 | 1.44 | 0.13 | 0 | -3 | 0 | -0.9379 | 0.4904 | 3.7 | 22 |  |  | 0.2388 | 0.3079 | -0.0251 | 1 |  |  |  |  | 133 |  | 0 | [286 empty] | 0
P | 2018-09-06T12:37:36+08:00 | 2018-09-06T19:39:02+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP654G | CPSM Final Test | FAIL | [27 empty] | 0.16378
P | 2018-09-05T23:04:34+08:00 | 2018-09-07T09:12:04+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP654G | CPSM Final Test | FAIL | [27 empty] | 0.07739
P | 2018-09-06T00:15:51+08:00 | 2018-09-07T09:48:28+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP645G | CPSM Final Test | FAIL | [13 empty] | 26.8
P | 2018-09-06T16:08:19+08:00 | 2018-09-07T10:15:08+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP654G | CPSM Final Test | PASS |  | 168.62 | -0.02 | -0.77 | 56.21 | 50.46 | 1.44 | 0.13 | -1 | -1 | -2 | -0.9039 | 0.54 | 7.5 | -8 |  |  | 0.1917 | 0.2104 | 0.017 | 1 |  |  |  |  | 133 |  | 0 | [287 empty] | 0
P | 2018-09-07T07:40:47+08:00 | 2018-09-07T11:33:29+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP655G | CPSM Final Test | FAIL
P | 2018-09-07T07:38:44+08:00 | 2018-09-07T11:27:27+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP658G | CPSM Final Test | FAIL | [316 empty] | 0
P | 2018-09-07T09:21:29+08:00 | 2018-09-07T12:31:22+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP659G | CPSM Final Test | PASS |  | 162.41 | -0.02 | -0.92 | 56.42 | 49.11 | 1.36 | 0.13 | -2 | -2 | 0 | -0.9178 | 0.525 | 6 | -17 |  |  | 0.2775 | 0.5966 | 0.0264 | 0.99991 |  |  |  |  | 133 |  | 0.00865 | [289 empty] | 0
P | 2018-09-07T09:40:45+08:00 | 2018-09-07T12:37:46+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP965G | CPSM Final Test | PASS | 129 | 175.46 | -0.02 | -0.99 | 56.11 | 48.85 | 1.36 | 0.13 | 0 | -2 | 0 | -0.846 | 0.522 | 4 | 7 |  |  | 0.1623 | 0.2623 | -0.0104 | 0.99998 | [21 empty] | 0 | [270 empty] | 0
P | 2018-09-07T09:31:38+08:00 | 2018-09-07T12:44:24+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | J18FNP657G | CPSM Final Test | PASS |  | 158.27 | -0.02 | -0.92 | 56.41 | 54.78 | 1.51 | 0.13 | -1 | 2 | 0 | -0.9179 | 0.5148 | 3.5 | -4 |  |  | 0.2651 | 0.5468 | 0.0287 | 0.99997 |  |  |  |  | 133 |  | 0 | [290 empty] | 0
P | 2018-09-07T09:45:36+08:00 | 2018-09-07T12:50:31+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | J18FNP656G | CPSM Final Test | PASS |  | 167.54 | -0.02 | -0.31 | 56.21 | 51.56 | 0.98 | 0.13 | 0 | 2 | 0 | -0.9078 | 0.5268 | 4.9 | 22 |  |  | 0.1335 | 0.1349 | -0.0068 | 1 |  |  |  |  | 133 |  | 0 | [291 empty] | 0
P | 2018-09-07T09:33:06+08:00 | 2018-09-07T12:57:05+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP658G | CPSM Final Test | PASS |  | 164.93 | 0.17 | -0.92 | 56.23 | 49.02 | 1.13 | 0.13 | 1 | -2 | 0 | -0.9239 | 0.4938 | 4.8 | 9 |  |  | 0.112 | 0.7294 | -0.0106 | 1 |  |  |  |  | 133 |  | 0 | [288 empty] | 0
P | 2018-09-07T09:37:08+08:00 | 2018-09-07T13:21:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP655G | CPSM Final Test | PASS |  | 164.75 | 0.37 | -0.92 | 56.19 | 50.21 | 1.59 | 0.13 | -1 | -2 | 0 | -0.8521 | 0.519 | 4.3 | 13 |  |  | 0.1131 | 0.4545 | 0.0289 | 0.99998 |  |  |  |  | 133 |  | 0.00982 | [292 empty] | 0
P | 2018-09-07T15:43:15+08:00 | 2018-09-07T16:00:30+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP970G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-09-05T16:43:53+08:00 | 2018-09-07T17:20:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP645G | CPSM Final Test | FAIL | [301 empty] | 0
P | 2018-09-07T16:20:04+08:00 | 2018-09-07T17:22:25+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP970G | CPSM Final Test | FAIL
P | 2018-09-07T13:36:57+08:00 | 2018-09-07T17:43:31+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP966G | CPSM Final Test | FAIL
P | 2018-09-06T17:02:10+08:00 | 2018-09-07T18:26:20+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP645G | CPSM Final Test | PASS |  | 158.36 | -0.02 | -0.38 | 56.52 | 51.99 | 1.67 | 0.13 | -1 | 1 | 1 | -1.0157 | 0.5128 | 2.9 | 16 |  |  | 0.0833 | 0.1862 | -0.0246 | 1 |  |  |  |  | 133 |  | 0 | [273 empty] | 0
P | 2018-09-07T15:30:10+08:00 | 2018-09-07T19:02:18+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP971G | CPSM Final Test | PASS | 129 | 163.58 | 0.27 | -1.07 | 56.21 | 52.32 | 1.59 | 0.13 | 1 | -3 | 0 | -0.9159 | 0.521 | 7.7 | 0 |  |  | 0.1 | 0.2343 | -0.0409 | 0.99998 | [21 empty] | 0 | [278 empty] | 0
P | 2018-09-07T15:49:55+08:00 | 2018-09-07T19:17:34+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP968G | CPSM Final Test | PASS | 129 | 169.52 | -0.02 | -0.69 | 56.43 | 55.2 | 1.97 | 0.13 | -1 | -2 | 0 | -0.9657 | 0.4898 | 3.3 | -3 |  |  | 0.0657 | 0.2341 | 0.021 | 0.99999 | [21 empty] | 0.00312 | [279 empty] | 0
P | 2018-09-07T15:48:47+08:00 | 2018-09-07T19:15:19+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP969G | CPSM Final Test | PASS | 129 | 166.37 | -0.02 | -0.61 | 56.32 | 48.18 | 1.74 | 0.13 | 1 | -2 | 0 | -0.8938 | 0.5124 | 3.7 | -14 |  |  | 0.0592 | 0.1259 | -0.0343 | 1 | [21 empty] | 0 | [280 empty] | 0
P | 2018-09-07T15:40:14+08:00 | 2018-09-07T19:28:01+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP966G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.99 | 56.19 | 51.31 | 1.36 | 0.13 | -1 | 4 | 0 | -0.9418 | 0.5328 | 7.8 | -1 |  |  | 0.0867 | 0.2714 | 0.0318 | 0.99997 | [21 empty] | 0 | [273 empty] | 0
P | 2018-09-07T14:06:23+08:00 | 2018-09-07T19:41:29+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP967G | CPSM Final Test | FAIL
P | 2018-09-07T16:51:15+08:00 | 2018-09-07T20:34:09+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP970G | CPSM Final Test | PASS | 129 | 170.24 | 1.06 | -1.15 | 56.19 | 52.32 | 1.21 | 0.13 | -1 | 3 | 0 | -0.7782 | 0.499 | 5.4 | 57 |  |  | 0.0363 | 0.7081 | 0.1224 | 0.99992 | [21 empty] | 0.00791 | [281 empty] | 0
P | 2018-09-07T16:53:56+08:00 | 2018-09-07T21:25:35+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP967G | CPSM Final Test | PASS | 129 | 166.91 | -0.02 | -0.77 | 56.55 | 51.22 | 1.44 | 0.13 | 1 | 2 | 0 | -0.9397 | 0.5092 | 5.8 | -1 |  |  | 0.0831 | 0.0942 | -0.017 | 0.99999 | [21 empty] | 0.00771 | [282 empty] | 0
P | 2018-09-08T09:32:41+08:00 | 2018-09-08T10:19:05+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP972G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-09-08T06:26:43+08:00 | 2018-09-08T12:48:29+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP979G | CPSM Final Test | FAIL
P | 2018-09-08T09:40:47+08:00 | 2018-09-08T12:52:02+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP980G | CPSM Final Test | PASS | 129 | 164.93 | 0.17 | -0.69 | 56.42 | 50.29 | 1.06 | 0.13 | -1 | 2 | 0 | -0.8559 | 0.549 | 6.4 | -15 |  |  | 0.1212 | 0.2982 | -0.0334 | 0.99998 | [21 empty] | 0 | [258 empty] | 0
P | 2018-09-08T09:58:15+08:00 | 2018-09-08T13:03:30+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP974G | CPSM Final Test | PASS | 129 | 172.49 | -0.02 | -1.22 | 56.02 | 49.11 | 1.06 | 0.13 | -1 | -2 | 0 | -0.8759 | 0.5078 | 3.2 | 63 |  |  | 0.1127 | 0.2966 | 0.0305 | 0.99999 | [21 empty] | 0 | [169 empty] | 0
P | 2018-09-08T06:54:43+08:00 | 2018-09-08T12:59:31+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP975G | CPSM Final Test | FAIL
P | 2018-09-08T09:57:07+08:00 | 2018-09-08T13:11:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP975G | CPSM Final Test | PASS | 129 | 159.53 | -0.02 | -0.99 | 56.34 | 49.61 | 1.44 | 0.13 | 1 | -2 | 0 | -0.854 | 0.4994 | 0.1 | 33 |  |  | 0.1282 | 0.5932 | 0.08 | 0.99983 | [21 empty] | 0.00983 | [283 empty] | 0
P | 2018-09-08T09:37:36+08:00 | 2018-09-08T14:33:23+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP979G | CPSM Final Test | PASS | 129 | 163.49 | 0.37 | -0.77 | 56.04 | 49.36 | 1.21 | 0.13 | 1 | 1 | 1 | -0.862 | 0.5464 | 0.3 | -7 |  |  | 0.1084 | 0.1565 | -0.0012 | 1 | [21 empty] | 0 | [284 empty] | 0
P | 2018-09-08T04:41:07+08:00 | 2018-09-08T14:56:03+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP977G | CPSM Final Test | FAIL
P | 2018-09-08T11:05:23+08:00 | 2018-09-08T15:15:40+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP976G | CPSM Final Test | PASS | 129 | 159.17 | 0.46 | -0.38 | 56.48 | 52.32 | 1.36 | 0.13 | 0 | 1 | 0 | -0.8819 | 0.51 | 3.6 | 1 |  |  | 0.0924 | 0.2892 | 0.0037 | 0.99998 | [21 empty] | 0 | [285 empty] | 0
P | 2018-09-08T14:10:37+08:00 | 2018-09-08T15:37:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP984G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-09-08T09:48:35+08:00 | 2018-09-08T15:59:05+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP977G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.92 | 56.3 | 50.55 | 1.21 | 0.13 | -1 | 1 | 0 | -0.868 | 0.531 | 7 | 5 |  |  | 0.0562 | 0.501 | 0.0258 | 0.99999 | [21 empty] | 0 | [286 empty] | 0
P | 2018-09-08T12:46:40+08:00 | 2018-09-08T16:32:07+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP982G | CPSM Final Test | FAIL | [13 empty] | 18
P | 2018-09-06T14:22:34+08:00 | 2018-09-08T16:50:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP973G | CPSM Final Test | FAIL
P | 2018-09-08T05:06:25+08:00 | 2018-09-08T17:01:07+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP972G | CPSM Final Test | FAIL | [14 empty] | 813
P | 2018-09-07T15:36:44+08:00 | 2018-09-08T17:54:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP973G | CPSM Final Test | PASS | 129 | 168.17 | 0.17 | -0.99 | 56.22 | 49.11 | 1.29 | 0.13 | 0 | -2 | 0 | -0.9019 | 0.5102 | 5.3 | 2 |  |  | 0.1772 | 0.2578 | 0.0394 | 0.99999 | [21 empty] | 0 | [287 empty] | 0
P | 2018-09-08T14:32:02+08:00 | 2018-09-08T17:57:04+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP981G | CPSM Final Test | PASS | 129 | 162.77 | 0.17 | -0.99 | 56.22 | 49.7 | 1.44 | 0.13 | 0 | 2 | 0 | -0.8599 | 0.496 | 8.1 | 14 |  |  | 0.138 | 0.1167 | -0.0493 | 1 | [21 empty] | 0 | [288 empty] | 0
P | 2018-09-08T14:41:31+08:00 | 2018-09-08T18:04:16+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP983G | CPSM Final Test | PASS | 129 | 159.26 | 0.17 | -0.77 | 56.21 | 50.8 | 1.21 | 0.13 | 0 | -2 | 1 | -0.9258 | 0.5008 | 3.7 | 12 |  |  | 0.1273 | 0.2773 | -0.0445 | 0.99997 | [21 empty] | 0.00798 | [289 empty] | 0
P | 2018-09-08T14:39:24+08:00 | 2018-09-08T18:32:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP982G | CPSM Final Test | PASS | 129 | 161.15 | -0.02 | -1.22 | 56.1 | 51.65 | 1.44 | 0.13 | 0 | 3 | 0 | -0.84 | 0.5486 | 4.1 | 8 |  |  | 0.226 | 0.4142 | 0.0953 | 0.99971 | [21 empty] | 0.00405 | [290 empty] | 0
P | 2018-09-08T11:03:46+08:00 | 2018-09-08T18:45:13+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP972G | CPSM Final Test | PASS | 129 | 166.64 | 0.37 | -1.15 | 56.1 | 49.95 | 1.36 | 0.13 | -2 | -3 | -1 | -0.8599 | 0.5204 | 5.9 | -8 |  |  | 0.0968 | 0.0823 | 0.0131 | 0.99999 | [21 empty] | 0 | [291 empty] | 0
P | 2018-09-06T19:12:03+08:00 | 2018-09-10T10:36:13+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP984G | CPSM Final Test | FAIL
P | 2018-09-06T18:08:17+08:00 | 2018-09-10T11:39:59+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP984G | CPSM Final Test | FAIL
P | 2018-09-10T08:01:32+08:00 | 2018-09-10T11:50:53+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP985G | CPSM Final Test | FAIL | [13 empty] | 19
P | 2018-09-10T07:30:39+08:00 | 2018-09-10T12:02:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP987G | CPSM Final Test | FAIL
P | 2018-09-10T07:45:27+08:00 | 2018-09-10T12:31:53+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP991G | CPSM Final Test | FAIL
P | 2018-09-10T09:45:39+08:00 | 2018-09-10T13:18:54+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP988G | CPSM Final Test | PASS | 129 | 160.16 | -0.02 | -0.69 | 56.41 | 50.63 | 1.21 | 0.13 | 1 | -3 | 0 | -0.892 | 0.55 | 3.7 | 12 |  |  | 0.0694 | 0.2466 | -0.015 | 0.99999 | [21 empty] | 0 | [292 empty] | 0
P | 2018-09-10T09:51:34+08:00 | 2018-09-10T13:20:25+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP986G | CPSM Final Test | PASS | 129 | 168.44 | 0.37 | -0.99 | 56.31 | 51.31 | 1.13 | 0.13 | 1 | -3 | 0 | -0.9138 | 0.5456 | 8.9 | 13 |  |  | 0.1577 | 0.1712 | 0.0126 | 0.99999 | [21 empty] | 0 | [286 empty] | 0
P | 2018-09-10T09:46:46+08:00 | 2018-09-10T13:48:02+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP987G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -1.15 | 56.2 | 51.31 | 1.51 | 0.13 | 1 | 1 | 0 | -0.846 | 0.516 | 3.6 | 3 |  |  | 0.158 | 0.0862 | 0.0073 | 0.99997 | [21 empty] | 0 | [293 empty] | 0
P | 2018-09-10T09:56:13+08:00 | 2018-09-10T13:51:09+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP985G | CPSM Final Test | PASS | 129 | 160.25 | -0.02 | -0.38 | 56.51 | 51.73 | 1.36 | 0.13 | 1 | 3 | 0 | -0.876 | 0.508 | -1.9 | 11 |  |  | 0.0444 | 0.092 | 0.0016 | 0.99999 | [21 empty] | 0 | [294 empty] | 0
P | 2018-09-10T10:08:40+08:00 | 2018-09-10T14:16:15+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP991G | CPSM Final Test | PASS | 129 | 158.27 | -0.02 | 0.45 | 56.48 | 48.51 | 1.21 | 0.13 | -1 | 3 | 1 | -0.8101 | 0.525 | 9.7 | 16 |  |  | 0.0959 | 0.4623 | 0.085 | 0.99985 | [21 empty] | 0.00661 | [295 empty] | 0
P | 2018-09-06T13:19:33+08:00 | 2018-09-10T16:28:43+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP984G | CPSM Final Test | FAIL | [339 empty] | 0
P | 2018-09-08T14:54:08+08:00 | 2018-09-10T17:34:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP984G | CPSM Final Test | PASS | 129 | 168.62 | 0.17 | -0.31 | 56.4 | 48.6 | 1.82 | 0.13 | 0 | -3 | -1 | -0.8879 | 0.4984 | 5.5 | 7 |  |  | 0.0511 | 0.1799 | -0.0077 | 1 | [21 empty] | 0 | [296 empty] | 0
P | 2018-09-10T15:29:18+08:00 | 2018-09-10T18:40:51+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP990G | CPSM Final Test | PASS | 129 | 163.4 | 0.17 | -0.69 | 56.4 | 50.97 | 1.44 | 0.13 | -1 | 1 | 0 | -0.9856 | 0.493 | 5.8 | 1 |  |  | 0.1787 | 0.1364 | 0.0422 | 0.99998 | [21 empty] | 0 | [127 empty] | 0
P | 2018-09-10T15:48:54+08:00 | 2018-09-10T18:45:02+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP992G | CPSM Final Test | PASS | 129 | 168.53 | -0.02 | -0.92 | 56.3 | 51.48 | 1.36 | 0.13 | 1 | -2 | 0 | -0.884 | 0.5238 | 0.6 | 79 |  |  | 0.1144 | 0.1959 | 0.0422 | 0.99994 | [21 empty] | 0 | [297 empty] | 0
P | 2018-09-10T15:53:38+08:00 | 2018-09-10T18:47:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP989G | CPSM Final Test | PASS | 129 | 161.78 | -0.02 | -0.31 | 56.49 | 51.22 | 1.36 | 0.13 | 1 | -3 | 0 | -0.9259 | 0.5202 | 2.1 | 6 |  |  | 0.0764 | 0.1857 | 0.0064 | 0.99999 | [21 empty] | 0 | [298 empty] | 0
P | 2018-09-09T23:39:40+08:00 | 2018-09-10T20:27:20+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP978G | CPSM Final Test | FAIL | [342 empty] | 0
P | 2018-09-10T10:03:30+08:00 | 2018-09-10T21:32:44+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP978G | CPSM Final Test | PASS | 129 | 168.26 | -0.02 | -0.61 | 56.33 | 50.72 | 1.21 | 0.13 | 0 | 3 | 0 | -0.9458 | 0.5214 | 2.4 | 60 |  |  | 0.1299 | 0.1312 | 0.0342 | 0.99996 | [21 empty] | 0.04315 | [299 empty] | 0
P | 2018-09-10T19:42:51+08:00 | 2018-09-10T22:54:21+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP998G | CPSM Final Test | PASS | 129 | 160.16 | 0.27 | -0.54 | 56.49 | 51.31 | 1.82 | 0.13 | 0 | 1 | 0 | -0.9277 | 0.5174 | 9.9 | 21 |  |  | 0.1489 | 0.2258 | 0.0397 | 0.99997 | [21 empty] | 0.01295 | [300 empty] | 0
P | 2018-09-10T19:38:59+08:00 | 2018-09-10T22:56:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP993G | CPSM Final Test | PASS | 129 | 173.48 | -0.02 | -0.54 | 56.31 | 50.8 | 0.91 | 0.13 | 1 | 3 | 0 | -0.8839 | 0.518 | 2.6 | -18 |  |  | 0.1143 | 0.2254 | 0.0337 | 0.99999 | [21 empty] | 0 | [301 empty] | 0
P | 2018-09-10T19:46:38+08:00 | 2018-09-10T22:58:24+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP994G | CPSM Final Test | PASS | 129 | 167 | 0.27 | -0.46 | 56.51 | 51.05 | 1.74 | 0.13 | 1 | -4 | 0 | -0.9538 | 0.519 | 12.3 | 3 |  |  | 0.0516 | 0.1613 | -0.0053 | 1 | [21 empty] | 0.01378 | [302 empty] | 0
P | 2018-09-10T19:56:41+08:00 | 2018-09-10T23:03:37+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP997G | CPSM Final Test | PASS | 129 | 157.55 | 0.4 | 0.23 | 56.32 | 52.66 | 1.51 | 0.13 | -1 | 1 | 0 | -0.9198 | 0.5492 | 0.6 | -5 |  |  | 0.1484 | 0.173 | 0.0085 | 0.99999 | [21 empty] | 0 | [303 empty] | 0
P | 2018-09-10T17:34:08+08:00 | 2018-09-10T22:26:21+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP002G | CPSM Final Test | FAIL
P | 2018-09-10T07:01:48+08:00 | 2018-09-11T08:40:59+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP995G | CPSM Final Test | FAIL
P | 2018-09-10T19:51:24+08:00 | 2018-09-11T09:47:05+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP995G | CPSM Final Test | PASS | 129 | 168.98 | -0.02 | -0.84 | 56.5 | 52.49 | 1.13 | 0.13 | 0 | -4 | 0 | -0.9279 | 0.507 | 5.6 | 11 |  |  | 0.0981 | 0.322 | 0.0028 | 0.99999 | [21 empty] | 0 | [115 empty] | 0
P | 2018-09-10T04:34:14+08:00 | 2018-09-11T11:26:15+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP002G | CPSM Final Test | FAIL
P | 2018-09-11T09:02:22+08:00 | 2018-09-11T12:03:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP999G | CPSM Final Test | FAIL
P | 2018-09-11T08:51:54+08:00 | 2018-09-11T12:22:18+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP001G | CPSM Final Test | FAIL
P | 2018-09-10T20:00:15+08:00 | 2018-09-11T12:29:17+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP002G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.84 | 56.41 | 53.93 | 1.29 | 0.13 | -1 | 3 | 0 | -0.8939 | 0.5406 | 3.5 | 33 |  |  | 0.1309 | 0.4416 | 0.0615 | 0.99992 | [21 empty] | 0 | [277 empty] | 0
P | 2018-09-11T10:40:14+08:00 | 2018-09-11T13:30:04+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP003G | CPSM Final Test | PASS | 129 | 154.67 | -0.02 | -0.46 | 56.22 | 49.78 | 1.44 | 0.13 | -2 | 2 | 0 | -0.9238 | 0.514 | 6.7 | -6 |  |  | 0.0736 | 0.1996 | -0.0181 | 0.99999 | [17 empty] | 0 |  |  |  | 0
P | 2018-09-11T10:43:23+08:00 | 2018-09-11T13:42:17+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP005G | CPSM Final Test | PASS | 129 | 169.52 | 0.17 | -1.22 | 56.23 | 50.46 | 0.91 | 0.13 | 0 | -1 | 0 | -0.8959 | 0.543 | 4.8 | 25 |  |  | 0.3855 | 0.4147 | -0.0379 | 0.99992 | [21 empty] | 0 | [115 empty] | 0
P | 2018-09-11T10:33:10+08:00 | 2018-09-11T13:48:39+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP999G | CPSM Final Test | PASS | 129 | 167.45 | 0.17 | -0.99 | 56.32 | 48.68 | 1.36 | 0.13 | -2 | -2 | 0 | -0.87 | 0.5428 | 6 | 5 |  |  | 0.1868 | 0.3803 | 0.027 | 0.99998 | [21 empty] | 0.00813 |  |  |  | 0
P | 2018-09-11T10:37:06+08:00 | 2018-09-11T14:06:38+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP001G | CPSM Final Test | PASS | 129 | 159.44 | -0.02 | 0 | 56.34 | 50.63 | 1.29 | 0.13 | -1 | -1 | 0 | -0.7961 | 0.489 | 3.7 | 11 |  |  | 0.122 | 0.2811 | 0.0292 | 0.99998 | [21 empty] | 0 | [304 empty] | 0
P | 2018-09-11T09:02:35+08:00 | 2018-09-11T15:50:15+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP004G | CPSM Final Test | FAIL
P | 2018-09-11T08:49:07+08:00 | 2018-09-11T16:27:34+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP006G | CPSM Final Test | FAIL | [55 empty] | 0
P | 2018-09-11T15:56:15+08:00 | 2018-09-11T16:55:47+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP011G | CPSM Final Test | FAIL | [348 empty] | FALSE
P | 2018-09-11T15:11:27+08:00 | 2018-09-11T17:25:47+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP010G | CPSM Final Test | FAIL | [13 empty] | 30.1
P | 2018-09-11T12:38:21+08:00 | 2018-09-11T17:33:04+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP006G | CPSM Final Test | PASS | 129 | 161.6 | -0.02 | -0.92 | 56.12 | 50.97 | 1.51 | 0.13 | 0 | -1 | 0 | -0.9378 | 0.543 | 7.6 | 2 |  |  | 0.0887 | 0.2282 | -0.0155 | 1 | [21 empty] | 0.03127 | [299 empty] | 0
P | 2018-09-11T15:00:36+08:00 | 2018-09-11T17:31:09+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP007G | CPSM Final Test | FAIL | [13 empty] | 24.6
P | 2018-09-11T15:41:07+08:00 | 2018-09-11T17:32:26+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP996G | CPSM Final Test | FAIL
P | 2018-09-11T12:26:25+08:00 | 2018-09-11T17:34:51+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP004G | CPSM Final Test | PASS | 129 | 157.46 | 0.17 | -1.22 | 56.23 | 49.7 | 0.98 | 0.13 | 1 | 3 | 0 | -0.9038 | 0.546 | 4.5 | 1 |  |  | 0.0902 | 0.1761 | -0.0378 | 1 | [21 empty] | 0 | [306 empty] | 0
P | 2018-09-11T15:44:17+08:00 | 2018-09-11T18:37:36+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP009G | CPSM Final Test | PASS | 129 | 164.93 | 0.27 | 0.3 | 56.6 | 50.04 | 1.74 | 0.13 | -1 | 3 | 0 | -0.8681 | 0.5368 | 3.8 | -2 |  |  | 0.076 | 0.1077 | -0.0186 | 0.99999 | [21 empty] | 0 | [307 empty] | 0
P | 2018-09-11T12:24:17+08:00 | 2018-09-11T18:57:51+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP008G | CPSM Final Test | FAIL
P | 2018-09-11T16:26:01+08:00 | 2018-09-11T19:23:39+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP011G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.92 | 56.62 | 50.72 | 1.59 | 0.13 | -2 | -2 | 0 | -0.842 | 0.5338 | 7.3 | 4 |  |  | 0.2648 | 0.4112 | 0.0493 | 0.99991 | [21 empty] | 0.00409 | [308 empty] | 0
P | 2018-09-11T16:18:37+08:00 | 2018-09-11T19:25:46+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP010G | CPSM Final Test | PASS | 129 | 168.98 | 0.27 | -1.07 | 56.32 | 52.07 | 1.36 | 0.13 | 0 | 3 | 1 | -0.9417 | 0.549 | 0.1 | 6 |  |  | 0.0486 | 0.276 | 0.0118 | 0.99999 | [21 empty] | 0.00628 | [309 empty] | 0
P | 2018-09-11T16:15:52+08:00 | 2018-09-11T19:31:21+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP007G | CPSM Final Test | PASS | 129 | 165.56 | -0.21 | -0.77 | 56.24 | 49.78 | 1.51 | 0.13 | -2 | 2 | 1 | -0.9001 | 0.498 | 4.2 | -5 |  |  | 0.0858 | 0.132 | 0.0285 | 0.99999 | [21 empty] | 0 | [310 empty] | 0
P | 2018-09-11T16:36:47+08:00 | 2018-09-11T19:32:50+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP996G | CPSM Final Test | PASS | 129 | 158.45 | 0.65 | -0.99 | 56.32 | 51.14 | 0.98 | 0.13 | -1 | -2 | 1 | -0.9379 | 0.541 | 3.3 | 23 |  |  | 0.0958 | 0.4245 | -0.0035 | 0.99998 | [21 empty] | 0 | [296 empty] | 0
P | 2018-09-11T15:41:04+08:00 | 2018-09-11T20:00:53+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP008G | CPSM Final Test | PASS | 129 | 151.34 | 0.37 | -1.15 | 56.11 | 50.46 | 1.21 | 0.13 | 1 | -5 | 0 | -0.8919 | 0.4968 | 8 | -41 |  |  | 0.1494 | 0.1844 | 0.022 | 0.99999 | [21 empty] | 0 | [311 empty] | 0
P | 2018-09-12T13:32:23+08:00 | 2018-09-12T16:38:14+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP013G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -0.38 | 56.5 | 48.26 | 1.44 | 0.13 | 0 | -2 | 0 | -0.8819 | 0.483 | 4.6 | 2 |  |  | 0.0939 | 0.1851 | 0.0009 | 1 | [21 empty] | 0 | [312 empty] | 0
P | 2018-09-12T13:21:56+08:00 | 2018-09-12T16:38:25+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP012G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.38 | 56.76 | 51.56 | 1.74 | 0.13 | -1 | 1 | 0 | -0.868 | 0.5198 | 6.2 | 14 |  |  | 0.2037 | 0.281 | -0.1921 | 0.99999 | [21 empty] | 0 | [51 empty] | 0
P | 2018-09-12T13:35:59+08:00 | 2018-09-12T16:38:27+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP015G | CPSM Final Test | PASS | 129 | 166.28 | -0.02 | -0.46 | 56.74 | 56.9 | 1.97 | 0.13 | 1 | -2 | 0 | -0.9239 | 0.519 | 5.9 | -3 |  |  | 0.0771 | 0.2566 | 0.0313 | 0.99999 | [21 empty] | 0 | [188 empty] | 0
P | 2018-09-12T13:25:03+08:00 | 2018-09-12T16:39:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP016G | CPSM Final Test | PASS | 129 | 158.45 | -0.02 | -0.38 | 56.52 | 49.7 | 1.82 | 0.13 | 1 | 3 | 0 | -0.9078 | 0.5082 | 7.5 | 25 |  |  | 0.1534 | 0.2843 | 0.0021 | 0.99998 | [21 empty] | 0 | [311 empty] | 0
P | 2018-09-12T15:35:19+08:00 | 2018-09-12T18:17:43+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP019G | CPSM Final Test | PASS | 129 | 159.71 | 0.46 | -0.61 | 56.41 | 51.9 | 1.13 | 0.13 | -1 | -4 | 0 | -0.85 | 0.5456 | 10.6 | 11 |  |  | 0.2509 | 0.4179 | 0.0425 | 0.99979 | [21 empty] | 0.00593 | [313 empty] | 0
P | 2018-09-12T15:24:19+08:00 | 2018-09-12T18:27:24+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP018G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -0.69 | 56.4 | 53.51 | 1.36 | 0.13 | 1 | 4 | 0 | -0.9159 | 0.5088 | 8.1 | 17 |  |  | 0.1176 | 0.2404 | 0.0257 | 0.99998 | [21 empty] | 0 | [299 empty] | 0
P | 2018-09-12T15:14:42+08:00 | 2018-09-12T18:22:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP014G | CPSM Final Test | PASS | 129 | 162.5 | -0.02 | -0.61 | 56.41 | 51.14 | 1.44 | 0.13 | -2 | -2 | 0 | -0.9997 | 0.5022 | 3.3 | -2 |  |  | 0.1165 | 0.2559 | 0.0141 | 1 | [21 empty] | 0 | [314 empty] | 0
P | 2018-09-11T22:35:46+08:00 | 2018-09-13T08:22:19+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP017G | CPSM Final Test | FAIL
P | 2018-09-12T15:29:02+08:00 | 2018-09-13T08:22:57+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP017G | CPSM Final Test | PASS | 129 | 175.37 | -0.02 | -0.38 | 56.44 | 55.03 | 1.59 | 0.13 | -1 | 2 | 0 | -0.9657 | 0.5236 | -2.3 | 14 |  |  | 0.0838 | 0.0985 | -0.0568 | 0.99999 | [21 empty] | 0 | [315 empty] | 0
P | 2018-09-13T07:46:33+08:00 | 2018-09-13T11:33:28+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | J18FTP020G | CPSM Final Test | FAIL | [13 empty] | -9.1
P | 2018-09-13T08:28:58+08:00 | 2018-09-13T11:34:57+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | J18FTP022G | CPSM Final Test | FAIL | [13 empty] | -6.2
P | 2018-09-13T07:02:21+08:00 | 2018-09-13T11:36:03+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | J18FTP021G | CPSM Final Test | FAIL | [13 empty] | -7.8
P | 2018-09-13T08:00:40+08:00 | 2018-09-13T12:52:59+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | J18FTP026G | CPSM Final Test | FAIL
P | 2018-09-13T07:58:10+08:00 | 2018-09-13T12:55:29+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | J18FTP026G | CPSM Final Test | FAIL
P | 2018-09-13T09:26:18+08:00 | 2018-09-13T13:07:17+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | J18FTP025G | CPSM Final Test | PASS | 129 | 172.49 | 0.08 | -0.99 | 56.42 | 55.2 | 1.29 | 0.13 | -1 | -2 | 0 | -0.8979 | 0.547 | 2.3 | 18 |  |  | 0.1293 | 0.0961 | 0.0372 | 1 | [21 empty] | 0 | [316 empty] | 0
P | 2018-09-13T07:46:26+08:00 | 2018-09-13T13:07:13+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | J18FTP026G | CPSM Final Test | FAIL | [41 empty] | 0
P | 2018-09-13T09:41:56+08:00 | 2018-09-13T13:12:23+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | J18FTP024G | CPSM Final Test | PASS | 129 | 164.03 | 0.27 | -1.15 | 56.32 | 49.28 | 0.91 | 0.13 | 0 | -2 | 0 | -0.9298 | 0.539 | 7.9 | 8 |  |  | 0.155 | 0.144 | -0.0137 | 0.99999 | [21 empty] | 0 | [317 empty] | 0
P | 2018-09-13T10:12:43+08:00 | 2018-09-13T13:18:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | J18FTP023G | CPSM Final Test | PASS | 129 | 166.1 | -0.3 | -0.77 | 56.72 | 54.61 | 1.36 | 0.13 | 0 | 3 | -1 | -0.852 | 0.5314 | -1 | 2 |  |  | 0.2175 | 0.2316 | 0.0279 | 0.99999 | [21 empty] | 0 | [318 empty] | 0
P | 2018-09-13T09:40:01+08:00 | 2018-09-13T13:33:30+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | J18FTP020G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -1.07 | 56.43 | 55.37 | 1.21 | 0.13 | -1 | -1 | 0 | -0.87 | 0.496 | 1.3 | 4 |  |  | 0.1402 | 0.14 | 0.0293 | 0.99998 | [21 empty] | 0 | [319 empty] | 0
P | 2018-09-13T10:01:58+08:00 | 2018-09-13T13:35:03+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | J18FTP022G | CPSM Final Test | PASS | 129 | 166.46 | 0.27 | -0.99 | 56.44 | 52.24 | 1.06 | 0.13 | -2 | -3 | 0 | -0.8959 | 0.4534 | 8.3 | 16 |  |  | 0.0512 | 0.1627 | -0.0153 | 1 | [21 empty] | 0 | [320 empty] | 0
P | 2018-09-13T09:19:12+08:00 | 2018-09-13T13:36:17+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | J18FTP021G | CPSM Final Test | PASS | 129 | 160.25 | 0.08 | -0.77 | 56.63 | 50.72 | 1.29 | 0.13 | -1 | -1 | 0 | -0.826 | 0.5416 | 3.2 | 3 |  |  | 0.2291 | 0.5628 | 0.0403 | 0.99984 | [21 empty] | 0.01103 | [321 empty] | 0
P | 2018-09-13T13:10:25+08:00 | 2018-09-13T16:05:54+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | J18FTP026G | CPSM Final Test | PASS | 129 | 166.19 | 0.17 | -0.69 | 56.86 | 51.82 | 0.75 | 0.13 | 1 | -2 | 0 | -0.9538 | 0.4758 | -3.4 | 11 |  |  | 0.0543 | 0.2687 | -0.0465 | 1 | [20 empty] | 0 | 0
P | 2018-09-13T14:31:53+08:00 | 2018-09-13T17:29:58+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | J18USM076G | CPSM Final Test | FAIL | [13 empty] | -5.4
P | 2018-09-13T18:08:45+08:00 | 2018-09-13T18:43:03+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | J18USM082G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-09-13T13:06:45+08:00 | 2018-09-13T18:55:06+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | J18USM076G | CPSM Final Test | FAIL
P | 2018-09-13T16:05:23+08:00 | 2018-09-13T19:06:49+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | J18USM078G | CPSM Final Test | PASS |  | 161.33 | -0.02 | 0 | 56.61 | 51.48 | 1.29 | 0.13 | -1 | 1 | 0 | -0.8719 | 0.5016 | 1.2 | 20 |  |  | 0.0632 | 0.3484 | -0.0063 | 0.99997 |  |  |  |  | 133 |  | 0 | [337 empty] | 0
P | 2018-09-13T15:45:42+08:00 | 2018-09-13T19:10:31+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | J18USM075G | CPSM Final Test | PASS |  | 163.04 | 0.17 | 0.15 | 56.81 | 51.48 | 1.21 | 0.13 | -1 | -2 | 0 | -0.9398 | 0.537 | -2.1 | 22 |  |  | 0.0443 | 0.1219 | -0.0063 | 1 |  |  |  |  | 133 |  | 0 | [338 empty] | 0
P | 2018-09-13T15:34:48+08:00 | 2018-09-13T19:12:59+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | J18USM077G | CPSM Final Test | PASS |  | 160.7 | -0.02 | -0.99 | 56.76 | 51.48 | 1.44 | 0.13 | 1 | -2 | 1 | -0.9238 | 0.4706 | 14.1 | -19 |  |  | 0.1537 | 0.3632 | 0.0725 | 0.99971 |  |  |  |  | 133 |  | 0.0057 | [299 empty] | 0
P | 2018-09-13T17:12:41+08:00 | 2018-09-13T19:19:59+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | J18USM081G | CPSM Final Test | FAIL | [13 empty] | -6.7
P | 2018-09-13T17:13:59+08:00 | 2018-09-13T19:21:15+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | J18USM080G | CPSM Final Test | FAIL | [13 empty] | -5.1
P | 2018-09-13T17:12:47+08:00 | 2018-09-13T19:25:02+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | J18USM079G | CPSM Final Test | FAIL | [13 empty] | -6.9
P | 2018-09-13T16:00:55+08:00 | 2018-09-13T20:43:14+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | J18USM076G | CPSM Final Test | PASS |  | 167.99 | -0.02 | -0.84 | 56.62 | 50.88 | 1.36 | 0.13 | 1 | 4 | 1 | -0.9397 | 0.5152 | -2.3 | 9 |  |  | 0.0817 | 0.3311 | -0.0189 | 0.99998 |  |  |  |  | 133 |  | 0 | [339 empty] | 0
P | 2018-09-13T18:25:54+08:00 | 2018-09-13T21:18:59+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | J18USM082G | CPSM Final Test | PASS |  | 160.43 | 0.46 | -1.38 | 56.28 | 48.6 | 0.52 | 0.22 | 1 | -3 | 1 | -0.8859 | 0.544 | -1.4 | 11 |  |  | 0.1134 | 0.1549 | 0.0212 | 0.99999 |  |  |  |  | 133 |  | 0 | [45 empty] | 0
P | 2018-09-13T18:16:20+08:00 | 2018-09-13T21:23:29+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | J18USM081G | CPSM Final Test | PASS |  | 161.6 | 0.37 | -1.3 | 56.56 | 51.14 | 1.21 | 0.13 | -2 | -2 | 0 | -0.8121 | 0.5018 | -0.9 | 26 |  |  | 0.2135 | 0.2406 | -0.0043 | 0.99998 |  |  |  |  | 133 |  | 0 | [340 empty] | 0
P | 2018-09-13T18:17:37+08:00 | 2018-09-13T21:25:23+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | J18USM080G | CPSM Final Test | PASS |  | 161.51 | 0.27 | -0.99 | 56.93 | 52.41 | 1.13 | 0.13 | 1 | 3 | 0 | -0.9458 | 0.5198 | 1 | -8 |  |  | 0.0423 | 0.2206 | 0.0026 | 0.99999 |  |  |  |  | 133 |  | 0 | [341 empty] | 0
P | 2018-09-13T18:18:54+08:00 | 2018-09-13T21:29:02+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | J18USM079G | CPSM Final Test | PASS |  | 159.62 | 0.17 | -0.23 | 56.6 | 50.38 | 1.29 | 0.13 | -2 | -3 | 0 | -0.9557 | 0.462 | 0.8 | -5 |  |  | 0.048 | 0.1687 | -0.0173 | 0.99999 |  |  |  |  | 133 |  | 0 | [342 empty] | 0
P | 2018-09-14T09:39:00+08:00 | 2018-09-14T12:20:45+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI230G | CPSM Final Test | FAIL | [13 empty] | -7.1
P | 2018-09-14T09:49:47+08:00 | 2018-09-14T12:33:42+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI233G | CPSM Final Test | FAIL | [13 empty] | -7.5
P | 2018-09-14T10:30:33+08:00 | 2018-09-14T12:37:45+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI234G | CPSM Final Test | FAIL | [13 empty] | -17.3
P | 2018-09-14T10:41:06+08:00 | 2018-09-14T13:59:33+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI231G | CPSM Final Test | PASS | 129 | 162.14 | 0.27 | -0.84 | 56.41 | 54.61 | 1.51 | 0.13 | -1 | 1 | 0 | -0.8261 | 0.4658 | -3.7 | 11 |  | 0 | 0.2317 | 0.3622 | 0.024 | 0.99995 | [350 empty] | 0
P | 2018-09-14T10:51:37+08:00 | 2018-09-14T14:02:05+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI232G | CPSM Final Test | PASS | 129 | 159.62 | 0.27 | -0.69 | 56.5 | 49.28 | 1.67 | 0.13 | 1 | -2 | -1 | -0.8062 | 0.508 | -4.3 | 30 |  | 0.02134 | 0.1577 | 0.2886 | 0.0062 | 0.99996 | [351 empty] | 0
P | 2018-09-14T10:59:52+08:00 | 2018-09-14T14:20:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI230G | CPSM Final Test | PASS | 129 | 164.48 | 0.08 | -0.77 | 56.51 | 49.78 | 1.44 | 0.13 | 1 | -2 | -1 | -0.8321 | 0.518 | 3.3 | 14 |  | 0 | 0.0674 | 0.2352 | -0.0096 | 0.99999 | [89 empty] | 0
P | 2018-09-14T11:11:45+08:00 | 2018-09-14T14:34:14+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI233G | CPSM Final Test | PASS | 129 | 168.53 | 0.56 | 0.38 | 56.61 | 52.49 | 1.13 | 0.13 | -1 | -2 | 0 | -0.8739 | 0.533 | 4.1 | 1 |  | 0 | 0.1531 | 0.3275 | -0.0048 | 1 | [352 empty] | 0
P | 2018-09-14T11:34:09+08:00 | 2018-09-14T14:38:45+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI234G | CPSM Final Test | PASS | 129 | 161.51 | 0.17 | -0.54 | 56.8 | 51.48 | 1.82 | 0.13 | 1 | 3 | 0 | -0.8201 | 0.4958 | 7.1 | -2 |  | 0 | 0.0711 | 0.2889 | 0.0298 | 0.99998 | [353 empty] | 0
P | 2018-09-14T16:25:24+08:00 | 2018-09-14T17:48:32+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI240G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-09-14T15:43:44+08:00 | 2018-09-14T18:30:12+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI240G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-09-14T12:58:33+08:00 | 2018-09-14T19:07:23+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI237G | CPSM Final Test | FAIL | [13 empty] | -7.8
P | 2018-09-14T14:48:40+08:00 | 2018-09-14T19:25:16+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI240G | CPSM Final Test | FAIL | [13 empty] | -7.6
P | 2018-09-14T15:00:23+08:00 | 2018-09-14T19:19:11+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI242G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-09-14T12:38:10+08:00 | 2018-09-14T19:52:35+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI235G | CPSM Final Test | FAIL | [13 empty] | 29.5
P | 2018-09-14T14:32:08+08:00 | 2018-09-14T20:07:37+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | J18SDI241G | CPSM Final Test | FAIL
P | 2018-09-14T14:30:28+08:00 | 2018-09-14T20:09:17+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | J18SDI241G | CPSM Final Test | FAIL
P | 2018-09-14T13:28:17+08:00 | 2018-09-14T20:10:59+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI238G | CPSM Final Test | FAIL
P | 2018-09-14T16:25:04+08:00 | 2018-09-14T20:51:07+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI236G | CPSM Final Test | PASS | 129 | 162.77 | 0.46 | -1.15 | 56.43 | 56.05 | 1.74 | 0.22 | 1 | -2 | 0 | -0.9379 | 0.5368 | 3 | 1 |  | 0.0178 | 0.0652 | 0.7883 | -0.0081 | 0.99999 | [354 empty] | 0
P | 2018-09-14T16:37:00+08:00 | 2018-09-14T20:54:12+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI239G | CPSM Final Test | PASS | 129 | 169.07 | 0.17 | -0.84 | 56.63 | 49.95 | 1.74 | 0.13 | 0 | 1 | 0 | -0.8181 | 0.4972 | 0.1 | 7 |  | 0 | 0.061 | 0.157 | 0.0099 | 0.99999 | [18 empty] | 0
P | 2018-09-14T16:49:38+08:00 | 2018-09-14T21:16:39+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI238G | CPSM Final Test | PASS | 129 | 162.23 | 0.17 | -0.69 | 56.52 | 51.56 | 1.59 | 0.13 | -2 | 3 | 0 | -0.9698 | 0.5352 | 2.7 | 52 |  | 0 | 0.0329 | 0.2313 | -0.0129 | 0.99999 | [355 empty] | 0
P | 2018-09-14T17:06:58+08:00 | 2018-09-14T21:25:24+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI240G | CPSM Final Test | PASS | 129 | 156.74 | 0.17 | -0.46 | 56.66 | 48.68 | 1.51 | 0.13 | -2 | -2 | 0 | -0.9477 | 0.481 | 2.3 | -3 |  | 0 | 0.0783 | 0.2576 | 0.0205 | 0.99999 | [356 empty] | 0
P | 2018-09-14T17:09:47+08:00 | 2018-09-14T21:21:29+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI242G | CPSM Final Test | PASS | 129 | 166.1 | 0.17 | -0.99 | 56.48 | 49.7 | 1.59 | 0.13 | 1 | -4 | 0 | -0.8399 | 0.5286 | -2.1 | -8 |  | 0 | 0.0601 | 0.0428 | 0.0034 | 1 | [357 empty] | 0
P | 2018-09-13T23:38:54+08:00 | 2018-09-15T08:51:51+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI235G | CPSM Final Test | FAIL | [14 empty] | 1318
P | 2018-09-13T23:11:17+08:00 | 2018-09-15T08:54:39+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI237G | CPSM Final Test | FAIL | [379 empty] | 6424
P | 2018-09-14T01:16:16+08:00 | 2018-09-15T09:23:29+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | J18SDI241G | CPSM Final Test | FAIL
P | 2018-09-14T16:02:58+08:00 | 2018-09-15T10:00:19+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI237G | CPSM Final Test | PASS | 129 | 161.51 | 0.27 | -0.61 | 56.22 | 48.09 | 1.59 | 0.13 | 1 | -3 | 1 | -0.9339 | 0.4652 | 7.1 | -68 |  | 0.01692 | 0.1402 | 0.1713 | -0.0337 | 0.99997 | [358 empty] | 0
P | 2018-09-14T16:15:23+08:00 | 2018-09-15T10:36:47+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI235G | CPSM Final Test | PASS | 129 | 163.31 | 0.17 | -1.07 | 56.47 | 49.95 | 1.36 | 0.13 | 0 | -1 | 0 | -0.8321 | 0.5192 | 3.5 | 15 |  | 0.016 | 0.1476 | 0.2422 | 0.0569 | 0.99999 | [359 empty] | 0
P | 2018-09-13T23:35:58+08:00 | 2018-09-15T11:03:47+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | J18SDI241G | CPSM Final Test | FAIL | [19 empty] | 0.248
P | 2018-09-13T21:38:40+08:00 | 2018-09-15T13:01:05+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | J18SDI241G | CPSM Final Test | FAIL | [19 empty] | 0.2732
P | 2018-09-15T10:16:39+08:00 | 2018-09-15T15:08:07+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI245G | CPSM Final Test | FAIL
P | 2018-09-15T07:45:43+08:00 | 2018-09-15T15:12:23+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI248G | CPSM Final Test | FAIL | [14 empty] | 2514
P | 2018-09-15T13:08:51+08:00 | 2018-09-15T15:30:32+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI241G | CPSM Final Test | FAIL | [13 empty] | -6.1
P | 2018-09-15T11:52:43+08:00 | 2018-09-15T15:45:32+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI244G | CPSM Final Test | PASS | 129 | 166.82 | 0.37 | -0.08 | 56.52 | 49.45 | 0.98 | 0.13 | 1 | -4 | 0 | -0.9477 | 0.531 | -3.2 | -7 |  | 0.00923 | 0.0699 | 0.2784 | 0.0035 | 1 | [45 empty] | 0
P | 2018-09-15T12:51:32+08:00 | 2018-09-15T15:49:37+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI246G | CPSM Final Test | PASS | 129 | 159.35 | 0.08 | -1.07 | 56.42 | 48.94 | 1.36 | 0.13 | 1 | 2 | 0 | -0.9019 | 0.5316 | -3.9 | -23 |  | 0 | 0.1518 | 0.1805 | -0.0065 | 0.99999 | [360 empty] | 0
P | 2018-09-15T07:29:01+08:00 | 2018-09-15T15:56:58+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI247G | CPSM Final Test | FAIL
P | 2018-09-15T11:43:00+08:00 | 2018-09-15T15:57:54+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI247G | CPSM Final Test | PASS | 129 | 158.9 | -0.02 | -0.54 | 56.73 | 52.83 | 1.74 | 0.13 | 1 | 2 | 0 | -0.9756 | 0.5232 | -2.4 | 23 |  | 0 | 0.0562 | 0.0833 | 0.0371 | 0.99995 | [361 empty] | 0
P | 2018-09-15T11:43:59+08:00 | 2018-09-15T15:52:41+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI243G | CPSM Final Test | PASS | 129 | 160.34 | -0.02 | -0.46 | 56.72 | 53.51 | 1.74 | 0.13 | -1 | 4 | 0 | -1.1294 | 0.5026 | 3.4 | 11 |  | 0 | 0.0614 | 0.3187 | 0.0232 | 0.99999 | [355 empty] | 0
P | 2018-09-15T09:08:23+08:00 | 2018-09-15T16:16:23+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI245G | CPSM Final Test | FAIL
P | 2018-09-15T11:29:03+08:00 | 2018-09-15T16:56:51+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI248G | CPSM Final Test | PASS | 129 | 161.06 | 0.17 | -0.92 | 56.61 | 49.28 | 1.29 | 0.13 | 1 | 3 | 0 | -0.858 | 0.492 | -4.6 | 20 |  | 0 | 0.0367 | 0.0886 | -0.0254 | 1 | [362 empty] | 0
P | 2018-09-15T12:42:23+08:00 | 2018-09-15T17:21:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI245G | CPSM Final Test | PASS | 129 | 159.89 | -0.02 | -0.38 | 56.53 | 48.01 | 1.51 | 0.13 | 0 | -2 | 0 | -0.9478 | 0.5156 | 1.8 | -10 |  | 0 | 0.0384 | 0.1373 | 0.0207 | 1 | [133 empty] | 0
P | 2018-09-15T14:19:42+08:00 | 2018-09-15T17:29:58+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI241G | CPSM Final Test | PASS | 129 | 167.72 | 0.17 | -0.99 | 56.78 | 52.32 | 0.98 | 0.13 | -2 | -2 | 0 | -0.9178 | 0.5042 | 9.3 | 30 |  | 0 | 0.1452 | 0.1815 | 0.0275 | 0.99999 | [358 empty] | 0
P | 2018-09-17T10:06:22+08:00 | 2018-09-17T14:15:13+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI250G | CPSM Final Test | FAIL | [5 empty] | 39.03
P | 2018-09-17T10:26:54+08:00 | 2018-09-17T14:15:14+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI251G | CPSM Final Test | FAIL | [5 empty] | 45.47
P | 2018-09-17T09:54:38+08:00 | 2018-09-17T14:26:57+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI250G | CPSM Final Test | FAIL | [5 empty] | 44.37
P | 2018-09-17T11:53:42+08:00 | 2018-09-17T14:57:31+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI256G | CPSM Final Test | FAIL | [13 empty] | -9.3
P | 2018-09-17T12:10:53+08:00 | 2018-09-17T15:00:31+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI255G | CPSM Final Test | FAIL | [13 empty] | -10
P | 2018-09-17T11:54:28+08:00 | 2018-09-17T15:39:05+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI249G | CPSM Final Test | FAIL | [14 empty] | 1550
P | 2018-09-17T10:54:27+08:00 | 2018-09-17T16:16:57+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI255G | CPSM Final Test | FAIL
P | 2018-09-17T13:45:44+08:00 | 2018-09-17T16:43:02+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI254G | CPSM Final Test | PASS | 129 | 157.73 | -0.02 | -0.69 | 56.7 | 51.73 | 0.91 | 0.13 | 1 | -2 | 0 | -0.9038 | 0.4992 | -4.3 | -16 |  | 0 | 0.0993 | 0.2264 | 0.0388 | 0.99996 | [363 empty] | 0
P | 2018-09-17T11:48:49+08:00 | 2018-09-17T16:45:36+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI252G | CPSM Final Test | PASS | 129 | 159.44 | 0.37 | -0.84 | 56.54 | 50.97 | 1.06 | 0.13 | 0 | -3 | 0 | -0.872 | 0.5 | -5 | 9 |  | 0.00978 | 0.1571 | 0.1836 | 0.035 | 0.99998 | [20 empty] | 0
P | 2018-09-17T12:21:04+08:00 | 2018-09-17T16:47:46+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI251G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -1.15 | 56.55 | 48.6 | 1.67 | 0.13 | -1 | 3 | 1 | -0.9238 | 0.5138 | -2.5 | 11 |  | 0 | 0.123 | 0.4939 | -0.0253 | 1 | [364 empty] | 0
P | 2018-09-17T13:25:36+08:00 | 2018-09-17T16:57:43+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI256G | CPSM Final Test | PASS | 129 | 165.56 | 0.17 | -0.69 | 56.17 | 51.82 | 1.13 | 0.13 | -1 | 1 | 0 | -0.9118 | 0.5188 | -1.8 | 8 |  | 0 | 0.1161 | 0.1872 | 0.0436 | 0.99996 | [46 empty] | 0
P | 2018-09-17T12:10:47+08:00 | 2018-09-17T17:00:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI250G | CPSM Final Test | PASS | 129 | 160.88 | 0.08 | -0.54 | 56.6 | 50.55 | 1.44 | 0.13 | 1 | -4 | 0 | -0.9437 | 0.5192 | 1.2 | 0 |  | 0 | 0.1016 | 0.2677 | 0.022 | 1 | [133 empty] | 0
P | 2018-09-17T13:35:42+08:00 | 2018-09-17T17:23:29+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI255G | CPSM Final Test | PASS | 129 | 164.03 | -0.02 | -0.08 | 56.65 | 52.24 | 1.51 | 0.13 | 0 | -4 | 0 | -0.84 | 0.5094 | 2.8 | 64 |  | 0 | 0.0511 | 0.2846 | -0.0157 | 0.99999 | [295 empty] | 0
P | 2018-09-17T13:46:46+08:00 | 2018-09-17T17:23:41+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI249G | CPSM Final Test | PASS | 129 | 163.22 | 0.37 | -0.92 | 56.71 | 54.61 | 1.21 | 0.18 | -1 | -2 | 0 | -0.9698 | 0.5188 | -2.6 | 28 |  | 0 | 0.0285 | 0.1823 | -0.036 | 1 | [365 empty] | 0
P | 2018-09-17T15:28:55+08:00 | 2018-09-17T18:40:40+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI258G | CPSM Final Test | FAIL | [13 empty] | -7.8
P | 2018-09-17T15:34:57+08:00 | 2018-09-17T19:22:19+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI253G | CPSM Final Test | FAIL
P | 2018-09-17T14:26:04+08:00 | 2018-09-17T20:10:39+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI257G | CPSM Final Test | FAIL | [387 empty] | 0
P | 2018-09-17T17:36:33+08:00 | 2018-09-17T20:11:01+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI261G | CPSM Final Test | FAIL | [13 empty] | -8.9
P | 2018-09-17T16:56:05+08:00 | 2018-09-17T20:22:02+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI262G | CPSM Final Test | PASS | 129 | 167.18 | 0.08 | -0.77 | 56.33 | 50.72 | 1.29 | 0.13 | -1 | 1 | -1 | -0.9637 | 0.4758 | -3.8 | 14 |  | 0 | 0.0609 | 0.1339 | -0.005 | 1 | [204 empty] | 0
P | 2018-09-17T16:48:31+08:00 | 2018-09-17T20:29:50+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI259G | CPSM Final Test | PASS | 129 | 168.71 | -0.02 | -1.15 | 56.47 | 48.77 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9757 | 0.5262 | -4.4 | 1 |  | 0.03973 | 0.143 | 0.0961 | 0.0118 | 0.99999 | [125 empty] | 0
P | 2018-09-17T17:04:48+08:00 | 2018-09-17T20:41:30+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI258G | CPSM Final Test | PASS | 129 | 165.56 | 0.46 | 0.15 | 56.75 | 49.61 | 1.67 | 0.13 | 1 | 2 | 0 | -0.9637 | 0.5284 | -2.3 | 11 |  | 0.0343 | 0.0681 | 0.2467 | 0.0354 | 0.99998 | [367 empty] | 0
P | 2018-09-17T16:46:35+08:00 | 2018-09-17T21:00:59+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI261G | CPSM Final Test | FAIL | [13 empty] | -8.4
P | 2018-09-17T13:51:35+08:00 | 2018-09-17T21:05:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI253G | CPSM Final Test | FAIL
P | 2018-09-17T18:30:46+08:00 | 2018-09-17T21:15:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI260G | CPSM Final Test | PASS | 129 | 163.49 | 0.27 | -0.99 | 56.6 | 52.66 | 1.13 | 0.13 | 1 | 2 | 0 | -0.9338 | 0.5206 | -0.5 | -5 |  | 0.02603 | 0.1458 | 0.2482 | 0.017 | 0.99998 | [295 empty] | 0
P | 2018-09-17T17:18:22+08:00 | 2018-09-17T21:16:11+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI257G | CPSM Final Test | PASS | 129 | 164.84 | 0.46 | -0.61 | 56.42 | 55.37 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9337 | 0.514 | 2.5 | 9 |  | 0 | 0.0444 | 0.1232 | 0.0322 | 0.99999 | [366 empty] | 0
P | 2018-09-17T17:28:38+08:00 | 2018-09-17T21:18:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI253G | CPSM Final Test | PASS | 129 | 172.13 | 0.37 | -0.84 | 56.7 | 54.61 | 0.98 | 0.13 | 0 | 1 | 0 | -0.9597 | 0.5272 | 1.4 | 9 |  | 0 | 0.044 | 0.196 | 0.0056 | 1 | [368 empty] | 0
P | 2018-09-17T20:37:45+08:00 | 2018-09-17T23:26:46+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI267G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -0.84 | 56.18 | 50.63 | 0.83 | 0.13 | -1 | -2 | 0 | -0.8859 | 0.5162 | 2.2 | 3 |  | 0 | 0.0956 | 0.2587 | 0.046 | 0.99997 | [369 empty] | 0
P | 2018-09-17T05:13:01+08:00 | 2018-09-18T08:34:33+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI261G | CPSM Final Test | FAIL | [14 empty] | 3166
P | 2018-09-17T08:16:16+08:00 | 2018-09-18T09:17:32+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | J18SDI264G | CPSM Final Test | FAIL | [14 empty] | 1015
P | 2018-09-17T20:34:57+08:00 | 2018-09-18T09:41:34+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI268G | CPSM Final Test | PASS | 129 | 168.44 | -0.02 | -0.54 | 56.16 | 51.22 | 1.06 | 0.13 | 0 | -2 | 0 | -0.9179 | 0.538 | -2.8 | 5 |  | 0 | 0.0455 | 0.1619 | -0.0441 | 1 | [188 empty] | 0
P | 2018-09-17T18:53:47+08:00 | 2018-09-18T10:19:23+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI261G | CPSM Final Test | PASS | 129 | 165.11 | 0.17 | -0.61 | 56.51 | 50.38 | 1.51 | 0.13 | 1 | -4 | 0 | -0.9658 | 0.5336 | -3.8 | -3 |  | 0 | 0.1413 | 0.1336 | 0.0087 | 1 | [370 empty] | 0
P | 2018-09-17T21:18:11+08:00 | 2018-09-18T10:59:05+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI263G | CPSM Final Test | PASS | 129 | 164.03 | -0.02 | -0.61 | 56.31 | 51.31 | 1.21 | 0.13 | -1 | -1 | 1 | -0.7603 | 0.5092 | -3.5 | 25 |  | 0 | 0.0613 | 0.212 | 0.0101 | 0.99996 | [371 empty] | 0
P | 2018-09-17T20:46:54+08:00 | 2018-09-18T11:02:22+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | J18SDI264G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.16 | 56.82 | 52.24 | 1.59 | 0.13 | 0 | -1 | 1 | -0.9138 | 0.4974 | 5.6 | 1 |  | 0 | 0.0749 | 0.0859 | 0.0358 | 1 | [372 empty] | 0
P | 2018-09-18T12:11:06+08:00 | 2018-09-18T14:46:03+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI266G | CPSM Final Test | FAIL | [13 empty] | -7.1
P | 2018-09-18T11:40:19+08:00 | 2018-09-18T15:42:30+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI269G | CPSM Final Test | PASS | 129 | 165.56 | -0.02 | -0.31 | 56.55 | 51.82 | 1.59 | 0.13 | -2 | -2 | 0 | -0.9518 | 0.5176 | 3.1 | -3 |  | 0 | 0.1471 | 0.469 | 0.0319 | 0.99999 | [302 empty] | 0
P | 2018-09-18T13:28:34+08:00 | 2018-09-18T16:46:05+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | J18SDI266G | CPSM Final Test | PASS | 129 | 170.87 | 0.17 | -0.08 | 57.01 | 54.86 | 0.91 | 0.13 | -1 | 5 | 0 | -0.9479 | 0.5206 | 1.7 | 7 |  | 0 | 0.1128 | 0.2491 | 0.0739 | 0.99997 | [366 empty] | 0
P | 2018-09-18T15:43:29+08:00 | 2018-09-18T18:38:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI270G | CPSM Final Test | FAIL | [13 empty] | -9.1
P | 2018-09-18T15:37:32+08:00 | 2018-09-18T18:41:31+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI271G | CPSM Final Test | FAIL | [13 empty] | -6.2
P | 2018-09-18T15:33:39+08:00 | 2018-09-18T18:42:49+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI273G | CPSM Final Test | FAIL | [13 empty] | -8.4
P | 2018-09-18T17:12:05+08:00 | 2018-09-18T20:17:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | J18SDI265G | CPSM Final Test | PASS | 129 | 159.53 | 0.17 | -0.99 | 56.51 | 51.22 | 0.98 | 0.13 | 0 | 1 | 0 | -0.9139 | 0.4862 | -3 | 7 |  | 0 | 0.0874 | 0.2187 | -0.0495 | 0.99999 | [373 empty] | 0
P | 2018-09-18T17:21:14+08:00 | 2018-09-18T20:35:10+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | J18SDI274G | CPSM Final Test | PASS | 129 | 163.76 | 0.27 | -0.23 | 56.57 | 54.02 | 0.98 | 0.13 | -2 | 3 | 0 | -0.9059 | 0.528 | -4.3 | 8 |  | 0.08384 | 0.2056 | 0.1634 | 0.0049 | 0.99998 | [374 empty] | 0
P | 2018-09-18T17:20:37+08:00 | 2018-09-18T20:31:48+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | J18SDI275G | CPSM Final Test | PASS | 129 | 168.26 | 0.46 | -1.07 | 56.24 | 52.83 | 0.83 | 0.13 | -1 | 3 | 1 | -0.9598 | 0.5228 | -3.3 | -21 |  | 0 | 0.1349 | 0.148 | -0.0058 | 1 | [375 empty] | 0
P | 2018-09-18T17:10:53+08:00 | 2018-09-18T20:38:07+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | J18SDI270G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -0.61 | 56.94 | 51.14 | 1.44 | 0.13 | 1 | -3 | 0 | -0.9419 | 0.5066 | 0.8 | 22 |  | 0 | 0.1185 | 0.1243 | -0.0293 | 0.99998 | [8 empty] | 0
P | 2018-09-18T17:09:31+08:00 | 2018-09-18T20:41:03+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | J18SDI271G | CPSM Final Test | PASS | 129 | 165.02 | 0.37 | -0.23 | 56.94 | 50.29 | 0.91 | 0.13 | -1 | 1 | 0 | -0.9637 | 0.515 | 3.9 | 37 |  | 0 | 0.1458 | 0.1738 | -0.0337 | 0.99998 | [376 empty] | 0
P | 2018-09-18T17:08:14+08:00 | 2018-09-18T20:42:39+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | J18SDI273G | CPSM Final Test | PASS | 129 | 161.51 | 0.08 | 0.15 | 56.42 | 51.73 | 1.51 | 0.13 | 0 | -3 | 0 | -0.8599 | 0.5166 | 4.7 | -1 |  | 0.01109 | 0.1267 | 0.0736 | 0.0283 | 0.99993 | [377 empty] | 0
P | 2018-09-18T02:12:44+08:00 | 2018-09-19T08:30:33+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | J18SDI272G | CPSM Final Test | FAIL
P | 2018-09-18T17:21:39+08:00 | 2018-09-19T08:43:31+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | J18SDI272G | CPSM Final Test | PASS | 129 | 157.1 | 0.17 | -0.84 | 56.62 | 51.05 | 1.21 | 0.13 | 1 | 2 | 1 | -0.7862 | 0.541 | -3.8 | -21 |  | 0.00745 | 0.1684 | 0.5334 | 0.1523 | 0.99972 | [378 empty] | 0
P | 2018-09-19T12:32:40+08:00 | 2018-09-19T14:07:48+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP660G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-09-19T11:54:26+08:00 | 2018-09-19T14:46:02+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP660G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-09-19T10:45:44+08:00 | 2018-09-19T15:54:44+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP660G | CPSM Final Test | FAIL | [5 empty] | 45.21
P | 2018-09-19T10:31:01+08:00 | 2018-09-19T16:19:51+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP663G | CPSM Final Test | FAIL | [400 empty] | FALSE
P | 2018-09-19T13:23:58+08:00 | 2018-09-19T17:24:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP666G | CPSM Final Test | PASS |  | 25.88 | 0.37 | 0.15 | 56.75 | 50.8 | 1.13 | 0.13 | 0 | 2 | 0 | -0.9876 | 0.5112 | 6.3 | -1 |  |  | 0.1844 | 0.1981 | 0.0142 | 0.99997 |  |  |  |  | 133 |  | 0 | [373 empty] | 0
P | 2018-09-19T13:22:30+08:00 | 2018-09-19T17:26:30+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP667G | CPSM Final Test | PASS |  | 162.77 | -0.02 | -0.54 | 56.77 | 51.48 | 1.21 | 0.13 | 0 | 1 | 0 | -1.0056 | 0.518 | 3.5 | 3 |  |  | 0.0661 | 0.2676 | -0.0033 | 1 |  |  |  |  | 133 |  | 0.0169 | [374 empty] | 0
P | 2018-09-19T14:07:57+08:00 | 2018-09-19T17:30:20+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP668G | CPSM Final Test | PASS |  | 162.32 | 0.27 | -0.84 | 56.56 | 50.46 | 1.06 | 0.13 | 1 | -2 | 0 | -0.9299 | 0.52 | -2.1 | 2 |  |  | 0.0742 | 0.1551 | -0.0459 | 1 |  |  |  |  | 133 |  | 0 | [115 empty] | 0
P | 2018-09-19T14:10:25+08:00 | 2018-09-19T17:37:54+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | J18FNP661G | CPSM Final Test | PASS |  | 165.47 | 0.27 | -0.38 | 56.55 | 51.65 | 1.51 | 0.13 | 0 | 1 | 0 | -0.9298 | 0.5118 | 10.1 | 5 |  |  | 0.1824 | 0.3088 | 0.0414 | 0.99998 |  |  |  |  | 133 |  | 0 | [375 empty] | 0
P | 2018-09-19T13:20:14+08:00 | 2018-09-19T18:46:28+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP660G | CPSM Final Test | PASS |  | 162.5 | 0.17 | -0.99 | 56.25 | 51.05 | 1.06 | 0.13 | -1 | 3 | 0 | -0.8957 | 0.5166 | 6.2 | 2 |  |  | 0.0989 | 0.1816 | 0.0494 | 0.99994 |  |  |  |  | 133 |  | 0 | [376 empty] | 0
P | 2018-09-19T13:25:26+08:00 | 2018-09-19T19:23:27+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP663G | CPSM Final Test | PASS |  | 162.5 | 0.17 | 0 | 56.49 | 52.07 | 1.21 | -0.14 | -1 | -1 | 0 | -0.9238 | 0.5282 | 9.2 | 36 |  |  | 0.1896 | 0.1158 | 0.0242 | 0.99997 |  |  |  |  | 133 |  | 0 | [377 empty] | 0
P | 2018-09-19T08:18:13+08:00 | 2018-09-19T19:46:09+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP662G | CPSM Final Test | FAIL | [19 empty] | 0.2333
P | 2018-09-19T07:13:20+08:00 | 2018-09-20T09:04:39+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP665G | CPSM Final Test | FAIL | [348 empty] | FALSE
P | 2018-09-18T18:58:45+08:00 | 2018-09-20T09:05:37+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP662G | CPSM Final Test | FAIL | [19 empty] | 0.2327
P | 2018-09-19T06:16:44+08:00 | 2018-09-20T10:01:15+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | J18FNP665G | CPSM Final Test | FAIL | [13 empty] | 26.3
P | 2018-09-18T17:35:13+08:00 | 2018-09-20T10:29:09+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP662G | CPSM Final Test | FAIL
P | 2018-09-19T19:38:21+08:00 | 2018-09-20T11:15:37+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | J18FNP664G | CPSM Final Test | PASS |  | 158.18 | 0.87 | -0.61 | 56.77 | 53.09 | 1.21 | 0.13 | -1 | -2 | 0 | -0.892 | 0.4758 | 4.6 | 38 |  |  | 0.4074 | 0.4261 | -0.0395 | 0.99999 |  |  |  |  | 133 |  | 0 | [378 empty] | 0
P | 2018-09-19T20:31:27+08:00 | 2018-09-20T11:24:20+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | J18FNP671G | CPSM Final Test | PASS |  | 123.8 | 0.17 | -0.16 | 56.33 | 49.95 | 0.98 | 0.13 | -1 | 1 | 0 | -0.858 | 0.489 | 5.8 | 14 |  |  | 0.1115 | 0.1926 | 0.0171 | 0.99999 |  |  |  |  | 133 |  | 0 | [379 empty] | 0
P | 2018-09-19T20:21:39+08:00 | 2018-09-20T11:18:38+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP670G | CPSM Final Test | PASS |  | 162.41 | 0.27 | -0.99 | 56.5 | 48.26 | 1.06 | 0.13 | 0 | 2 | 1 | -0.838 | 0.4972 | 5.4 | 40 |  |  | 0.1438 | 0.2653 | 0.0263 | 0.99996 |  |  |  |  | 133 |  | 0 | [145 empty] | 0
P | 2018-09-19T20:17:42+08:00 | 2018-09-20T11:21:10+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP669G | CPSM Final Test | PASS |  | 158.45 | -0.02 | -0.46 | 56.61 | 50.97 | 1.21 | 0.13 | -1 | -2 | 0 | -0.87 | 0.5342 | 3.3 | 6 |  |  | 0.0584 | 0.2205 | -0.0071 | 1 |  |  |  |  | 133 |  | 0 | [380 empty] | 0
P | 2018-09-19T20:14:02+08:00 | 2018-09-20T11:24:42+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP674G | CPSM Final Test | PASS |  | 164.66 | 0.17 | -0.92 | 56.6 | 52.66 | 1.21 | 0.13 | 0 | 3 | -1 | -0.7642 | 0.52 | 6 | -4 |  |  | 0.0383 | 0.1369 | -0.0531 | 0.99999 |  |  |  |  | 133 |  | 0 | [381 empty] | 0
P | 2018-09-19T20:05:43+08:00 | 2018-09-20T11:38:11+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP672G | CPSM Final Test | PASS |  | 158.72 | 0.27 | -0.46 | 56.61 | 49.36 | 1.29 | 0.13 | -1 | -2 | -1 | -0.866 | 0.4912 | 2 | 41 |  |  | 0.0924 | 0.1544 | 0.0242 | 0.99998 |  |  |  |  | 133 |  | 0 | [52 empty] | 0
P | 2018-09-20T11:20:43+08:00 | 2018-09-20T13:44:20+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP677G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-09-20T11:18:25+08:00 | 2018-09-20T13:46:38+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP677G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-09-20T11:18:34+08:00 | 2018-09-20T13:52:42+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | J18FNP675G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-09-18T13:58:01+08:00 | 2018-09-20T14:06:21+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP662G | CPSM Final Test | FAIL | [19 empty] | 0.261
P | 2018-09-20T10:13:00+08:00 | 2018-09-20T14:38:08+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP680G | CPSM Final Test | FAIL
P | 2018-09-20T12:36:07+08:00 | 2018-09-20T16:19:05+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | J18FNP676G | CPSM Final Test | PASS |  | 164.12 | 0.17 | -0.61 | 56.51 | 50.88 | 1.36 | 0.13 | 1 | 3 | 0 | -0.9498 | 0.4912 | 0.3 | 18 |  |  | 0.0987 | 0.0998 | -0.0346 | 0.99999 |  |  |  |  | 133 |  | 0 | [382 empty] | 0
P | 2018-09-20T12:32:32+08:00 | 2018-09-20T16:19:12+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | J18FNP677G | CPSM Final Test | PASS |  | 169.61 | 0.97 | -0.84 | 56.64 | 48.18 | 0.37 | 0.13 | 1 | 1 | 0 | -1.0097 | 0.5042 | 7.9 | 18 |  |  | 0.2093 | 0.4592 | 0.0229 | 0.99999 |  |  |  |  | 133 |  | 0 | [383 empty] | 0
P | 2018-09-20T12:22:41+08:00 | 2018-09-20T16:19:40+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP678G | CPSM Final Test | PASS |  | 164.21 | 0.17 | -0.99 | 56.51 | 50.55 | 1.29 | 0.13 | 1 | 2 | 0 | -0.9378 | 0.472 | 2.8 | 1 |  |  | 0.1575 | 0.3682 | 0.021 | 1 |  |  |  |  | 133 |  | 0 | [384 empty] | 0
P | 2018-09-20T12:18:43+08:00 | 2018-09-20T16:19:45+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP673G | CPSM Final Test | PASS |  | 157.91 | 0.27 | -1.15 | 56.49 | 50.63 | 1.21 | 0.13 | 1 | -4 | 0 | -0.9377 | 0.5394 | 5.8 | 3 |  |  | 0.1174 | 0.1693 | -0.0017 | 1 |  |  |  |  | 133 |  | 0 | [385 empty] | 0
P | 2018-09-20T12:35:38+08:00 | 2018-09-20T16:31:14+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | J18FNP675G | CPSM Final Test | PASS |  | 159.44 | 0.17 | -1.15 | 56.47 | 48.51 | 1.13 | 0.13 | -1 | 1 | 0 | -0.9637 | 0.51 | 3.8 | 43 |  |  | 0.0872 | 0.2615 | 0.0296 | 0.99996 |  |  |  |  | 133 |  | 0 | [386 empty] | 0
P | 2018-09-20T08:03:20+08:00 | 2018-09-20T16:47:48+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP680G | CPSM Final Test | FAIL
P | 2018-09-20T05:45:54+08:00 | 2018-09-20T19:05:14+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP680G | CPSM Final Test | FAIL
P | 2018-09-18T08:58:31+08:00 | 2018-09-20T19:05:51+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP662G | CPSM Final Test | FAIL | [19 empty] | 0.2596
P | 2018-09-20T15:46:22+08:00 | 2018-09-20T19:10:39+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP031G | CPSM Final Test | FAIL | [13 empty] | 27.3
P | 2018-09-20T14:56:35+08:00 | 2018-09-20T19:16:19+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP027G | CPSM Final Test | FAIL | [13 empty] | 21
P | 2018-09-20T15:06:54+08:00 | 2018-09-20T19:50:07+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP031G | CPSM Final Test | FAIL | [13 empty] | 15.5
P | 2018-09-19T14:02:11+08:00 | 2018-09-20T20:10:35+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | J18FNP662G | CPSM Final Test | PASS |  | 158.09 | 0.37 | -0.84 | 57.05 | 49.28 | 1.36 | 0.13 | 0 | 1 | 0 | -0.9319 | 0.5052 | 0.8 | 7 |  |  | 0.185 | 0.0585 | 0.1268 | 0.99982 |  |  |  |  | 133 |  | 0.01688 | [341 empty] | 0
P | 2018-09-20T12:25:34+08:00 | 2018-09-20T20:10:42+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | J18FNP680G | CPSM Final Test | PASS |  | 163.31 | 0.27 | -0.84 | 56.49 | 50.12 | 1.06 | 0.13 | 0 | 2 | 0 | -0.9498 | 0.527 | 3.4 | 32 |  |  | 0.0755 | 0.2345 | -0.1082 | 0.99999 |  |  |  |  | 133 |  | 0 | [387 empty] | 0
P | 2018-09-20T14:33:04+08:00 | 2018-09-20T20:23:57+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP031G | CPSM Final Test | FAIL | [14 empty] | 1186
P | 2018-09-20T17:14:00+08:00 | 2018-09-20T20:54:06+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP028G | CPSM Final Test | PASS | 129 | 162.68 | -0.02 | -1.15 | 56.46 | 50.29 | 1.06 | 0.13 | 1 | 2 | 0 | -0.9339 | 0.549 | 5.3 | 7 |  |  | 0.0951 | 0.1297 | 0.0173 | 0.99999 | [21 empty] | 0 | [279 empty] | 0
P | 2018-09-20T17:30:38+08:00 | 2018-09-20T21:00:58+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP029G | CPSM Final Test | PASS | 129 | 169.07 | -0.02 | -0.77 | 56.51 | 49.11 | 1.36 | 0.13 | 1 | -4 | 1 | -0.872 | 0.5264 | 2.9 | 2 |  |  | 0.0678 | 0.1473 | 0.0706 | 0.99995 | [21 empty] | 0 | [373 empty] | 0
P | 2018-09-20T17:09:56+08:00 | 2018-09-20T20:55:12+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | J18FNP679G | CPSM Final Test | PASS |  | 155.12 | 0.27 | -0.54 | 56.5 | 49.61 | 1.06 | 0.13 | 0 | 2 | 0 | -0.9437 | 0.5336 | 10.2 | 0 |  |  | 0.1019 | 0.232 | 0.003 | 0.99999 |  |  |  |  | 133 |  | 0 | [389 empty] | 0
P | 2018-09-20T01:34:07+08:00 | 2018-09-21T08:38:47+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP027G | CPSM Final Test | FAIL
P | 2018-09-20T17:06:27+08:00 | 2018-09-21T08:45:51+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP027G | CPSM Final Test | PASS | 129 | 161.6 | -0.02 | -0.69 | 56.74 | 51.56 | 1.59 | 0.13 | 1 | -2 | 0 | -0.838 | 0.5388 | 10.5 | -47 |  |  | 0.0887 | 0.0499 | 0.0013 | 0.99996 | [21 empty] | 0.005 | [302 empty] | 0
P | 2018-09-20T01:32:00+08:00 | 2018-09-21T09:25:01+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP031G | CPSM Final Test | FAIL | [14 empty] | 860
P | 2018-09-21T10:03:36+08:00 | 2018-09-21T10:55:52+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP032G | CPSM Final Test | FAIL |  |  |  | -9.33
P | 2018-09-20T17:28:31+08:00 | 2018-09-21T11:09:17+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP031G | CPSM Final Test | PASS | 129 | 162.68 | 0.37 | -0.31 | 56.51 | 48.26 | 1.36 | 0.13 | 0 | -1 | 0 | -0.7802 | 0.4968 | 12.4 | 32 |  |  | 0.0364 | 0.1181 | 0.0103 | 0.99999 | [21 empty] | 0 | [375 empty] | 0
P | 2018-09-21T09:44:58+08:00 | 2018-09-21T11:14:30+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP032G | CPSM Final Test | FAIL
P | 2018-09-21T08:09:24+08:00 | 2018-09-21T12:41:34+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP037G | CPSM Final Test | FAIL
P | 2018-09-21T10:14:53+08:00 | 2018-09-21T13:46:43+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | J18FNP665G | CPSM Final Test | PASS |  | 155.57 | 0.17 | -0.99 | 56.55 | 51.39 | 1.36 | 0.13 | -2 | -2 | 1 | -0.8681 | 0.4852 | 1.4 | -1 |  |  | 0.0639 | 0.2281 | -0.0211 | 0.99998 |  |  |  |  | 133 |  | 0 | [338 empty] | 0
P | 2018-09-21T10:41:08+08:00 | 2018-09-21T13:53:08+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP038G | CPSM Final Test | PASS | 129 | 170.42 | -0.02 | -0.31 | 56.34 | 50.04 | 0.91 | 0.13 | 1 | 2 | 0 | -0.9259 | 0.521 | 5.1 | -14 |  |  | 0.0894 | 0.3654 | -0.0209 | 0.99997 | [21 empty] | 0 | [376 empty] | 0
P | 2018-09-21T10:38:40+08:00 | 2018-09-21T13:49:09+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP034G | CPSM Final Test | PASS | 129 | 168.89 | -0.02 | -1.15 | 56.48 | 48.18 | 0.91 | 0.13 | 1 | 4 | 0 | -0.8261 | 0.5298 | 2.1 | 17 |  |  | 0.236 | 0.0891 | 0.0621 | 0.99989 | [21 empty] | 0.00711 | [377 empty] | 0
P | 2018-09-21T10:29:44+08:00 | 2018-09-21T13:53:58+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP032G | CPSM Final Test | PASS | 129 | 158.72 | 0.37 | -1.45 | 56.12 | 50.46 | 1.21 | 0.13 | 1 | -13 | 1 | -0.9497 | 0.5308 | 5.5 | -14 |  |  | 0.0993 | 0.0542 | 0.03 | 0.99999 | [21 empty] | 0 | [378 empty] | 0
P | 2018-09-21T10:25:29+08:00 | 2018-09-21T14:27:40+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP037G | CPSM Final Test | PASS | 129 | 157.1 | 0.17 | -1.07 | 56.41 | 51.65 | 0.83 | 0.13 | 0 | 3 | 1 | -0.854 | 0.4806 | 5.9 | 18 |  |  | 0.0636 | 0.2688 | 0.0005 | 0.99998 | [21 empty] | 0 | [372 empty] | 0
P | 2018-09-21T11:15:45+08:00 | 2018-09-21T14:33:29+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP036G | CPSM Final Test | PASS | 129 | 156.02 | -0.02 | -0.61 | 56.46 | 49.7 | 1.29 | 0.13 | -1 | -1 | 1 | -0.9099 | 0.528 | 3.2 | 10 |  |  | 0.1025 | 0.4434 | 0.0541 | 0.99999 | [21 empty] | 0 | [280 empty] | 0
P | 2018-09-21T14:22:30+08:00 | 2018-09-21T16:29:05+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP042G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-09-21T14:05:08+08:00 | 2018-09-21T17:23:08+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP033G | CPSM Final Test | FAIL | [13 empty] | 17.7
P | 2018-09-21T13:50:59+08:00 | 2018-09-21T17:32:33+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP035G | CPSM Final Test | FAIL
P | 2018-09-21T15:22:12+08:00 | 2018-09-21T18:36:01+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP040G | CPSM Final Test | PASS | 129 | 161.69 | 0.17 | -0.77 | 56.77 | 48.68 | 1.29 | 0.13 | 0 | -1 | 0 | -0.9538 | 0.45 | 6.4 | 7 |  |  | 0.0821 | 0.097 | 0.0013 | 0.99997 | [21 empty] | 0 | [379 empty] | 0
P | 2018-09-21T14:58:29+08:00 | 2018-09-21T18:38:24+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP030G | CPSM Final Test | PASS | 129 | 164.84 | 0.27 | -0.84 | 56.48 | 49.78 | 0.6 | 0.13 | -2 | 1 | 0 | -0.9298 | 0.4756 | 4.3 | 0 |  |  | 0.1359 | 0.2758 | 0.0139 | 0.99999 | [21 empty] | 0 | [380 empty] | 0
P | 2018-09-21T15:29:24+08:00 | 2018-09-21T18:39:52+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP041G | CPSM Final Test | PASS | 129 | 169.97 | -0.02 | -0.77 | 56.52 | 50.21 | 1.44 | 0.13 | 1 | -2 | 1 | -0.9338 | 0.516 | 4 | -6 |  |  | 0.0669 | 0.1431 | -0.0101 | 1 | [21 empty] | 0 | [372 empty] | 0
P | 2018-09-21T15:32:42+08:00 | 2018-09-21T18:41:44+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP043G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | 0.07 | 56.71 | 55.88 | 1.51 | 0.13 | -2 | 3 | 1 | -0.946 | 0.5132 | 6.3 | 10 |  |  | 0.0792 | 0.1093 | 0.0317 | 0.99995 | [21 empty] | 0 | [378 empty] | 0
P | 2018-09-21T15:44:08+08:00 | 2018-09-21T19:23:24+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP033G | CPSM Final Test | PASS | 129 | 172.13 | -0.02 | -0.69 | 56.77 | 49.11 | 1.67 | 0.13 | 0 | -2 | 1 | -0.9658 | 0.5454 | -3.1 | -5 |  |  | 0.2464 | 0.2631 | 0.0229 | 0.99999 | [21 empty] | 0 | [381 empty] | 0
P | 2018-09-21T12:01:51+08:00 | 2018-09-21T19:21:41+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP035G | CPSM Final Test | FAIL | [14 empty] | 4008
P | 2018-09-21T15:25:48+08:00 | 2018-09-21T19:28:15+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP042G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -1.07 | 56.48 | 50.12 | 0.83 | 0.13 | -1 | -1 | 0 | -0.844 | 0.5382 | 5.1 | 6 |  |  | 0.0964 | 0.3286 | 0.0456 | 0.99996 | [21 empty] | 0 | [382 empty] | 0
P | 2018-09-21T15:41:46+08:00 | 2018-09-21T21:05:59+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP035G | CPSM Final Test | PASS | 129 | 166.73 | 0.27 | -1.3 | 56.46 | 50.63 | 0.68 | 0.13 | -1 | 3 | 1 | -0.8759 | 0.536 | -0.8 | 22 |  |  | 0.2336 | 0.1062 | 0.03 | 0.99997 | [21 empty] | 0.0053 | [383 empty] | 0
P | 2018-09-22T09:47:03+08:00 | 2018-09-22T10:38:22+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP044G | CPSM Final Test | FAIL | [5 empty] | 45.72
P | 2018-09-22T10:03:37+08:00 | 2018-09-22T10:44:23+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP048G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-09-22T08:48:15+08:00 | 2018-09-22T11:25:21+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP051G | CPSM Final Test | FAIL | [13 empty] | 15.7
P | 2018-09-22T08:32:41+08:00 | 2018-09-22T11:46:16+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP050G | CPSM Final Test | FAIL
P | 2018-09-22T10:16:32+08:00 | 2018-09-22T13:14:46+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP046G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -0.92 | 56.72 | 54.69 | 1.13 | 0.13 | 1 | -2 | -1 | -0.9139 | 0.5028 | 9.5 | 8 |  |  | 0.0764 | 0.3188 | 0.0006 | 0.99998 | [21 empty] | 0 | [384 empty] | 0
P | 2018-09-22T10:26:32+08:00 | 2018-09-22T13:22:02+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP049G | CPSM Final Test | PASS | 129 | 160.97 | -0.02 | -0.61 | 56.73 | 54.61 | 0.98 | 0.13 | -1 | 3 | 0 | -0.9019 | 0.519 | 4.4 | -1 |  |  | 0.1079 | 0.2382 | -0.0057 | 0.99997 | [21 empty] | 0 | [385 empty] | 0
P | 2018-09-22T10:12:42+08:00 | 2018-09-22T13:18:14+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP044G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.54 | 57.11 | 48.01 | 0.91 | 0.13 | -1 | 1 | -1 | -0.844 | 0.5244 | 2 | 47 |  |  | 0.0781 | 0.2245 | 0.0267 | 0.99999 | [21 empty] | 0 | [286 empty] | 0
P | 2018-09-22T10:06:48+08:00 | 2018-09-22T13:25:37+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP051G | CPSM Final Test | PASS | 129 | 167.72 | -0.02 | -0.84 | 56.5 | 50.63 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9339 | 0.5476 | 1.4 | 3 |  |  | 0.1992 | 0.3524 | 0.0295 | 0.99978 | [21 empty] | 0.02327 | [386 empty] | 0
P | 2018-09-22T10:24:00+08:00 | 2018-09-22T13:28:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP048G | CPSM Final Test | PASS | 129 | 159.89 | -0.02 | -0.99 | 56.95 | 50.29 | 1.29 | 0.13 | -1 | 2 | 0 | -0.838 | 0.5022 | 3.1 | 12 |  |  | 0.1906 | 0.2684 | 0.0783 | 0.99998 | [21 empty] | 0.00748 | [387 empty] | 0
P | 2018-09-22T10:09:28+08:00 | 2018-09-22T13:44:56+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP050G | CPSM Final Test | PASS | 129 | 159.62 | 0.17 | -0.54 | 56.52 | 49.87 | 1.13 | 0.13 | -1 | 1 | 1 | -0.9179 | 0.548 | 8.2 | 0 |  |  | 0.1337 | 0.1767 | 0.0317 | 0.99999 | [21 empty] | 0 | [330 empty] | 0
P | 2018-09-22T05:57:39+08:00 | 2018-09-22T14:08:54+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP045G | CPSM Final Test | FAIL | [431 empty] | 6752
P | 2018-09-22T13:44:40+08:00 | 2018-09-22T15:21:52+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP047G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-09-22T10:03:17+08:00 | 2018-09-22T15:28:56+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP045G | CPSM Final Test | PASS | 129 | 158.99 | -0.02 | -1.07 | 56.41 | 54.69 | 0.75 | 0.13 | -1 | -3 | 0 | -0.9061 | 0.516 | 4.7 | 8 |  |  | 0.4694 | 0.3789 | 0.0097 | 0.99994 | [21 empty] | 0 | [388 empty] | 0
P | 2018-09-22T12:44:53+08:00 | 2018-09-22T16:34:22+08:00 | CPSM Final Test TE83064-16 | 25703003 | Autosampler | 186015085 | J18FTP056G | CPSM Final Test | FAIL
P | 2018-09-22T14:41:22+08:00 | 2018-09-22T17:57:15+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | J18FTP039G | CPSM Final Test | PASS | 129 | 162.41 | -0.02 | -0.08 | 56.88 | 49.61 | 1.59 | 0.13 | 0 | -3 | -1 | -0.9358 | 0.512 | -3.8 | 110 |  |  | 0.0337 | 0.2534 | 0.0004 | 0.99996 | [21 empty] | 0 | [375 empty] | 0
P | 2018-09-22T14:46:27+08:00 | 2018-09-22T17:57:43+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | J18FTP053G | CPSM Final Test | PASS | 129 | 158.81 | -0.02 | -1.07 | 56.71 | 54.61 | 1.29 | 0.13 | 0 | -2 | 1 | -0.876 | 0.5384 | -0.8 | 10 |  |  | 0.1971 | 0.1004 | 0.0829 | 0.99998 | [21 empty] | 0.01227 | [389 empty] | 0
P | 2018-09-22T14:33:16+08:00 | 2018-09-22T17:58:14+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP047G | CPSM Final Test | PASS | 129 | 160.97 | -0.02 | -0.92 | 56.43 | 48.01 | 1.06 | 0.13 | 1 | 2 | 0 | -0.9278 | 0.5162 | 3.2 | 14 |  |  | 0.15 | 0.2229 | -0.0051 | 1 | [21 empty] | 0 | [390 empty] | 0
P | 2018-09-22T14:36:31+08:00 | 2018-09-22T18:10:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP057G | CPSM Final Test | PASS | 129 | 166.37 | -0.02 | -1.15 | 56.77 | 54.44 | 1.06 | 0.13 | 1 | -2 | -1 | -0.9278 | 0.548 | 8.5 | -13 |  |  | 0.0636 | 0.309 | -0.0198 | 1 | [21 empty] | 0 | [391 empty] | 0
P | 2018-09-22T14:39:38+08:00 | 2018-09-22T18:18:54+08:00 | CPSM Final Test TE83064-16 | 25703003 | Autosampler | 186015085 | J18FTP056G | CPSM Final Test | PASS | 129 | 169.88 | -0.02 | -1.3 | 56.63 | 49.95 | 0.75 | 0.13 | 1 | -3 | 0 | -0.9539 | 0.504 | 11.3 | 29 |  |  | 0.2509 | 0.484 | 0.0327 | 0.99998 | [21 empty] | 0 | [392 empty] | 0
P | 2018-09-20T18:42:10+08:00 | 2018-09-24T10:18:57+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP052G | CPSM Final Test | FAIL | [227 empty] | 0
P | 2018-09-24T08:25:04+08:00 | 2018-09-24T10:50:34+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP058G | CPSM Final Test | FAIL | [13 empty] | 22
P | 2018-09-22T14:30:34+08:00 | 2018-09-24T11:24:33+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | J18FTP052G | CPSM Final Test | PASS | 129 | 161.15 | -0.02 | -0.46 | 56.76 | 51.82 | 2.05 | 0.13 | 1 | -5 | 0 | -0.9617 | 0.479 | 5.6 | 14 |  |  | 0.1734 | 0.576 | -0.0124 | 1 | [21 empty] | 0 | [184 empty] | 0
P | 2018-09-22T14:50:07+08:00 | 2018-09-24T11:55:46+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP055G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.61 | 56.62 | 48.51 | 1.06 | 0.13 | 0 | 0 | 0 | -0.888 | 0.5182 | -1.9 | -3 |  |  | 0.3041 | 0.3272 | -0.0289 | 1 | [21 empty] | 0 | [393 empty] | 0
P | 2018-09-24T09:58:44+08:00 | 2018-09-24T12:40:30+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | J18FTP054G | CPSM Final Test | PASS | 129 | 159.89 | -0.21 | -0.99 | 56.48 | 48.94 | 1.36 | 0.13 | 1 | -3 | 1 | -0.9038 | 0.4906 | 4.1 | 12 |  |  | 0.0695 | 0.2368 | 0.0124 | 0.99999 | [21 empty] | 0 | [394 empty] | 0
P | 2018-09-24T05:07:22+08:00 | 2018-09-24T14:08:16+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP058G | CPSM Final Test | FAIL | [431 empty] | 1337
P | 2018-09-24T04:54:14+08:00 | 2018-09-24T14:21:24+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP058G | CPSM Final Test | FAIL | [431 empty] | 0
P | 2018-09-24T09:37:49+08:00 | 2018-09-24T15:27:52+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | J18FTP058G | CPSM Final Test | PASS | 129 | 163.31 | 0.17 | 0.07 | 56.6 | 51.05 | 1.13 | 0.13 | 1 | -2 | 0 | -0.9138 | 0.5144 | 1.8 | 7 |  |  | 0.0614 | 0.1612 | 0.0002 | 0.99999 | [21 empty] | 0 | [388 empty] | 0
P | 2018-09-24T02:41:47+08:00 | 2018-09-25T08:34:16+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | J18FTP062G | CPSM Final Test | FAIL | [14 empty] | 1974
P | 2018-09-24T02:18:42+08:00 | 2018-09-25T09:16:00+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | J18FTP065G | CPSM Final Test | FAIL | [438 empty] | 0
P | 2018-09-24T01:57:30+08:00 | 2018-09-25T09:11:30+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | J18FTP059G | CPSM Final Test | FAIL | [378 empty] | 0
P | 2018-06-01T09:42:51+08:00 | 2018-06-01T10:35:39+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP545G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-06-01T10:05:45+08:00 | 2018-06-01T13:03:21+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP544G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -0.99 | 56.17 | 48.43 | 1.51 | 0.13 | 1 | -2 | 0 | -0.8501 | 0.5202 | 0.3 | 0 |  |  | 0.1406 | 0.2012 | -0.0394 | 0.99999 | [21 empty] | 0 | [275 empty] | 0
P | 2018-06-01T10:01:21+08:00 | 2018-06-01T13:04:57+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP548G | CPSM Final Test | PASS | 129 | 164.39 | -0.02 | -0.46 | 56.48 | 49.11 | 1.67 | 0.13 | 0 | -1 | 0 | -0.8839 | 0.5422 | 4.2 | 35 |  |  | 0.0773 | 0.0537 | -0.0267 | 1 | [21 empty] | 0.03894 | [396 empty] | 0
P | 2018-06-01T13:37:30+08:00 | 2018-06-01T14:49:43+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP549G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-06-01T04:36:15+08:00 | 2018-06-01T15:13:08+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP547G | CPSM Final Test | FAIL
P | 2018-06-01T12:39:12+08:00 | 2018-06-01T16:21:07+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP551G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-06-01T11:28:58+08:00 | 2018-06-01T16:58:15+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP549G | CPSM Final Test | FAIL | [14 empty] | 661
P | 2018-06-01T09:54:42+08:00 | 2018-06-01T17:11:46+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP547G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.77 | 56.24 | 51.05 | 1.36 | 0.13 | 0 | -1 | 0 | -0.85 | 0.538 | 2.3 | 19 |  |  | 0.0613 | 0.3299 | -0.0278 | 1 | [21 empty] | 0.04737 | [397 empty] | 0
P | 2018-06-01T14:17:49+08:00 | 2018-06-01T17:49:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP552G | CPSM Final Test | PASS | 129 | 157.82 | -0.02 | -0.99 | 56.12 | 49.19 | 1.29 | 0.13 | 0 | 1 | 1 | -0.8859 | 0.4936 | 0.9 | -5 |  |  | 0.046 | 0.0749 | -0.0018 | 1 | [21 empty] | 0 | [398 empty] | 0
P | 2018-06-01T14:21:32+08:00 | 2018-06-01T17:49:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP550G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.92 | 56.03 | 51.39 | 0.98 | 0.13 | 0 | -1 | -1 | -0.85 | 0.4974 | -0.2 | -15 |  |  | 0.0756 | 0.1042 | -0.0411 | 0.99999 | [21 empty] | 0 | [399 empty] | 0
P | 2018-06-01T14:29:31+08:00 | 2018-06-01T17:56:13+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP553G | CPSM Final Test | PASS | 129 | 172.85 | -0.02 | -0.99 | 56.03 | 50.04 | 1.06 | 0.13 | -1 | -5 | 0 | -0.8779 | 0.4916 | 1.4 | -13 |  |  | 0.0959 | 0.3196 | -0.0265 | 1 | [21 empty] | 0 | [400 empty] | 0
P | 2018-06-01T11:25:42+08:00 | 2018-06-01T18:09:41+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP545G | CPSM Final Test | PASS | 129 | 167.99 | -0.02 | -0.92 | 56.03 | 50.21 | 1.29 | 0.13 | -1 | -4 | 0 | -0.8281 | 0.5206 | 3.6 | 8 |  |  | 0.4023 | 0.8666 | 0.0908 | 0.99996 | [21 empty] | 0 | [401 empty] | 0
P | 2018-06-01T14:30:10+08:00 | 2018-06-01T18:21:47+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP551G | CPSM Final Test | PASS | 129 | 162.77 | -0.02 | -0.92 | 56.32 | 49.78 | 1.44 | 0.13 | 0 | -2 | 0 | -0.9079 | 0.5312 | 1.2 | 79 |  |  | 0.2593 | 0.582 | -0.0203 | 0.99998 | [21 empty] | 0.00808 | [402 empty] | 0
P | 2018-06-01T14:13:37+08:00 | 2018-06-01T18:42:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP549G | CPSM Final Test | PASS | 129 | 157.91 | -0.21 | -1.3 | 56.31 | 48.6 | 1.21 | 0.13 | -2 | -2 | 0 | -0.8899 | 0.521 | -0.8 | 1 |  |  | 0.0818 | 0.1537 | -0.023 | 1 | [21 empty] | 0.02023 | [329 empty] | 0
P | 2018-06-01T18:57:10+08:00 | 2018-06-01T22:05:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP500G | CPSM Final Test | PASS | 129 | 168.17 | -0.21 | -0.84 | 56.14 | 50.8 | 1.13 | 0.13 | -2 | 1 | 0 | -0.898 | 0.48 | -0.3 | 7 |  |  | 0.0497 | 0.3813 | 0.0137 | 0.99999 | [21 empty] | 0 | [403 empty] | 0
P | 2018-06-01T18:50:31+08:00 | 2018-06-01T22:05:53+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP497G | CPSM Final Test | PASS | 129 | 163.4 | -0.11 | -0.46 | 56.35 | 49.7 | 1.06 | 0.13 | 1 | 2 | 0 | -0.9359 | 0.5164 | -0.9 | 4 |  |  | 0.0986 | 0.2381 | -0.0387 | 1 | [21 empty] | 0 | [331 empty] | 0
P | 2018-06-01T19:05:14+08:00 | 2018-06-01T22:11:17+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP558G | CPSM Final Test | PASS | 129 | 169.52 | -0.02 | -0.84 | 56.22 | 50.12 | 1.59 | 0.13 | -1 | -2 | 0 | -0.9179 | 0.4904 | 4.4 | 46 |  |  | 0.107 | 0.311 | 0.008 | 0.99999 | [21 empty] | 0.01743 | [404 empty] | 0
P | 2018-06-01T19:12:26+08:00 | 2018-06-01T22:11:47+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP557G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.46 | 56.16 | 50.29 | 1.21 | 0.13 | -1 | -1 | 1 | -0.9218 | 0.532 | 5 | -7 |  |  | 0.0989 | 0.1317 | -0.0229 | 1 | [21 empty] | 0 | [405 empty] | 0
P | 2018-06-01T19:05:10+08:00 | 2018-06-01T22:14:13+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP556G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -0.69 | 56.01 | 50.04 | 1.29 | 0.13 | -1 | -3 | 0 | -0.9278 | 0.487 | 6.8 | 3 |  |  | 0.0684 | 0.1269 | -0.0186 | 1 | [21 empty] | 0 | [205 empty] | 0
P | 2018-06-01T03:08:38+08:00 | 2018-06-02T11:52:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP555G | CPSM Final Test | FAIL
P | 2018-06-02T09:45:27+08:00 | 2018-06-02T12:57:59+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP496G | CPSM Final Test | PASS | 129 | 160.25 | -0.02 | -0.46 | 56.38 | 48.68 | 1.21 | 0.13 | 1 | -2 | 0 | -0.8341 | 0.5306 | 3.5 | 8 |  |  | 0.3452 | 0.4173 | 0.0451 | 0.99994 | [21 empty] | 0 | [406 empty] | 0
P | 2018-06-02T09:39:46+08:00 | 2018-06-02T12:57:37+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP512G | CPSM Final Test | PASS | 129 | 160.07 | -0.02 | -0.92 | 56.24 | 51.99 | 1.21 | 0.13 | 1 | -2 | 0 | -0.8281 | 0.529 | 7.4 | 4 |  |  | 0.0161 | 0.2588 | 0.0161 | 1 | [21 empty] | 0 | [143 empty] | 0
P | 2018-06-02T09:58:50+08:00 | 2018-06-02T12:59:03+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP561G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -0.46 | 56.7 | 54.27 | 1.36 | 0.13 | -1 | -1 | 1 | -0.8379 | 0.5264 | -3.5 | -4 |  |  | 0.0328 | 0.1272 | -0.0005 | 1 | [21 empty] | 0 | [407 empty] | 0
P | 2018-06-02T09:54:20+08:00 | 2018-06-02T13:10:33+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP560G | CPSM Final Test | PASS | 129 | 171.59 | -0.02 | -1.3 | 55.91 | 49.28 | 0.91 | 0.13 | 0 | 1 | 0 | -0.864 | 0.505 | 0.8 | 4 |  |  | 0.0679 | 0.1986 | -0.0064 | 1 | [21 empty] | 0 | [408 empty] | 0
P | 2018-06-02T13:50:54+08:00 | 2018-06-02T14:49:49+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP563G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-06-02T12:08:47+08:00 | 2018-06-02T16:09:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP554G | CPSM Final Test | FAIL
P | 2018-06-02T02:56:08+08:00 | 2018-06-02T16:53:01+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP559G | CPSM Final Test | FAIL | [19 empty] | -0.544
P | 2018-06-01T19:30:30+08:00 | 2018-06-02T17:00:10+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP555G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.92 | 56.12 | 50.46 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9159 | 0.4994 | 0.3 | 9 |  |  | 0.0285 | 0.1002 | -0.0128 | 1 | [21 empty] | 0 | [409 empty] | 0
P | 2018-06-02T14:22:43+08:00 | 2018-06-02T17:18:57+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP503G | CPSM Final Test | PASS | 129 | 161.42 | -0.02 | -0.77 | 56.5 | 48.51 | 0.98 | 0.13 | 1 | 2 | 0 | -0.864 | 0.518 | -0.2 | 20 |  |  | 0.1941 | 0.1886 | 0.0944 | 0.99993 | [21 empty] | 0 | [410 empty] | 0
P | 2018-06-02T14:26:43+08:00 | 2018-06-02T17:18:49+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP509G | CPSM Final Test | PASS | 129 | 165.11 | -1.1 | -0.38 | 56.81 | 50.55 | 1.36 | 0.13 | -1 | 1 | 0 | -0.7842 | 0.5152 | 4.5 | 9 |  |  | 0.0875 | 0.0849 | -0.0504 | 1 | [21 empty] | 0 | [258 empty] | 0
P | 2018-06-02T14:12:41+08:00 | 2018-06-02T17:21:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP562G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -0.84 | 56.33 | 48.94 | 0.68 | 0.13 | 1 | 1 | -1 | -0.8222 | 0.505 | -1.8 | 18 |  |  | 0.064 | 0.1455 | 0.0232 | 1 | [21 empty] | 0 | [368 empty] | 0
P | 2018-06-02T14:08:59+08:00 | 2018-06-02T17:53:51+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP554G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -0.84 | 55.91 | 49.45 | 1.51 | 0.13 | 0 | 1 | 0 | -0.842 | 0.502 | 0.9 | 117 |  |  | 0.0732 | 0.0394 | -0.0028 | 0.99999 | [21 empty] | 0 | [411 empty] | 0
P | 2018-06-02T09:54:35+08:00 | 2018-06-02T17:57:05+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP559G | CPSM Final Test | PASS | 129 | 161.51 | -0.02 | -0.77 | 56.31 | 49.95 | 1.51 | 0.13 | -1 | -2 | 0 | -0.8879 | 0.5186 | -4.3 | -3 |  |  | 0.2093 | 0.3475 | -0.0179 | 0.99999 | [21 empty] | 0 | [150 empty] | 0
P | 2018-06-01T19:55:32+08:00 | 2018-06-03T08:45:11+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP563G | CPSM Final Test | FAIL
P | 2018-06-02T14:20:22+08:00 | 2018-06-03T08:45:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP563G | CPSM Final Test | PASS | 129 | 161.06 | -0.3 | -0.77 | 56.33 | 50.72 | 1.29 | 0.13 | 1 | 1 | -1 | -0.8779 | 0.5042 | -3.3 | -2 |  |  | 0.1266 | 0.1728 | -0.0164 | 0.99999 | [21 empty] | 0 | [188 empty] | 0
P | 2018-06-03T09:24:53+08:00 | 2018-06-03T09:50:10+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP565G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-06-03T08:35:17+08:00 | 2018-06-03T10:22:07+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP508G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-06-03T08:36:34+08:00 | 2018-06-03T10:28:35+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP495G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-06-03T09:00:38+08:00 | 2018-06-03T10:35:31+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP566G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-06-03T08:51:18+08:00 | 2018-06-03T10:44:51+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP566G | CPSM Final Test | FAIL | [9 empty] | 39
P | 2018-06-03T09:40:43+08:00 | 2018-06-03T12:51:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP567G | CPSM Final Test | PASS | 129 | 164.21 | -0.21 | -0.84 | 56.3 | 48.6 | 1.21 | 0.13 | -1 | 1 | 0 | -0.876 | 0.5496 | -3 | 0 |  |  | 0.0527 | 0.1823 | -0.0178 | 0.99999 | [21 empty] | 0 | [412 empty] | 0
P | 2018-06-03T09:44:50+08:00 | 2018-06-03T12:53:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP564G | CPSM Final Test | PASS | 129 | 170.96 | -0.02 | -0.99 | 56.01 | 50.63 | 1.06 | 0.13 | 1 | 1 | -1 | -0.844 | 0.5 | -1.7 | 1 |  |  | 0.0668 | 0.18 | -0.0096 | 1 | [21 empty] | 0 | [368 empty] | 0
P | 2018-06-03T09:57:16+08:00 | 2018-06-03T13:01:01+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP568G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -0.84 | 56.1 | 49.78 | 1.51 | 0.13 | -1 | 4 | 0 | -0.882 | 0.4872 | -1.7 | 8 |  |  | 0.091 | 0.3666 | 0.0258 | 0.99999 | [21 empty] | 0.00749 | [188 empty] | 0
P | 2018-06-03T09:28:42+08:00 | 2018-06-03T12:59:31+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP508G | CPSM Final Test | PASS | 129 | 162.5 | -0.02 | -0.54 | 56.5 | 48.26 | 0.98 | 0.13 | -1 | 4 | 0 | -0.8361 | 0.533 | -1.3 | 36 |  |  | 0.0271 | 0.2641 | -0.0104 | 1 | [21 empty] | 0.00563 | [134 empty] | 0
P | 2018-06-03T09:37:31+08:00 | 2018-06-03T13:01:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP565G | CPSM Final Test | PASS | 129 | 167.63 | -0.59 | -0.77 | 56.46 | 49.36 | 1.51 | 0.13 | -1 | -1 | 0 | -0.8061 | 0.548 | -1.5 | 3 |  |  | 0.0777 | 0.2718 | -0.0298 | 1 | [21 empty] | 0 | [409 empty] | 0
P | 2018-06-03T09:32:35+08:00 | 2018-06-03T13:08:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP495G | CPSM Final Test | PASS | 129 | 157.37 | -0.21 | -0.92 | 56.19 | 49.78 | 1.13 | 0.13 | 1 | -2 | 0 | -0.7782 | 0.5252 | 3.4 | -6 |  |  | 0.0561 | 0.2803 | 0.0299 | 0.99997 | [21 empty] | 0 | [141 empty] | 0
P | 2018-06-03T09:48:05+08:00 | 2018-06-03T13:09:01+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP566G | CPSM Final Test | PASS | 129 | 159.62 | -0.11 | -0.46 | 56.54 | 49.28 | 1.21 | 0.13 | 0 | -1 | 0 | -0.88 | 0.543 | 2.8 | 8 |  |  | 0.1059 | 0.4362 | 0.0054 | 0.99999 | [21 empty] | 0 | [413 empty] | 0
P | 2018-06-03T13:00:54+08:00 | 2018-06-03T15:48:55+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP571G | CPSM Final Test | FAIL | [13 empty] | -5.4
P | 2018-06-01T09:58:27+08:00 | 2018-06-03T15:54:00+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP546G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.84 | 56.24 | 48.85 | 0.91 | 0.13 | 0 | 1 | 0 | -0.862 | 0.4982 | 1.4 | -1 |  |  | 0.1723 | 0.3509 | -0.0082 | 0.99997 | [21 empty] | 0 | [414 empty] | 0
P | 2018-06-03T11:20:02+08:00 | 2018-06-03T17:09:45+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP575G | CPSM Final Test | FAIL | [458 empty] | FALSE
P | 2018-06-03T14:18:15+08:00 | 2018-06-03T17:21:07+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP569G | CPSM Final Test | PASS | 129 | 169.88 | -0.02 | -1.07 | 56.21 | 51.56 | 1.13 | 0.13 | 0 | -2 | 0 | -0.9279 | 0.4826 | -0.3 | 27 |  |  | 0.1885 | 0.3512 | -0.0057 | 0.99997 | [21 empty] | 0 | [416 empty] | 0
P | 2018-06-03T14:11:37+08:00 | 2018-06-03T17:20:49+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP573G | CPSM Final Test | PASS | 129 | 164.3 | 0.17 | -0.99 | 56.41 | 54.95 | 1.36 | 0.13 | 1 | -4 | 0 | -0.9098 | 0.5272 | 3.1 | -2 |  |  | 0.0939 | 0.0704 | -0.0411 | 1 | [21 empty] | 0 | [417 empty] | 0
P | 2018-06-03T10:49:35+08:00 | 2018-06-03T17:23:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP574G | CPSM Final Test | FAIL
P | 2018-06-03T14:06:28+08:00 | 2018-06-03T17:23:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP574G | CPSM Final Test | PASS | 129 | 167.81 | 0.46 | -0.92 | 56.31 | 49.19 | 0.83 | 0.13 | -1 | 1 | 0 | -0.8819 | 0.4948 | 4.3 | 5 |  |  | 0.1819 | 0.4902 | 0.0073 | 1 | [21 empty] | 0 | [418 empty] | 0
P | 2018-06-03T14:29:19+08:00 | 2018-06-03T17:30:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP572G | CPSM Final Test | PASS | 129 | 161.69 | 0.27 | -1.22 | 56.12 | 48.85 | 0.75 | 0.13 | 1 | -1 | 1 | -0.864 | 0.489 | 10.3 | -4 |  |  | 0.1229 | 0.1746 | 0.0044 | 0.99997 | [21 empty] | 0 | [419 empty] | 0
P | 2018-06-03T11:10:30+08:00 | 2018-06-03T17:40:53+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP570G | CPSM Final Test | FAIL
P | 2018-06-03T14:24:55+08:00 | 2018-06-03T17:51:09+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP571G | CPSM Final Test | PASS | 129 | 161.42 | -0.02 | -0.46 | 56.2 | 49.61 | 1.21 | 0.13 | 1 | 1 | 0 | -0.9758 | 0.4946 | 3.2 | 23 |  |  | 0.0311 | 0.2836 | -0.0286 | 1 | [21 empty] | 0 | [182 empty] | 0
P | 2018-06-03T14:14:53+08:00 | 2018-06-03T19:34:27+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP575G | CPSM Final Test | PASS | 129 | 162.95 | 0.27 | -0.69 | 56.49 | 48.51 | 1.44 | 0.13 | 0 | 1 | 0 | -0.9237 | 0.516 | 7.3 | 33 |  |  | 0.0549 | 0.1446 | 0.0102 | 0.99998 | [21 empty] | 0 | [420 empty] | 0
P | 2018-06-02T20:12:04+08:00 | 2018-06-04T08:39:19+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP570G | CPSM Final Test | FAIL
P | 2018-06-02T19:04:40+08:00 | 2018-06-04T09:46:43+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP570G | CPSM Final Test | FAIL
P | 2018-06-02T18:25:40+08:00 | 2018-06-04T10:25:43+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP570G | CPSM Final Test | FAIL | [42 empty] | 0
P | 2018-06-03T14:25:42+08:00 | 2018-06-04T11:28:37+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP570G | CPSM Final Test | PASS | 129 | 169.7 | -0.02 | -1.3 | 56.11 | 51.73 | 1.06 | 0.13 | 1 | 3 | 0 | -0.9758 | 0.529 | 8 | 3 |  |  | 0.3373 | 0.1332 | 0.0166 | 0.99998 | [21 empty] | 0 | [75 empty] | 0
P | 2018-06-04T13:32:09+08:00 | 2018-06-04T17:22:22+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP470G | CPSM Final Test | FAIL
P | 2018-06-04T13:50:46+08:00 | 2018-06-04T17:25:56+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP471G | CPSM Final Test | FAIL
P | 2018-06-04T14:11:02+08:00 | 2018-06-04T17:32:53+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | F18FNP472G | CPSM Final Test | FAIL
P | 2018-06-04T15:35:23+08:00 | 2018-06-04T18:38:32+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP474G | CPSM Final Test | FAIL | [13 empty] | -6.3
P | 2018-06-04T16:07:49+08:00 | 2018-06-04T19:04:43+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP473G | CPSM Final Test | PASS |  | 168.17 | -0.02 | -0.99 | 56.13 | 49.78 | 1.13 | 0.13 | 0 | -2 | 0 | -0.9398 | 0.532 | -0.3 | 7 |  |  | 0.1042 | 0.113 | -0.0153 | 1 |  |  |  |  | 133 |  | 0 | [436 empty] | 0
P | 2018-06-04T15:27:15+08:00 | 2018-06-04T19:10:48+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP470G | CPSM Final Test | PASS |  | 165.74 | -0.02 | -0.77 | 56.13 | 48.18 | 1.29 | 0.13 | 0 | -1 | 0 | -0.8999 | 0.5258 | 1.7 | 6 |  |  | 0.0449 | 0.0948 | -0.0222 | 0.99999 |  |  |  |  | 133 |  | 0 | [356 empty] | 0
P | 2018-06-04T15:38:21+08:00 | 2018-06-04T19:13:46+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP471G | CPSM Final Test | PASS |  | 164.3 | -0.02 | -0.99 | 56.17 | 48.43 | 0.98 | 0.13 | 1 | -2 | 0 | -0.8281 | 0.5332 | 8.7 | -10 |  |  | 0.0516 | 0.169 | -0.0118 | 0.99999 |  |  |  |  | 133 |  | 0 | [437 empty] | 0
P | 2018-06-04T15:51:57+08:00 | 2018-06-04T19:21:43+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | F18FNP472G | CPSM Final Test | PASS |  | 165.65 | 0.17 | -0.54 | 56.71 | 55.54 | 1.44 | 0.13 | 0 | 2 | 0 | -0.8957 | 0.5076 | 1.8 | 5 |  |  | 0.1757 | 0.1427 | 0.0556 | 0.99999 |  |  |  |  | 133 |  | 0 | [438 empty] | 0
P | 2018-06-04T18:29:40+08:00 | 2018-06-04T19:31:22+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP476G | CPSM Final Test | FAIL | [5 empty] | 46.65
P | 2018-06-04T19:14:45+08:00 | 2018-06-04T20:02:00+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP476G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-06-04T17:06:57+08:00 | 2018-06-04T20:42:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP474G | CPSM Final Test | PASS |  | 162.95 | 0.27 | -0.92 | 56.24 | 49.02 | 1.36 | 0.13 | 0 | -3 | -1 | -0.9198 | 0.4984 | 9.7 | 7 |  |  | 0.1074 | 0.2297 | -0.0421 | 0.99999 |  |  |  |  | 133 |  | 0 | [439 empty] | 0
P | 2018-06-04T18:29:27+08:00 | 2018-06-04T21:23:24+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP475G | CPSM Final Test | PASS |  | 165.2 | 0.27 | -0.69 | 56.34 | 50.29 | 1.74 | 0.13 | 0 | -2 | 0 | -0.9877 | 0.5198 | -2.8 | 3 |  |  | 0.1158 | 0.0416 | 0.0456 | 0.99997 |  |  |  |  | 133 |  | 0 | [440 empty] | 0
P | 2018-06-04T07:12:22+08:00 | 2018-06-05T09:07:40+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP476G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-06-04T06:19:26+08:00 | 2018-06-05T10:00:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP476G | CPSM Final Test | FAIL | [13 empty] | -5.2
P | 2018-06-04T19:49:33+08:00 | 2018-06-05T11:17:20+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | F18FNP477G | CPSM Final Test | PASS |  | 162.86 | 0.17 | -0.77 | 56.72 | 53.68 | 1.44 | 0.13 | 1 | -2 | 1 | -0.824 | 0.4652 | 3.2 | 5 |  |  | 0.1564 | 0.1783 | 0.0236 | 0.99998 |  |  |  |  | 133 |  | 0 | [441 empty] | 0
P | 2018-06-04T20:10:01+08:00 | 2018-06-05T12:04:02+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP476G | CPSM Final Test | PASS |  | 163.76 | 0.46 | -0.99 | 56.2 | 48.51 | 1.29 | 0.13 | 1 | 3 | 1 | -0.8081 | 0.5268 | 1.9 | 51 |  |  | 0.0964 | 0.1592 | -0.0287 | 1 |  |  |  |  | 133 |  | 0 | [442 empty] | 0
P | 2018-06-05T11:09:07+08:00 | 2018-06-05T13:18:26+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP479G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-06-05T10:45:54+08:00 | 2018-06-05T13:34:40+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP478G | CPSM Final Test | PASS |  | 155.84 | -0.4 | -1.15 | 56.42 | 54.44 | 1.06 | 0.13 | 0 | -2 | -1 | -0.8341 | 0.533 | -0.5 | 0 |  |  | 0.0851 | 0.1466 | -0.0099 | 0.99999 |  |  |  |  | 133 |  | 0 | [443 empty] | 0
P | 2018-06-05T10:51:39+08:00 | 2018-06-05T13:35:54+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP479G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-06-05T10:38:05+08:00 | 2018-06-05T13:49:28+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP479G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-06-05T13:55:28+08:00 | 2018-06-05T14:33:28+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP481G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-06-05T14:17:56+08:00 | 2018-06-05T14:53:58+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP482G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-06-05T09:29:20+08:00 | 2018-06-05T15:19:57+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP480G | CPSM Final Test | FAIL
P | 2018-06-05T12:24:38+08:00 | 2018-06-05T16:25:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP480G | CPSM Final Test | PASS |  | 162.23 | -0.02 | -0.77 | 56.41 | 54.78 | 0.83 | 0.13 | -1 | 2 | -1 | -0.8979 | 0.547 | -1.3 | 21 |  |  | 0.0501 | 0.0844 | -0.0099 | 0.99999 |  |  |  |  | 133 |  | 0 | [444 empty] | 0
P | 2018-06-05T07:14:27+08:00 | 2018-06-05T17:13:06+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP479G | CPSM Final Test | FAIL | [27 empty] | 0.18148
P | 2018-06-05T14:14:28+08:00 | 2018-06-05T17:17:30+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP481G | CPSM Final Test | PASS |  | 161.78 | 0.37 | -1.07 | 56.16 | 49.02 | 1.21 | 0.13 | -1 | -1 | 0 | -0.894 | 0.462 | 6 | 209 |  |  | 0.1199 | 0.1102 | 0.0022 | 0.99999 |  |  |  |  | 133 |  | 0 | [445 empty] | 0
P | 2018-06-05T14:35:57+08:00 | 2018-06-05T17:38:20+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP482G | CPSM Final Test | PASS |  | 161.6 | -0.4 | -0.84 | 56.4 | 48.68 | 1.06 | 0.13 | -1 | 2 | 0 | -0.8321 | 0.5328 | 4.2 | 45 |  |  | 0.2167 | 0.2324 | -0.0055 | 0.99998 |  |  |  |  | 133 |  | 0 | [446 empty] | 0
P | 2018-06-05T15:23:20+08:00 | 2018-06-05T18:08:07+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | F18FNP483G | CPSM Final Test | PASS |  | 170.42 | -0.02 | -0.99 | 56.33 | 51.22 | 0.83 | 0.13 | 1 | -2 | 0 | -0.9219 | 0.547 | -2 | 0 |  |  | 0.0372 | 0.2563 | 0.0283 | 0.99998 |  |  |  |  | 133 |  | 0 | [447 empty] | 0
P | 2018-06-05T06:13:11+08:00 | 2018-06-05T18:14:22+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP479G | CPSM Final Test | FAIL | [27 empty] | 0.04882
P | 2018-06-05T12:13:47+08:00 | 2018-06-05T19:17:16+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP479G | CPSM Final Test | PASS |  | 164.39 | -0.02 | -0.61 | 56.33 | 48.68 | 1.06 | 0.13 | 2 | 2 | 0 | -0.9038 | 0.5186 | 3.1 | 17 |  |  | 0.0898 | 0.2352 | -0.0577 | 1 |  |  |  |  | 133 |  | 0 | [448 empty] | 0
P | 2018-06-05T16:59:31+08:00 | 2018-06-05T19:44:10+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP485G | CPSM Final Test | PASS |  | 162.68 | -0.02 | -0.92 | 56.45 | 48.01 | 1.29 | 0.13 | 0 | -1 | 1 | -0.9019 | 0.5438 | -0.3 | 63 |  |  | 0.2134 | 0.4293 | -0.0247 | 0.99999 |  |  |  |  | 133 |  | 0 | [449 empty] | 0
P | 2018-06-05T16:49:05+08:00 | 2018-06-05T19:51:44+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | F18FNP484G | CPSM Final Test | PASS |  | 161.06 | -0.49 | -1.15 | 56.3 | 49.7 | 0.91 | 0.13 | -1 | -3 | -1 | -0.8181 | 0.5442 | 6.9 | -9 |  |  | 0.0591 | 0.3072 | 0.0191 | 0.99997 |  |  |  |  | 133 |  | 0 | [450 empty] | 0
P | 2018-06-05T18:23:48+08:00 | 2018-06-05T21:12:46+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP486G | CPSM Final Test | PASS |  | 164.03 | 0.08 | -0.92 | 56.13 | 51.9 | 1.36 | 0.13 | 1 | 1 | 0 | -0.9458 | 0.5086 | 3.6 | 3 |  |  | 0.0953 | 0.2868 | -0.0159 | 0.99999 |  |  |  |  | 133 |  | 0 | [451 empty] | 0
P | 2018-06-06T09:41:42+08:00 | 2018-06-06T10:26:16+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP577G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-06-06T09:08:22+08:00 | 2018-06-06T10:26:47+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP576G | CPSM Final Test | FAIL | [5 empty] | 45.21
P | 2018-06-06T07:45:48+08:00 | 2018-06-06T11:49:21+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP576G | CPSM Final Test | FAIL | [5 empty] | 46.82
P | 2018-06-06T10:03:59+08:00 | 2018-06-06T13:04:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP577G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | -0.69 | 56.42 | 48.09 | 1.29 | 0.13 | -2 | 2 | 0 | -0.8101 | 0.493 | 1.3 | -83 |  |  | 0.1158 | 0.2048 | -0.0082 | 0.99999 | [21 empty] | 0 | [437 empty] | 0
P | 2018-06-06T14:10:34+08:00 | 2018-06-06T15:13:22+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP580G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-06-06T04:08:48+08:00 | 2018-06-06T15:26:21+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP576G | CPSM Final Test | FAIL | [14 empty] | 1112
P | 2018-06-06T06:40:14+08:00 | 2018-06-06T16:35:38+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP578G | CPSM Final Test | FAIL
P | 2018-06-06T11:37:56+08:00 | 2018-06-06T16:42:20+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP578G | CPSM Final Test | PASS | 129 | 161.24 | 0.17 | -0.99 | 56.14 | 49.61 | 0.98 | 0.13 | 0 | 1 | 1 | -0.8979 | 0.5218 | -3 | -10 |  |  | 0.0548 | 0.0697 | -0.0645 | 0.99998 | [21 empty] | 0 | [262 empty] | 0
P | 2018-06-06T09:47:34+08:00 | 2018-06-06T17:10:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP576G | CPSM Final Test | PASS | 129 | 163.58 | -0.11 | -0.61 | 56.31 | 49.19 | 1.13 | 0.13 | -1 | -2 | -1 | -0.852 | 0.514 | 0.7 | 0 |  |  | 0.0783 | 0.1804 | 0.0384 | 0.99998 | [21 empty] | 0 | [438 empty] | 0
P | 2018-06-06T11:56:13+08:00 | 2018-06-06T17:11:19+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP579G | CPSM Final Test | FAIL | [472 empty] | 3332
P | 2018-06-06T14:41:58+08:00 | 2018-06-06T17:50:34+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP580G | CPSM Final Test | PASS | 129 | 167.45 | -0.02 | -1.07 | 55.99 | 49.19 | 1.29 | 0.13 | 0 | 2 | 0 | -0.9119 | 0.4892 | 6.9 | 190 |  |  | 0.0685 | 0.1839 | -0.008 | 0.99999 | [21 empty] | 0 | [439 empty] | 0
P | 2018-06-06T15:02:42+08:00 | 2018-06-06T18:00:18+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP581G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.84 | 56.03 | 51.22 | 1.06 | 0.13 | 0 | 1 | 1 | -0.8341 | 0.534 | -0.5 | 18 |  |  | 0.1081 | 0.1893 | -0.0077 | 0.99999 | [21 empty] | 0 | [440 empty] | 0
P | 2018-06-06T14:33:46+08:00 | 2018-06-06T18:17:03+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP579G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.54 | 56.34 | 50.8 | 1.13 | 0.13 | 0 | -2 | 0 | -0.9199 | 0.5254 | 2.8 | -13 |  |  | 0.1057 | 0.1581 | -0.0081 | 1 | [21 empty] | 0 | [429 empty] | 0
P | 2018-06-06T15:41:25+08:00 | 2018-06-06T18:26:22+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP583G | CPSM Final Test | PASS | 129 | 153.41 | -0.02 | -1.15 | 56.01 | 50.63 | 1.21 | 0.13 | 0 | 1 | 0 | -0.87 | 0.516 | -1.2 | -8 |  |  | 0.0302 | 0.3003 | -0.0423 | 0.99999 | [21 empty] | 0 | [441 empty] | 0
P | 2018-06-06T15:56:26+08:00 | 2018-06-06T18:28:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP584G | CPSM Final Test | FAIL
P | 2018-06-06T15:59:37+08:00 | 2018-06-06T18:38:52+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP582G | CPSM Final Test | PASS | 129 | 162.5 | -0.02 | -0.69 | 56.13 | 48.94 | 0.98 | 0.13 | -1 | 2 | 0 | -0.828 | 0.544 | 1.3 | 4 |  |  | 0.0283 | 0.1058 | 0.0297 | 0.99994 | [21 empty] | 0.01734 | [442 empty] | 0
P | 2018-06-06T18:04:39+08:00 | 2018-06-06T19:19:29+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP585G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-06-06T17:12:17+08:00 | 2018-06-06T21:03:37+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP584G | CPSM Final Test | PASS | 129 | 168.8 | 0.08 | -0.84 | 56.2 | 49.61 | 1.36 | 0.13 | 1 | -2 | -1 | -0.9498 | 0.545 | -2 | 45 |  |  | 0.1087 | 0.3157 | -0.0257 | 0.99999 | [21 empty] | 0 | [443 empty] | 0
P | 2018-06-06T18:28:40+08:00 | 2018-06-06T21:11:56+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP586G | CPSM Final Test | PASS | 129 | 173.12 | 0.27 | -0.92 | 56.49 | 49.87 | 0.75 | 0.13 | 1 | -2 | 0 | -0.9199 | 0.5302 | -3.7 | -19 |  |  | 0.0428 | 0.0735 | -0.0315 | 1 | [21 empty] | 0 | [444 empty] | 0
P | 2018-06-06T18:42:04+08:00 | 2018-06-06T21:49:15+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP585G | CPSM Final Test | PASS | 129 | 165.56 | -0.02 | -0.38 | 56.46 | 49.87 | 1.74 | 0.13 | 0 | 2 | -1 | -0.9378 | 0.516 | 5 | 42 |  |  | 0.0804 | 0.0987 | -0.0373 | 1 | [21 empty] | 0 | [445 empty] | 0
P | 2018-06-07T09:28:05+08:00 | 2018-06-07T10:35:52+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP589G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-06-07T10:14:00+08:00 | 2018-06-07T11:07:32+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP590G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-06-07T09:49:59+08:00 | 2018-06-07T12:48:50+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP587G | CPSM Final Test | PASS | 129 | 162.23 | 0.08 | -1.22 | 56.2 | 50.04 | 0.98 | 0.13 | 0 | -2 | -1 | -0.7901 | 0.5398 | -3.4 | -6 |  |  | 0.1097 | 0.1434 | 0.0106 | 0.99996 | [21 empty] | 0 | [446 empty] | 0
P | 2018-06-07T10:22:42+08:00 | 2018-06-07T13:11:08+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP588G | CPSM Final Test | PASS | 129 | 160.88 | -0.02 | -0.84 | 56.42 | 53.09 | 0.91 | 0.13 | -1 | -5 | 0 | -0.8441 | 0.539 | -0.9 | 5 |  |  | 0.1335 | 0.1671 | 0.0032 | 0.99997 | [21 empty] | 0 | [447 empty] | 0
P | 2018-06-07T10:40:46+08:00 | 2018-06-07T13:38:54+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP590G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.84 | 56.09 | 48.68 | 1.29 | 0.13 | -2 | 3 | 1 | -0.9359 | 0.52 | 5.1 | 3 |  |  | 0.1002 | 0.3576 | -0.0197 | 0.99999 | [21 empty] | 0 | [448 empty] | 0
P | 2018-06-07T06:05:15+08:00 | 2018-06-07T13:58:42+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP589G | CPSM Final Test | FAIL | [309 empty] | 1575
P | 2018-06-07T13:11:54+08:00 | 2018-06-07T14:08:46+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP592G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-06-07T10:01:59+08:00 | 2018-06-07T15:04:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP589G | CPSM Final Test | PASS | 129 | 159.98 | -0.02 | -0.99 | 55.88 | 49.53 | 0.98 | 0.15 | 0 | -2 | 0 | -0.9019 | 0.5158 | -3.3 | -31 |  |  | 0.0874 | 0.2593 | -0.0245 | 0.99998 | [21 empty] | 0 | [266 empty] | 0
P | 2018-06-07T13:22:27+08:00 | 2018-06-07T16:15:33+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP591G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -0.84 | 56.46 | 49.61 | 0.91 | 0.13 | -1 | 1 | 1 | -0.8281 | 0.5164 | 3.7 | 0 |  |  | 0.0679 | 0.2346 | 0.0118 | 0.99999 | [21 empty] | 0 | [429 empty] | 0
P | 2018-06-07T13:40:20+08:00 | 2018-06-07T16:41:04+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP592G | CPSM Final Test | PASS | 129 | 165.83 | -0.02 | -0.77 | 56.2 | 48.51 | 1.36 | 0.13 | 1 | 1 | 0 | -0.9219 | 0.5182 | 5.5 | -12 |  |  | 0.103 | 0.1038 | -0.0145 | 1 | [21 empty] | 0 | [449 empty] | 0
P | 2018-06-07T14:21:45+08:00 | 2018-06-07T17:11:55+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP593G | CPSM Final Test | PASS | 129 | 163.31 | -0.21 | -0.46 | 56.18 | 50.12 | 1.36 | 0.13 | -1 | -3 | -1 | -0.826 | 0.546 | 3.9 | 6 |  |  | 0.0821 | 0.3468 | -0.0056 | 0.99992 | [21 empty] | 0 | [450 empty] | 0
P | 2018-06-07T14:38:57+08:00 | 2018-06-07T17:25:25+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP508G | CPSM Final Test | PASS | 129 | 162.5 | -0.02 | -0.61 | 56.5 | 49.78 | 0.98 | 0.13 | -1 | -2 | 0 | -0.89 | 0.5192 | -0.3 | 46 |  |  | 0.1595 | 0.1091 | 0.0087 | 0.99997 | [21 empty] | 0 | [451 empty] | 0
P | 2018-06-07T15:23:16+08:00 | 2018-06-07T18:38:56+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP595G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-06-07T13:59:38+08:00 | 2018-06-07T20:02:34+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP595G | CPSM Final Test | FAIL | [495 empty] | 5007
P | 2018-06-07T17:14:48+08:00 | 2018-06-07T20:33:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP594G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.54 | 57.07 | 51.73 | 1.21 | 0.13 | -1 | 1 | 0 | -0.8599 | 0.5388 | 0.2 | -5 |  |  | 0.0218 | 0.3214 | -0.0339 | 1 | [21 empty] | 0 | [266 empty] | 0
P | 2018-06-07T17:35:02+08:00 | 2018-06-07T20:58:34+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP596G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | -0.61 | 56.2 | 50.38 | 1.36 | 0.13 | -1 | -2 | 0 | -0.9598 | 0.4966 | -3.4 | 38 |  |  | 0.1501 | 0.2076 | -0.0187 | 0.99999 | [21 empty] | 0 | [453 empty] | 0
P | 2018-06-07T17:01:06+08:00 | 2018-06-07T21:07:58+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP595G | CPSM Final Test | PASS | 129 | 167.72 | -0.11 | -0.99 | 56.23 | 49.53 | 1.06 | 0.13 | -1 | 1 | 1 | -0.854 | 0.5146 | 11.2 | -18 |  |  | 0.0922 | 0.3896 | 0.0731 | 0.99996 | [21 empty] | 0.03524 | [452 empty] | 0
P | 2018-06-07T18:25:49+08:00 | 2018-06-07T21:24:52+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP597G | CPSM Final Test | PASS | 129 | 164.39 | -0.02 | -1.07 | 56.12 | 50.38 | 0.68 | 0.13 | 1 | 1 | 0 | -0.9617 | 0.468 | 2.8 | 11 |  |  | 0.0893 | 0.2039 | -0.0358 | 0.99999 | [21 empty] | 0 | [405 empty] | 0
P | 2018-06-07T18:45:57+08:00 | 2018-06-07T21:36:23+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP598G | CPSM Final Test | PASS | 129 | 161.42 | 0.17 | -0.84 | 56.54 | 51.14 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9477 | 0.5138 | 6.3 | -3 |  |  | 0.0752 | 0.163 | -0.0144 | 1 | [21 empty] | 0 | [454 empty] | 0
P | 2018-06-08T09:58:30+08:00 | 2018-06-08T10:55:23+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP602G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-06-08T10:31:30+08:00 | 2018-06-08T11:16:40+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP600G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-06-08T10:52:26+08:00 | 2018-06-08T11:27:13+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP601G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-06-08T11:29:08+08:00 | 2018-06-08T13:30:55+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP604G | CPSM Final Test | FAIL | [5 empty] | 46.82
P | 2018-06-08T10:26:56+08:00 | 2018-06-08T13:31:29+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP602G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -0.84 | 56.5 | 49.7 | 1.13 | 0.13 | -1 | 1 | 0 | -0.852 | 0.5376 | -1.6 | 33 |  |  | 0.0866 | 0.1188 | 0.0261 | 0.99999 | [21 empty] | 0 | [411 empty] | 0
P | 2018-06-08T11:28:47+08:00 | 2018-06-08T13:35:32+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP603G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-06-08T08:09:58+08:00 | 2018-06-08T13:38:12+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP600G | CPSM Final Test | FAIL | [251 empty] | 0
P | 2018-06-08T07:21:39+08:00 | 2018-06-08T13:51:33+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP599G | CPSM Final Test | FAIL | [42 empty] | 0.09593
P | 2018-06-08T11:09:50+08:00 | 2018-06-08T13:56:03+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP601G | CPSM Final Test | PASS | 129 | 167.27 | 0.27 | -0.54 | 56.39 | 48.6 | 0.75 | 0.13 | 0 | 2 | 0 | -0.7941 | 0.4856 | 4.3 | 29 |  |  | 0.0746 | 0.2298 | -0.0136 | 0.99999 | [21 empty] | 0 | [455 empty] | 0
P | 2018-06-08T12:52:22+08:00 | 2018-06-08T14:06:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP605G | CPSM Final Test | FAIL | [5 empty] | 45.89
P | 2018-06-08T10:54:05+08:00 | 2018-06-08T14:43:56+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP600G | CPSM Final Test | PASS | 129 | 164.21 | 0.37 | -0.54 | 55.9 | 48.34 | 1.13 | 0.13 | 1 | 2 | 0 | -0.854 | 0.51 | 2.2 | -27 |  |  | 0.1128 | 0.1663 | -0.0485 | 1 | [21 empty] | 0 | [208 empty] | 0
P | 2018-06-08T10:36:36+08:00 | 2018-06-08T14:54:33+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP599G | CPSM Final Test | PASS | 129 | 156.47 | 0.37 | -0.61 | 56.02 | 50.21 | 0.91 | 0.13 | -1 | 1 | 0 | -0.8839 | 0.4966 | 6.2 | 10 |  |  | 0.1259 | 0.1891 | -0.0456 | 0.99999 | [21 empty] | 0 | [456 empty] | 0
P | 2018-06-08T12:32:10+08:00 | 2018-06-08T16:11:48+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP603G | CPSM Final Test | PASS | 129 | 159.44 | 0.08 | -0.99 | 56.27 | 50.12 | 1.21 | 0.13 | -1 | 1 | 0 | -0.8779 | 0.5394 | 1.3 | 2 |  |  | 0.0507 | 0.1876 | 0.0509 | 0.99997 | [21 empty] | 0 | [457 empty] | 0
P | 2018-06-08T08:44:38+08:00 | 2018-06-08T16:15:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP604G | CPSM Final Test | FAIL
P | 2018-06-08T12:30:02+08:00 | 2018-06-08T16:21:59+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP604G | CPSM Final Test | PASS | 129 | 162.86 | 0.56 | -0.08 | 56.28 | 48.43 | 1.13 | 0.13 | 1 | 3 | 0 | -0.864 | 0.4734 | -1.3 | 19 |  |  | 0.1429 | 0.1932 | 0.0055 | 0.99999 | [21 empty] | 0 | [147 empty] | 0
P | 2018-06-08T13:29:29+08:00 | 2018-06-08T16:36:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP605G | CPSM Final Test | PASS | 129 | 165.2 | 0.37 | -0.92 | 56.16 | 48.26 | 0.98 | 0.13 | 1 | 2 | -1 | -0.844 | 0.523 | -3.6 | -9 |  |  | 0.0648 | 0.2548 | -0.0211 | 0.99999 | [21 empty] | 0.03407 | [368 empty] | 0
P | 2018-06-08T13:12:29+08:00 | 2018-06-08T17:38:12+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP607G | CPSM Final Test | FAIL | [495 empty] | 19230
P | 2018-06-08T15:25:20+08:00 | 2018-06-08T18:57:50+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP607G | CPSM Final Test | PASS | 129 | 160.43 | -0.11 | -0.54 | 55.89 | 48.34 | 1.13 | 0.13 | -1 | 2 | 0 | -0.7143 | 0.499 | 4.6 | 5 |  |  | 0.0772 | 0.1033 | 0.0248 | 0.99999 | [21 empty] | 0 | [452 empty] | 0
P | 2018-06-08T15:12:27+08:00 | 2018-06-08T19:07:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP606G | CPSM Final Test | FAIL | [5 empty] | 46.57
P | 2018-06-08T18:37:49+08:00 | 2018-06-08T19:43:22+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP615G | CPSM Final Test | FAIL | [5 empty] | 43.52
P | 2018-06-08T18:41:57+08:00 | 2018-06-08T21:29:12+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP608G | CPSM Final Test | PASS | 129 | 161.87 | 0.46 | -0.38 | 56.21 | 49.28 | 1.51 | 0.13 | 1 | -2 | 0 | -0.9317 | 0.5378 | 1.5 | 41 |  |  | 0.1172 | 0.1818 | 0.0166 | 0.99999 | [21 empty] | 0 | [458 empty] | 0
P | 2018-06-08T18:25:09+08:00 | 2018-06-08T21:28:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP610G | CPSM Final Test | PASS | 129 | 176.09 | 0.27 | -0.84 | 56.41 | 52.24 | 1.29 | 0.13 | 0 | 1 | 0 | -0.9456 | 0.543 | -0.2 | 12 |  |  | 0.0648 | 0.2103 | 0.0062 | 1 | [21 empty] | 0.04059 | [457 empty] | 0
P | 2018-06-08T18:18:41+08:00 | 2018-06-08T21:31:11+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP611G | CPSM Final Test | PASS | 129 | 164.57 | 0.46 | -1.38 | 56.13 | 49.28 | 1.06 | 0.13 | 1 | -2 | 0 | -0.9197 | 0.52 | 1.1 | -66 |  |  | 0.3064 | 0.2032 | -0.0632 | 1 | [21 empty] | 0 | [459 empty] | 0
P | 2018-06-08T17:31:31+08:00 | 2018-06-08T21:33:10+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP613G | CPSM Final Test | PASS | 129 | 160.97 | 0.46 | -0.23 | 56.23 | 50.55 | 0.83 | 0.13 | 0 | -2 | -1 | -0.9339 | 0.5418 | 6.9 | -3 |  |  | 0.1736 | 0.2315 | -0.0307 | 0.99999 | [21 empty] | 0 | [208 empty] | 0
P | 2018-06-08T17:19:59+08:00 | 2018-06-08T21:35:37+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP609G | CPSM Final Test | PASS | 129 | 156.74 | -0.02 | -0.69 | 56.24 | 50.21 | 1.44 | 0.13 | 1 | -2 | -1 | -0.9398 | 0.4972 | 0.3 | 8 |  |  | 0.1113 | 0.4017 | -0.0387 | 0.99998 | [21 empty] | 0 | [460 empty] | 0
P | 2018-06-08T19:10:36+08:00 | 2018-06-08T22:10:44+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP615G | CPSM Final Test | PASS | 129 | 172.22 | 0.56 | -0.92 | 56.48 | 48.34 | 1.13 | 0.13 | 0 | -1 | 0 | -0.9078 | 0.5436 | 0.3 | -13 |  |  | 0.1126 | 0.2918 | -0.0142 | 1 | [21 empty] | 0 | [19 empty] | 0
P | 2018-06-08T01:31:09+08:00 | 2018-06-09T08:49:15+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP606G | CPSM Final Test | FAIL | [14 empty] | 1658
P | 2018-06-08T17:10:12+08:00 | 2018-06-09T10:34:05+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP606G | CPSM Final Test | PASS | 129 | 158.72 | 1.03 | -0.61 | 56.41 | 49.78 | 0.98 | 0.13 | 1 | -3 | 0 | -0.8719 | 0.522 | -0.9 | -9 |  |  | 0.1312 | 0.2189 | -0.0043 | 1 | [21 empty] | 0 | [461 empty] | 0
P | 2018-06-09T11:30:21+08:00 | 2018-06-09T14:53:15+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP612G | CPSM Final Test | PASS | 129 | 155.57 | 0.27 | -1.07 | 56.22 | 49.11 | 1.06 | 0.13 | 0 | -3 | 0 | -0.866 | 0.5068 | 5.3 | 8 |  |  | 0.0632 | 0.071 | -0.0168 | 0.99997 | [21 empty] | 0 | [462 empty] | 0
P | 2018-06-09T11:41:37+08:00 | 2018-06-09T15:05:47+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP614G | CPSM Final Test | PASS | 129 | 164.75 | 0.37 | -0.38 | 56.43 | 54.53 | 1.13 | 0.13 | 0 | -2 | 0 | -0.8442 | 0.5328 | -1.4 | -1 |  |  | 0.256 | 0.0888 | -0.0252 | 0.99999 | [21 empty] | 0 | [463 empty] | 0
P | 2018-06-09T14:23:34+08:00 | 2018-06-09T15:21:18+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP621G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-06-09T12:42:22+08:00 | 2018-06-09T15:35:08+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP618G | CPSM Final Test | PASS | 129 | 160.7 | -0.02 | -0.99 | 56.14 | 49.02 | 1.13 | 0.13 | 0 | 2 | 0 | -0.8959 | 0.5006 | 4.3 | -4 |  |  | 0.0199 | 0.0837 | -0.0303 | 1 | [21 empty] | 0 | [464 empty] | 0
P | 2018-06-09T12:56:50+08:00 | 2018-06-09T15:48:07+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP616G | CPSM Final Test | PASS | 129 | 162.32 | -0.11 | -1.3 | 56.23 | 48.26 | 0.83 | 0.13 | -1 | 1 | 0 | -0.9259 | 0.4674 | -3.4 | 19 |  |  | 0.1017 | 0.249 | -0.0227 | 0.99999 | [21 empty] | 0 | [399 empty] | 0
P | 2018-06-09T13:18:46+08:00 | 2018-06-09T16:13:47+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP619G | CPSM Final Test | PASS | 129 | 160.7 | 0.27 | -1.07 | 56.18 | 49.53 | 0.83 | 0.13 | 0 | -2 | 0 | -0.9019 | 0.4948 | -1 | 3 |  |  | 0.0814 | 0.3374 | -0.0198 | 1 | [21 empty] | 0.01839 | [465 empty] | 0
P | 2018-06-09T13:53:41+08:00 | 2018-06-09T16:46:28+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP617G | CPSM Final Test | PASS | 129 | 159.71 | 0.27 | -0.69 | 56.45 | 50.12 | 0.91 | 0.13 | 1 | -2 | 0 | -0.9219 | 0.511 | 1.2 | 12 |  |  | 0.1112 | 0.1277 | 0.0028 | 0.99999 | [21 empty] | 0 | [466 empty] | 0
P | 2018-06-09T14:02:15+08:00 | 2018-06-09T16:45:30+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP620G | CPSM Final Test | PASS | 129 | 160.97 | 0.84 | -0.84 | 56.32 | 48.77 | 1.44 | 0.13 | 0 | 1 | 0 | -0.9338 | 0.5 | -2.4 | 3 |  |  | 0.1059 | 0.4394 | -0.013 | 0.99999 | [21 empty] | 0 | [467 empty] | 0
P | 2018-06-09T14:52:26+08:00 | 2018-06-09T17:54:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP621G | CPSM Final Test | PASS | 129 | 165.83 | 0.37 | -0.84 | 56.16 | 49.28 | 1.06 | 0.13 | 1 | -2 | -1 | -0.8041 | 0.5328 | -0.7 | 80 |  |  | 0.0354 | 0.1681 | -0.0005 | 0.99998 | [21 empty] | 0 | [452 empty] | 0
P | 2018-06-10T09:38:29+08:00 | 2018-06-10T11:05:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | F18FNP489G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-06-10T08:16:29+08:00 | 2018-06-10T12:27:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | F18FNP489G | CPSM Final Test | FAIL
P | 2018-06-10T10:28:53+08:00 | 2018-06-10T13:39:39+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | F18FTP622G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.23 | 56.23 | 48.09 | 1.29 | 0.13 | 1 | -2 | 1 | -0.8759 | 0.466 | 6.2 | -3 |  |  | 0.0753 | 0.1948 | -0.0048 | 0.99997 | [21 empty] | 0 | [468 empty] | 0
P | 2018-06-10T10:25:23+08:00 | 2018-06-10T13:40:58+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | F18FNP487G | CPSM Final Test | PASS |  | 163.04 | -0.02 | 0.23 | 56.82 | 48.01 | 2.12 | 0.13 | 0 | -2 | 0 | -0.89 | 0.5488 | 1.9 | 8 |  |  | 0.0764 | 0.1524 | -0.002 | 1 |  |  |  |  | 133 |  | 0 | [484 empty] | 0
P | 2018-06-10T10:14:18+08:00 | 2018-06-10T13:46:13+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | F18FNP488G | CPSM Final Test | PASS |  | 165.47 | -0.02 | -0.38 | 56.43 | 51.31 | 1.21 | 0.13 | 0 | -2 | 0 | -0.8581 | 0.516 | 2.5 | 30 |  |  | 0.1336 | 0.1682 | -0.0271 | 0.99997 |  |  |  |  | 133 |  | 0 | [485 empty] | 0
P | 2018-06-10T10:48:50+08:00 | 2018-06-10T13:48:43+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP623G | CPSM Final Test | PASS | 129 | 166.28 | 0.37 | -0.61 | 56.4 | 51.65 | 1.29 | 0.13 | 1 | -2 | 0 | -0.9078 | 0.4834 | 1.6 | 0 |  |  | 0.2507 | 0.2108 | -0.0479 | 0.99999 | [21 empty] | 0 | [471 empty] | 0
P | 2018-06-10T10:56:58+08:00 | 2018-06-10T13:55:30+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP625G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -0.77 | 56.2 | 51.22 | 1.21 | 0.13 | 0 | -2 | 0 | -0.8959 | 0.4958 | 5.8 | -19 |  |  | 0.1602 | 0.1636 | 0.0332 | 0.99999 | [21 empty] | 0 | [472 empty] | 0
P | 2018-06-10T10:21:48+08:00 | 2018-06-10T14:15:35+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | F18FNP489G | CPSM Final Test | PASS |  | 167.09 | 0.56 | -0.84 | 56.62 | 48.77 | 1.44 | 0.13 | 1 | 1 | 0 | -0.8779 | 0.5496 | -1 | 2 |  |  | 0.0985 | 0.0915 | -0.0066 | 0.99999 |  |  |  |  | 133 |  | 0 | [488 empty] | 0
P | 2018-06-10T13:29:47+08:00 | 2018-06-10T17:33:56+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP624G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -0.46 | 56.44 | 53.34 | 1.06 | 0.13 | -1 | -3 | 0 | -0.8121 | 0.541 | 8.3 | 1 |  |  | 0.1339 | 0.2317 | 0.011 | 0.99996 | [21 empty] | 0 | [273 empty] | 0
P | 2018-06-10T11:53:59+08:00 | 2018-06-10T17:34:38+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP629G | CPSM Final Test | FAIL | [13 empty] | -6.3
P | 2018-06-10T14:29:55+08:00 | 2018-06-10T17:51:34+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP627G | CPSM Final Test | PASS | 129 | 177.08 | 0.17 | -0.69 | 56.53 | 50.38 | 1.97 | 0.13 | -1 | 1 | -1 | -0.9178 | 0.489 | 7.8 | 12 |  |  | 0.1277 | 0.2002 | -0.0216 | 0.99998 | [21 empty] | 0 | [474 empty] | 0
P | 2018-06-09T20:37:05+08:00 | 2018-06-11T08:51:32+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP629G | CPSM Final Test | FAIL
P | 2018-06-10T14:44:19+08:00 | 2018-06-11T08:58:06+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP629G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | -0.54 | 56.42 | 52.24 | 1.21 | 0.13 | 0 | -4 | 1 | -0.9038 | 0.516 | 8.5 | 3 |  |  | 0.0979 | 0.1149 | -0.031 | 0.99999 | [21 empty] | 0 | [475 empty] | 0
P | 2018-06-09T19:01:04+08:00 | 2018-06-11T09:57:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP626G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-06-10T17:21:01+08:00 | 2018-06-11T11:42:14+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | F18FTP633G | CPSM Final Test | PASS | 129 | 162.14 | 0.97 | -0.31 | 56.7 | 54.53 | 1.21 | 0.13 | 0 | -1 | 0 | -0.876 | 0.539 | 4.2 | 341 |  |  | 0.052 | 0.0649 | -0.0071 | 1 | [21 empty] | 0 | [476 empty] | 0
P | 2018-06-10T17:23:56+08:00 | 2018-06-11T11:44:39+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP628G | CPSM Final Test | PASS | 129 | 164.21 | 0.78 | -0.69 | 56.32 | 50.38 | 1.06 | 0.13 | 1 | -4 | 0 | -0.862 | 0.481 | -3.6 | 127 |  |  | 0.0865 | 0.1998 | -0.0083 | 0.99998 | [21 empty] | 0 | [477 empty] | 0
P | 2018-06-10T17:27:50+08:00 | 2018-06-11T11:47:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | F18FTP631G | CPSM Final Test | PASS | 129 | 162.77 | 0.46 | -0.61 | 56.43 | 53.09 | 1.51 | 0.13 | -1 | -1 | 0 | -0.8421 | 0.4998 | 2.7 | 20 |  |  | 0.1021 | 0.3109 | 0.1031 | 0.99999 | [21 empty] | 0 | [478 empty] | 0
P | 2018-06-10T14:29:19+08:00 | 2018-06-11T11:58:53+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP626G | CPSM Final Test | PASS | 129 | 161.15 | 0.78 | -0.77 | 55.93 | 49.11 | 1.21 | 0.13 | 1 | -2 | 0 | -0.8879 | 0.4848 | 9 | 20 |  |  | 0.0487 | 0.2209 | -0.0091 | 0.99999 | [21 empty] | 0 | [479 empty] | 0
P | 2018-06-11T10:31:07+08:00 | 2018-06-11T13:23:50+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP632G | CPSM Final Test | PASS | 129 | 160.16 | 0.97 | -0.16 | 56.79 | 50.29 | 1.13 | 0.13 | 0 | -4 | -1 | -0.884 | 0.5272 | -1.9 | -19 |  |  | 0.0243 | 0.2236 | 0.0115 | 0.99999 | [21 empty] | 0 | [480 empty] | 0
P | 2018-06-11T10:38:57+08:00 | 2018-06-11T13:28:56+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP634G | CPSM Final Test | PASS | 129 | 161.42 | -0.02 | -0.38 | 56.71 | 52.15 | 0.98 | 0.13 | -1 | 1 | 1 | -0.85 | 0.545 | 6.3 | -6 |  |  | 0.1651 | 0.1038 | 0.0032 | 0.99999 | [21 empty] | 0 | [481 empty] | 0
P | 2018-06-11T10:27:56+08:00 | 2018-06-11T13:26:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP630G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -0.46 | 56.62 | 49.95 | 1.74 | 0.13 | 0 | -2 | 0 | -0.7642 | 0.5484 | -0.3 | 25 |  |  | 0.119 | 0.26 | 0.0615 | 0.99998 | [21 empty] | 0 | [482 empty] | 0
P | 2018-06-11T10:15:28+08:00 | 2018-06-11T14:11:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | F18FTP635G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-06-11T13:05:54+08:00 | 2018-06-11T14:28:01+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP638G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-06-11T08:34:06+08:00 | 2018-06-11T15:52:41+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | F18FTP635G | CPSM Final Test | FAIL | [13 empty] | 15.8
P | 2018-06-11T13:57:05+08:00 | 2018-06-11T17:22:24+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | F18FTP637G | CPSM Final Test | PASS | 129 | 164.75 | 0.27 | -0.46 | 56.14 | 48.18 | 1.29 | 0.13 | 0 | -1 | -1 | -0.8979 | 0.5416 | -3.3 | -4 |  |  | 0.0983 | 0.1442 | -0.0097 | 1 | [21 empty] | 0 | [82 empty] | 0
P | 2018-06-11T13:46:58+08:00 | 2018-06-11T17:31:23+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP638G | CPSM Final Test | PASS | 129 | 166.19 | 0.27 | -0.38 | 56.35 | 48.26 | 0.98 | 0.13 | -1 | 1 | -1 | -0.8241 | 0.472 | 1.7 | 15 |  |  | 0.058 | 0.3052 | -0.0295 | 0.99999 | [21 empty] | 0 | [301 empty] | 0
P | 2018-06-11T13:54:43+08:00 | 2018-06-11T17:35:55+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | F18FTP636G | CPSM Final Test | PASS | 129 | 166.46 | 0.27 | 0.38 | 56.22 | 51.14 | 1.36 | 0.13 | 1 | 3 | -1 | -0.9077 | 0.4858 | 10 | 80 |  |  | 0.047 | 0.5025 | -0.0305 | 0.99999 | [21 empty] | 0 | [483 empty] | 0
P | 2018-06-11T16:23:17+08:00 | 2018-06-11T16:37:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP640G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-06-11T12:13:23+08:00 | 2018-06-11T18:01:11+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | F18FTP635G | CPSM Final Test | PASS | 129 | 164.57 | 0.37 | -0.77 | 56.08 | 48.85 | 1.06 | 0.13 | -1 | -4 | 0 | -0.8939 | 0.4898 | 6.3 | 51 |  |  | 0.0335 | 0.0997 | -0.0256 | 1 | [21 empty] | 0 | [484 empty] | 0
P | 2018-06-11T14:14:51+08:00 | 2018-06-11T18:37:56+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP639G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-06-11T16:26:58+08:00 | 2018-06-11T21:21:50+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP642G | CPSM Final Test | PASS | 129 | 167.9 | -0.02 | -0.54 | 56.43 | 52.58 | 0.98 | 0.13 | 1 | -2 | 0 | -0.9278 | 0.498 | -1.8 | 34 |  |  | 0.0901 | 0.2218 | 0.0039 | 0.99998 | [21 empty] | 0 | [485 empty] | 0
P | 2018-06-11T16:26:23+08:00 | 2018-06-11T21:23:26+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP639G | CPSM Final Test | PASS | 129 | 167.45 | -0.02 | -0.38 | 56.25 | 48.18 | 0.75 | 0.13 | -1 | 1 | 0 | -0.9299 | 0.4982 | 6.3 | 2 |  |  | 0.0925 | 0.3415 | -0.0319 | 1 | [21 empty] | 0 | [486 empty] | 0
P | 2018-06-11T16:19:43+08:00 | 2018-06-11T21:25:19+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP643G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.69 | 56.72 | 54.69 | 1.21 | 0.13 | 1 | 2 | 1 | -0.9118 | 0.521 | 6.3 | 2 |  |  | 0.1227 | 0.3401 | -0.0266 | 0.99999 | [21 empty] | 0 | [487 empty] | 0
P | 2018-06-11T18:39:24+08:00 | 2018-06-11T21:27:22+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP640G | CPSM Final Test | PASS | 129 | 165.02 | 0.97 | -0.46 | 56.79 | 50.21 | 1.67 | 0.13 | 1 | -1 | -1 | -0.872 | 0.5354 | 4.3 | -6 |  |  | 0.0923 | 0.1152 | 0.043 | 0.99993 | [21 empty] | 0 | [256 empty] | 0
P | 2018-06-12T09:21:00+08:00 | 2018-06-12T10:31:50+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP645G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-06-12T08:24:02+08:00 | 2018-06-12T11:28:48+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP645G | CPSM Final Test | FAIL | [13 empty] | 17.3
P | 2018-06-11T19:43:54+08:00 | 2018-06-12T11:57:17+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP644G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -0.77 | 56.18 | 49.45 | 0.98 | 0.13 | 1 | -4 | 0 | -0.8501 | 0.492 | 1 | 51 |  |  | 0.1372 | 0.2158 | 0.0009 | 1 | [21 empty] | 0 | [488 empty] | 0
P | 2018-06-12T09:56:25+08:00 | 2018-06-12T13:29:00+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP645G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.46 | 56.35 | 48.34 | 1.29 | 0.13 | 1 | 3 | 0 | -0.8321 | 0.5114 | 14.1 | 32 |  |  | 0.0554 | 0.2347 | 0.0932 | 0.99996 | [21 empty] | 0 | [489 empty] | 0
P | 2018-06-12T09:58:40+08:00 | 2018-06-12T13:28:24+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP646G | CPSM Final Test | PASS | 129 | 167 | 0.27 | -0.54 | 56.5 | 51.73 | 1.21 | 0.13 | 1 | -2 | 0 | -0.8979 | 0.5358 | 3 | -5 |  |  | 0.0973 | 0.219 | -0.0031 | 0.99998 | [21 empty] | 0 | [490 empty] | 0
P | 2018-06-11T00:09:07+08:00 | 2018-06-12T15:27:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | F18FTP641G | CPSM Final Test | FAIL | [18 empty] | 3.3528
P | 2018-06-12T09:19:25+08:00 | 2018-06-12T15:52:54+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP650G | CPSM Final Test | FAIL | [14 empty] | 1425
P | 2018-06-12T08:42:29+08:00 | 2018-06-12T16:27:57+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP649G | CPSM Final Test | FAIL
P | 2018-06-11T19:48:10+08:00 | 2018-06-12T16:30:19+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | F18FTP641G | CPSM Final Test | PASS | 129 | 168.26 | -0.02 | -0.61 | 56.41 | 54.53 | 1.36 | 0.13 | -1 | 3 | 0 | -0.88 | 0.5432 | 12.1 | 6 |  |  | 0.096 | 0.2567 | -0.0485 | 1 | [21 empty] | 0 | [91 empty] | 0
P | 2018-06-12T12:26:12+08:00 | 2018-06-12T16:41:03+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | F18FTP648G | CPSM Final Test | PASS | 129 | 165.11 | -0.02 | 0.68 | 56.82 | 48.6 | 1.29 | 0.13 | 1 | 3 | 0 | -0.896 | 0.546 | 4.8 | 9 |  |  | 0.1459 | 0.0807 | -0.0245 | 1 | [21 empty] | 0 | [201 empty] | 0
P | 2018-06-12T08:25:51+08:00 | 2018-06-12T16:46:28+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP650G | CPSM Final Test | FAIL
P | 2018-06-12T14:51:16+08:00 | 2018-06-12T18:09:10+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP652G | CPSM Final Test | PASS | 129 | 165.29 | -0.02 | -0.46 | 56.62 | 51.14 | 1.82 | 0.13 | -1 | 2 | 1 | -0.8301 | 0.5296 | 1.2 | 19 |  |  | 0.0662 | 0.2805 | 0.0339 | 0.99995 | [21 empty] | 0 | [491 empty] | 0
P | 2018-06-12T12:35:13+08:00 | 2018-06-12T18:12:09+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP649G | CPSM Final Test | PASS | 129 | 160.97 | 0.17 | -0.61 | 56.23 | 48.43 | 1.06 | 0.13 | -1 | 1 | -1 | -0.9239 | 0.5298 | 9.6 | 0 |  |  | 0.0956 | 0.4197 | -0.0108 | 1 | [21 empty] | 0 | [492 empty] | 0
P | 2018-06-12T14:53:19+08:00 | 2018-06-12T18:11:34+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP653G | CPSM Final Test | PASS | 129 | 160.34 | 0.37 | -0.54 | 56.14 | 50.29 | 1.29 | 0.13 | 0 | 1 | 0 | -0.892 | 0.513 | -4.5 | 6 |  |  | 0.0619 | 0.5344 | -0.0499 | 1 | [21 empty] | 0 | [493 empty] | 0
P | 2018-06-12T14:44:35+08:00 | 2018-06-12T18:13:37+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP647G | CPSM Final Test | PASS | 129 | 165.56 | 0.17 | -0.69 | 56 | 48.26 | 1.29 | 0.13 | 0 | 1 | 0 | -0.892 | 0.5344 | -0.3 | 7 |  |  | 0.0707 | 0.1669 | -0.0137 | 0.99999 | [21 empty] | 0 | [494 empty] | 0
P | 2018-06-12T12:36:09+08:00 | 2018-06-12T18:30:44+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP650G | CPSM Final Test | PASS | 129 | 162.68 | 0.17 | -0.38 | 56.42 | 51.05 | 1.29 | 0.13 | 1 | -3 | 1 | -0.8939 | 0.4886 | 1.3 | -4 |  |  | 0.0406 | 0.0751 | 0.0355 | 0.99999 | [21 empty] | 0 | [495 empty] | 0
P | 2018-06-12T13:14:42+08:00 | 2018-06-12T19:26:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | F18FTP654G | CPSM Final Test | FAIL | [539 empty] | 0
P | 2018-06-12T16:20:43+08:00 | 2018-06-12T20:32:07+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | F18FTP654G | CPSM Final Test | PASS | 129 | 167 | -0.02 | -0.23 | 56.47 | 49.45 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9197 | 0.4984 | 9.3 | -19 |  |  | 0.1076 | 0.1521 | -0.0371 | 0.99999 | [21 empty] | 0 | [496 empty] | 0
P | 2018-06-12T18:53:43+08:00 | 2018-06-12T22:13:15+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | F18FTP655G | CPSM Final Test | PASS | 129 | 162.32 | 0.17 | -0.31 | 56.2 | 49.7 | 1.59 | 0.13 | 0 | -3 | 0 | -0.9179 | 0.547 | 7.7 | -29 |  |  | 0.0534 | 0.2338 | 0.0344 | 0.99998 | [21 empty] | 0 | [149 empty] | 0
P | 2018-06-12T18:57:17+08:00 | 2018-06-12T22:20:28+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP656G | CPSM Final Test | PASS | 129 | 171.59 | 0.27 | -0.23 | 56.4 | 54.27 | 1.21 | 0.13 | 0 | -2 | 0 | -0.898 | 0.4878 | 4 | -16 |  |  | 0.1674 | 0.6071 | -0.012 | 1 | [21 empty] | 0 | [497 empty] | 0
P | 2018-06-12T18:45:34+08:00 | 2018-06-12T22:26:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | F18FTP658G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -0.99 | 56.12 | 49.87 | 0.52 | 0.13 | 0 | 1 | 0 | -0.9477 | 0.527 | 5 | -14 |  |  | 0.1029 | 0.3769 | 0.0077 | 0.99999 | [21 empty] | 0 | [498 empty] | 0
P | 2018-06-12T03:08:42+08:00 | 2018-06-13T10:29:03+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP651G | CPSM Final Test | FAIL | [508 empty] | 0
P | 2018-06-13T08:46:36+08:00 | 2018-06-13T11:19:12+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP657G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-06-13T10:18:39+08:00 | 2018-06-13T11:24:12+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP661G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-06-13T08:18:55+08:00 | 2018-06-13T11:31:28+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP659G | CPSM Final Test | FAIL | [13 empty] | -7.4
P | 2018-06-12T18:48:52+08:00 | 2018-06-13T11:34:41+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP651G | CPSM Final Test | PASS | 129 | 166.19 | -0.02 | -0.61 | 56.18 | 48.94 | 1.29 | 0.13 | 1 | 3 | 0 | -0.86 | 0.4906 | 1.2 | 3 |  |  | 0.102 | 0.2674 | -0.0034 | 0.99993 | [21 empty] | 0 | [465 empty] | 0
P | 2018-06-13T11:05:00+08:00 | 2018-06-13T13:47:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | F18FTP662G | CPSM Final Test | PASS | 129 | 162.14 | -0.02 | -0.08 | 56.24 | 48.68 | 1.51 | 0.13 | 0 | -2 | 1 | -0.8739 | 0.5158 | -1.7 | 7 |  |  | 0.0427 | 0.1452 | 0.0243 | 1 | [21 empty] | 0.02618 | [499 empty] | 0
P | 2018-06-13T10:02:54+08:00 | 2018-06-13T13:53:14+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | F18FTP657G | CPSM Final Test | PASS | 129 | 158.18 | -0.02 | -0.46 | 55.85 | 49.11 | 0.83 | 0.13 | -2 | 2 | -1 | -0.88 | 0.5324 | 5.3 | 3 |  |  | 0.0805 | 0.1539 | 0.0567 | 0.99997 | [21 empty] | 0 | [500 empty] | 0
P | 2018-06-13T10:51:25+08:00 | 2018-06-13T13:56:18+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | F18FTP661G | CPSM Final Test | PASS | 129 | 157.73 | 0.46 | -0.54 | 55.99 | 49.95 | 1.29 | 0.13 | 0 | 1 | -1 | -0.8242 | 0.5372 | -1.1 | -33 |  |  | 0.0832 | 0.2597 | 0.0542 | 0.99993 | [21 empty] | 0 | [501 empty] | 0
P | 2018-06-13T12:17:28+08:00 | 2018-06-13T14:55:43+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP663G | CPSM Final Test | FAIL | [13 empty] | -9.4
P | 2018-06-13T05:06:35+08:00 | 2018-06-13T14:59:07+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP660G | CPSM Final Test | FAIL
P | 2018-06-13T04:17:09+08:00 | 2018-06-13T15:33:14+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP659G | CPSM Final Test | FAIL
P | 2018-06-13T11:31:32+08:00 | 2018-06-13T15:49:01+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | F18FTP664G | CPSM Final Test | FAIL
P | 2018-06-13T13:33:20+08:00 | 2018-06-13T16:30:27+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | F18FTP665G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.54 | 56.42 | 56.39 | 1.44 | 0.13 | 0 | -2 | 0 | -0.8978 | 0.5 | 7.7 | -20 |  |  | 0.1004 | 0.2118 | -0.0731 | 0.99998 | [21 empty] | 0.01082 | [496 empty] | 0
P | 2018-06-13T15:56:53+08:00 | 2018-06-13T16:39:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP667G | CPSM Final Test | FAIL | [5 empty] | 47.16
P | 2018-06-13T10:02:51+08:00 | 2018-06-13T16:43:31+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | F18FTP660G | CPSM Final Test | PASS | 129 | 162.68 | 0.78 | -0.38 | 56.49 | 48.26 | 1.21 | 0.13 | 0 | 1 | 0 | -0.87 | 0.542 | 2.3 | -14 |  |  | 0.0457 | 0.2118 | -0.0485 | 1 | [21 empty] | 0 | [502 empty] | 0
P | 2018-06-13T13:36:35+08:00 | 2018-06-13T16:55:43+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | F18FTP663G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.23 | 55.92 | 48.77 | 1.36 | 0.13 | 1 | 1 | 0 | -0.8341 | 0.4848 | 1.9 | 22 |  |  | 0.1005 | 0.1036 | 0.0048 | 0.99999 | [21 empty] | 0 | [465 empty] | 0
P | 2018-06-13T15:03:01+08:00 | 2018-06-13T17:07:48+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP668G | CPSM Final Test | FAIL | [13 empty] | -11.2
P | 2018-06-13T09:55:11+08:00 | 2018-06-13T17:18:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | F18FTP659G | CPSM Final Test | PASS | 129 | 165.83 | 0.37 | -0.23 | 55.91 | 50.38 | 1.9 | 0.13 | 0 | -2 | 1 | -0.8361 | 0.5046 | 0.2 | -15 |  |  | 0.1124 | 0.1125 | 0.0286 | 1 | [21 empty] | 0 | [503 empty] | 0
P | 2018-06-13T14:42:53+08:00 | 2018-06-13T17:27:56+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP668G | CPSM Final Test | FAIL | [13 empty] | -7.3
P | 2018-06-13T13:40:16+08:00 | 2018-06-13T17:33:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | F18FTP664G | CPSM Final Test | PASS | 129 | 168.71 | -0.02 | -0.46 | 56.4 | 54.86 | 1.21 | 0.13 | 0 | -2 | 0 | -0.8978 | 0.474 | 5.8 | 12 |  |  | 0.094 | 0.1407 | -0.0269 | 0.99999 | [21 empty] | 0 | [504 empty] | 0
P | 2018-06-13T15:49:35+08:00 | 2018-06-13T18:44:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP666G | CPSM Final Test | PASS | 129 | 163.67 | 0.27 | -0.99 | 56.46 | 51.73 | 1.67 | 0.13 | 0 | 1 | 0 | -0.9398 | 0.5458 | -0.9 | 134 |  |  | 0.1078 | 0.1853 | -0.0468 | 1 | [21 empty] | 0 | [505 empty] | 0
P | 2018-06-13T16:18:04+08:00 | 2018-06-13T19:10:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP667G | CPSM Final Test | PASS | 129 | 163.22 | 0.56 | -1.07 | 56.27 | 48.68 | 0.83 | 0.13 | -1 | -1 | 0 | -0.8779 | 0.502 | 4.4 | 10 |  |  | 0.0799 | 0.1686 | 0.0505 | 0.99996 | [21 empty] | 0 | [506 empty] | 0
P | 2018-06-13T18:21:10+08:00 | 2018-06-13T19:24:56+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP673G | CPSM Final Test | FAIL | [5 empty] | 42.5
P | 2018-06-13T16:05:24+08:00 | 2018-06-13T19:30:54+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP668G | CPSM Final Test | PASS | 129 | 162.68 | -0.11 | -0.99 | 56.32 | 49.78 | 1.51 | 0.13 | 0 | -2 | 0 | -0.9139 | 0.527 | 1.3 | -5 |  |  | 0.114 | 0.0965 | -0.0361 | 0.99995 | [21 empty] | 0 | [507 empty] | 0
P | 2018-06-13T17:42:40+08:00 | 2018-06-13T21:18:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP671G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.77 | 56.17 | 48.77 | 1.36 | 0.13 | 0 | 1 | 0 | -0.8959 | 0.531 | 6.4 | -11 |  |  | 0.0493 | 0.0376 | -0.0619 | 1 | [21 empty] | 0 | [508 empty] | 0
P | 2018-06-13T18:25:18+08:00 | 2018-06-13T21:32:13+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP670G | CPSM Final Test | PASS | 129 | 165.47 | -0.3 | -0.99 | 56.2 | 49.28 | 1.36 | 0.13 | 0 | 1 | 0 | -0.9278 | 0.5082 | 1 | 12 |  |  | 0.0869 | 0.2391 | -0.0209 | 0.99997 | [21 empty] | 0 | [509 empty] | 0
P | 2018-06-13T18:53:03+08:00 | 2018-06-13T21:56:40+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP673G | CPSM Final Test | PASS | 129 | 160.7 | 0.75 | -0.46 | 56.72 | 48.68 | 1.44 | 0.13 | -1 | -2 | -1 | -0.904 | 0.4918 | 7.5 | -3 |  |  | 0.0677 | 0.3809 | -0.0167 | 1 | [21 empty] | 0 | [510 empty] | 0
P | 2018-06-13T04:52:54+08:00 | 2018-06-14T08:31:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP669G | CPSM Final Test | FAIL
P | 2018-06-13T18:42:25+08:00 | 2018-06-14T08:32:25+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP669G | CPSM Final Test | PASS | 129 | 165.47 | 0.27 | -0.99 | 56.21 | 48.6 | 1.06 | 0.13 | 1 | 1 | 1 | -0.9378 | 0.4766 | 0 | 4 |  |  | 0.0865 | 0.1436 | 0.0101 | 0.99999 | [21 empty] | 0 | [511 empty] | 0
P | 2018-06-13T00:25:48+08:00 | 2018-06-14T10:45:19+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP672G | CPSM Final Test | FAIL
P | 2018-06-14T09:40:47+08:00 | 2018-06-14T11:05:58+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP674G | CPSM Final Test | FAIL | [5 empty] | 46.48
P | 2018-06-14T10:38:32+08:00 | 2018-06-14T11:13:30+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP676G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-06-13T17:35:33+08:00 | 2018-06-14T12:29:43+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP672G | CPSM Final Test | PASS | 129 | 167.27 | 0.84 | -0.99 | 56.17 | 49.28 | 0.75 | 0.13 | 0 | 1 | 0 | -0.9039 | 0.5274 | 6.8 | -9 |  |  | 0.1224 | 0.1702 | 0.0173 | 0.99999 | [21 empty] | 0 | [512 empty] | 0
P | 2018-06-14T10:34:24+08:00 | 2018-06-14T13:16:38+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP675G | CPSM Final Test | PASS | 129 | 160.25 | 0.17 | -0.92 | 56.51 | 49.02 | 1.21 | 0.13 | 1 | -2 | 0 | -0.878 | 0.5318 | 2.3 | 44 |  |  | 0.0371 | 0.369 | 0.0189 | 1 | [21 empty] | 0 | [513 empty] | 0
P | 2018-06-14T10:23:22+08:00 | 2018-06-14T13:38:54+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP674G | CPSM Final Test | PASS | 129 | 166.55 | 0.46 | -0.99 | 56.78 | 48.6 | 1.21 | 0.13 | -1 | 1 | -1 | -0.8062 | 0.5284 | 11.8 | -1 |  |  | 0.0464 | 0.4634 | 0.0244 | 0.99996 | [21 empty] | 0 | [514 empty] | 0
P | 2018-06-14T10:56:01+08:00 | 2018-06-14T13:44:04+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP676G | CPSM Final Test | PASS | 129 | 169.88 | 0.08 | -0.46 | 56.94 | 48.85 | 1.67 | 0.13 | 0 | -1 | 1 | -0.8161 | 0.5266 | -1.5 | -3 |  |  | 0.0995 | 0.0934 | 0.0267 | 0.99998 | [21 empty] | 0 | [507 empty] | 0
P | 2018-06-14T13:10:28+08:00 | 2018-06-14T14:30:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP681G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-06-14T13:33:43+08:00 | 2018-06-14T15:12:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP677G | CPSM Final Test | FAIL | [558 empty] | FALSE
P | 2018-06-14T14:07:52+08:00 | 2018-06-14T16:54:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP680G | CPSM Final Test | PASS | 129 | 166.19 | 0.56 | -0.69 | 56.41 | 48.94 | 1.59 | 0.13 | 0 | -2 | 0 | -0.8142 | 0.5216 | 12 | -1 |  |  | 0.1829 | 0.0519 | -0.0309 | 0.99998 | [21 empty] | 0 | [83 empty] | 0
P | 2018-06-14T13:50:23+08:00 | 2018-06-14T17:14:57+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP681G | CPSM Final Test | PASS | 129 | 164.93 | 0.56 | -0.46 | 56.76 | 48.18 | 1.21 | 0.13 | 0 | -3 | 0 | -0.9218 | 0.508 | 6.9 | -12 |  |  | 0.0786 | 0.1165 | -0.0567 | 1 | [21 empty] | 0.00825 | [516 empty] | 0
P | 2018-06-14T09:29:52+08:00 | 2018-06-14T17:19:48+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP678G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-06-14T13:24:50+08:00 | 2018-06-14T19:23:10+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP678G | CPSM Final Test | PASS | 129 | 164.39 | 0.27 | -1.3 | 56.18 | 49.45 | 1.13 | 0.13 | 0 | -2 | 0 | -0.876 | 0.5238 | 10.5 | -72 |  |  | 0.0567 | 0.1297 | 0.0143 | 1 | [21 empty] | 0 | [517 empty] | 0
P | 2018-06-14T08:53:49+08:00 | 2018-06-14T19:52:21+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP677G | CPSM Final Test | FAIL | [558 empty] | FALSE
P | 2018-06-14T17:20:55+08:00 | 2018-06-14T20:52:53+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP679G | CPSM Final Test | PASS | 129 | 163.04 | 0.97 | -0.77 | 56.42 | 51.65 | 0.91 | 0.13 | -1 | -1 | 0 | -0.8918 | 0.521 | -0.1 | -4 |  |  | 0.0354 | 0.0608 | 0.0436 | 1 | [21 empty] | 0 | [518 empty] | 0
P | 2018-06-14T17:42:26+08:00 | 2018-06-14T21:00:44+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP682G | CPSM Final Test | PASS | 129 | 164.66 | 0.68 | -0.61 | 56.43 | 50.97 | 1.36 | 0.13 | 0 | -1 | 0 | -0.854 | 0.5386 | 1.6 | -31 |  |  | 0.1624 | 0.0696 | -0.0022 | 0.99998 | [21 empty] | 0.01167 | [519 empty] | 0
P | 2018-06-14T18:25:59+08:00 | 2018-06-14T21:06:28+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP684G | CPSM Final Test | PASS | 129 | 163.76 | -0.21 | -0.69 | 56.21 | 49.28 | 1.36 | 0.13 | 0 | 1 | -1 | -0.8759 | 0.5392 | 8.8 | 14 |  |  | 0.1097 | 0.451 | -0.0008 | 0.99999 | [21 empty] | 0 | [520 empty] | 0
P | 2018-06-14T18:43:02+08:00 | 2018-06-14T21:26:32+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP683G | CPSM Final Test | PASS | 129 | 158.9 | 0.4 | -0.69 | 56.33 | 50.21 | 0.98 | 0.13 | 1 | -2 | 0 | -0.892 | 0.539 | 5.6 | 70 |  |  | 0.051 | 0.2489 | 0.027 | 0.99999 | [21 empty] | 0 | [521 empty] | 0
P | 2018-06-13T20:12:55+08:00 | 2018-06-15T08:33:15+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP677G | CPSM Final Test | FAIL
P | 2018-06-14T14:23:05+08:00 | 2018-06-15T08:33:47+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP677G | CPSM Final Test | PASS | 129 | 163.49 | 0.27 | -0.38 | 56.1 | 50.72 | 1.29 | 0.13 | 1 | 3 | 0 | -0.9258 | 0.494 | -0.8 | 36 |  |  | 0.118 | 0.1739 | -0.0228 | 0.99999 | [21 empty] | 0 | [110 empty] | 0
P | 2018-06-15T09:28:14+08:00 | 2018-06-15T10:31:49+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM979G | CPSM Final Test | FAIL | [5 empty] | 46.91
P | 2018-06-15T07:56:47+08:00 | 2018-06-15T11:15:50+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM975G | CPSM Final Test | FAIL | [14 empty] | 1024
P | 2018-06-15T08:51:03+08:00 | 2018-06-15T11:49:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM977G | CPSM Final Test | PASS |  | 159.44 | 0.27 | -0.84 | 56.11 | 50.21 | 1.29 | 0.13 | 1 | -2 | 0 | -0.8002 | 0.4982 | 3.5 | -2 |  |  | 0.0572 | 0.2899 | -0.0191 | 0.99998 |  |  |  |  | 133 |  | 0 | [537 empty] | 0
P | 2018-06-15T09:01:51+08:00 | 2018-06-15T11:51:05+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM976G | CPSM Final Test | PASS |  | 166.19 | 0.17 | -1.07 | 56.17 | 49.28 | 1.29 | 0.13 | 1 | -2 | 0 | -0.864 | 0.5332 | 2.5 | 63 |  |  | 0.1191 | 0.3382 | -0.0158 | 0.99996 |  |  |  |  | 133 |  | 0 | [538 empty] | 0
P | 2018-06-15T09:22:08+08:00 | 2018-06-15T12:17:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM978G | CPSM Final Test | PASS |  | 163.76 | -0.02 | -0.69 | 56.43 | 55.71 | 1.36 | 0.13 | 0 | 1 | 0 | -0.858 | 0.533 | 0.4 | 3 |  |  | 0.2937 | 0.2701 | -0.0384 | 0.99999 |  |  |  |  | 133 |  | 0 | [13 empty] | 0
P | 2018-06-15T09:36:18+08:00 | 2018-06-15T13:03:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM975G | CPSM Final Test | PASS |  | 163.94 | 0.27 | -0.99 | 56.43 | 51.48 | 0.98 | 0.13 | 0 | 1 | 0 | -0.7921 | 0.5302 | 15 | -8 |  |  | 0.0787 | 0.1292 | -0.0346 | 0.99999 |  |  |  |  | 133 |  | 0 | [539 empty] | 0
P | 2018-06-15T10:00:02+08:00 | 2018-06-15T13:06:43+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM979G | CPSM Final Test | PASS |  | 165.92 | -0.02 | -0.61 | 56.35 | 49.02 | 1.44 | 0.13 | -1 | -1 | -1 | -0.848 | 0.5154 | 1.9 | 26 |  |  | 0.0723 | 0.1436 | -0.0397 | 1 |  |  |  |  | 133 |  | 0 | [540 empty] | 0
P | 2018-06-15T13:26:04+08:00 | 2018-06-15T14:18:54+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM982G | CPSM Final Test | FAIL | [5 empty] | 47.16
P | 2018-06-15T12:03:42+08:00 | 2018-06-15T15:41:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM982G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-06-15T13:23:57+08:00 | 2018-06-15T16:19:46+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM985G | CPSM Final Test | FAIL | [13 empty] | -5.5
P | 2018-06-15T15:41:19+08:00 | 2018-06-15T16:23:53+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM986G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-06-15T11:14:50+08:00 | 2018-06-15T16:30:08+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM982G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-06-15T13:39:05+08:00 | 2018-06-15T16:44:48+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM983G | CPSM Final Test | PASS |  | 168.71 | 0.27 | -0.99 | 56.47 | 48.01 | 0.83 | 0.13 | 1 | 1 | 0 | -0.7543 | 0.4882 | -0.6 | 8 |  |  | 0.0943 | 0.2676 | 0.0907 | 0.99992 |  |  |  |  | 133 |  | 0 | [541 empty] | 0
P | 2018-06-15T16:10:59+08:00 | 2018-06-15T17:01:07+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM987G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-06-15T16:04:53+08:00 | 2018-06-15T17:07:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM987G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-06-15T14:32:21+08:00 | 2018-06-15T17:19:39+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM984G | CPSM Final Test | PASS |  | 160.79 | -0.02 | -0.84 | 56.22 | 51.48 | 1.29 | 0.13 | -1 | -3 | 0 | -0.8899 | 0.52 | 0.9 | 14 |  |  | 0.0519 | 0.2104 | -0.0254 | 1 |  |  |  |  | 133 |  | 0.02629 | [542 empty] | 0
P | 2018-06-15T14:13:37+08:00 | 2018-06-15T17:27:42+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM981G | CPSM Final Test | PASS |  | 163.94 | 0.37 | -0.69 | 56.44 | 52.66 | 1.44 | 0.13 | 0 | -3 | 0 | -0.8441 | 0.5438 | -5 | 1 |  |  | 0.0731 | 0.1526 | 0.0256 | 0.99998 |  |  |  |  | 133 |  | 0 | [543 empty] | 0
P | 2018-06-15T14:37:37+08:00 | 2018-06-15T17:27:35+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM986G | CPSM Final Test | FAIL | [13 empty] | -5.3
P | 2018-06-15T15:30:03+08:00 | 2018-06-15T18:17:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM980G | CPSM Final Test | PASS |  | 172.13 | 0.27 | -0.46 | 56.32 | 48.26 | 1.44 | 0.13 | 0 | -1 | 0 | -0.804 | 0.5198 | 2.8 | 11 |  |  | 0.0806 | 0.0675 | -0.0536 | 1 |  |  |  |  | 133 |  | 0 | [339 empty] | 0
P | 2018-06-15T14:51:51+08:00 | 2018-06-15T18:23:48+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM985G | CPSM Final Test | PASS |  | 155.3 | 0.27 | -0.84 | 56.47 | 49.45 | 1.67 | 0.13 | 1 | 3 | 0 | -0.868 | 0.522 | 9.3 | 23 |  |  | 0.0818 | 0.1453 | -0.012 | 1 |  |  |  |  | 133 |  | 0 | [544 empty] | 0
P | 2018-06-15T13:52:29+08:00 | 2018-06-15T18:34:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM982G | CPSM Final Test | PASS |  | 167.27 | 0.46 | -0.77 | 56.28 | 48.85 | 1.51 | 0.13 | -1 | -1 | 0 | -0.7862 | 0.521 | -4.5 | 40 |  |  | 0.1203 | 0.4274 | -0.0565 | 1 |  |  |  |  | 133 |  | 0 | [39 empty] | 0
P | 2018-06-15T16:02:36+08:00 | 2018-06-15T19:31:23+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM986G | CPSM Final Test | PASS |  | 160.7 | 0.78 | -0.92 | 56.4 | 48.51 | 1.59 | 0.13 | -1 | 1 | 0 | -0.8161 | 0.531 | 0.8 | 15 |  |  | 0.029 | 0.4161 | -0.0025 | 0.99997 |  |  |  |  | 133 |  | 0 | [545 empty] | 0
P | 2018-06-15T16:36:03+08:00 | 2018-06-15T19:54:49+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM987G | CPSM Final Test | PASS |  | 158.27 | 0.46 | -0.84 | 56.32 | 48.51 | 0.91 | 0.13 | -2 | 0 | -1 | -0.858 | 0.512 | -2.3 | 37 |  |  | 0.0203 | 0.1698 | -0.0087 | 0.99998 |  |  |  |  | 133 |  | 0 | [546 empty] | 0
P | 2018-06-16T12:21:03+08:00 | 2018-06-16T14:14:04+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | F18USM988G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-06-16T11:55:09+08:00 | 2018-06-16T14:11:17+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | F18USM991G | CPSM Final Test | FAIL | [5 empty] | 42.25
P | 2018-06-16T11:40:53+08:00 | 2018-06-16T14:13:03+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | F18USM993G | CPSM Final Test | FAIL | [5 empty] | 45.47
P | 2018-06-16T10:48:57+08:00 | 2018-06-16T15:46:10+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | F18USM988G | CPSM Final Test | FAIL | [13 empty] | -6.1
P | 2018-06-16T10:23:45+08:00 | 2018-06-16T15:42:41+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | F18USM991G | CPSM Final Test | FAIL | [13 empty] | -8.4
P | 2018-06-16T13:03:32+08:00 | 2018-06-16T15:52:08+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | F18USM996G | CPSM Final Test | FAIL | [5 empty] | 46.48
P | 2018-06-16T13:07:06+08:00 | 2018-06-16T16:22:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | F18USM990G | CPSM Final Test | PASS |  | 163.13 | 0.17 | -0.69 | 56.33 | 50.8 | 1.44 | 0.13 | 1 | -2 | 1 | -0.8658 | 0.531 | 0.8 | -3 |  |  | 0.0914 | 0.1078 | 0.0016 | 0.99999 |  |  |  |  | 133 |  | 0 | [547 empty] | 0
P | 2018-06-16T13:10:47+08:00 | 2018-06-16T16:22:37+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | F18USM989G | CPSM Final Test | PASS |  | 163.94 | 0.68 | -0.92 | 56.31 | 50.04 | 0.91 | 0.13 | 1 | 1 | 0 | -0.8579 | 0.5024 | 6.7 | 32 |  |  | 0.0587 | 0.1707 | -0.0404 | 0.99999 |  |  |  |  | 133 |  | 0 | [548 empty] | 0
P | 2018-06-16T12:59:28+08:00 | 2018-06-16T16:22:13+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | F18USM992G | CPSM Final Test | PASS |  | 161.42 | -0.02 | -0.54 | 56.42 | 53.34 | 1.44 | 0.13 | 0 | 1 | 0 | -0.8281 | 0.518 | 3.8 | 30 |  |  | 0.1277 | 0.1508 | -0.0551 | 1 |  |  |  |  | 133 |  | 0 | [510 empty] | 0
P | 2018-06-16T13:17:34+08:00 | 2018-06-16T17:50:12+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | F18USM988G | CPSM Final Test | PASS |  | 159.44 | 0.27 | 0.38 | 56.63 | 48.43 | 1.21 | 0.13 | -2 | 1 | 0 | -0.874 | 0.5184 | -0.8 | 93 |  |  | 0.0497 | 0.1156 | 0.0422 | 0.99998 |  |  |  |  | 133 |  | 0 | [549 empty] | 0
P | 2018-06-16T13:03:13+08:00 | 2018-06-16T17:46:49+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | F18USM991G | CPSM Final Test | PASS |  | 163.67 | 0.46 | -0.46 | 56.26 | 48.26 | 1.29 | -0.16 | 1 | 1 | 0 | -0.874 | 0.528 | 4.2 | 12 |  |  | 0.1218 | 0.1962 | -0.0328 | 0.99999 |  |  |  |  | 133 |  | 0 | [550 empty] | 0
P | 2018-06-16T08:05:41+08:00 | 2018-06-16T17:48:15+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | F18USM993G | CPSM Final Test | FAIL
P | 2018-06-16T12:56:58+08:00 | 2018-06-16T17:48:49+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | F18USM993G | CPSM Final Test | PASS |  | 163.31 | 0.78 | -0.61 | 56.19 | 48.85 | 1.36 | 0.13 | 1 | -2 | 1 | -0.8839 | 0.5276 | -4.6 | 0 |  |  | 0.1035 | 0.1927 | -0.0151 | 0.99999 |  |  |  |  | 133 |  | 0 | [551 empty] | 0
P | 2018-06-16T11:01:12+08:00 | 2018-06-16T17:49:54+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | F18USM994G | CPSM Final Test | FAIL | [13 empty] | -5.1
P | 2018-06-16T14:27:50+08:00 | 2018-06-16T18:52:06+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | F18USM996G | CPSM Final Test | PASS |  | 166.1 | -0.02 | -0.61 | 55.99 | 48.68 | 1.13 | 0.13 | 1 | -2 | 1 | -0.8859 | 0.5052 | -2.6 | 29 |  |  | 0.0366 | 0.2926 | -0.0295 | 0.99999 |  |  |  |  | 133 |  | 0 | [552 empty] | 0
P | 2018-06-16T17:43:49+08:00 | 2018-06-16T19:12:37+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | F18USM999G | CPSM Final Test | FAIL | [5 empty] | 47.16
P | 2018-06-15T19:43:16+08:00 | 2018-06-17T09:07:50+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | F18USM994G | CPSM Final Test | FAIL | [17 empty] | 59.8294
P | 2018-06-16T14:25:33+08:00 | 2018-06-17T10:11:02+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | F18USM994G | CPSM Final Test | PASS |  | 164.57 | -0.02 | -0.77 | 56.2 | 48.77 | 1.21 | 0.13 | -1 | 4 | 0 | -1.2391 | 0.5 | -2.5 | 25 |  |  | 0.1041 | 0.2338 | 0.0522 | 0.99993 |  |  |  |  | 133 |  | 0.00934 | [553 empty] | 0
P | 2018-06-16T18:28:13+08:00 | 2018-06-17T10:50:55+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | F18USM999G | CPSM Final Test | PASS |  | 158.09 | 0.37 | -0.54 | 56.82 | 48.09 | 1.74 | 0.13 | 0 | 1 | 1 | -0.83 | 0.525 | 0.3 | 23 |  |  | 0.1021 | 0.3591 | -0.0305 | 0.99999 |  |  |  |  | 133 |  | 0 | [554 empty] | 0
P | 2018-06-17T09:22:30+08:00 | 2018-06-17T10:56:26+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI943G | CPSM Final Test | FAIL | [5 empty] | 46.14
P | 2018-06-16T18:36:45+08:00 | 2018-06-17T10:58:24+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | F18USM998G | CPSM Final Test | PASS |  | 160.52 | 0.46 | -0.61 | 56.47 | 48.01 | 1.29 | 0.13 | -1 | -3 | 0 | -0.9059 | 0.496 | -4.2 | 14 |  |  | 0.0638 | 0.14 | -0.0154 | 1 |  |  |  |  | 133 |  | 0 | [555 empty] | 0
P | 2018-06-16T18:52:36+08:00 | 2018-06-17T10:59:23+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | F18USM995G | CPSM Final Test | PASS |  | 162.95 | 0.17 | -0.99 | 56.34 | 50.88 | 1.06 | 0.13 | 1 | -2 | 1 | -0.9158 | 0.5046 | 2.3 | 10 |  |  | 0.0523 | 0.0543 | -0.0417 | 0.99999 |  |  |  |  | 133 |  | 0 | [556 empty] | 0
P | 2018-06-16T19:11:03+08:00 | 2018-06-17T11:01:05+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | F18USM997G | CPSM Final Test | PASS |  | 166.37 | 0.17 | -0.92 | 56.23 | 50.21 | 0.98 | 0.13 | 0 | -1 | 1 | -0.9319 | 0.532 | 2.4 | 8 |  |  | 0.1183 | 0.2393 | 0.0018 | 0.99997 |  |  |  |  | 133 |  | 0 | [557 empty] | 0
P | 2018-06-17T13:13:43+08:00 | 2018-06-17T15:21:15+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI949G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-06-17T04:32:42+08:00 | 2018-06-17T15:46:14+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI943G | CPSM Final Test | FAIL | [14 empty] | -80
P | 2018-06-17T12:52:49+08:00 | 2018-06-17T15:55:18+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI947G | CPSM Final Test | FAIL | [13 empty] | -9.8
P | 2018-06-17T10:08:06+08:00 | 2018-06-17T16:36:40+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI939G | CPSM Final Test | PASS | 129 | 159.26 | 0.17 | -0.46 | 56.43 | 53.17 | 1.06 | 0.13 | 1 | 3 | 0 | -0.8661 | 0.4898 | -3.8 | 25 |  | 0 | 0.066 | 0.4584 | -0.0151 | 0.99999 | [565 empty] | 0
P | 2018-06-17T10:16:24+08:00 | 2018-06-17T16:43:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI942G | CPSM Final Test | PASS | 129 | 164.84 | 0.27 | -0.54 | 56.21 | 49.02 | 1.9 | 0.13 | 0 | 1 | 0 | -0.8959 | 0.544 | 8.3 | 5 |  | 0 | 0.141 | 0.4153 | -0.0079 | 0.99999 | [566 empty] | 0
P | 2018-06-17T11:35:09+08:00 | 2018-06-17T16:59:49+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI949G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-06-17T10:09:28+08:00 | 2018-06-17T17:31:00+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI943G | CPSM Final Test | PASS | 129 | 157.01 | -0.02 | -0.77 | 56.43 | 49.61 | 0.91 | 0.13 | -2 | 1 | 0 | -0.6804 | 0.5128 | -2.3 | 26 |  | 0 | 0.1446 | 0.1979 | 0.0209 | 0.99999 | [567 empty] | 0
P | 2018-06-17T14:34:31+08:00 | 2018-06-17T17:29:42+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI940G | CPSM Final Test | PASS | 129 | 162.23 | 0.68 | -0.61 | 56.5 | 48.85 | 1.21 | 0.13 | -1 | -2 | 0 | -0.7542 | 0.524 | -2.8 | 11 |  | 0 | 0.0652 | 0.2818 | 0.0286 | 0.99993 | [568 empty] | 0
P | 2018-06-17T14:30:38+08:00 | 2018-06-17T17:37:27+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI946G | CPSM Final Test | PASS | 129 | 162.86 | 0.17 | -0.77 | 56.22 | 49.78 | 1.21 | 0.13 | 0 | -1 | 1 | -0.8998 | 0.4978 | -0.2 | 12 |  | 0 | 0.0436 | 0.4056 | -0.0483 | 0.99999 | [569 empty] | 0
P | 2018-06-17T14:20:39+08:00 | 2018-06-17T17:39:15+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI944G | CPSM Final Test | PASS | 129 | 158 | -0.02 | -0.16 | 56.43 | 50.46 | 0.98 | 0.13 | 1 | 2 | 0 | -0.87 | 0.55 | -3.8 | -10 |  | 0 | 0.4063 | 0.2347 | -0.0185 | 1 | [534 empty] | 0
P | 2018-06-17T14:24:04+08:00 | 2018-06-17T17:55:54+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI947G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.54 | 56.61 | 50.21 | 1.36 | 0.13 | 0 | 2 | 0 | -0.87 | 0.515 | -2.8 | -2 |  | 0 | 0.0743 | 0.3318 | -0.0409 | 0.99999 | [570 empty] | 0
P | 2018-06-17T14:17:29+08:00 | 2018-06-17T19:34:17+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI949G | CPSM Final Test | PASS | 129 | 164.93 | -0.02 | -0.92 | 55.94 | 49.87 | 1.06 | 0.13 | -2 | -2 | -1 | -0.836 | 0.4856 | -2.9 | 10 |  | 0 | 0.1267 | 0.1715 | -0.0265 | 0.99999 | [57 empty] | 0
P | 2018-06-18T10:49:56+08:00 | 2018-06-18T11:56:14+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI951G | CPSM Final Test | FAIL | [5 empty] | 45.55
P | 2018-06-18T10:32:58+08:00 | 2018-06-18T12:04:50+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI941G | CPSM Final Test | FAIL | [5 empty] | 46.82
P | 2018-06-18T08:57:01+08:00 | 2018-06-18T13:26:41+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI956G | CPSM Final Test | FAIL | [14 empty] | -130
P | 2018-06-18T09:06:30+08:00 | 2018-06-18T13:31:18+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI941G | CPSM Final Test | FAIL | [13 empty] | -8.5
P | 2018-06-18T08:12:23+08:00 | 2018-06-18T13:43:47+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI950G | CPSM Final Test | FAIL | [14 empty] | 1173
P | 2018-06-18T08:10:48+08:00 | 2018-06-18T14:04:42+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI945G | CPSM Final Test | FAIL
P | 2018-06-18T11:18:26+08:00 | 2018-06-18T14:14:52+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI955G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.31 | 56.46 | 48.94 | 1.44 | 0.13 | -1 | 1 | 0 | -0.892 | 0.523 | -1.6 | 59 |  | 0 | 0.0791 | 0.3758 | -0.0038 | 1 | [451 empty] | 0
P | 2018-06-18T11:04:19+08:00 | 2018-06-18T14:22:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI957G | CPSM Final Test | PASS | 129 | 166.82 | 0.17 | -0.46 | 56.42 | 50.46 | 1.44 | 0.13 | 0 | -1 | 0 | -0.8879 | 0.544 | 8.3 | -14 |  | 0 | 0.12 | 0.1145 | -0.0245 | 0.99999 | [571 empty] | 0
P | 2018-06-18T11:00:35+08:00 | 2018-06-18T14:24:59+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI953G | CPSM Final Test | PASS | 129 | 164.57 | -0.02 | -0.99 | 56.42 | 53.26 | 0.83 | 0.13 | 0 | 1 | 0 | -0.8622 | 0.541 | 4.4 | 11 |  | 0 | 0.2193 | 0.1844 | 0.0079 | 0.99999 | [56 empty] | 0
P | 2018-06-18T11:23:05+08:00 | 2018-06-18T15:05:22+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI951G | CPSM Final Test | PASS | 129 | 163.94 | -0.21 | -0.38 | 56.73 | 49.78 | 1.82 | 0.13 | 1 | -4 | 0 | -0.9737 | 0.518 | -4.1 | 8 |  | 0 | 0.1423 | 0.1234 | 0.022 | 0.99998 | [572 empty] | 0
P | 2018-06-18T11:11:51+08:00 | 2018-06-18T15:12:43+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI956G | CPSM Final Test | PASS | 129 | 165.38 | 0.27 | 0.15 | 56.81 | 49.28 | 1.36 | 0.13 | 1 | -2 | 0 | -0.8798 | 0.5304 | 3.2 | 8 |  | 0 | 0.13 | 0.1804 | -0.0361 | 1 | [573 empty] | 0
P | 2018-06-18T11:18:54+08:00 | 2018-06-18T15:31:34+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI941G | CPSM Final Test | PASS | 129 | 160.79 | 0.17 | -0.46 | 56.32 | 48.6 | 1.44 | 0.13 | -1 | -1 | 1 | -0.9458 | 0.48 | 5.5 | 67 |  | 0 | 0.0481 | 0.196 | -0.0273 | 0.99998 | [574 empty] | 0
P | 2018-06-18T10:58:05+08:00 | 2018-06-18T15:28:39+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI950G | CPSM Final Test | PASS | 129 | 162.23 | 0.17 | -0.69 | 56.46 | 48.85 | 1.21 | 0.13 | 1 | -2 | 0 | -0.8559 | 0.529 | -4.2 | 4 |  | 0 | 0.1224 | 0.2011 | 0.0202 | 0.99999 | [575 empty] | 0
P | 2018-06-18T15:33:07+08:00 | 2018-06-18T16:20:20+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI948G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-06-18T15:09:16+08:00 | 2018-06-18T16:22:05+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI952G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-06-18T05:53:04+08:00 | 2018-06-18T16:22:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI945G | CPSM Final Test | FAIL | [597 empty] | 0
P | 2018-06-18T13:16:56+08:00 | 2018-06-18T18:22:37+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI954G | CPSM Final Test | FAIL | [13 empty] | -7.2
P | 2018-06-18T13:00:32+08:00 | 2018-06-18T18:22:53+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI958G | CPSM Final Test | FAIL | [13 empty] | -6.5
P | 2018-06-18T13:14:59+08:00 | 2018-06-18T18:38:28+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI948G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-06-18T12:38:39+08:00 | 2018-06-18T18:39:17+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI959G | CPSM Final Test | FAIL
P | 2018-06-18T16:06:52+08:00 | 2018-06-18T19:24:00+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI960G | CPSM Final Test | PASS | 129 | 170.78 | 0.17 | -0.54 | 56.31 | 50.04 | 0.83 | 0.13 | 1 | 3 | 0 | -0.856 | 0.5346 | -3.4 | 0 |  | 0 | 0.0611 | 0.2028 | 0.0367 | 0.99995 | [577 empty] | 0
P | 2018-06-18T16:09:09+08:00 | 2018-06-18T19:23:24+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI961G | CPSM Final Test | PASS | 129 | 171.68 | -0.02 | -0.31 | 56.47 | 49.7 | 1.44 | 0.13 | 0 | 3 | 1 | -0.912 | 0.478 | -4.5 | -5 |  | 0 | 0.134 | 0.5596 | -0.0681 | 1 | [578 empty] | 0
P | 2018-06-18T15:45:40+08:00 | 2018-06-18T19:24:13+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI952G | CPSM Final Test | PASS | 129 | 165.47 | -0.02 | -0.38 | 56.5 | 49.28 | 1.06 | 0.13 | -1 | 1 | -1 | -0.8998 | 0.527 | -1.9 | -28 |  | 0 | 0.0687 | 0.8892 | -0.0352 | 0.99999 | [579 empty] | 0
P | 2018-06-18T15:41:42+08:00 | 2018-06-18T20:23:21+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI958G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -0.38 | 56.22 | 49.02 | 1.21 | 0.13 | 0 | -2 | 0 | -0.864 | 0.5498 | 5.4 | 38 |  | 0 | 0.1114 | 0.1839 | -0.0062 | 0.99999 | [580 empty] | 0
P | 2018-06-18T15:38:58+08:00 | 2018-06-18T20:24:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI959G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.61 | 56.47 | 48.01 | 1.36 | 0.13 | 1 | 2 | 1 | -0.8998 | 0.5326 | -1.3 | -12 |  | 0 | 0.0978 | 0.241 | -0.0265 | 1 | [581 empty] | 0
P | 2018-06-18T15:56:43+08:00 | 2018-06-18T21:23:14+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI948G | CPSM Final Test | PASS | 129 | 163.76 | 0.27 | -0.84 | 56.43 | 50.55 | 0.83 | 0.13 | -2 | -2 | 0 | -0.9079 | 0.546 | -0.1 | 39 |  | 0 | 0.1521 | 0.1346 | -0.0182 | 0.99998 | [582 empty] | 0
P | 2018-06-18T15:49:47+08:00 | 2018-06-19T08:23:09+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI954G | CPSM Final Test | PASS | 129 | 162.32 | 0.27 | -0.31 | 56.5 | 49.02 | 1.67 | 0.13 | 0 | 1 | 0 | -0.9059 | 0.5206 | 6.8 | 72 |  | 0 | 0.1562 | 0.1412 | -0.0542 | 0.99999 | [583 empty] | 0
P | 2018-06-19T09:43:38+08:00 | 2018-06-19T10:34:54+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI963G | CPSM Final Test | FAIL
P | 2018-06-19T08:00:14+08:00 | 2018-06-19T12:18:18+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI963G | CPSM Final Test | FAIL | [13 empty] | 18.6
P | 2018-06-19T09:59:32+08:00 | 2018-06-19T12:49:47+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI965G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -0.61 | 56.34 | 49.45 | 1.13 | 0.13 | 0 | -1 | 0 | -0.816 | 0.489 | -4.5 | -12 |  | 0 | 0.0476 | 0.0497 | -0.0624 | 0.99995 | [584 empty] | 0
P | 2018-06-19T10:13:54+08:00 | 2018-06-19T13:06:31+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI964G | CPSM Final Test | PASS | 129 | 159.98 | -0.02 | -0.77 | 56.11 | 49.36 | 1.51 | 0.13 | 0 | -1 | -1 | -0.8799 | 0.5176 | 3.1 | -20 |  | 0 | 0.0661 | 0.2928 | -0.046 | 0.99983 | [160 empty] | 0
P | 2018-06-19T06:12:25+08:00 | 2018-06-19T13:31:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI967G | CPSM Final Test | FAIL
P | 2018-06-19T10:09:16+08:00 | 2018-06-19T14:18:20+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI963G | CPSM Final Test | PASS | 129 | 167.54 | -0.02 | 0 | 56.42 | 54.27 | 1.74 | 0.13 | -1 | -2 | 0 | -0.8639 | 0.5346 | -4.4 | -7 |  | 0 | 0.1275 | 0.5175 | -0.0072 | 0.99997 | [582 empty] | 0
P | 2018-06-19T09:48:41+08:00 | 2018-06-19T14:35:58+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI966G | CPSM Final Test | FAIL
P | 2018-06-19T12:17:09+08:00 | 2018-06-19T15:33:18+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI970G | CPSM Final Test | PASS | 129 | 156.56 | -0.02 | -0.54 | 56.62 | 50.12 | 0.91 | 0.13 | -1 | 1 | 1 | -0.8042 | 0.535 | -1.2 | -6 |  | 0 | 0.1129 | 0.1735 | 0.0687 | 0.99996 | [585 empty] | 0
P | 2018-06-19T12:12:19+08:00 | 2018-06-19T16:21:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI966G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -0.69 | 56.4 | 51.39 | 1.67 | 0.13 | -1 | 2 | 0 | -0.936 | 0.4898 | -0.8 | 19 |  | 0 | 0.0383 | 0.3092 | -0.0196 | 0.99999 | [586 empty] | 0
P | 2018-06-19T13:40:47+08:00 | 2018-06-19T16:26:36+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI968G | CPSM Final Test | FAIL | [13 empty] | 19.9
P | 2018-06-19T02:49:52+08:00 | 2018-06-19T16:59:51+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI962G | CPSM Final Test | FAIL | [17 empty] | 56.0005
P | 2018-06-19T02:47:50+08:00 | 2018-06-19T17:01:53+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI962G | CPSM Final Test | FAIL
P | 2018-06-19T14:21:03+08:00 | 2018-06-19T17:04:14+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI975G | CPSM Final Test | FAIL | [13 empty] | -7.5
P | 2018-06-19T02:10:18+08:00 | 2018-06-19T17:39:25+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI962G | CPSM Final Test | FAIL
P | 2018-06-19T14:55:26+08:00 | 2018-06-19T17:48:17+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI973G | CPSM Final Test | PASS | 129 | 165.65 | 0.75 | -1.38 | 56.18 | 49.11 | 1.21 | 0.13 | -1 | 2 | 1 | -0.9039 | 0.5288 | -1.3 | -18 |  | 0 | 0.049 | 0.0607 | -0.0018 | 1 | [587 empty] | 0
P | 2018-06-19T15:01:35+08:00 | 2018-06-19T17:56:47+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI972G | CPSM Final Test | PASS | 129 | 164.75 | 0.17 | -0.84 | 56.23 | 48.6 | 0.91 | 0.13 | 1 | 1 | 0 | -0.9338 | 0.5022 | 1.6 | 19 |  | 0 | 0.0659 | 0.1543 | -0.0533 | 1 | [446 empty] | 0
P | 2018-06-19T01:28:09+08:00 | 2018-06-19T18:15:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI967G | CPSM Final Test | FAIL
P | 2018-06-19T15:32:46+08:00 | 2018-06-19T18:21:42+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI945G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-06-19T15:03:42+08:00 | 2018-06-19T18:32:02+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI968G | CPSM Final Test | PASS | 129 | 165.2 | -0.02 | -0.69 | 56.52 | 48.34 | 0.83 | 0.13 | 0 | 1 | 0 | -0.8979 | 0.534 | 3.4 | 16 |  | 0 | 0.1207 | 0.1023 | 0.0055 | 0.99999 | [374 empty] | 0
P | 2018-06-19T09:54:52+08:00 | 2018-06-19T18:42:43+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI962G | CPSM Final Test | PASS | 129 | 164.84 | 0.17 | -0.54 | 56.33 | 49.36 | 1.67 | 0.13 | 0 | 2 | 0 | -0.8381 | 0.491 | -2.3 | -2 |  | 0 | 0.1828 | 0.1343 | -0.0217 | 0.99999 | [588 empty] | 0
P | 2018-06-19T11:53:45+08:00 | 2018-06-19T19:31:32+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI975G | CPSM Final Test | FAIL | [610 empty] | 34577
P | 2018-06-19T12:45:35+08:00 | 2018-06-19T21:20:21+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI976G | CPSM Final Test | FAIL
P | 2018-06-18T22:19:17+08:00 | 2018-06-19T21:24:21+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI967G | CPSM Final Test | FAIL
P | 2018-06-19T09:57:01+08:00 | 2018-06-19T21:28:16+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI975G | CPSM Final Test | FAIL | [16 empty] | 0.32206
P | 2018-06-19T16:57:14+08:00 | 2018-06-19T22:22:16+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI945G | CPSM Final Test | PASS | 129 | 165.11 | 0.17 | -0.31 | 56.63 | 48.68 | 1.29 | 0.13 | 1 | 1 | 0 | -0.9258 | 0.4894 | -0.8 | 3 |  | 0 | 0.303 | 0.1885 | -0.0219 | 0.99999 | [590 empty] | 0
P | 2018-06-19T09:51:49+08:00 | 2018-06-19T22:27:31+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI967G | CPSM Final Test | PASS | 129 | 161.69 | 0.17 | -0.61 | 56.33 | 51.31 | 1.9 | 0.13 | 0 | -2 | 1 | -0.7922 | 0.5154 | -3.1 | 3 |  | 0 | 0.0845 | 0.0561 | 0.0018 | 0.99999 | [591 empty] | 0
P | 2018-06-19T19:52:13+08:00 | 2018-06-19T22:50:47+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI969G | CPSM Final Test | PASS | 129 | 161.78 | 0.37 | -0.46 | 56.48 | 48.77 | 1.13 | 0.13 | 0 | 1 | 0 | -0.9338 | 0.546 | 0.4 | -15 |  | 0 | 0.123 | 0.2382 | -0.0258 | 1 | [592 empty] | 0
P | 2018-06-19T19:59:48+08:00 | 2018-06-19T22:59:40+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI971G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.77 | 56.41 | 53.34 | 1.29 | 0.13 | 0 | -2 | 1 | -0.9617 | 0.5138 | 5 | -29 |  | 0 | 0.1074 | 0.2093 | 0.0381 | 0.99998 | [593 empty] | 0
P | 2018-06-19T17:02:58+08:00 | 2018-06-19T23:04:21+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI976G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.84 | 56.16 | 50.55 | 1.51 | 0.13 | 0 | 2 | 1 | -0.9258 | 0.5358 | 8 | -7 |  | 0 | 0.0454 | 0.1528 | 0.0215 | 0.99999 | [576 empty] | 0
P | 2018-06-20T09:39:21+08:00 | 2018-06-20T10:55:59+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI980G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-06-18T20:21:51+08:00 | 2018-06-20T11:03:26+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI975G | CPSM Final Test | FAIL
P | 2018-06-20T08:15:28+08:00 | 2018-06-20T11:21:15+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI974G | CPSM Final Test | FAIL | [13 empty] | -5.4
P | 2018-06-20T08:24:40+08:00 | 2018-06-20T11:27:08+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI979G | CPSM Final Test | FAIL | [13 empty] | -6.2
P | 2018-06-19T15:42:39+08:00 | 2018-06-20T12:06:28+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI975G | CPSM Final Test | PASS | 129 | 171.5 | -0.02 | -0.38 | 56.43 | 48.34 | 1.74 | 0.13 | 1 | 1 | -1 | -0.8821 | 0.489 | -4.3 | -17 |  | 0 | 0.2717 | 0.3129 | -0.0469 | 0.99996 | [589 empty] | 0
P | 2018-06-20T07:34:36+08:00 | 2018-06-20T12:17:12+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI979G | CPSM Final Test | FAIL | [13 empty] | -6.3
P | 2018-06-20T09:41:47+08:00 | 2018-06-20T12:46:17+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI981G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | -0.69 | 56.18 | 50.38 | 1.59 | 0.13 | 1 | -2 | -1 | -0.836 | 0.528 | -1.5 | 37 |  | 0 | 0.0761 | 0.0822 | -0.0027 | 0.99998 | [591 empty] | 0
P | 2018-06-20T09:56:43+08:00 | 2018-06-20T12:59:11+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI978G | CPSM Final Test | PASS | 129 | 165.56 | 0.56 | -0.69 | 56.62 | 50.63 | 1.29 | 0.13 | 0 | 2 | 0 | -0.8061 | 0.5134 | 1.6 | 23 |  | 0 | 0.065 | 0.2091 | -0.0443 | 0.99998 | [594 empty] | 0
P | 2018-06-20T09:48:21+08:00 | 2018-06-20T13:21:37+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI974G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -0.31 | 56.44 | 53.09 | 1.36 | 0.13 | 0 | -1 | 0 | -0.832 | 0.489 | -1 | -16 |  | 0 | 0.0758 | 0.1754 | 0.0015 | 1 | [143 empty] | 0
P | 2018-06-20T10:17:40+08:00 | 2018-06-20T13:36:41+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI980G | CPSM Final Test | PASS | 129 | 163.4 | 0.37 | -0.61 | 56.09 | 48.26 | 0.91 | 0.13 | 0 | -3 | 0 | -0.846 | 0.5418 | -0.3 | -17 |  | 0.00791 | 0.0819 | 0.0746 | -0.0053 | 1 | [595 empty] | 0
P | 2018-06-19T01:57:47+08:00 | 2018-06-20T13:39:29+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI977G | CPSM Final Test | FAIL | [16 empty] | 0.13604
P | 2018-06-20T09:55:54+08:00 | 2018-06-20T14:17:16+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI979G | CPSM Final Test | PASS | 129 | 171.14 | 0.08 | -0.54 | 56.42 | 51.22 | 1.9 | 0.13 | 0 | 3 | 0 | -0.8599 | 0.5112 | -0.3 | -11 |  | 0 | 0.1096 | 0.1164 | 0.0038 | 0.99998 | [596 empty] | 0
P | 2018-06-19T19:48:38+08:00 | 2018-06-20T14:42:37+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI977G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -0.77 | 56.23 | 51.48 | 1.44 | 0.13 | 0 | 1 | 0 | -0.8998 | 0.475 | -2.2 | 102 |  | 0 | 0.142 | 0.1229 | -0.0284 | 1 | [597 empty] | 0
P | 2018-06-20T12:19:17+08:00 | 2018-06-20T15:39:36+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI982G | CPSM Final Test | PASS | 129 | 165.2 | 0.27 | -0.77 | 56.52 | 49.61 | 1.36 | 0.13 | 1 | -2 | 0 | -0.8202 | 0.5292 | -4.8 | 30 |  | 0 | 0.0512 | 0.369 | -0.0301 | 0.99999 | [33 empty] | 0
P | 2018-06-20T12:17:28+08:00 | 2018-06-20T15:43:14+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI983G | CPSM Final Test | PASS | 129 | 166.1 | 0.27 | -0.92 | 56.32 | 50.29 | 2.12 | 0.13 | 1 | 2 | 0 | -0.8879 | 0.5134 | 5 | 8 |  | 0 | 0.0638 | 0.3211 | 0.009 | 0.99997 | [589 empty] | 0
P | 2018-06-20T14:02:10+08:00 | 2018-06-20T17:18:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM002G | CPSM Final Test | FAIL
P | 2018-06-20T14:45:34+08:00 | 2018-06-20T17:38:05+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM004G | CPSM Final Test | FAIL | [13 empty] | -8.6
P | 2018-06-20T14:02:59+08:00 | 2018-06-20T17:42:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM003G | CPSM Final Test | FAIL | [13 empty] | 17.8
P | 2018-06-20T13:17:17+08:00 | 2018-06-20T17:42:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM001G | CPSM Final Test | FAIL
P | 2018-06-20T13:22:27+08:00 | 2018-06-20T18:22:33+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM003G | CPSM Final Test | FAIL | [13 empty] | 25.3
P | 2018-06-20T15:22:09+08:00 | 2018-06-20T18:32:00+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM006G | CPSM Final Test | PASS |  | 163.22 | 0.68 | -0.61 | 56.46 | 48.18 | 1.74 | 0.13 | 0 | -4 | 0 | -0.9616 | 0.5402 | 0.1 | 61 |  |  | 0.0843 | 0.1765 | -0.0209 | 1 |  |  |  |  | 133 |  | 0 | [82 empty] | 0
P | 2018-06-20T15:40:12+08:00 | 2018-06-20T19:07:17+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM002G | CPSM Final Test | PASS |  | 170.15 | 0.27 | -0.84 | 56.41 | 52.07 | 1.51 | 0.13 | 0 | 2 | -1 | -0.9159 | 0.4842 | 3.5 | 39 |  |  | 0.3707 | 0.439 | -0.051 | 1 |  |  |  |  | 133 |  | 0 | [591 empty] | 0
P | 2018-06-20T17:57:31+08:00 | 2018-06-20T19:14:42+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM007G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-06-20T15:30:03+08:00 | 2018-06-20T19:30:55+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM001G | CPSM Final Test | PASS |  | 163.85 | 0.08 | -0.84 | 56.44 | 51.73 | 1.74 | 0.13 | 0 | -2 | 0 | -0.8978 | 0.5218 | 6.5 | 1 |  |  | 0.0765 | 0.0394 | -0.0301 | 0.99999 |  |  |  |  | 133 |  | 0 | [592 empty] | 0
P | 2018-06-20T16:43:27+08:00 | 2018-06-20T19:42:06+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM005G | CPSM Final Test | PASS |  | 165.47 | -0.02 | -0.54 | 56.7 | 49.61 | 1.82 | 0.13 | 0 | -2 | 0 | -0.9098 | 0.4944 | 1.1 | 57 |  |  | 0.0632 | 0.3292 | -0.0405 | 1 |  |  |  |  | 133 |  | 0 | [593 empty] | 0
P | 2018-06-20T11:46:25+08:00 | 2018-06-20T19:58:35+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM003G | CPSM Final Test | FAIL | [13 empty] | 24
P | 2018-06-20T18:20:27+08:00 | 2018-06-20T21:28:44+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM008G | CPSM Final Test | PASS |  | 167.45 | 0.37 | -0.61 | 56.54 | 48.6 | 1.51 | 0.13 | 0 | -1 | 0 | -0.9099 | 0.49 | -2.3 | 2 |  |  | 0.1048 | 0.3041 | -0.0253 | 1 |  |  |  |  | 133 |  | 0 | [217 empty] | 0
P | 2018-06-20T18:36:07+08:00 | 2018-06-20T21:59:32+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM007G | CPSM Final Test | PASS |  | 168.26 | -0.11 | -0.61 | 56.55 | 48.94 | 1.82 | 0.13 | 0 | -1 | 0 | -0.8022 | 0.5092 | 1.4 | -10 |  |  | 0.0607 | 0.3556 | 0.0187 | 0.99998 |  |  |  |  | 133 |  | 0 | [594 empty] | 0
P | 2018-06-19T22:35:46+08:00 | 2018-06-21T09:47:53+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM004G | CPSM Final Test | FAIL | [623 empty] | 12552
P | 2018-06-21T08:40:08+08:00 | 2018-06-21T09:46:50+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM009G | CPSM Final Test | FAIL | [5 empty] | 46.06
P | 2018-06-19T21:51:39+08:00 | 2018-06-21T09:53:21+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM003G | CPSM Final Test | FAIL
P | 2018-06-20T16:11:50+08:00 | 2018-06-21T10:53:25+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM004G | CPSM Final Test | PASS |  | 166.64 | -0.02 | -0.77 | 56.49 | 49.19 | 1.74 | 0.13 | -1 | -1 | 0 | -0.9098 | 0.5408 | -3.8 | 2 |  |  | 0.4574 | 0.4195 | 0.0032 | 0.99996 |  |  |  |  | 133 |  | 0 | [595 empty] | 0
P | 2018-06-21T07:58:16+08:00 | 2018-06-21T10:49:54+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM010G | CPSM Final Test | FAIL | [13 empty] | -6.4
P | 2018-06-21T07:43:19+08:00 | 2018-06-21T11:28:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM012G | CPSM Final Test | FAIL
P | 2018-06-20T15:52:30+08:00 | 2018-06-21T11:40:49+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM003G | CPSM Final Test | PASS |  | 160.07 | 0.37 | -0.77 | 56.61 | 48.51 | 1.67 | 0.13 | -1 | -1 | -1 | -0.9198 | 0.458 | 6.1 | 22 |  |  | 0.1476 | 0.2986 | -0.0059 | 1 |  |  |  |  | 133 |  | 0 | [596 empty] | 0
P | 2018-06-21T09:45:37+08:00 | 2018-06-21T12:27:27+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM011G | CPSM Final Test | FAIL
P | 2018-06-21T10:53:30+08:00 | 2018-06-21T12:31:08+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM014G | CPSM Final Test | FAIL | [5 empty] | 46.65
P | 2018-06-21T09:13:29+08:00 | 2018-06-21T12:34:18+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM009G | CPSM Final Test | PASS |  | 164.48 | 0.27 | 0.45 | 56.64 | 48.94 | 1.97 | 0.13 | -1 | -1 | 0 | -0.9058 | 0.535 | 5.8 | -6 |  |  | 0.0835 | 0.1281 | -0.025 | 0.99999 |  |  |  |  | 133 |  | 0 | [410 empty] | 0
P | 2018-06-21T09:35:59+08:00 | 2018-06-21T12:41:59+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM012G | CPSM Final Test | PASS |  | 163.67 | -0.02 | -0.54 | 56.23 | 50.38 | 1.59 | 0.13 | 0 | 1 | 1 | -0.9199 | 0.5454 | -2.7 | 5 |  |  | 0.1054 | 0.0717 | -0.01 | 1 |  |  |  |  | 133 |  | 0 | [597 empty] | 0
P | 2018-06-21T09:24:05+08:00 | 2018-06-21T12:55:18+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM010G | CPSM Final Test | PASS |  | 165.74 | -0.11 | -0.69 | 56.51 | 48.85 | 1.36 | 0.13 | 0 | -2 | 0 | -0.874 | 0.537 | 8.4 | 96 |  |  | 0.1177 | 0.37 | -0.0059 | 1 |  |  |  |  | 133 |  | 0 | [598 empty] | 0
P | 2018-06-21T07:59:35+08:00 | 2018-06-21T14:13:29+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM011G | CPSM Final Test | FAIL
P | 2018-06-21T08:29:06+08:00 | 2018-06-21T14:09:14+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM013G | CPSM Final Test | FAIL
P | 2018-06-21T09:31:37+08:00 | 2018-06-21T14:24:47+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM015G | CPSM Final Test | FAIL
P | 2018-06-21T11:42:19+08:00 | 2018-06-21T15:50:30+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM014G | CPSM Final Test | PASS |  | 170.87 | -0.3 | -0.46 | 56.96 | 48.68 | 2.05 | 0.13 | 0 | -1 | 0 | -0.818 | 0.4692 | 0.8 | 18 |  |  | 0.0918 | 0.1668 | -0.018 | 0.99993 |  |  |  |  | 133 |  | 0 | [599 empty] | 0
P | 2018-06-21T11:06:32+08:00 | 2018-06-21T16:01:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM011G | CPSM Final Test | PASS |  | 167.27 | 0.17 | -0.46 | 56.62 | 49.78 | 1.82 | 0.13 | 1 | -2 | -1 | -0.8441 | 0.524 | -2.8 | 14 |  |  | 0.0722 | 0.1917 | -0.0098 | 0.99997 |  |  |  |  | 133 |  | 0 | [595 empty] | 0
P | 2018-06-21T11:19:10+08:00 | 2018-06-21T15:56:54+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM013G | CPSM Final Test | PASS |  | 160.97 | 0.17 | -0.84 | 56.46 | 49.11 | 1.59 | 0.13 | 1 | 10 | 1 | -0.9278 | 0.5294 | -4.2 | -8 |  |  | 0.0651 | 0.4858 | -0.1085 | 1 |  |  |  |  | 133 |  | 0 | [301 empty] | 0
P | 2018-06-21T11:58:12+08:00 | 2018-06-21T16:12:17+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM015G | CPSM Final Test | PASS |  | 159.89 | -0.02 | -0.69 | 56.22 | 50.38 | 1.67 | 0.13 | 0 | -3 | -1 | -0.8899 | 0.532 | 7 | -3 |  |  | 0.1292 | 0.1885 | -0.012 | 1 |  |  |  |  | 133 |  | 0 | [600 empty] | 0
P | 2018-06-21T15:19:52+08:00 | 2018-06-21T16:30:57+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM019G | CPSM Final Test | FAIL | [5 empty] | 45.72
P | 2018-06-21T14:21:44+08:00 | 2018-06-21T16:31:28+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM016G | CPSM Final Test | FAIL | [13 empty] | -9.3
P | 2018-06-21T13:51:00+08:00 | 2018-06-21T16:41:32+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM017G | CPSM Final Test | FAIL | [13 empty] | -6.4
P | 2018-06-21T14:31:21+08:00 | 2018-06-21T16:48:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM018G | CPSM Final Test | FAIL | [13 empty] | -6.6
P | 2018-06-21T13:10:54+08:00 | 2018-06-21T17:42:18+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM016G | CPSM Final Test | FAIL
P | 2018-06-21T15:16:16+08:00 | 2018-06-21T18:50:20+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM017G | CPSM Final Test | PASS |  | 166.37 | 0.46 | -0.61 | 56.41 | 54.95 | 2.51 | 0.13 | 0 | -1 | 0 | -0.8939 | 0.5196 | 1.3 | 6 |  |  | 0.0442 | 0.2679 | 0.0283 | 0.99998 |  |  |  |  | 133 |  | 0 | [82 empty] | 0
P | 2018-06-21T15:40:05+08:00 | 2018-06-21T18:53:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM018G | CPSM Final Test | PASS |  | 163.22 | 0.75 | -0.84 | 56.62 | 49.61 | 1.36 | 0.13 | -2 | -2 | 0 | -0.9278 | 0.5382 | -1.4 | 15 |  |  | 0.0758 | 0.0911 | -0.0549 | 0.99998 |  |  |  |  | 133 |  | 0 | [601 empty] | 0
P | 2018-06-21T14:10:42+08:00 | 2018-06-21T19:49:52+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM022G | CPSM Final Test | FAIL | [13 empty] | -7.8
P | 2018-06-21T15:26:36+08:00 | 2018-06-21T19:59:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM016G | CPSM Final Test | PASS |  | 159.89 | 0.27 | -0.46 | 56.53 | 48.94 | 1.82 | 0.13 | 1 | 1 | 0 | -0.9199 | 0.521 | 4 | 42 |  |  | 0.1315 | 0.1047 | -0.0661 | 0.99999 |  |  |  |  | 133 |  | 0 | [602 empty] | 0
P | 2018-06-21T15:55:24+08:00 | 2018-06-21T20:21:39+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM019G | CPSM Final Test | PASS |  | 162.41 | -0.11 | -0.46 | 56.74 | 48.01 | 1.59 | 0.13 | -2 | 1 | 0 | -0.9319 | 0.5208 | -0.1 | 21 |  |  | 0.0414 | 0.1553 | -0.0372 | 0.99996 |  |  |  |  | 133 |  | 0 | [469 empty] | 0
P | 2018-06-21T16:37:26+08:00 | 2018-06-21T20:23:15+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM021G | CPSM Final Test | PASS |  | 164.93 | 0.27 | -0.31 | 56.49 | 48.6 | 1.44 | 0.13 | 1 | -3 | 0 | -0.9318 | 0.5194 | 1.7 | 29 |  |  | 0.0611 | 0.1803 | -0.0181 | 0.99999 |  |  |  |  | 133 |  | 0 | [603 empty] | 0
P | 2018-06-21T17:00:17+08:00 | 2018-06-21T21:53:32+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM022G | CPSM Final Test | PASS |  | 167.9 | 0.27 | -1.15 | 56.22 | 49.78 | 1.13 | 0.13 | -1 | -2 | 1 | -0.88 | 0.4886 | 1.3 | 8 |  |  | 0.065 | 0.2247 | -0.023 | 0.99996 |  |  |  |  | 133 |  | 0 | [604 empty] | 0
P | 2018-06-21T02:51:45+08:00 | 2018-06-22T08:27:17+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM023G | CPSM Final Test | FAIL
P | 2018-06-21T17:39:31+08:00 | 2018-06-22T08:27:55+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM023G | CPSM Final Test | PASS |  | 162.77 | -0.02 | -0.92 | 56.3 | 48.43 | 1.29 | 0.13 | 1 | -4 | 0 | -0.8321 | 0.5132 | 8.3 | 14 |  |  | 0.1367 | 0.376 | -0.0558 | 0.99999 |  |  |  |  | 133 |  | 0 | [605 empty] | 0
P | 2018-06-21T00:12:43+08:00 | 2018-06-22T08:32:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM020G | CPSM Final Test | FAIL
P | 2018-06-21T16:22:42+08:00 | 2018-06-22T08:33:17+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM020G | CPSM Final Test | PASS |  | 160.25 | -0.11 | 0 | 56.44 | 53.42 | 1.9 | 0.13 | 0 | -1 | 1 | -0.9977 | 0.5264 | -4.1 | 44 |  |  | 0.1275 | 0.1096 | -0.0238 | 0.99999 |  |  |  |  | 133 |  | 0 | [606 empty] | 0
P | 2018-06-22T09:56:26+08:00 | 2018-06-22T12:23:53+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM026G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-06-22T08:44:59+08:00 | 2018-06-22T13:51:02+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM027G | CPSM Final Test | FAIL | [5 empty] | 42.67
P | 2018-06-22T10:55:42+08:00 | 2018-06-22T13:53:45+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM025G | CPSM Final Test | PASS |  | 166.73 | 0.37 | -0.92 | 56.31 | 49.19 | 1.59 | 0.13 | 0 | -3 | 0 | -0.9119 | 0.5364 | -1.3 | -6 |  |  | 0.0993 | 0.268 | -0.035 | 0.99999 |  |  |  |  | 133 |  | 0 | [607 empty] | 0
P | 2018-06-22T07:55:09+08:00 | 2018-06-22T14:03:18+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM024G | CPSM Final Test | FAIL
P | 2018-06-22T07:51:00+08:00 | 2018-06-22T14:29:19+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM026G | CPSM Final Test | FAIL
P | 2018-06-22T11:10:10+08:00 | 2018-06-22T14:29:53+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM026G | CPSM Final Test | PASS |  | 163.49 | -0.11 | -0.99 | 56.27 | 48.68 | 1.9 | 0.13 | -1 | 1 | 0 | -0.9517 | 0.5132 | 1.1 | -1 |  |  | 0.2013 | 0.1815 | 0.0215 | 0.99998 |  |  |  |  | 133 |  | 0 | [608 empty] | 0
P | 2018-06-22T10:59:13+08:00 | 2018-06-22T15:08:54+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM024G | CPSM Final Test | PASS |  | 168.53 | 0.37 | -0.92 | 56.62 | 51.48 | 1.21 | 0.13 | -1 | -3 | 0 | -0.9058 | 0.523 | -2.6 | 18 |  |  | 0.0975 | 0.0987 | -0.0151 | 0.99998 |  |  |  |  | 133 |  | 0 | [609 empty] | 0
P | 2018-06-22T11:52:23+08:00 | 2018-06-22T15:32:11+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM029G | CPSM Final Test | PASS |  | 164.39 | 0.84 | -0.84 | 56.31 | 49.61 | 1.36 | 0.13 | -2 | -2 | 1 | -0.9437 | 0.5188 | -3 | 23 |  |  | 0.0742 | 0.0967 | -0.0344 | 0.99999 |  |  |  |  | 133 |  | 0 | [506 empty] | 0
P | 2018-06-22T08:26:54+08:00 | 2018-06-22T15:36:18+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM030G | CPSM Final Test | FAIL | [27 empty] | 5.4958
P | 2018-06-22T11:18:01+08:00 | 2018-06-22T16:34:54+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM027G | CPSM Final Test | PASS |  | 164.3 | 0.75 | -1.22 | 56.09 | 48.01 | 0.98 | 0.13 | 1 | -1 | -1 | -0.7921 | 0.541 | -1.9 | 16 |  |  | 0.0731 | 0.6945 | -0.0239 | 0.99999 |  |  |  |  | 133 |  | 0 | [610 empty] | 0
P | 2018-06-22T12:01:36+08:00 | 2018-06-22T16:39:30+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM030G | CPSM Final Test | PASS |  | 160.43 | 0.68 | -0.84 | 56.59 | 51.31 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9399 | 0.512 | -2.6 | 8 |  |  | 0.0919 | 0.2358 | -0.0552 | 0.99999 |  |  |  |  | 133 |  | 0 | [611 empty] | 0
P | 2018-06-22T14:11:57+08:00 | 2018-06-22T16:59:17+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM033G | CPSM Final Test | PASS |  | 169.79 | 0.46 | -0.99 | 56.17 | 49.02 | 1.29 | 0.13 | -1 | -2 | 0 | -0.8879 | 0.5282 | -3.4 | 18 |  |  | 0.1486 | 0.0751 | -0.0128 | 1 |  |  |  |  | 133 |  | 0 | [90 empty] | 0
P | 2018-06-22T11:03:02+08:00 | 2018-06-22T16:59:47+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM031G | CPSM Final Test | FAIL | [640 empty] | 0
P | 2018-06-22T14:49:05+08:00 | 2018-06-22T17:46:13+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM032G | CPSM Final Test | PASS |  | 164.57 | 0.37 | -1.15 | 56.31 | 48.68 | 1.44 | 0.13 | 1 | 4 | 1 | -0.9079 | 0.5048 | -1.2 | 3 |  |  | 0.0594 | 0.1964 | 0.0238 | 0.99999 |  |  |  |  | 133 |  | 0 | [613 empty] | 0
P | 2018-06-22T09:31:10+08:00 | 2018-06-22T17:47:01+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM028G | CPSM Final Test | FAIL | [642 empty] | 9890
P | 2018-06-22T13:39:06+08:00 | 2018-06-22T18:52:37+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | F18USM028G | CPSM Final Test | PASS |  | 157.28 | 0.17 | -0.38 | 56.42 | 51.31 | 0.98 | 0.13 | -1 | 4 | 0 | -0.9259 | 0.516 | -2.3 | 16 |  |  | 0.0611 | 0.2459 | -0.0083 | 0.99996 |  |  |  |  | 133 |  | 0 | [614 empty] | 0
P | 2018-06-22T08:53:54+08:00 | 2018-06-22T19:08:55+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM031G | CPSM Final Test | FAIL | [17 empty] | 2.7043
P | 2018-06-22T12:54:25+08:00 | 2018-06-22T20:09:18+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM035G | CPSM Final Test | FAIL | [637 empty] | 2430
P | 2018-06-22T14:01:24+08:00 | 2018-06-22T20:12:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | F18USM031G | CPSM Final Test | PASS |  | 162.5 | 0.17 | -0.54 | 56.24 | 50.38 | 1.59 | 0.13 | 1 | -3 | 0 | -0.9617 | 0.5168 | -3.8 | 17 |  |  | 0.0494 | 0.2346 | -0.0451 | 0.99997 |  |  |  |  | 133 |  | 0 | [612 empty] | 0
P | 2018-06-22T16:48:19+08:00 | 2018-06-22T20:30:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | F18USM036G | CPSM Final Test | PASS |  | 171.59 | 0.17 | -0.16 | 57.02 | 50.63 | 1.82 | 0.13 | 0 | 1 | -1 | -0.8998 | 0.509 | -2.3 | 23 |  |  | 0.1392 | 0.1527 | -0.0293 | 0.99999 |  |  |  |  | 133 |  | 0 | [346 empty] | 0
P | 2018-06-22T17:04:24+08:00 | 2018-06-22T20:31:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | F18USM034G | CPSM Final Test | PASS |  | 162.77 | 0.17 | -0.77 | 56.48 | 48.51 | 1.29 | 0.13 | 0 | -2 | 0 | -0.8919 | 0.502 | 3.5 | 58 |  |  | 0.1397 | 0.2318 | -0.0214 | 1 |  |  |  |  | 133 |  | 0 | [615 empty] | 0
P | 2018-06-22T18:47:08+08:00 | 2018-06-22T21:33:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | F18USM039G | CPSM Final Test | PASS |  | 165.02 | -0.11 | -0.99 | 56.49 | 50.8 | 1.44 | 0.13 | 0 | 1 | 1 | -0.9697 | 0.4672 | 1.1 | 36 |  |  | 0.0743 | 0.2503 | -0.0092 | 0.99998 |  |  |  |  | 133 |  | 0 | [616 empty] | 0
P | 2018-06-22T00:28:15+08:00 | 2018-06-23T08:35:28+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM035G | CPSM Final Test | FAIL
P | 2018-06-22T16:31:51+08:00 | 2018-06-23T08:36:02+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM035G | CPSM Final Test | PASS |  | 166.73 | 0.27 | -0.38 | 56.51 | 50.21 | 1.51 | 0.13 | -1 | 3 | -1 | -0.9138 | 0.5324 | 4.4 | 2 |  |  | 0.135 | 0.3072 | -0.0656 | 1 |  |  |  |  | 133 |  | 0 | [609 empty] | 0
P | 2018-06-22T01:41:49+08:00 | 2018-06-23T09:01:11+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM037G | CPSM Final Test | FAIL | [645 empty] | 0
P | 2018-06-22T02:20:25+08:00 | 2018-06-23T09:07:20+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM038G | CPSM Final Test | FAIL | [221 empty] | 0
P | 2018-06-22T17:21:30+08:00 | 2018-06-23T10:06:55+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | F18USM037G | CPSM Final Test | PASS |  | 163.22 | -0.3 | -0.61 | 56.51 | 50.88 | 1.9 | 0.13 | 1 | 1 | 0 | -0.906 | 0.498 | -0.7 | -1 |  |  | 0.0261 | 0.1143 | -0.0475 | 1 |  |  |  |  | 133 |  | 0 | [617 empty] | 0
P | 2018-06-22T17:43:53+08:00 | 2018-06-23T10:12:46+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | F18USM038G | CPSM Final Test | PASS |  | 162.95 | -0.02 | -0.77 | 56.5 | 49.45 | 1.59 | 0.13 | -1 | -2 | 0 | -0.9019 | 0.532 | -0.1 | 8 |  |  | 0.1423 | 0.272 | -0.0182 | 0.99995 |  |  |  |  | 133 |  | 0 | [193 empty] | 0
P | 2018-06-23T09:03:22+08:00 | 2018-06-23T12:49:20+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI984G | CPSM Final Test | FAIL | [14 empty] | 2588
P | 2018-06-23T10:38:28+08:00 | 2018-06-23T13:28:36+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | F18USM040G | CPSM Final Test | PASS |  | 165.47 | 0.27 | -1.07 | 56.2 | 50.46 | 1.36 | 0.13 | 0 | -2 | 0 | -0.9299 | 0.462 | 4.9 | 21 |  |  | 0.1015 | 0.1428 | -0.0118 | 0.99999 |  |  |  |  | 133 |  | 0 | [609 empty] | 0
P | 2018-06-23T11:13:52+08:00 | 2018-06-23T13:57:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI985G | CPSM Final Test | PASS | 129 | 159.89 | -0.02 | -0.84 | 56.51 | 49.28 | 1.06 | 0.13 | 0 | -2 | -1 | -0.8681 | 0.5454 | 2.5 | -11 |  | 0 | 0.0896 | 0.1058 | -0.0243 | 1 | [625 empty] | 0
P | 2018-06-23T07:00:24+08:00 | 2018-06-23T14:52:18+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI984G | CPSM Final Test | FAIL
P | 2018-06-23T10:56:21+08:00 | 2018-06-23T14:52:42+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI984G | CPSM Final Test | PASS | 129 | 168.62 | -0.21 | -0.61 | 56.32 | 49.87 | 1.21 | 0.13 | 0 | 2 | 0 | -0.8241 | 0.5192 | -2.8 | 0 |  | 0 | 0.1531 | 0.2797 | -0.0377 | 0.99999 | [488 empty] | 0
P | 2018-06-23T11:33:26+08:00 | 2018-06-23T14:56:51+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI986G | CPSM Final Test | PASS | 129 | 169.34 | -0.21 | -0.61 | 56.62 | 49.28 | 1.67 | 0.13 | -1 | 1 | 0 | -0.854 | 0.507 | -4.9 | -17 |  | 0 | 0.0436 | 0.2707 | -0.0355 | 1 | [624 empty] | 0
P | 2018-06-23T14:10:36+08:00 | 2018-06-23T15:06:13+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI988G | CPSM Final Test | FAIL | [5 empty] | 46.31
P | 2018-06-23T11:49:55+08:00 | 2018-06-23T15:22:55+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI987G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -0.61 | 56.32 | 48.09 | 1.13 | 0.13 | -1 | -1 | 0 | -0.814 | 0.5338 | -0.6 | 8 |  | 0 | 0.061 | 0.0773 | -0.0426 | 1 | [619 empty] | 0
P | 2018-06-23T15:02:14+08:00 | 2018-06-23T15:45:47+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI990G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-06-23T14:54:52+08:00 | 2018-06-23T17:46:58+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI989G | CPSM Final Test | PASS | 129 | 164.48 | 0.17 | -1.22 | 56.17 | 51.9 | 0.98 | 0.13 | 0 | -2 | 0 | -0.9339 | 0.513 | -4.5 | 10 |  | 0 | 0.0989 | 0.3548 | -0.0142 | 0.99999 | [200 empty] | 0
P | 2018-06-23T14:38:24+08:00 | 2018-06-23T17:57:37+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI988G | CPSM Final Test | PASS | 129 | 166.28 | 0.17 | -0.92 | 56.35 | 48.68 | 1.21 | 0.13 | -1 | -2 | 1 | -0.9417 | 0.5432 | 3.6 | 2 |  | 0 | 0.0822 | 0.1646 | 0.0177 | 0.99999 | [626 empty] | 0
P | 2018-06-23T15:24:00+08:00 | 2018-06-23T18:19:01+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI990G | CPSM Final Test | PASS | 129 | 170.42 | 0.08 | -0.77 | 56.4 | 48.6 | 1.06 | 0.13 | 0 | -2 | 0 | -0.8461 | 0.4966 | 8 | 19 |  | 0 | 0.1604 | 0.093 | 0.0526 | 0.99998 | [621 empty] | 0
P | 2018-06-23T17:34:30+08:00 | 2018-06-23T18:18:57+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI992G | CPSM Final Test | FAIL | [5 empty] | 45.47
P | 2018-06-23T16:25:28+08:00 | 2018-06-23T19:27:59+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI992G | CPSM Final Test | FAIL | [13 empty] | -8.6
P | 2018-06-23T16:13:28+08:00 | 2018-06-23T20:05:39+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI993G | CPSM Final Test | FAIL
P | 2018-06-23T17:02:50+08:00 | 2018-06-23T20:28:02+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI994G | CPSM Final Test | PASS | 129 | 168.26 | 0.37 | -0.61 | 56.32 | 49.53 | 1.67 | 0.13 | 0 | -1 | 0 | -0.9397 | 0.525 | 0.8 | 6 |  | 0 | 0.0505 | 0.0739 | -0.0386 | 1 | [627 empty] | 0
P | 2018-06-23T17:56:44+08:00 | 2018-06-23T21:28:23+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI992G | CPSM Final Test | PASS | 129 | 165.92 | 0.27 | -1.22 | 56.43 | 48.34 | 1.36 | 0.13 | -1 | -1 | 0 | -0.88 | 0.535 | 2.6 | 3 |  | 0 | 0.0442 | 0.0949 | -0.0283 | 1 | [628 empty] | 0
P | 2018-06-23T18:09:33+08:00 | 2018-06-23T21:50:27+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI993G | CPSM Final Test | PASS | 129 | 166.82 | 0.17 | -0.84 | 56.32 | 50.55 | 1.59 | 0.13 | -1 | 1 | 0 | -0.936 | 0.5338 | 11.2 | 8 |  | 0 | 0.1022 | 0.1098 | -0.0332 | 0.99999 | [629 empty] | 0
P | 2018-06-24T09:51:58+08:00 | 2018-06-24T12:03:38+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI998G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-06-24T09:26:02+08:00 | 2018-06-24T12:29:34+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI998G | CPSM Final Test | FAIL | [13 empty] | -5.2
P | 2018-06-24T10:04:50+08:00 | 2018-06-24T13:30:21+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI999G | CPSM Final Test | FAIL | [14 empty] | 3760
P | 2018-06-24T11:11:47+08:00 | 2018-06-24T13:59:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI997G | CPSM Final Test | PASS | 129 | 162.59 | 0.17 | -0.84 | 56.43 | 54.44 | 1.67 | 0.13 | -1 | 1 | 0 | -0.7942 | 0.5266 | 2.2 | 11 |  | 0 | 0.2113 | 0.1889 | 0.0703 | 0.99996 | [424 empty] | 0
P | 2018-06-24T07:28:53+08:00 | 2018-06-24T14:05:39+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI001G | CPSM Final Test | FAIL | [13 empty] | 30.9
P | 2018-06-24T10:57:48+08:00 | 2018-06-24T14:30:46+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI998G | CPSM Final Test | PASS | 129 | 158.99 | -0.02 | -1.07 | 56.42 | 50.38 | 0.98 | 0.13 | 0 | -1 | 1 | -0.8858 | 0.4956 | -2.2 | -12 |  | 0 | 0.065 | 0.3385 | -0.0211 | 1 | [630 empty] | 0
P | 2018-06-24T06:54:57+08:00 | 2018-06-24T14:39:35+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI001G | CPSM Final Test | FAIL | [13 empty] | 21.6
P | 2018-06-24T11:47:36+08:00 | 2018-06-24T15:14:33+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI999G | CPSM Final Test | PASS | 129 | 163.94 | 0.17 | -0.92 | 56.33 | 48.68 | 1.06 | 0.13 | -1 | -1 | 0 | -0.84 | 0.5124 | -1.4 | 10 |  | 0 | 0.1107 | 0.1227 | -0.0211 | 0.99999 | [631 empty] | 0
P | 2018-06-24T06:11:47+08:00 | 2018-06-24T15:22:45+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI001G | CPSM Final Test | FAIL | [13 empty] | 35.3
P | 2018-06-24T12:15:13+08:00 | 2018-06-24T15:37:37+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI996G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | -0.77 | 56.41 | 51.22 | 0.98 | 0.13 | 1 | 2 | 0 | -0.8121 | 0.522 | -4.6 | 8 |  | 0 | 0.0657 | 0.1228 | -0.0001 | 1 | [632 empty] | 0
P | 2018-06-24T13:27:11+08:00 | 2018-06-24T16:13:48+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI003G | CPSM Final Test | PASS | 129 | 162.23 | 0.27 | -1.15 | 55.87 | 49.36 | 1.44 | 0.13 | 0 | -3 | 0 | -0.848 | 0.5418 | 5.8 | 17 |  | 0 | 0.0599 | 0.0955 | -0.0183 | 0.99998 | [633 empty] | 0
P | 2018-06-24T16:11:36+08:00 | 2018-06-24T16:58:33+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI008G | CPSM Final Test | FAIL | [5 empty] | 46.31
P | 2018-06-24T10:47:16+08:00 | 2018-06-24T17:23:11+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI001G | CPSM Final Test | PASS | 129 | 166.01 | 0.27 | -0.69 | 56.44 | 54.78 | 1.67 | 0.13 | 0 | -2 | 0 | -0.8959 | 0.5028 | -4.2 | 11 |  | 0 | 0.1595 | 0.0685 | 0.0189 | 0.99998 | [634 empty] | 0
P | 2018-06-24T16:22:11+08:00 | 2018-06-24T19:21:38+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI007G | CPSM Final Test | PASS | 129 | 165.65 | 0.78 | -0.99 | 56.21 | 50.63 | 1.06 | 0.13 | 1 | -2 | 1 | -0.894 | 0.5062 | 3.4 | 2 |  | 0 | 0.0609 | 0.1021 | -0.011 | 0.99999 | [635 empty] | 0
P | 2018-06-24T00:31:19+08:00 | 2018-06-25T09:12:09+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI002G | CPSM Final Test | FAIL | [14 empty] | 937
P | 2018-06-23T23:46:24+08:00 | 2018-06-25T09:23:45+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI008G | CPSM Final Test | FAIL | [14 empty] | 1399
P | 2018-06-22T03:01:01+08:00 | 2018-06-25T10:34:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI995G | CPSM Final Test | FAIL
P | 2018-06-21T22:44:11+08:00 | 2018-06-25T10:53:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI991G | CPSM Final Test | FAIL | [19 empty] | -0.8185
P | 2018-06-24T16:51:44+08:00 | 2018-06-25T10:56:41+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI002G | CPSM Final Test | PASS | 129 | 160.97 | -0.11 | -0.77 | 56.59 | 51.82 | 1.67 | 0.13 | 1 | 3 | -1 | -0.876 | 0.5232 | 0.7 | 11 |  | 0 | 0.1235 | 0.1809 | 0.0012 | 1 | [636 empty] | 0
P | 2018-06-23T22:54:55+08:00 | 2018-06-25T11:54:01+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI005G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-06-23T16:48:55+08:00 | 2018-06-25T11:57:16+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI991G | CPSM Final Test | PASS | 129 | 166.19 | 0.08 | -0.99 | 56.21 | 48.51 | 1.51 | 0.13 | -1 | 1 | -1 | -0.9158 | 0.534 | -2.6 | -6 |  | 0 | 0.229 | 0.1689 | -0.0348 | 1 | [637 empty] | 0
P | 2018-06-23T21:09:22+08:00 | 2018-06-25T12:00:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI008G | CPSM Final Test | FAIL | [659 empty] | 0
P | 2018-06-25T12:06:50+08:00 | 2018-06-25T12:15:13+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI005G | CPSM Final Test | FAIL | [87 empty] | FALSE
P | 2018-06-23T18:47:31+08:00 | 2018-06-25T12:18:53+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI995G | CPSM Final Test | PASS | 129 | 168.44 | -0.02 | -0.92 | 56.32 | 50.04 | 1.36 | 0.13 | 0 | -2 | 0 | -0.9357 | 0.487 | 1.1 | -6 |  | 0 | 0.1474 | 0.2214 | -0.0369 | 0.99997 | [639 empty] | 0
P | 2018-06-25T06:03:36+08:00 | 2018-06-25T12:27:30+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI009G | CPSM Final Test | FAIL
P | 2018-06-25T07:51:28+08:00 | 2018-06-25T12:38:06+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI010G | CPSM Final Test | FAIL
P | 2018-06-25T09:15:33+08:00 | 2018-06-25T12:39:58+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI009G | CPSM Final Test | PASS | 129 | 161.42 | -0.11 | -0.84 | 56.22 | 49.61 | 1.44 | 0.13 | 1 | 1 | 0 | -0.8999 | 0.5372 | 5.1 | 11 |  | 0 | 0.0764 | 0.0679 | -0.0046 | 1 | [640 empty] | 0
P | 2018-06-24T16:35:05+08:00 | 2018-06-25T13:06:25+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI008G | CPSM Final Test | PASS | 129 | 161.51 | -0.02 | -0.46 | 56.24 | 50.63 | 1.21 | 0.13 | 1 | 1 | 0 | -0.8979 | 0.5376 | 0.7 | -4 |  | 0 | 0.1026 | 0.1556 | -0.0561 | 0.99999 | [638 empty] | 0
P | 2018-06-25T10:14:47+08:00 | 2018-06-25T14:23:04+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI010G | CPSM Final Test | PASS | 129 | 163.85 | 0.17 | -1.3 | 56.3 | 51.73 | 1.36 | 0.13 | -1 | 1 | 1 | -0.9537 | 0.5192 | 0.1 | 17 |  | 0 | 0.0584 | 0.1825 | 0.0044 | 0.99998 | [544 empty] | 0
P | 2018-06-25T11:47:36+08:00 | 2018-06-25T14:29:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI014G | CPSM Final Test | PASS | 129 | 166.64 | 0.08 | -0.92 | 56.17 | 51.73 | 1.29 | 0.13 | 1 | 2 | 0 | -0.8321 | 0.502 | 0.7 | -2 |  | 0 | 0.0335 | 0.1014 | -0.038 | 1 | [641 empty] | 0
P | 2018-06-25T11:54:18+08:00 | 2018-06-25T14:57:46+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI006G | CPSM Final Test | FAIL | [13 empty] | -9.3
P | 2018-06-25T03:36:37+08:00 | 2018-06-25T15:36:31+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI011G | CPSM Final Test | FAIL | [16 empty] | 3.1617
P | 2018-06-25T12:11:01+08:00 | 2018-06-25T15:39:59+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI005G | CPSM Final Test | PASS | 129 | 165.83 | 0.27 | -0.92 | 56.4 | 50.88 | 0.91 | 0.13 | -2 | 2 | 0 | -0.8302 | 0.504 | -3 | -1 |  | 0 | 0.0525 | 0.305 | -0.0087 | 0.99999 | [642 empty] | 0
P | 2018-06-25T14:58:36+08:00 | 2018-06-25T16:11:28+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI018G | CPSM Final Test | FAIL | [5 empty] | 46.23
P | 2018-06-25T15:23:39+08:00 | 2018-06-25T16:36:27+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI015G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-06-25T09:36:34+08:00 | 2018-06-25T16:39:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI011G | CPSM Final Test | PASS | 129 | 164.93 | -0.11 | -1.15 | 56.2 | 49.11 | 1.21 | 0.13 | -1 | -2 | 0 | -0.7901 | 0.531 | 0 | 22 |  | 0 | 0.2205 | 0.4148 | 0.0464 | 0.99999 | [152 empty] | 0
P | 2018-06-25T13:26:02+08:00 | 2018-06-25T17:11:24+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI006G | CPSM Final Test | PASS | 129 | 167.63 | 0.84 | -1.07 | 56.23 | 48.77 | 1.29 | 0.13 | 0 | -1 | 0 | -0.9597 | 0.4912 | 6.1 | 0 |  | 0 | 0.038 | 0.0816 | 0.002 | 0.99999 | [643 empty] | 0
P | 2018-06-25T14:14:06+08:00 | 2018-06-25T17:13:02+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI013G | CPSM Final Test | PASS | 129 | 163.49 | 0.37 | -1.3 | 56.52 | 49.36 | 1.36 | 0.13 | -2 | 2 | 0 | -0.8999 | 0.524 | 13.6 | 24 |  | 0 | 0.2568 | 0.2113 | 0.023 | 0.99999 | [644 empty] | 0
P | 2018-06-25T14:19:23+08:00 | 2018-06-25T17:16:36+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI012G | CPSM Final Test | PASS | 129 | 166.19 | 0.56 | -1.3 | 56.12 | 51.99 | 1.13 | 0.13 | -1 | 3 | 0 | -0.9816 | 0.464 | 2.3 | 0 |  | 0 | 0.1269 | 0.2583 | -0.04 | 0.99999 | [638 empty] | 0
P | 2018-06-25T14:30:11+08:00 | 2018-06-25T17:42:34+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI004G | CPSM Final Test | PASS | 129 | 163.85 | 0.97 | -0.84 | 56.33 | 48.77 | 1.21 | 0.13 | 0 | -2 | 0 | -0.9857 | 0.5154 | 1.1 | 13 |  | 0 | 0.0469 | 0.2846 | -0.0142 | 1 | [208 empty] | 0
P | 2018-06-25T15:35:02+08:00 | 2018-06-25T18:44:44+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI018G | CPSM Final Test | PASS | 129 | 166.01 | 0.3 | -0.69 | 56.05 | 49.61 | 1.21 | 0.13 | -1 | 1 | 0 | -0.8859 | 0.5162 | -2.2 | 4 |  | 0 | 0.0667 | 0.1681 | 0.0025 | 0.99999 | [645 empty] | 0
P | 2018-06-25T14:09:22+08:00 | 2018-06-25T18:58:40+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI017G | CPSM Final Test | FAIL | [14 empty] | 772
P | 2018-06-25T16:43:45+08:00 | 2018-06-25T19:29:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI021G | CPSM Final Test | FAIL | [13 empty] | -7.7
P | 2018-06-25T16:19:52+08:00 | 2018-06-25T19:43:26+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI022G | CPSM Final Test | FAIL
P | 2018-06-25T16:18:28+08:00 | 2018-06-25T19:44:50+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI022G | CPSM Final Test | FAIL
P | 2018-06-25T16:00:03+08:00 | 2018-06-25T20:17:59+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI015G | CPSM Final Test | PASS | 129 | 165.47 | 0.46 | -1.07 | 56.01 | 49.45 | 1.29 | 0.13 | 1 | 1 | 0 | -0.8939 | 0.5426 | -0.3 | -13 |  | 0 | 0.1992 | 0.3017 | -0.0252 | 0.99999 | [159 empty] | 0
P | 2018-06-25T16:34:01+08:00 | 2018-06-25T20:42:40+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI017G | CPSM Final Test | PASS | 129 | 162.86 | 0.27 | -0.77 | 56.29 | 51.48 | 1.06 | 0.13 | 1 | 1 | 0 | -0.9278 | 0.466 | 2.8 | 15 |  | 0 | 0.0875 | 0.3061 | -0.0067 | 1 | [646 empty] | 0
P | 2018-06-25T18:06:41+08:00 | 2018-06-25T21:29:26+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI021G | CPSM Final Test | PASS | 129 | 165.38 | 0.17 | -0.99 | 56.11 | 50.72 | 1.44 | 0.13 | 0 | 1 | 0 | -0.9358 | 0.513 | 1.2 | 14 |  | 0 | 0.1029 | 0.2078 | -0.0348 | 0.99998 | [647 empty] | 0
P | 2018-06-25T18:48:56+08:00 | 2018-06-25T21:37:12+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI019G | CPSM Final Test | PASS | 129 | 158.36 | 0.08 | -0.84 | 56.62 | 52.24 | 1.13 | 0.13 | 1 | 2 | 0 | -0.9537 | 0.521 | 0.5 | 10 |  | 0 | 0.1327 | 0.1151 | -0.0337 | 0.99999 | [648 empty] | 0
P | 2018-06-25T18:48:03+08:00 | 2018-06-25T21:46:22+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI016G | CPSM Final Test | PASS | 129 | 160.43 | 0.08 | -0.46 | 56.58 | 50.29 | 1.13 | 0.13 | -1 | -2 | 0 | -0.9497 | 0.5044 | 1.3 | -5 |  | 0 | 0.0731 | 0.084 | 0.0001 | 1 | [649 empty] | 0
P | 2018-06-25T18:54:56+08:00 | 2018-06-25T21:48:22+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI020G | CPSM Final Test | PASS | 129 | 166.1 | 0.08 | -1.38 | 55.78 | 50.12 | 1.13 | 0.13 | 0 | -1 | 0 | -0.9558 | 0.511 | 1.5 | 13 |  | 0 | 0.2434 | 0.2249 | -0.0081 | 0.99999 | [650 empty] | 0
P | 2018-06-25T00:54:08+08:00 | 2018-06-26T11:09:10+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI022G | CPSM Final Test | FAIL
P | 2018-06-24T22:31:32+08:00 | 2018-06-26T13:31:46+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI022G | CPSM Final Test | FAIL
P | 2018-06-26T10:29:58+08:00 | 2018-06-26T14:25:16+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI024G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-06-24T21:27:48+08:00 | 2018-06-26T14:35:30+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI022G | CPSM Final Test | FAIL
P | 2018-06-25T18:01:39+08:00 | 2018-06-26T14:36:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | F18SDI022G | CPSM Final Test | PASS | 129 | 166.1 | 0.08 | -0.84 | 56.19 | 49.02 | 1.21 | 0.13 | 1 | -3 | 1 | -0.8501 | 0.5366 | -1.9 | -2 |  | 0.0052 | 0.0361 | 0.2667 | 0.0141 | 0.99999 | [651 empty] | 0
P | 2018-06-26T13:13:12+08:00 | 2018-06-26T15:47:19+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI028G | CPSM Final Test | FAIL | [13 empty] | -5.3
P | 2018-06-26T12:28:28+08:00 | 2018-06-26T16:12:04+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI023G | CPSM Final Test | PASS | 129 | 164.03 | 0.68 | -1.3 | 55.9 | 49.11 | 0.6 | 0.13 | 1 | 3 | -1 | -0.9119 | 0.4828 | 3.9 | -8 |  | 0 | 0.085 | 0.4065 | -0.0136 | 1 | [652 empty] | 0
P | 2018-06-26T12:34:15+08:00 | 2018-06-26T16:14:18+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI025G | CPSM Final Test | PASS | 129 | 172.4 | 0.08 | -0.69 | 56.32 | 49.95 | 1.36 | 0.13 | -1 | -4 | 0 | -0.888 | 0.5078 | 5.4 | 9 |  | 0 | 0.0867 | 0.1225 | -0.057 | 0.99998 | [653 empty] | 0
P | 2018-06-26T12:31:36+08:00 | 2018-06-26T16:18:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI026G | CPSM Final Test | PASS | 129 | 167.9 | -0.3 | -0.61 | 56.49 | 50.04 | 1.29 | 0.13 | -2 | -1 | -1 | -0.9059 | 0.4748 | 6.9 | -6 |  | 0 | 0.0777 | 0.1386 | -0.0253 | 0.99999 | [654 empty] | 0
P | 2018-06-26T12:35:51+08:00 | 2018-06-26T16:25:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI027G | CPSM Final Test | PASS | 129 | 165.74 | 0.17 | -0.84 | 56.55 | 50.04 | 1.36 | 0.13 | 1 | 3 | 1 | -0.9217 | 0.5432 | 3.6 | 5 |  | 0 | 0.1085 | 0.1951 | -0.0195 | 1 | [41 empty] | 0
P | 2018-06-26T12:27:37+08:00 | 2018-06-26T16:29:24+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI024G | CPSM Final Test | PASS | 129 | 165.38 | 0.27 | -0.92 | 56.21 | 49.02 | 0.6 | 0.13 | -1 | -1 | 0 | -0.854 | 0.5438 | 1.5 | 3 |  | 0 | 0.1169 | 0.0382 | -0.0031 | 0.99999 | [655 empty] | 0
P | 2018-06-26T14:33:53+08:00 | 2018-06-26T17:35:12+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI030G | CPSM Final Test | PASS | 129 | 167.9 | 0.27 | -1.15 | 56.03 | 50.04 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9478 | 0.504 | -3.3 | -1 |  | 0 | 0.0786 | 0.2834 | 0.0487 | 0.99997 | [656 empty] | 0
P | 2018-06-26T14:30:15+08:00 | 2018-06-26T17:47:01+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI028G | CPSM Final Test | PASS | 129 | 161.6 | 0.56 | -1.07 | 56.2 | 50.04 | 1.13 | 0.13 | 0 | -1 | 1 | -0.9539 | 0.5208 | -0.9 | 3 |  | 0 | 0.0876 | 0.2468 | 0.0441 | 1 | [657 empty] | 0
P | 2018-06-26T18:34:51+08:00 | 2018-06-26T22:01:52+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI029G | CPSM Final Test | PASS | 129 | 174.74 | -0.02 | 0.23 | 56.63 | 49.36 | 1.36 | 0.13 | 1 | 3 | 0 | -0.9776 | 0.5196 | 3.6 | 5 |  | 0 | 0.0712 | 0.0531 | 0.0282 | 0.99999 | [655 empty] | 0
P | 2018-06-26T18:45:20+08:00 | 2018-06-26T22:03:06+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI032G | CPSM Final Test | PASS | 129 | 168.35 | 0.3 | -0.99 | 56.47 | 51.65 | 1.29 | 0.13 | -1 | -3 | 0 | -0.9238 | 0.5 | -0.8 | 41 |  | 0 | 0.0853 | 0.3091 | -0.0458 | 1 | [658 empty] | 0
P | 2018-06-26T04:51:32+08:00 | 2018-06-27T08:31:06+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI031G | CPSM Final Test | FAIL
P | 2018-06-26T18:41:19+08:00 | 2018-06-27T08:31:40+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI031G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -0.54 | 56.53 | 51.65 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9199 | 0.5308 | 4.1 | -6 |  | 0 | 0.076 | 0.0968 | -0.0288 | 1 | [659 empty] | 0
P | 2018-06-27T09:59:52+08:00 | 2018-06-27T12:39:34+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI035G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-06-27T11:17:56+08:00 | 2018-06-27T14:21:16+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | F18SDI037G | CPSM Final Test | PASS | 129 | 166.19 | 0.37 | -1.38 | 56.23 | 51.14 | 1.21 | 0.13 | -2 | 1 | 1 | -0.8839 | 0.532 | -2.3 | 34 |  | 0 | 0.0913 | 0.2186 | 0.0244 | 0.99999 | [660 empty] | 0
P | 2018-06-27T11:21:13+08:00 | 2018-06-27T14:27:05+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | F18SDI039G | CPSM Final Test | PASS | 129 | 168.08 | 0.49 | -1.22 | 56.07 | 49.95 | 1.29 | 0.13 | 0 | -4 | 0 | -0.9517 | 0.5098 | -4.7 | 10 |  | 0 | 0.1248 | 0.184 | 0.003 | 1 | [661 empty] | 0
P | 2018-06-27T11:22:59+08:00 | 2018-06-27T14:29:08+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | F18SDI036G | CPSM Final Test | PASS | 129 | 162.14 | 0.75 | -0.77 | 56.5 | 48.6 | 1.67 | 0.13 | -2 | -1 | 0 | -0.9298 | 0.5314 | -4.1 | 8 |  | 0 | 0.1434 | 0.2115 | 0.0084 | 0.99999 | [662 empty] | 0
P | 2018-06-27T11:19:43+08:00 | 2018-06-27T14:39:36+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | F18SDI035G | CPSM Final Test | PASS | 129 | 160.16 | -0.02 | -0.69 | 56.41 | 49.95 | 1.82 | 0.13 | -1 | -3 | 0 | -0.9298 | 0.518 | -2.5 | 2 |  | 0 | 0.078 | 0.098 | -0.0093 | 1 | [663 empty] | 0
P | 2018-06-27T12:01:39+08:00 | 2018-06-27T15:25:00+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | F18SDI033G | CPSM Final Test | PASS | 129 | 169.97 | -0.02 | -0.92 | 56.49 | 52.15 | 1.59 | 0.13 | 1 | 1 | 0 | -0.9499 | 0.4928 | -3.2 | -13 |  | 0 | 0.1324 | 0.1799 | 0.0207 | 0.99999 | [664 empty] | 0
P | 2018-06-27T11:54:54+08:00 | 2018-06-27T15:21:55+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | F18SDI034G | CPSM Final Test | PASS | 129 | 170.6 | 0.27 | -0.84 | 56.3 | 50.55 | 1.36 | 0.13 | 0 | -2 | 0 | -0.9637 | 0.5126 | -0.4 | 11 |  | 0.01986 | 0.1093 | 0.254 | 0.0076 | 0.99999 | [665 empty] | 0
P | 2018-06-27T18:14:18+08:00 | 2018-06-27T19:08:53+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | F18USM042G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-06-27T02:04:30+08:00 | 2018-06-27T20:29:24+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI038G | CPSM Final Test | FAIL | [17 empty] | 3.475
P | 2018-06-27T11:16:57+08:00 | 2018-06-27T21:32:20+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI038G | CPSM Final Test | PASS | 129 | 167.72 | -0.11 | -0.77 | 56.28 | 53.34 | 1.82 | 0.13 | -1 | 1 | 0 | -0.9018 | 0.5372 | 1.4 | 16 |  | 0 | 0.0772 | 0.1783 | -0.024 | 1 | [666 empty] | 0
P | 2018-06-27T18:41:36+08:00 | 2018-06-27T21:47:23+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | F18USM042G | CPSM Final Test | PASS |  | 161.78 | -0.02 | -0.92 | 56.71 | 48.77 | 0.98 | 0.13 | -2 | 1 | 0 | -0.9299 | 0.5022 | -2.5 | 54 |  |  | 0.2077 | 0.2002 | -0.0168 | 0.99999 |  |  |  |  | 133 |  | 0 | [473 empty] | 0
P | 2018-06-27T18:45:16+08:00 | 2018-06-27T21:51:22+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | F18USM044G | CPSM Final Test | PASS |  | 162.68 | -0.02 | -0.92 | 56.24 | 50.38 | 1.51 | 0.13 | 1 | -2 | 1 | -0.9019 | 0.493 | -1.6 | 4 |  |  | 0.1067 | 0.0868 | -0.0127 | 1 |  |  |  |  | 133 |  | 0 | [660 empty] | 0
P | 2018-06-27T19:01:31+08:00 | 2018-06-27T21:55:05+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | F18USM045G | CPSM Final Test | PASS |  | 163.04 | -0.02 | -0.16 | 56.5 | 49.36 | 1.29 | 0.13 | 1 | -2 | -1 | -0.8499 | 0.524 | -0.8 | -9 |  |  | 0.0449 | 0.3061 | -0.0184 | 1 |  |  |  |  | 133 |  | 0 | [74 empty] | 0
P | 2018-06-27T19:02:23+08:00 | 2018-06-27T22:02:58+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | F18USM046G | CPSM Final Test | PASS |  | 164.66 | 0.37 | -0.61 | 56.23 | 50.38 | 1.44 | 0.13 | 0 | -2 | 0 | -0.9018 | 0.5476 | 0.3 | 49 |  |  | 0.0676 | 0.1215 | 0.0014 | 0.99997 |  |  |  |  | 133 |  | 0 | [661 empty] | 0
P | 2018-06-27T19:15:07+08:00 | 2018-06-27T21:59:56+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | F18USM043G | CPSM Final Test | PASS |  | 165.47 | -0.02 | 0.07 | 56.21 | 51.9 | 1.13 | 0.13 | -1 | 1 | -1 | -0.9457 | 0.488 | 3.5 | 16 |  |  | 0.0236 | 0.202 | -0.0556 | 1 |  |  |  |  | 133 |  | 0 | [11 empty] | 0
P | 2018-06-28T08:47:53+08:00 | 2018-06-28T11:01:37+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | F18USM049G | CPSM Final Test | FAIL | [13 empty] | -9.4
P | 2018-06-28T07:45:00+08:00 | 2018-06-28T11:48:48+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | F18USM047G | CPSM Final Test | FAIL
P | 2018-06-27T19:59:43+08:00 | 2018-06-28T12:04:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | F18USM041G | CPSM Final Test | PASS |  | 162.05 | 0.27 | -1.22 | 56.28 | 50.88 | 1.06 | 0.13 | -1 | -1 | 0 | -0.8899 | 0.503 | 0.4 | 0 |  |  | 0.0881 | 0.1907 | -0.0479 | 0.99999 |  |  |  |  | 133 |  | 0 | [662 empty] | 0
P | 2018-06-28T09:51:19+08:00 | 2018-06-28T13:04:40+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | F18USM048G | CPSM Final Test | PASS |  | 164.57 | -0.3 | -0.16 | 56.52 | 48.85 | 1.44 | 0.13 | -1 | -4 | 0 | -0.8821 | 0.5184 | -3.7 | 19 |  |  | 0.0179 | 0.0954 | -0.0231 | 1 |  |  |  |  | 133 |  | 0 | [663 empty] | 0
P | 2018-06-28T06:46:26+08:00 | 2018-06-28T12:47:22+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | F18USM047G | CPSM Final Test | FAIL | [672 empty] | 2431
P | 2018-06-28T09:54:45+08:00 | 2018-06-28T13:25:17+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | F18USM049G | CPSM Final Test | PASS |  | 165.29 | 0.37 | -0.92 | 56.61 | 51.99 | 0.98 | 0.13 | 0 | -2 | 0 | -0.8819 | 0.522 | -1.8 | 64 |  |  | 0.1711 | 0.1793 | 0.0024 | 0.99998 |  |  |  |  | 133 |  | 0.02032 | [664 empty] | 0
P | 2018-06-28T10:10:09+08:00 | 2018-06-28T13:52:28+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | F18USM050G | CPSM Final Test | FAIL | [13 empty] | -6.2
P | 2018-06-28T04:58:14+08:00 | 2018-06-28T14:35:34+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | F18USM047G | CPSM Final Test | FAIL | [672 empty] | 117847
P | 2018-06-28T08:45:21+08:00 | 2018-06-28T15:17:16+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | F18USM050G | CPSM Final Test | FAIL
P | 2018-06-28T09:46:54+08:00 | 2018-06-28T15:41:06+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | F18USM047G | CPSM Final Test | PASS |  | 161.78 | -0.02 | -0.46 | 56.92 | 48.68 | 1.36 | 0.13 | 1 | 1 | 1 | -0.7961 | 0.4852 | -1.5 | -15 |  |  | 0.0637 | 0.1846 | 0.0433 | 0.99997 |  |  |  |  | 133 |  | 0 | [644 empty] | 0
P | 2018-06-28T13:31:31+08:00 | 2018-06-28T16:19:00+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI040G | CPSM Final Test | FAIL | [13 empty] | -5.3
P | 2018-06-28T12:01:19+08:00 | 2018-06-28T17:05:54+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | F18USM050G | CPSM Final Test | PASS |  | 163.04 | 0.17 | -0.69 | 56.32 | 48.77 | 1.51 | 0.13 | -1 | -3 | -1 | -0.8819 | 0.496 | 5 | 11 |  |  | 0.0801 | 0.0604 | 0.0154 | 0.99994 |  |  |  |  | 133 |  | 0 | [665 empty] | 0
P | 2018-06-28T13:58:27+08:00 | 2018-06-28T17:14:52+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI041G | CPSM Final Test | PASS | 129 | 158 | -0.02 | -0.54 | 56.41 | 54.27 | 0.98 | 0.13 | 1 | -2 | 0 | -0.8959 | 0.5396 | -1.5 | 4 |  | 0.03867 | 0.0891 | 0.3248 | -0.0511 | 0.99999 | [219 empty] | 0
P | 2018-06-28T14:10:12+08:00 | 2018-06-28T17:20:16+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI042G | CPSM Final Test | PASS | 129 | 167 | -0.02 | -0.08 | 56.54 | 50.29 | 1.36 | 0.13 | 0 | 1 | 0 | -0.9139 | 0.5178 | 4.5 | 0 |  | 0 | 0.1134 | 0.1561 | 0.0111 | 0.99992 | [149 empty] | 0
P | 2018-06-28T14:55:16+08:00 | 2018-06-28T18:19:08+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI040G | CPSM Final Test | PASS | 129 | 165.74 | 0.59 | -0.31 | 56.5 | 51.39 | 0.98 | 0.13 | 1 | -2 | 0 | -0.9936 | 0.509 | -3.7 | 0 |  | 0.0073 | 0.0212 | 0.2166 | -0.0145 | 0.99999 | [673 empty] | 0
P | 2018-06-28T12:39:44+08:00 | 2018-06-28T18:59:29+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI045G | CPSM Final Test | FAIL
P | 2018-06-28T15:56:40+08:00 | 2018-06-28T20:34:09+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI044G | CPSM Final Test | PASS | 129 | 160.88 | -0.02 | -0.69 | 56.47 | 50.12 | 0.91 | 0.13 | 0 | -3 | 0 | -0.9238 | 0.541 | 5.2 | 14 |  | 0 | 0.0602 | 0.2137 | -0.0422 | 1 | [674 empty] | 0
P | 2018-06-28T16:42:14+08:00 | 2018-06-28T20:34:12+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI048G | CPSM Final Test | FAIL
P | 2018-06-28T15:53:03+08:00 | 2018-06-28T20:34:28+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI043G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -0.16 | 56.41 | 51.05 | 1.36 | 0.13 | 1 | 2 | 0 | -0.91 | 0.542 | 4 | 53 |  | 0.02357 | 0.0796 | 0.1592 | -0.0856 | 0.99999 | [675 empty] | 0
P | 2018-06-28T15:49:37+08:00 | 2018-06-28T20:45:37+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI045G | CPSM Final Test | PASS | 129 | 159.8 | -0.02 | -0.38 | 56.46 | 50.29 | 1.21 | 0.13 | -1 | -5 | 0 | -0.8261 | 0.503 | 7.8 | 9 |  | 0.01819 | 0.0971 | 0.1229 | -0.0186 | 1 | [676 empty] | 0
P | 2018-06-28T18:39:20+08:00 | 2018-06-28T21:46:44+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI047G | CPSM Final Test | PASS | 129 | 162.41 | -0.02 | -0.46 | 56.78 | 50.88 | 1.21 | 0.13 | -1 | -2 | 1 | -0.9598 | 0.5408 | -4.8 | 90 |  | 0 | 0.0784 | 0.2232 | 0.0331 | 0.99998 | [563 empty] | 0
P | 2018-06-28T18:38:13+08:00 | 2018-06-28T22:18:38+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI048G | CPSM Final Test | PASS | 129 | 162.5 | -0.02 | -0.31 | 56.3 | 48.77 | 1.51 | 0.13 | 0 | -3 | 0 | -0.9777 | 0.4976 | 1 | 11 |  | 0.01775 | 0.1179 | 0.1707 | -0.0242 | 0.99998 | [282 empty] | 0
P | 2018-06-28T03:19:06+08:00 | 2018-06-29T10:04:56+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI046G | CPSM Final Test | FAIL | [698 empty] | 1488
P | 2018-06-29T07:59:35+08:00 | 2018-06-29T10:59:56+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI051G | CPSM Final Test | FAIL | [13 empty] | -13.1
P | 2018-06-28T18:42:01+08:00 | 2018-06-29T11:11:38+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI046G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | 0.3 | 56.62 | 49.36 | 1.67 | 0.14 | -1 | -3 | 0 | -0.9179 | 0.538 | 1.2 | 15 |  | 0 | 0.1245 | 0.2415 | -0.0545 | 0.99999 | [677 empty] | 0
P | 2018-06-29T11:37:49+08:00 | 2018-06-29T12:22:14+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI054G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-06-29T09:19:30+08:00 | 2018-06-29T12:28:02+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI049G | CPSM Final Test | PASS | 129 | 157.82 | 0.37 | -0.38 | 56.24 | 48.09 | 1.51 | 0.13 | 1 | 1 | 0 | -0.7862 | 0.515 | -0.9 | 2 |  | 0 | 0.2205 | 0.1172 | 0.0546 | 0.99994 | [651 empty] | 0
P | 2018-06-29T09:27:43+08:00 | 2018-06-29T12:28:11+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI050G | CPSM Final Test | PASS | 129 | 158.72 | 0.17 | -0.31 | 56.21 | 51.39 | 1.44 | 0.13 | 1 | 11 | 0 | -0.87 | 0.477 | -3.3 | -2 |  | 0 | 0.0983 | 0.1242 | 0.035 | 0.99998 | [678 empty] | 0
P | 2018-06-29T11:42:04+08:00 | 2018-06-29T12:30:30+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI056G | CPSM Final Test | FAIL | [5 empty] | 46.57
P | 2018-06-29T09:40:48+08:00 | 2018-06-29T12:44:48+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI052G | CPSM Final Test | PASS | 129 | 156.47 | 0.37 | -0.31 | 56.42 | 56.05 | 1.29 | 0.13 | -1 | -1 | 0 | -0.9118 | 0.536 | -3.2 | 24 |  | 0.01019 | 0.0595 | 0.3459 | -0.0522 | 1 | [679 empty] | 0
P | 2018-06-29T09:29:46+08:00 | 2018-06-29T13:00:30+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI051G | CPSM Final Test | PASS | 129 | 163.22 | 0.37 | -0.69 | 56.21 | 49.28 | 1.67 | 0.13 | 1 | 2 | 0 | -0.9159 | 0.5128 | 4.9 | 8 |  | 0 | 0.1372 | 0.1527 | 0.0027 | 0.99997 | [312 empty] | 0
P | 2018-06-29T10:07:57+08:00 | 2018-06-29T13:52:06+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI054G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-06-29T09:53:28+08:00 | 2018-06-29T13:52:44+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI053G | CPSM Final Test | FAIL | [13 empty] | -12.3
P | 2018-06-29T09:54:15+08:00 | 2018-06-29T13:58:51+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI055G | CPSM Final Test | FAIL | [13 empty] | -8.5
P | 2018-06-29T12:00:01+08:00 | 2018-06-29T15:52:44+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI054G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.46 | 56.22 | 49.53 | 1.29 | 0.13 | 0 | -4 | 0 | -0.848 | 0.541 | 2.5 | -1 |  | 0 | 0.044 | 0.1197 | -0.0488 | 0.99999 | [680 empty] | 0
P | 2018-06-29T12:06:17+08:00 | 2018-06-29T15:54:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI056G | CPSM Final Test | PASS | 129 | 161.51 | 0.17 | -0.54 | 56.04 | 48.94 | 1.44 | 0.13 | -1 | 1 | 1 | -0.8101 | 0.522 | -0.3 | 10 |  | 0 | 0.0278 | 0.2189 | 0.0187 | 0.99999 | [677 empty] | 0
P | 2018-06-29T11:56:33+08:00 | 2018-06-29T16:09:35+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI055G | CPSM Final Test | PASS | 129 | 164.84 | 0.17 | -0.69 | 56.32 | 50.12 | 1.74 | 0.13 | 0 | -2 | 0 | -0.8321 | 0.5 | 0.8 | -1 |  | 0.01425 | 0.2575 | 0.1532 | 0.017 | 0.99998 | [681 empty] | 0
P | 2018-06-29T13:47:59+08:00 | 2018-06-29T17:29:12+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI058G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-06-29T12:35:07+08:00 | 2018-06-29T18:42:04+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI058G | CPSM Final Test | FAIL | [14 empty] | 605
P | 2018-06-29T15:42:41+08:00 | 2018-06-29T18:54:42+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI060G | CPSM Final Test | PASS | 129 | 170.96 | -0.02 | -0.31 | 56.41 | 54.69 | 1.36 | 0.13 | -1 | -2 | -1 | -0.9498 | 0.55 | 7.5 | 20 |  | 0 | 0.0956 | 0.165 | 0.0329 | 0.99997 | [682 empty] | 0
P | 2018-06-29T04:49:56+08:00 | 2018-06-29T18:56:16+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI053G | CPSM Final Test | FAIL | [17 empty] | 0.7495
P | 2018-06-29T15:36:55+08:00 | 2018-06-29T18:59:18+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI057G | CPSM Final Test | PASS | 129 | 166.19 | 0.17 | -0.31 | 56.61 | 50.21 | 1.29 | 0.13 | 0 | -1 | 0 | -0.88 | 0.5222 | 1.8 | 46 |  | 0.01327 | 0.1896 | 0.3716 | -0.0468 | 0.99999 | [683 empty] | 0
P | 2018-06-29T16:17:38+08:00 | 2018-06-29T19:17:48+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI059G | CPSM Final Test | PASS | 129 | 163.94 | 0.17 | -0.54 | 56.17 | 49.95 | 1.44 | 0.13 | -1 | 2 | 0 | -0.8421 | 0.537 | -1.9 | 21 |  | 0 | 0.152 | 0.3135 | -0.0382 | 1 | [684 empty] | 0
P | 2018-06-29T15:38:36+08:00 | 2018-06-29T20:26:58+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI058G | CPSM Final Test | PASS | 129 | 162.41 | -0.02 | -0.16 | 56.43 | 53.26 | 1.74 | 0.13 | 1 | -2 | 0 | -0.8859 | 0.4958 | 1.8 | 31 |  | 0 | 0.1174 | 0.2771 | -0.0506 | 0.99998 | [163 empty] | 0
P | 2018-06-29T11:53:06+08:00 | 2018-06-29T20:30:40+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI053G | CPSM Final Test | PASS | 129 | 161.42 | 0.17 | -0.38 | 56.01 | 51.31 | 1.82 | 0.13 | 0 | -2 | 0 | -0.8539 | 0.5268 | -2 | -2 |  | 0 | 0.4933 | 0.6134 | -0.1987 | 0.99998 | [685 empty] | 0
P | 2018-06-29T16:28:32+08:00 | 2018-06-29T20:37:38+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI063G | CPSM Final Test | FAIL
P | 2018-06-29T18:30:11+08:00 | 2018-06-29T21:32:08+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI062G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.31 | 56.33 | 49.02 | 1.29 | 0.13 | 0 | 3 | 0 | -0.9159 | 0.508 | 2.1 | 3 |  | 0 | 0.1231 | 0.0472 | -0.0083 | 0.99999 | [686 empty] | 0
P | 2018-06-29T18:27:03+08:00 | 2018-06-29T21:32:09+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI061G | CPSM Final Test | PASS | 129 | 160.16 | -0.02 | -0.46 | 56.31 | 48.68 | 1.44 | 0.13 | 1 | -2 | 0 | -0.894 | 0.5106 | 3.6 | 39 |  | 0.01126 | 0.0915 | 0.2117 | 0.027 | 0.99999 | [687 empty] | 0
P | 2018-06-29T18:33:05+08:00 | 2018-06-29T21:43:20+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI063G | CPSM Final Test | PASS | 129 | 157.91 | 0.27 | -0.46 | 56.21 | 53.17 | 1.29 | 0.13 | 0 | 1 | 0 | -0.9059 | 0.5466 | -3.6 | 12 |  | 0 | 0.0603 | 0.0277 | -0.0122 | 0.99999 | [688 empty] | 0
P | 2018-06-30T10:26:15+08:00 | 2018-06-30T13:31:24+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI065G | CPSM Final Test | PASS | 129 | 153.23 | -0.02 | -0.46 | 55.93 | 48.85 | 1.59 | 0.13 | 0 | -1 | 0 | -0.8918 | 0.5454 | 10.3 | 0 |  | 0 | 0.1561 | 0.1288 | 0.0123 | 0.99999 | [689 empty] | 0
P | 2018-06-30T10:30:01+08:00 | 2018-06-30T13:32:09+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI064G | CPSM Final Test | PASS | 129 | 166.82 | 0.17 | -0.46 | 56.12 | 50.55 | 1.36 | 0.13 | -1 | -1 | 0 | -0.8619 | 0.5218 | 7.8 | 8 |  | 0.01332 | 0.1722 | 0.3519 | 0.022 | 0.99998 | [690 empty] | 0
P | 2018-06-30T10:33:40+08:00 | 2018-06-30T13:33:34+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI066G | CPSM Final Test | PASS | 129 | 162.68 | 0.17 | -0.08 | 56.18 | 51.48 | 1.51 | 0.13 | 1 | 1 | 0 | -0.8819 | 0.496 | 0.4 | -2 |  | 0 | 0.1027 | 0.3147 | 0.001 | 0.99999 | [279 empty] | 0
P | 2018-06-30T10:37:25+08:00 | 2018-06-30T13:34:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI068G | CPSM Final Test | PASS | 129 | 160.52 | -0.02 | 0.53 | 56.2 | 50.38 | 1.13 | 0.13 | 1 | 2 | 0 | -0.9158 | 0.453 | 0.3 | -2 |  | 0.02589 | 0.0639 | 0.1849 | -0.0129 | 0.99999 | [688 empty] | 0
P | 2018-06-30T07:23:51+08:00 | 2018-06-30T13:58:18+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI067G | CPSM Final Test | FAIL
P | 2018-06-30T10:41:05+08:00 | 2018-06-30T13:58:48+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI067G | CPSM Final Test | PASS | 129 | 161.87 | -0.02 | -0.54 | 56.17 | 49.45 | 1.44 | 0.13 | 1 | -2 | 0 | -0.894 | 0.5428 | 6.6 | 6 |  | 0.01034 | 0.0581 | 0.3536 | -0.0065 | 0.99997 | [173 empty] | 0
P | 2018-06-30T11:46:20+08:00 | 2018-06-30T15:32:50+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI069G | CPSM Final Test | PASS | 129 | 163.22 | 0.17 | -0.08 | 56.2 | 48.68 | 1.74 | 0.13 | 0 | -2 | 0 | -0.8998 | 0.544 | -1.1 | -13 |  | 0 | 0.0812 | 0.2886 | 0.0094 | 0.99999 | [691 empty] | 0
P | 2018-06-30T12:53:04+08:00 | 2018-06-30T15:52:18+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI070G | CPSM Final Test | PASS | 129 | 159.26 | -0.02 | 0 | 56.5 | 50.29 | 1.13 | 0.13 | 1 | -2 | 0 | -0.8939 | 0.512 | 6.2 | -8 |  | 0 | 0.0311 | 0.2669 | -0.0229 | 0.99999 | [692 empty] | 0
P | 2018-06-30T15:42:00+08:00 | 2018-06-30T16:36:54+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI073G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-06-30T15:32:49+08:00 | 2018-06-30T16:39:02+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI074G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-06-30T15:42:37+08:00 | 2018-06-30T18:25:50+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI071G | CPSM Final Test | PASS | 129 | 153.86 | -0.02 | -0.38 | 56.5 | 50.21 | 1.21 | 0.13 | 0 | -2 | 0 | -0.9198 | 0.5008 | 0.4 | 33 |  | 0 | 0.0421 | 0.1528 | -0.0264 | 1 | [693 empty] | 0
P | 2018-06-30T16:09:27+08:00 | 2018-06-30T19:08:30+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI073G | CPSM Final Test | PASS | 129 | 164.12 | 0.27 | -0.54 | 56.34 | 49.02 | 1.36 | 0.13 | 0 | 1 | 0 | -0.8202 | 0.508 | 7 | 14 |  | 0 | 0.4424 | 0.425 | -0.0437 | 0.99993 | [694 empty] | 0
P | 2018-06-30T16:05:56+08:00 | 2018-06-30T19:10:54+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI074G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | 0.23 | 56.53 | 49.28 | 1.67 | 0.13 | 1 | 1 | -1 | -0.9119 | 0.4988 | 11.8 | -17 |  | 0 | 0.1988 | 0.1941 | 0.0022 | 1 | [335 empty] | 0
P | 2018-06-30T14:00:14+08:00 | 2018-06-30T17:19:40+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | F18SDI072G | CPSM Final Test | FAIL | [14 empty] | 659
P | 2018-07-01T10:00:26+08:00 | 2018-07-01T12:12:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI079G | CPSM Final Test | FAIL | [13 empty] | -6.2
P | 2018-07-01T09:48:59+08:00 | 2018-07-01T12:55:00+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI075G | CPSM Final Test | PASS | 129 | 166.01 | 0.08 | -0.46 | 56.3 | 50.55 | 1.59 | 0.13 | -1 | -4 | 0 | -0.846 | 0.544 | -3.2 | 12 |  | 0 | 0.2001 | 0.225 | -0.005 | 0.99999 | [695 empty] | 0
P | 2018-07-01T09:58:12+08:00 | 2018-07-01T12:57:38+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI078G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -0.61 | 56.42 | 52.41 | 2.05 | 0.13 | 1 | -2 | 0 | -0.8401 | 0.5314 | -3 | -13 |  | 0 | 0.1227 | 0.1759 | -0.0214 | 0.99998 | [696 empty] | 0
P | 2018-07-01T10:02:26+08:00 | 2018-07-01T13:00:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI077G | CPSM Final Test | PASS | 129 | 166.19 | -0.02 | -0.54 | 56.34 | 49.45 | 1.59 | 0.13 | -1 | 1 | -1 | -0.82 | 0.5354 | -0.8 | -1 |  | 0.10802 | 0.3649 | 0.3362 | -0.0211 | 1 | [410 empty] | 0
P | 2018-07-01T10:16:45+08:00 | 2018-07-01T13:07:52+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI072G | CPSM Final Test | PASS | 129 | 165.92 | 0.17 | 0.07 | 56.83 | 50.97 | 1.06 | 0.13 | -1 | -2 | 0 | -1.0196 | 0.521 | -2.3 | 44 |  | 0 | 0.0649 | 0.1503 | 0.002 | 0.99999 | [697 empty] | 0
P | 2018-07-01T09:50:56+08:00 | 2018-07-01T12:30:36+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI080G | CPSM Final Test | FAIL
P | 2018-07-01T05:46:12+08:00 | 2018-07-01T14:01:39+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI076G | CPSM Final Test | FAIL | [16 empty] | 0.37923
P | 2018-07-01T11:06:40+08:00 | 2018-07-01T14:12:52+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI079G | CPSM Final Test | PASS | 129 | 163.31 | -0.02 | -0.46 | 56.62 | 50.12 | 1.21 | 0.13 | -1 | -2 | 0 | -0.8501 | 0.518 | -1.3 | 89 |  | 0 | 0.04 | 0.1341 | -0.0369 | 1 | [698 empty] | 0
P | 2018-07-01T09:53:55+08:00 | 2018-07-01T15:04:45+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI076G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.16 | 56.62 | 50.21 | 1.9 | 0.13 | 1 | 2 | -1 | -0.84 | 0.53 | 4.5 | 18 |  | 0 | 0.1862 | 0.2098 | 0.036 | 0.99997 | [538 empty] | 0
P | 2018-07-01T11:10:46+08:00 | 2018-07-01T15:09:36+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI080G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | 0.3 | 56.73 | 55.54 | 1.29 | 0.13 | 1 | -3 | 0 | -0.8061 | 0.499 | 4.2 | 35 |  | 0 | 0.09 | 0.2137 | -0.0497 | 0.99999 | [699 empty] | 0
P | 2018-07-01T12:00:21+08:00 | 2018-07-01T17:25:58+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI081G | CPSM Final Test | FAIL | [160 empty] | 2561
P | 2018-07-01T11:56:21+08:00 | 2018-07-01T17:29:58+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI081G | CPSM Final Test | FAIL
P | 2018-07-01T14:39:05+08:00 | 2018-07-01T18:04:04+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI083G | CPSM Final Test | PASS | 129 | 162.14 | 0.17 | -0.31 | 56.4 | 53.26 | 0.91 | 0.13 | 1 | -2 | 0 | -0.9058 | 0.521 | 0 | 8 |  | 0 | 0.1954 | 0.215 | -0.0283 | 0.99998 | [85 empty] | 0
P | 2018-07-01T14:42:21+08:00 | 2018-07-01T18:06:16+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI082G | CPSM Final Test | PASS | 129 | 169.7 | 0.17 | -0.69 | 56.53 | 50.12 | 1.13 | 0.13 | 1 | 3 | 0 | -0.8681 | 0.5214 | 1.2 | -31 |  | 0 | 0.0952 | 0.0583 | -0.0268 | 1 | [700 empty] | 0
P | 2018-07-01T14:43:09+08:00 | 2018-07-01T18:35:52+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI081G | CPSM Final Test | PASS | 129 | 162.41 | 0.08 | -0.31 | 56.72 | 53.51 | 1.29 | 0.13 | 0 | -4 | -1 | -0.8101 | 0.536 | 2.7 | 2 |  | 0 | 0.0948 | 0.2353 | 0.0188 | 0.99995 | [139 empty] | 0
P | 2018-07-01T00:39:37+08:00 | 2018-07-02T09:20:08+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI086G | CPSM Final Test | FAIL | [722 empty] | 0
P | 2018-07-01T16:59:52+08:00 | 2018-07-02T10:25:40+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI086G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.31 | 56.64 | 50.88 | 1.51 | 0.13 | 0 | -1 | 0 | -0.8819 | 0.5052 | -4 | 3 |  | 0 | 0.1044 | 0.0948 | 0.0132 | 1 | [701 empty] | 0
P | 2018-07-02T10:26:07+08:00 | 2018-07-02T11:18:16+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI088G | CPSM Final Test | FAIL | [5 empty] | 46.4
P | 2018-07-02T10:21:28+08:00 | 2018-07-02T13:20:29+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | F18SDI084G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -0.38 | 56.8 | 50.63 | 1.13 | 0.13 | 0 | 3 | -1 | -0.9038 | 0.5366 | -0.2 | 25 |  | 0 | 0.1 | 0.1316 | -0.0008 | 0.99999 | [139 empty] | 0
P | 2018-07-02T10:16:38+08:00 | 2018-07-02T13:30:46+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | F18SDI087G | CPSM Final Test | PASS | 129 | 156.02 | -0.02 | -0.31 | 56.76 | 51.14 | 1.13 | 0.13 | 1 | 2 | 0 | -0.9259 | 0.548 | -1 | -8 |  | 0 | 0.2286 | 0.2245 | -0.0169 | 0.99999 | [702 empty] | 0
P | 2018-07-02T10:10:52+08:00 | 2018-07-02T13:34:00+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | F18SDI085G | CPSM Final Test | PASS | 129 | 164.75 | 0.27 | -0.38 | 56.61 | 49.53 | 1.44 | 0.13 | 1 | -2 | 1 | -0.854 | 0.487 | 12.8 | -16 |  | 0 | 0.1616 | 0.2882 | -0.0281 | 0.99998 | [699 empty] | 0
P | 2018-07-02T10:49:03+08:00 | 2018-07-02T13:37:08+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | F18SDI090G | CPSM Final Test | PASS | 129 | 161.96 | 0.27 | 0.61 | 56.93 | 50.38 | 1.67 | 0.13 | -1 | -3 | 0 | -0.8978 | 0.5462 | 8.6 | 15 |  | 0 | 0.1171 | 0.0722 | -0.0115 | 1 | [703 empty] | 0
P | 2018-07-02T10:45:04+08:00 | 2018-07-02T13:39:47+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI089G | CPSM Final Test | PASS | 129 | 159.89 | -0.02 | 0.15 | 56.31 | 50.63 | 2.12 | 0.13 | 0 | -2 | -1 | -0.9159 | 0.5488 | 3.8 | 10 |  | 0 | 0.1715 | 0.3167 | -0.044 | 0.99997 | [704 empty] | 0
P | 2018-07-02T10:52:11+08:00 | 2018-07-02T13:52:48+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI088G | CPSM Final Test | PASS | 129 | 158.63 | 0.17 | 0.07 | 56.47 | 51.05 | 1.74 | 0.13 | -2 | -2 | -1 | -0.9279 | 0.5246 | -4.7 | -17 |  | 0 | 0.0353 | 0.0842 | -0.0221 | 0.99996 | [705 empty] | 0
P | 2018-07-02T08:57:09+08:00 | 2018-07-02T15:03:38+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI091G | CPSM Final Test | FAIL | [16 empty] | 0.29977
P | 2018-07-02T12:00:24+08:00 | 2018-07-02T16:06:30+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | F18SDI091G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | 0.3 | 56.63 | 48.94 | 1.06 | 0.13 | 1 | 2 | 1 | -0.9099 | 0.5348 | 1.6 | 27 |  | 0 | 0.0864 | 0.1718 | -0.0335 | 0.99997 | [701 empty] | 0
P | 2018-07-02T16:47:00+08:00 | 2018-07-02T20:02:02+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP519G | CPSM Final Test | PASS |  | 164.66 | -0.02 | -0.23 | 56.34 | 50.88 | 1.67 | 0.13 | -1 | -3 | 0 | -0.9319 | 0.5322 | -0.5 | 5 |  |  | 0.1594 | 0.1598 | -0.0385 | 0.99999 |  |  |  |  | 133 |  | 0.01147 | [699 empty] | 0
P | 2018-07-04T09:37:58+08:00 | 2018-07-04T12:07:32+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP521G | CPSM Final Test | FAIL | [13 empty] | -7.7
P | 2018-07-04T10:29:29+08:00 | 2018-07-04T12:27:58+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | F18FNP523G | CPSM Final Test | FAIL | [13 empty] | 18.1
P | 2018-07-04T07:46:17+08:00 | 2018-07-04T13:17:13+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP524G | CPSM Final Test | FAIL
P | 2018-07-04T09:23:15+08:00 | 2018-07-04T12:57:35+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP520G | CPSM Final Test | FAIL
P | 2018-07-04T09:25:41+08:00 | 2018-07-04T13:31:46+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | F18FNP523G | CPSM Final Test | FAIL | [13 empty] | 15.2
P | 2018-07-04T10:41:26+08:00 | 2018-07-04T13:50:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP522G | CPSM Final Test | PASS |  | 162.5 | -0.59 | -0.99 | 56.19 | 51.99 | 1.44 | 0.13 | -1 | -2 | 1 | -0.824 | 0.517 | 3.2 | 32 |  |  | 0.0418 | 0.065 | -0.0084 | 0.99999 |  |  |  |  | 133 |  | 0 | [214 empty] | 0
P | 2018-07-04T07:41:20+08:00 | 2018-07-04T14:04:10+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP521G | CPSM Final Test | FAIL | [728 empty] | 0
P | 2018-07-04T11:10:25+08:00 | 2018-07-04T14:22:59+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP520G | CPSM Final Test | PASS |  | 158.54 | 0.17 | 0.23 | 56.45 | 50.38 | 0.98 | 0.13 | 1 | 3 | -1 | -0.8261 | 0.523 | 3.2 | 81 |  |  | 0.1473 | 0.1611 | -0.0206 | 0.99999 |  |  |  |  | 133 |  | 0.01916 | [701 empty] | 0
P | 2018-07-04T10:26:11+08:00 | 2018-07-04T14:30:30+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP525G | CPSM Final Test | PASS |  | 163.49 | 0.37 | -0.61 | 56.13 | 51.56 | 1.67 | 0.13 | 1 | 2 | 0 | -0.842 | 0.504 | 4.6 | 33 |  |  | 0.4521 | 0.4526 | -0.0193 | 0.99998 |  |  |  |  | 133 |  | 0.02411 | [702 empty] | 0
P | 2018-07-04T10:52:45+08:00 | 2018-07-04T15:09:36+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP521G | CPSM Final Test | PASS |  | 159.98 | -0.02 | -0.77 | 56.43 | 55.54 | 1.36 | 0.13 | 0 | 3 | 0 | -0.8002 | 0.4916 | 6.4 | 14 |  |  | 0.1536 | 0.2164 | -0.0517 | 1 |  |  |  |  | 133 |  | 0 | [700 empty] | 0
P | 2018-07-04T11:28:43+08:00 | 2018-07-04T15:36:54+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | F18FNP523G | CPSM Final Test | PASS |  | 160.34 | 0.27 | -0.46 | 56.51 | 48.94 | 1.36 | 0.13 | 1 | -2 | -1 | -0.83 | 0.463 | 9.8 | 22 |  |  | 0.0944 | 0.1089 | -0.0139 | 1 |  |  |  |  | 133 |  | 0 | [703 empty] | 0
P | 2018-07-04T04:36:13+08:00 | 2018-07-04T16:27:17+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP524G | CPSM Final Test | FAIL | [27 empty] | 0.04679
P | 2018-07-04T10:31:45+08:00 | 2018-07-04T17:30:17+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP524G | CPSM Final Test | PASS |  | 163.31 | -0.02 | -0.31 | 56.46 | 51.14 | 1.59 | 0.13 | 1 | -3 | 0 | -0.862 | 0.477 | -3.7 | 0 |  |  | 0.0911 | 0.1677 | -0.0396 | 0.99999 |  |  |  |  | 133 |  | 0 | [335 empty] | 0
P | 2018-07-09T18:01:40+08:00 | 2018-07-09T19:06:55+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP527G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-07-09T18:43:01+08:00 | 2018-07-09T21:43:09+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP532G | CPSM Final Test | PASS |  | 163.31 | 0.17 | -1.3 | 56.17 | 50.8 | 0.75 | 0.13 | 0 | 1 | 0 | -0.9558 | 0.5164 | 2.6 | 15 |  |  | 0.1222 | 0.1614 | -0.0457 | 0.99999 |  |  |  |  | 133 |  | 0 | [609 empty] | 0
P | 2018-07-09T18:32:58+08:00 | 2018-07-09T21:48:03+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | F18FNP529G | CPSM Final Test | PASS |  | 163.31 | -0.02 | -0.46 | 56.33 | 50.29 | 1.13 | 0.13 | 0 | 2 | 0 | -0.9219 | 0.544 | 5.4 | 15 |  |  | 0.105 | 0.1591 | -0.0359 | 1 |  |  |  |  | 133 |  | 0.02759 | [700 empty] | 0
P | 2018-07-09T18:34:17+08:00 | 2018-07-09T21:50:37+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | F18FNP527G | CPSM Final Test | PASS |  | 162.5 | 0.27 | -1.15 | 56.01 | 48.51 | 1.29 | 0.13 | -1 | 1 | 0 | -0.894 | 0.5034 | 5.9 | 42 |  |  | 0.2096 | 0.512 | 0.0094 | 1 |  |  |  |  | 133 |  | 0 | [701 empty] | 0
P | 2018-07-09T18:53:09+08:00 | 2018-07-09T21:55:10+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | F18FNP533G | CPSM Final Test | PASS |  | 167.36 | -0.02 | -0.54 | 56.24 | 49.87 | 1.51 | 0.13 | -1 | 1 | 0 | -0.9279 | 0.5368 | -4.3 | 55 |  |  | 0.1248 | 0.119 | -0.0005 | 0.99996 |  |  |  |  | 133 |  | 0 | [703 empty] | 0
P | 2018-07-09T19:13:21+08:00 | 2018-07-09T21:56:57+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP526G | CPSM Final Test | PASS |  | 163.4 | -0.02 | -0.46 | 56.31 | 50.63 | 1.67 | 0.13 | 1 | -2 | 0 | -0.9598 | 0.5048 | -1.8 | 7 |  |  | 0.1268 | 0.1902 | -0.021 | 1 |  |  |  |  | 133 |  | 0.01241 | [704 empty] | 0
P | 2018-07-09T19:32:42+08:00 | 2018-07-09T22:30:26+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP530G | CPSM Final Test | PASS |  | 161.78 | 0.27 | -0.08 | 56.13 | 48.94 | 0.98 | 0.13 | 1 | 3 | 0 | -0.9657 | 0.5422 | 6.2 | 5 |  |  | 0.1057 | 0.1555 | 0.011 | 0.99995 |  |  |  |  | 133 |  | 0.01052 | [705 empty] | 0
P | 2018-07-10T08:11:33+08:00 | 2018-07-10T11:29:45+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP535G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-07-10T07:28:24+08:00 | 2018-07-10T11:46:16+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP536G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-07-10T08:38:59+08:00 | 2018-07-10T11:56:54+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP742G | CPSM Final Test | FAIL | [5 empty] | 45.64
P | 2018-07-10T06:48:55+08:00 | 2018-07-10T12:34:03+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP531G | CPSM Final Test | FAIL | [13 empty] | 45.6
P | 2018-07-09T20:07:29+08:00 | 2018-07-10T13:25:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | F18FNP528G | CPSM Final Test | PASS |  | 162.77 | -0.21 | -0.08 | 56.43 | 55.2 | 1.51 | 0.13 | 1 | 3 | 0 | -0.904 | 0.514 | 8 | 29 |  |  | 0.3483 | 0.4394 | 0.0486 | 0.99998 |  |  |  |  | 133 |  | 0 | [249 empty] | 0
P | 2018-07-10T10:06:46+08:00 | 2018-07-10T13:49:25+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP703G | CPSM Final Test | PASS | 129 | 163.4 | 0.59 | -0.92 | 56.12 | 51.05 | 0.98 | 0.13 | -1 | -2 | 0 | -0.8102 | 0.491 | 10 | -17 |  |  | 0.0692 | 0.2892 | -0.0092 | 0.99999 | [21 empty] | 0 | [691 empty] | 0
P | 2018-07-10T09:57:30+08:00 | 2018-07-10T13:54:03+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP744G | CPSM Final Test | PASS | 129 | 165.74 | 0.37 | -0.46 | 56.21 | 50.63 | 1.36 | 0.13 | 1 | -2 | 1 | -0.8679 | 0.508 | 4.8 | -5 |  |  | 0.0536 | 0.1135 | -0.0318 | 1 | [21 empty] | 0 | [692 empty] | 0
P | 2018-07-10T05:15:18+08:00 | 2018-07-10T13:59:22+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP536G | CPSM Final Test | FAIL | [104 empty] | 1439
P | 2018-07-10T05:15:39+08:00 | 2018-07-10T14:07:19+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP531G | CPSM Final Test | FAIL | [736 empty] | 3046
P | 2018-07-10T05:30:41+08:00 | 2018-07-10T14:10:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP535G | CPSM Final Test | FAIL
P | 2018-07-10T09:50:39+08:00 | 2018-07-10T14:11:05+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | F18FNP535G | CPSM Final Test | PASS |  | 163.85 | 0.08 | -0.54 | 56.07 | 48.43 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9398 | 0.5034 | 0.3 | 29 |  |  | 0.3204 | 0.171 | -0.0253 | 1 |  |  |  |  | 133 |  | 0 | [709 empty] | 0
P | 2018-07-10T10:17:56+08:00 | 2018-07-10T14:29:48+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP742G | CPSM Final Test | PASS | 129 | 161.33 | 0.46 | 0.07 | 56.31 | 49.28 | 1.44 | 0.13 | 0 | 2 | -1 | -0.8439 | 0.4662 | 10.4 | 1 |  |  | 0.1367 | 0.7202 | -0.0107 | 0.99999 | [21 empty] | 0 | [184 empty] | 0
P | 2018-07-10T13:58:49+08:00 | 2018-07-10T14:39:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP748G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-07-10T09:37:20+08:00 | 2018-07-10T15:04:54+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | F18FNP536G | CPSM Final Test | PASS |  | 162.23 | -0.02 | 0 | 57 | 48.34 | 1.9 | 0.13 | -1 | -2 | 0 | -0.9119 | 0.543 | -2 | 23 |  |  | 0.1904 | 0.1902 | 0.0014 | 0.99996 |  |  |  |  | 133 |  | 0 | [76 empty] | 0
P | 2018-07-10T09:41:29+08:00 | 2018-07-10T15:12:57+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | F18FNP531G | CPSM Final Test | PASS |  | 163.49 | 0.17 | -0.46 | 56.11 | 49.87 | 1.36 | 0.13 | 0 | -1 | 0 | -0.8819 | 0.4912 | 7.5 | 16 |  |  | 0.0465 | 0.3613 | -0.0402 | 0.99999 |  |  |  |  | 133 |  | 0 | [708 empty] | 0
P | 2018-07-10T14:01:27+08:00 | 2018-07-10T16:53:45+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP706G | CPSM Final Test | PASS | 129 | 161.33 | 0.08 | 0.07 | 56.17 | 49.11 | 1.21 | 0.13 | 1 | 3 | 0 | -0.8521 | 0.5016 | 6.3 | -10 |  |  | 0.1744 | 0.3087 | -0.0122 | 0.99997 | [21 empty] | 0 | [695 empty] | 0
P | 2018-07-10T14:19:18+08:00 | 2018-07-10T17:34:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP748G | CPSM Final Test | PASS | 129 | 159.98 | 0.08 | -0.38 | 56.23 | 50.97 | 1.36 | 0.13 | -1 | 1 | -1 | -0.9159 | 0.543 | 6.3 | 3 |  |  | 0.2258 | 0.2908 | -0.0441 | 1 | [21 empty] | 0 | [696 empty] | 0
P | 2018-07-10T13:14:02+08:00 | 2018-07-10T17:36:08+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP746G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-07-10T12:01:14+08:00 | 2018-07-10T19:22:27+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP733G | CPSM Final Test | FAIL
P | 2018-07-10T09:56:07+08:00 | 2018-07-10T19:25:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP726G | CPSM Final Test | FAIL
P | 2018-07-10T15:23:59+08:00 | 2018-07-10T19:53:26+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP732G | CPSM Final Test | FAIL | [14 empty] | 638
P | 2018-07-08T19:48:24+08:00 | 2018-07-10T20:09:18+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP534G | CPSM Final Test | FAIL | [732 empty] | 0
P | 2018-07-10T15:55:31+08:00 | 2018-07-10T20:43:24+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP745G | CPSM Final Test | PASS | 129 | 160.88 | 0.27 | -0.54 | 56.61 | 50.63 | 1.44 | 0.13 | 1 | 11 | 1 | -0.9438 | 0.507 | 0.3 | 11 |  |  | 0.1023 | 0.1381 | 0.0052 | 0.99999 | [21 empty] | 0 | [61 empty] | 0
P | 2018-07-10T15:41:50+08:00 | 2018-07-10T21:06:27+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP733G | CPSM Final Test | PASS | 129 | 157.46 | 0.08 | 0.23 | 56.15 | 49.45 | 1.06 | 0.13 | -1 | 1 | 1 | -0.874 | 0.5448 | 3.5 | -8 |  |  | 0.0553 | 0.3426 | -0.0101 | 0.99998 | [21 empty] | 0.00748 | [693 empty] | 0
P | 2018-07-10T14:40:41+08:00 | 2018-07-10T21:09:31+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP726G | CPSM Final Test | PASS | 129 | 170.06 | 0.08 | -0.54 | 56.43 | 55.12 | 1.29 | 0.13 | 1 | -2 | 0 | -0.9318 | 0.513 | 8.9 | -3 |  |  | 0.1104 | 0.0979 | -0.0043 | 0.99999 | [21 empty] | 0 | [697 empty] | 0
P | 2018-07-08T06:53:36+08:00 | 2018-07-11T09:04:06+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP534G | CPSM Final Test | FAIL | [732 empty] | 0
P | 2018-07-11T08:52:00+08:00 | 2018-07-11T09:59:11+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP722G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-07-09T20:30:22+08:00 | 2018-07-11T10:19:48+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP746G | CPSM Final Test | FAIL
P | 2018-07-09T20:28:28+08:00 | 2018-07-11T10:21:42+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP746G | CPSM Final Test | FAIL | [42 empty] | 0
P | 2018-07-10T00:38:17+08:00 | 2018-07-11T10:39:08+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP732G | CPSM Final Test | FAIL | [741 empty] | 2619
P | 2018-07-10T00:34:35+08:00 | 2018-07-11T10:42:50+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP732G | CPSM Final Test | FAIL
P | 2018-07-11T08:48:53+08:00 | 2018-07-11T11:11:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP713G | CPSM Final Test | FAIL | [13 empty] | 38.6
P | 2018-07-08T04:36:32+08:00 | 2018-07-11T11:21:10+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP534G | CPSM Final Test | FAIL | [732 empty] | 2372
P | 2018-07-10T15:25:05+08:00 | 2018-07-11T11:24:52+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP746G | CPSM Final Test | PASS | 129 | 162.41 | 0.65 | -0.61 | 56.54 | 49.11 | 1.21 | 0.13 | 1 | 1 | 0 | -0.8839 | 0.4838 | -3.3 | 8 |  |  | 0.225 | 0.5456 | -0.008 | 1 | [21 empty] | 0 | [486 empty] | 0
P | 2018-07-10T17:38:43+08:00 | 2018-07-11T11:48:22+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP732G | CPSM Final Test | PASS | 129 | 169.97 | 0.08 | -0.46 | 56.34 | 48.18 | 1.44 | 0.13 | -1 | -2 | 1 | -0.7084 | 0.5418 | 7.2 | 20 |  |  | 0.1846 | 0.4713 | 0.0994 | 0.99997 | [21 empty] | 0.00637 | [698 empty] | 0
P | 2018-07-09T19:58:51+08:00 | 2018-07-11T12:27:38+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | F18FNP534G | CPSM Final Test | PASS |  | 160.43 | 0.17 | -1.07 | 56.13 | 50.97 | 1.36 | 0.13 | -1 | -1 | -1 | -0.7363 | 0.542 | -4.1 | -3 |  |  | 0.2575 | 0.3181 | -0.0624 | 0.99997 |  |  |  |  | 133 |  | 0 | [704 empty] | 0
P | 2018-07-11T08:12:37+08:00 | 2018-07-11T12:40:23+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP718G | CPSM Final Test | FAIL
P | 2018-07-11T09:14:49+08:00 | 2018-07-11T12:49:12+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP712G | CPSM Final Test | PASS | 129 | 156.2 | 0.68 | -0.99 | 56.04 | 49.95 | 1.29 | 0.13 | -1 | 2 | 0 | -0.9358 | 0.526 | -4.7 | -5 |  |  | 0.2096 | 0.2286 | -0.0248 | 0.99998 | [21 empty] | 0 | [699 empty] | 0
P | 2018-07-11T09:25:35+08:00 | 2018-07-11T12:45:53+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP722G | CPSM Final Test | PASS | 129 | 166.82 | 0.56 | -0.99 | 56.31 | 49.7 | 1.13 | 0.13 | -1 | 1 | 0 | -0.9059 | 0.5022 | 7.9 | 16 |  |  | 0.1551 | 0.1862 | -0.0313 | 0.99999 | [21 empty] | 0 | [700 empty] | 0
P | 2018-07-11T10:00:04+08:00 | 2018-07-11T13:14:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP713G | CPSM Final Test | PASS | 129 | 165.11 | 0.08 | -0.46 | 56.13 | 49.95 | 1.21 | 0.13 | -1 | -1 | 0 | -0.852 | 0.543 | 8.4 | 6 |  |  | 0.0525 | 0.2008 | -0.0194 | 0.99999 | [21 empty] | 0 | [701 empty] | 0
P | 2018-07-11T10:26:30+08:00 | 2018-07-11T13:45:57+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP718G | CPSM Final Test | PASS | 129 | 167.45 | 0.27 | -1.15 | 55.94 | 50.21 | 1.44 | 0.13 | 0 | -2 | 1 | -0.8721 | 0.4896 | 2.3 | 15 |  |  | 0.1549 | 0.186 | -0.013 | 0.99999 | [21 empty] | 0 | [347 empty] | 0
P | 2018-07-11T03:40:39+08:00 | 2018-07-11T16:44:11+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP710G | CPSM Final Test | FAIL | [24 empty] | 0
P | 2018-07-11T10:12:25+08:00 | 2018-07-11T17:49:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP710G | CPSM Final Test | PASS | 129 | 161.06 | 0.27 | -0.84 | 56.15 | 50.29 | 1.36 | 0.13 | 1 | -3 | 0 | -0.844 | 0.532 | 4.2 | 18 |  |  | 0.0557 | 0.0623 | 0.0002 | 1 |  |  |  | 0 | [17 empty] | 0
P | 2018-07-11T13:17:12+08:00 | 2018-07-11T18:22:27+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP714G | CPSM Final Test | FAIL | [5 empty] | 44.62
P | 2018-07-11T13:36:49+08:00 | 2018-07-11T19:14:00+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP716G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-07-11T15:55:20+08:00 | 2018-07-11T19:15:59+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP719G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-07-11T12:04:08+08:00 | 2018-07-11T19:35:31+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP714G | CPSM Final Test | FAIL
P | 2018-07-11T16:01:32+08:00 | 2018-07-11T20:36:50+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP711G | CPSM Final Test | PASS | 129 | 167.99 | 0.08 | -0.54 | 56.31 | 48.09 | 0.83 | 0.13 | 0 | -2 | 0 | -0.9337 | 0.5422 | 3.8 | 9 |  |  | 0.0614 | 0.3264 | -0.0022 | 1 | [21 empty] | 0.01329 | [169 empty] | 0
P | 2018-07-11T16:02:52+08:00 | 2018-07-11T20:33:24+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP717G | CPSM Final Test | PASS | 129 | 161.69 | 0.37 | -0.92 | 56.14 | 50.63 | 1.29 | 0.13 | 0 | -2 | 1 | -0.866 | 0.5064 | 1.9 | -11 |  |  | 0.0357 | 0.1145 | -0.0552 | 0.99999 | [21 empty] | 0 | [689 empty] | 0
P | 2018-07-11T16:13:10+08:00 | 2018-07-11T20:36:36+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP729G | CPSM Final Test | PASS | 129 | 161.87 | 0.17 | -0.31 | 55.92 | 51.56 | 1.13 | 0.13 | 1 | -2 | 0 | -0.832 | 0.479 | -0.7 | 31 |  |  | 0.1294 | 0.371 | 0.0378 | 0.99997 | [21 empty] | 0 | [698 empty] | 0
P | 2018-07-11T17:35:39+08:00 | 2018-07-11T21:19:39+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP719G | CPSM Final Test | PASS | 129 | 161.87 | 0.46 | -0.46 | 56.44 | 52.75 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9358 | 0.5246 | -0.4 | -1 |  |  | 0.0672 | 0.2816 | -0.0434 | 0.99999 | [21 empty] | 0 | [264 empty] | 0
P | 2018-07-11T16:25:24+08:00 | 2018-07-11T21:57:38+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP716G | CPSM Final Test | PASS | 129 | 163.76 | 0.17 | -0.46 | 56.48 | 50.21 | 1.06 | 0.13 | 1 | -4 | 1 | -0.9638 | 0.537 | 8.4 | -2 |  |  | 0.062 | 0.0556 | -0.0196 | 0.99998 | [21 empty] | 0 | [702 empty] | 0
P | 2018-07-10T22:32:12+08:00 | 2018-07-12T09:07:27+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP714G | CPSM Final Test | FAIL
P | 2018-07-11T01:01:39+08:00 | 2018-07-12T09:54:17+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP715G | CPSM Final Test | FAIL | [19 empty] | -0.9505
P | 2018-07-11T15:49:49+08:00 | 2018-07-12T10:51:37+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP714G | CPSM Final Test | PASS | 129 | 165.02 | 0.08 | -0.08 | 57 | 53.51 | 1.74 | 0.13 | 0 | 1 | 0 | -0.9399 | 0.537 | 3.4 | 0 |  |  | 0.1727 | 0.8331 | -0.0471 | 1 | [21 empty] | 0 | [703 empty] | 0
P | 2018-07-11T17:27:58+08:00 | 2018-07-12T10:57:27+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP715G | CPSM Final Test | PASS | 129 | 155.48 | 0.17 | -0.84 | 56.35 | 49.45 | 0.91 | 0.13 | -1 | -2 | 0 | -0.8839 | 0.547 | 5.6 | 6 |  |  | 0.0921 | 0.156 | -0.0041 | 0.99999 | [21 empty] | 0.01552 | [347 empty] | 0
P | 2018-07-12T09:57:27+08:00 | 2018-07-12T14:20:00+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP736G | CPSM Final Test | FAIL | [13 empty] | -5.3
P | 2018-07-12T09:05:27+08:00 | 2018-07-12T15:12:00+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP736G | CPSM Final Test | FAIL | [14 empty] | 620
P | 2018-07-12T08:26:03+08:00 | 2018-07-12T15:32:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP734G | CPSM Final Test | FAIL | [747 empty] | 2631
P | 2018-07-12T12:49:11+08:00 | 2018-07-12T15:55:35+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP741G | CPSM Final Test | FAIL | [13 empty] | -7.4
P | 2018-07-12T12:21:09+08:00 | 2018-07-12T16:16:18+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP724G | CPSM Final Test | PASS | 129 | 164.57 | 0.17 | -0.31 | 56.23 | 49.78 | 1.06 | 0.13 | 0 | -2 | 0 | -0.9318 | 0.5048 | 1.8 | 29 |  |  | 0.139 | 0.3284 | -0.0093 | 0.99999 | [21 empty] | 0 | [317 empty] | 0
P | 2018-07-12T12:25:37+08:00 | 2018-07-12T16:13:41+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP743G | CPSM Final Test | PASS | 129 | 163.31 | 0.37 | -0.38 | 56.32 | 50.29 | 1.36 | 0.13 | -1 | -4 | 0 | -0.8479 | 0.486 | -2.4 | 11 |  |  | 0.1462 | 0.3023 | 0.0026 | 0.99999 | [21 empty] | 0 | [705 empty] | 0
P | 2018-07-12T12:31:06+08:00 | 2018-07-12T16:31:01+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP740G | CPSM Final Test | FAIL
P | 2018-07-12T07:19:43+08:00 | 2018-07-12T16:38:36+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP734G | CPSM Final Test | FAIL
P | 2018-07-12T11:59:09+08:00 | 2018-07-12T16:39:08+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP734G | CPSM Final Test | PASS | 129 | 165.11 | -0.02 | -0.46 | 56.42 | 53.42 | 1.59 | 0.13 | -1 | 2 | 1 | -0.8979 | 0.5488 | 5.1 | 7 |  |  | 0.1565 | 0.3532 | -0.0032 | 0.99997 | [21 empty] | 0 | [704 empty] | 0
P | 2018-07-12T12:08:44+08:00 | 2018-07-12T16:56:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP736G | CPSM Final Test | PASS | 129 | 166.73 | 0.17 | -1.07 | 56.18 | 49.36 | 1.44 | 0.13 | 1 | -4 | 0 | -0.7922 | 0.512 | 8.1 | 10 |  |  | 0.0651 | 0.304 | -0.0198 | 1 | [21 empty] | 0 | [706 empty] | 0
P | 2018-07-12T14:39:28+08:00 | 2018-07-12T17:25:04+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP721G | CPSM Final Test | PASS | 129 | 160.16 | -0.11 | -0.23 | 56.43 | 55.12 | 1.82 | 0.13 | -1 | 1 | 0 | -0.9298 | 0.536 | 3.7 | -20 |  |  | 0.1026 | 0.1831 | -0.0266 | 0.99999 | [21 empty] | 0 | [707 empty] | 0
P | 2018-07-12T14:31:04+08:00 | 2018-07-12T18:15:49+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP740G | CPSM Final Test | PASS | 129 | 161.78 | 0.56 | -0.92 | 56.17 | 49.95 | 1.06 | 0.13 | 1 | -2 | 1 | -0.864 | 0.5464 | 4.1 | -3 |  |  | 0.2597 | 0.6452 | 0.0109 | 0.99997 | [21 empty] | 0 | [708 empty] | 0
P | 2018-07-12T14:04:13+08:00 | 2018-07-12T18:36:27+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP737G | CPSM Final Test | FAIL | [14 empty] | 1307
P | 2018-07-12T18:56:44+08:00 | 2018-07-12T19:31:35+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP730G | CPSM Final Test | FAIL | [5 empty] | 46.65
P | 2018-07-12T14:13:06+08:00 | 2018-07-12T20:03:17+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP741G | CPSM Final Test | FAIL
P | 2018-07-12T12:36:49+08:00 | 2018-07-12T20:03:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP737G | CPSM Final Test | FAIL
P | 2018-07-12T17:08:12+08:00 | 2018-07-12T21:47:45+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | F18FTP741G | CPSM Final Test | PASS | 129 | 161.24 | 0.37 | -0.61 | 56.54 | 51.82 | 1.13 | 0.13 | -1 | -3 | 0 | -0.9957 | 0.4906 | 5.1 | 8 |  |  | 0.1121 | 0.1667 | -0.0227 | 1 | [21 empty] | 0 | [709 empty] | 0
P | 2018-07-12T16:20:20+08:00 | 2018-07-12T21:48:21+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP737G | CPSM Final Test | PASS | 129 | 154.58 | 0.17 | -0.77 | 56.21 | 50.29 | 0.98 | 0.13 | 1 | -2 | 0 | -0.91 | 0.511 | -2.8 | 32 |  |  | 0.0608 | 0.069 | 0.0005 | 1 |  |  |  | 0 | [17 empty] | 0.01027
P | 2018-07-12T05:22:12+08:00 | 2018-07-13T08:53:38+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP731G | CPSM Final Test | FAIL | [78 empty] | 0
P | 2018-07-12T04:50:44+08:00 | 2018-07-13T09:37:35+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP730G | CPSM Final Test | FAIL
P | 2018-07-12T19:07:55+08:00 | 2018-07-13T09:59:06+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP731G | CPSM Final Test | PASS | 129 | 163.31 | 0.27 | -0.84 | 56.13 | 50.72 | 1.06 | 0.13 | 1 | 2 | 0 | -0.9578 | 0.522 | -0.5 | 18 |  |  | 0.0824 | 0.2425 | 0.0037 | 0.99997 | [21 empty] | 0 | [35 empty] | 0
P | 2018-07-13T10:48:48+08:00 | 2018-07-13T13:45:08+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP723G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-07-13T13:45:48+08:00 | 2018-07-13T14:57:34+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP747G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-07-13T09:46:13+08:00 | 2018-07-13T15:20:41+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP704G | CPSM Final Test | FAIL | [13 empty] | 15.7
P | 2018-07-13T12:06:44+08:00 | 2018-07-13T16:36:38+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP747G | CPSM Final Test | FAIL | [13 empty] | -10.2
P | 2018-07-13T13:47:50+08:00 | 2018-07-13T16:54:25+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP728G | CPSM Final Test | PASS | 129 | 169.88 | 0.17 | -0.77 | 56.13 | 50.55 | 0.83 | 0.13 | 0 | 1 | 0 | -0.848 | 0.5004 | -0.3 | 4 |  |  | 0.0341 | 0.1776 | 0.016 | 0.99996 | [21 empty] | 0.01187 | [418 empty] | 0
P | 2018-07-13T12:16:58+08:00 | 2018-07-13T16:58:48+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP723G | CPSM Final Test | PASS | 129 | 155.75 | 0.08 | -0.92 | 56.05 | 49.7 | 1.21 | 0.13 | -2 | -2 | 0 | -0.888 | 0.5336 | -1.2 | 5 |  |  | 0.0746 | 0.1135 | -0.0096 | 0.99999 | [21 empty] | 0 | [492 empty] | 0
P | 2018-07-13T12:33:27+08:00 | 2018-07-13T17:20:59+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP704G | CPSM Final Test | PASS | 129 | 159.71 | 0.17 | -0.77 | 56.23 | 49.19 | 0.75 | 0.13 | -1 | -4 | 0 | -0.9179 | 0.4664 | 6.5 | 15 |  |  | 0.106 | 0.2373 | -0.0059 | 1 | [21 empty] | 0.01177 | [710 empty] | 0
P | 2018-07-13T14:31:43+08:00 | 2018-07-13T17:54:18+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP708G | CPSM Final Test | PASS | 129 | 154.76 | 0.17 | -1.15 | 56.21 | 50.29 | 1.21 | 0.13 | 1 | -2 | 0 | -0.8481 | 0.5362 | -3.3 | 18 |  |  | 0.0759 | 0.2499 | -0.0165 | 0.99999 | [21 empty] | 0.01212 | [696 empty] | 0
P | 2018-07-11T20:28:54+08:00 | 2018-07-13T17:59:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP730G | CPSM Final Test | FAIL | [754 empty] | 0
P | 2018-07-13T14:43:50+08:00 | 2018-07-13T18:00:44+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP707G | CPSM Final Test | PASS | 129 | 160.25 | 0.27 | -1.15 | 56.33 | 48.77 | 1.51 | 0.13 | 0 | -2 | 0 | -0.7742 | 0.5326 | 4.4 | 25 |  |  | 0.0754 | 0.4145 | 0.0253 | 0.99991 | [21 empty] | 0 | [704 empty] | 0
P | 2018-07-13T14:21:41+08:00 | 2018-07-13T18:36:50+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP747G | CPSM Final Test | PASS | 129 | 162.5 | 0.37 | -0.69 | 56.18 | 50.55 | 1.21 | 0.13 | -1 | 1 | -1 | -0.902 | 0.518 | 7.4 | 16 |  |  | 0.0734 | 0.1851 | 0.0048 | 0.99997 | [21 empty] | 0 | [35 empty] | 0
P | 2018-07-12T19:14:09+08:00 | 2018-07-13T19:05:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP730G | CPSM Final Test | PASS | 129 | 159.71 | 0.37 | -0.61 | 56.18 | 48.94 | 1.29 | 0.13 | -2 | 1 | -1 | -0.9617 | 0.4956 | -1.7 | -2 |  |  | 0.1717 | 0.2794 | -0.0808 | 0.99999 | [21 empty] | 0 | [711 empty] | 0
P | 2018-07-14T07:39:07+08:00 | 2018-07-14T10:21:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP739G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-07-14T08:48:15+08:00 | 2018-07-14T11:51:32+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP725G | CPSM Final Test | PASS | 129 | 161.42 | -0.11 | -0.46 | 56.42 | 54.27 | 1.67 | 0.13 | 1 | -2 | 0 | -0.8021 | 0.4978 | 2.3 | 25 |  |  | 0.1495 | 0.2749 | 0.0316 | 0.99995 | [21 empty] | 0 | [712 empty] | 0
P | 2018-07-14T04:10:01+08:00 | 2018-07-14T13:50:10+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP739G | CPSM Final Test | FAIL | [42 empty] | 0.12283
P | 2018-07-14T12:51:43+08:00 | 2018-07-14T15:10:53+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP735G | CPSM Final Test | FAIL | [5 empty] | 44.96
P | 2018-07-14T13:59:07+08:00 | 2018-07-14T15:53:27+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP705G | CPSM Final Test | FAIL | [83 empty] | FALSE
P | 2018-07-14T13:53:52+08:00 | 2018-07-14T16:56:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | F18FTP727G | CPSM Final Test | PASS | 129 | 166.19 | 0.27 | -0.31 | 56.48 | 50.29 | 1.67 | 0.13 | 1 | 4 | 0 | -0.852 | 0.504 | 0.7 | 12 |  |  | 0.0462 | 0.2383 | -0.0274 | 1 | [21 empty] | 0 | [711 empty] | 0
P | 2018-07-14T01:03:07+08:00 | 2018-07-14T16:57:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP739G | CPSM Final Test | FAIL | [42 empty] | 0.07694
P | 2018-07-14T13:37:39+08:00 | 2018-07-14T17:00:42+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | F18FTP738G | CPSM Final Test | PASS | 129 | 167 | 0.27 | -0.38 | 56.62 | 49.02 | 1.51 | 0.13 | 0 | -3 | 0 | -0.9537 | 0.4798 | 0.8 | 13 |  |  | 0.0561 | 0.4727 | 0.0296 | 0.99995 | [21 empty] | 0 | [713 empty] | 0
P | 2018-07-14T14:24:24+08:00 | 2018-07-14T17:16:51+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | F18FTP720G | CPSM Final Test | PASS | 129 | 168.17 | 0.17 | -0.77 | 56.16 | 51.56 | 1.44 | 0.3 | 1 | -2 | -1 | -0.8421 | 0.528 | -4.8 | 10 |  |  | 0.1949 | 0.3853 | -0.0123 | 0.99997 | [21 empty] | 0.01396 | [6 empty] | 0
P | 2018-07-14T14:01:18+08:00 | 2018-07-14T17:50:07+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | F18FTP735G | CPSM Final Test | PASS | 129 | 165.56 | 0.75 | -0.99 | 56.07 | 48.18 | 0.98 | 0.13 | -2 | -2 | 0 | -0.9059 | 0.524 | 8.8 | 9 |  |  | 0.0741 | 0.0658 | 0.0044 | 0.99999 | [21 empty] | 0.00431 | [714 empty] | 0
P | 2018-07-14T14:40:00+08:00 | 2018-07-14T17:53:08+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | F18FTP709G | CPSM Final Test | PASS | 129 | 163.94 | -0.02 | -0.77 | 56.13 | 48.68 | 1.44 | 0.13 | 1 | -2 | -1 | -0.838 | 0.4874 | 5.3 | -5 |  |  | 0.0313 | 0.0797 | -0.0462 | 1 | [21 empty] | 0 | [715 empty] | 0
P | 2018-07-14T09:00:06+08:00 | 2018-07-14T18:00:52+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | F18FTP739G | CPSM Final Test | PASS | 129 | 161.69 | 0.17 | -0.77 | 56.2 | 50.46 | 1.44 | 0.13 | -1 | -2 | 0 | -0.9757 | 0.5088 | 2.3 | 91 |  |  | 0.0494 | 0.09 | -0.0124 | 1 | [21 empty] | 0 | [716 empty] | 0
P | 2018-07-12T18:55:21+08:00 | 2018-07-16T10:57:13+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP705G | CPSM Final Test | FAIL | [760 empty] | 0
P | 2018-07-14T14:56:17+08:00 | 2018-07-16T12:03:03+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | F18FTP705G | CPSM Final Test | PASS | 129 | 153.77 | 0.17 | -1.15 | 56 | 49.45 | 1.06 | 0.13 | 1 | -2 | 1 | -0.9418 | 0.534 | -0.4 | -5 |  |  | 0.0747 | 0.3143 | -0.0257 | 0.99995 | [21 empty] | 0 | [717 empty] | 0
P | 2018-07-16T13:33:42+08:00 | 2018-07-16T18:19:21+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP537G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-07-16T13:45:09+08:00 | 2018-07-16T18:19:59+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP538G | CPSM Final Test | FAIL
P | 2018-07-16T16:02:34+08:00 | 2018-07-17T10:03:49+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP538G | CPSM Final Test | PASS |  | 162.41 | 0.17 | -0.61 | 56.12 | 50.8 | 1.21 | 0.13 | 0 | -4 | 0 | -0.9397 | 0.5 | 2.8 | 6 |  |  | 0.0808 | 0.2361 | -0.0274 | 1 |  |  |  |  | 133 |  | 0 | [733 empty] | 0
P | 2018-07-16T17:08:51+08:00 | 2018-07-17T10:03:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | G18FNP539G | CPSM Final Test | PASS |  | 167.72 | 0.17 | -0.69 | 56.43 | 54.69 | 1.29 | 0.13 | -1 | 1 | 1 | -0.9896 | 0.51 | 9.2 | 6 |  |  | 0.0813 | 0.3059 | -0.0135 | 0.99998 |  |  |  |  | 133 |  | 0 | [38 empty] | 0
P | 2018-07-16T15:56:31+08:00 | 2018-07-17T10:04:01+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP537G | CPSM Final Test | PASS |  | 164.12 | 0.27 | -0.69 | 56.6 | 48.18 | 1.36 | 0.13 | 0 | 1 | 0 | -0.9179 | 0.512 | 6.6 | 30 |  |  | 0.0708 | 0.2964 | -0.0187 | 0.99999 |  |  |  |  | 133 |  | 0 | [676 empty] | 0
P | 2018-07-20T16:37:47+08:00 | 2018-07-20T18:44:37+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP547G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-07-20T18:06:42+08:00 | 2018-07-20T19:02:27+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | G18FNP543G | CPSM Final Test | FAIL | [5 empty] | 46.31
P | 2018-07-20T14:23:22+08:00 | 2018-07-20T20:34:20+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | G18FNP544G | CPSM Final Test | FAIL
P | 2018-07-20T17:36:23+08:00 | 2018-07-20T21:19:36+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP545G | CPSM Final Test | PASS |  | 163.04 | 0.17 | -0.46 | 56.53 | 51.22 | 0.75 | 0.14 | 1 | -2 | 1 | -0.9279 | 0.496 | 7.2 | 23 |  |  | 0.0407 | 0.3271 | -0.0216 | 0.99999 |  |  |  |  | 133 |  | 0 | [734 empty] | 0
P | 2018-07-20T18:25:42+08:00 | 2018-07-20T21:38:33+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | G18FNP540G | CPSM Final Test | PASS |  | 163.67 | -0.02 | -0.61 | 56.48 | 52.07 | 1.44 | 0.13 | 1 | 1 | 0 | -0.9198 | 0.5148 | 7.9 | 63 |  |  | 0.056 | 0.3675 | -0.0272 | 0.99999 |  |  |  |  | 133 |  | 0 | [732 empty] | 0
P | 2018-07-20T18:25:30+08:00 | 2018-07-20T21:34:08+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP541G | CPSM Final Test | PASS |  | 162.77 | -0.02 | -0.16 | 56.52 | 50.21 | 1.67 | 0.13 | 1 | 2 | 0 | -0.9358 | 0.458 | 5.9 | 31 |  |  | 0.2289 | 0.1872 | -0.0243 | 0.99999 |  |  |  |  | 133 |  | 0 | [735 empty] | 0
P | 2018-07-20T17:41:12+08:00 | 2018-07-20T21:36:41+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP547G | CPSM Final Test | PASS |  | 156.2 | -0.02 | -0.46 | 56.57 | 49.02 | 0.98 | 0.13 | -2 | -1 | 0 | -0.9439 | 0.5436 | 10.3 | -6 |  |  | 0.108 | 0.1627 | -0.0045 | 0.99999 |  |  |  |  | 133 |  | 0 | [432 empty] | 0
P | 2018-07-20T17:28:51+08:00 | 2018-07-20T22:22:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | G18FNP544G | CPSM Final Test | PASS |  | 162.95 | -0.02 | -0.46 | 56.42 | 56.39 | 1.59 | 0.13 | -1 | -3 | -1 | -0.9458 | 0.505 | 4.6 | 19 |  |  | 0.2245 | 0.2246 | -0.0041 | 1 |  |  |  |  | 133 |  | 0 | [736 empty] | 0
P | 2018-07-20T02:48:00+08:00 | 2018-07-21T08:16:59+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP542G | CPSM Final Test | FAIL
P | 2018-07-20T17:32:29+08:00 | 2018-07-21T08:17:53+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP542G | CPSM Final Test | PASS |  | 162.41 | -0.02 | -0.23 | 56.52 | 50.38 | 1.21 | 0.13 | -1 | -2 | 0 | -0.87 | 0.505 | 8.4 | 34 |  |  | 0.116 | 0.4767 | 0.0333 | 0.99999 |  |  |  |  | 133 |  | 0 | [737 empty] | 0
P | 2018-07-20T03:43:14+08:00 | 2018-07-21T09:25:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | G18FNP543G | CPSM Final Test | FAIL
P | 2018-07-20T02:37:25+08:00 | 2018-07-21T10:20:56+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | G18FNP546G | CPSM Final Test | FAIL
P | 2018-07-20T02:30:27+08:00 | 2018-07-21T10:27:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | G18FNP546G | CPSM Final Test | FAIL | [27 empty] | 0
P | 2018-07-20T18:34:34+08:00 | 2018-07-21T11:14:53+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | G18FNP543G | CPSM Final Test | PASS |  | 160.61 | 0.17 | -0.38 | 56.41 | 49.19 | 1.13 | 0.13 | -2 | -2 | -1 | -0.8121 | 0.5192 | 4.4 | -4 |  |  | 0.1179 | 0.4299 | -0.0071 | 0.99995 |  |  |  |  | 133 |  | 0 | [738 empty] | 0
P | 2018-07-21T07:10:57+08:00 | 2018-07-21T11:56:16+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP549G | CPSM Final Test | FAIL | [767 empty] | 0
P | 2018-07-21T09:46:49+08:00 | 2018-07-21T12:41:13+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | G18FNP548G | CPSM Final Test | PASS |  | 159.44 | -0.02 | -0.61 | 56.31 | 48.01 | 1.13 | 0.13 | 1 | 1 | 0 | -0.9118 | 0.5122 | -3.7 | 15 |  |  | 0.0773 | 0.2984 | -0.002 | 1 |  |  |  |  | 133 |  | 0 | [740 empty] | 0
P | 2018-07-21T09:29:41+08:00 | 2018-07-21T12:36:07+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP775G | CPSM Final Test | PASS | 129 | 162.41 | -0.02 | -0.54 | 56.43 | 53.51 | 1.21 | 0.13 | 1 | 1 | 0 | -0.7941 | 0.4916 | 0.9 | 0 |  |  | 0.1437 | 0.3384 | -0.0681 | 1 | [21 empty] | 0 | [101 empty] | 0
P | 2018-07-21T09:16:43+08:00 | 2018-07-21T12:39:15+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP778G | CPSM Final Test | PASS | 129 | 160.88 | 0.17 | -0.77 | 56.47 | 50.29 | 0.91 | 0.13 | 1 | 2 | -1 | -0.8998 | 0.5282 | 14.6 | 16 |  |  | 0.1503 | 0.1053 | -0.0095 | 0.99998 | [21 empty] | 0 | [544 empty] | 0
P | 2018-07-21T09:24:01+08:00 | 2018-07-21T12:40:12+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP550G | CPSM Final Test | PASS |  | 159.89 | 0.46 | -0.46 | 56.46 | 50.46 | 1.13 | 0.13 | -1 | -3 | 0 | -0.854 | 0.4864 | -4.1 | 16 |  |  | 0.0454 | 0.2531 | -0.0168 | 0.99999 |  |  |  |  | 133 |  | 0 | [370 empty] | 0
P | 2018-07-21T09:20:17+08:00 | 2018-07-21T12:42:20+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP784G | CPSM Final Test | PASS | 129 | 161.78 | -0.02 | -0.69 | 56.19 | 51.65 | 0.83 | 0.13 | 0 | 1 | -1 | -0.8819 | 0.5018 | 13 | 5 |  |  | 0.1611 | 0.3548 | -0.0204 | 0.99999 | [21 empty] | 0 | [726 empty] | 0
P | 2018-07-20T00:13:03+08:00 | 2018-07-21T12:45:18+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | G18FNP546G | CPSM Final Test | FAIL
P | 2018-07-21T09:33:37+08:00 | 2018-07-21T13:02:46+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP549G | CPSM Final Test | PASS |  | 163.85 | 0.37 | -0.31 | 56.49 | 50.29 | 1.13 | 0.13 | 0 | -4 | 0 | -0.9018 | 0.538 | 2.6 | 15 |  |  | 0.1173 | 0.5142 | -0.0271 | 0.99999 |  |  |  |  | 133 |  | 0 | [739 empty] | 0
P | 2018-07-20T18:29:10+08:00 | 2018-07-21T13:48:30+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | G18FNP546G | CPSM Final Test | PASS |  | 170.96 | 0.17 | -0.54 | 56.23 | 50.88 | 0.83 | 0.13 | -1 | -2 | -1 | -1.0077 | 0.5054 | 3.6 | -10 |  |  | 0.0714 | 0.131 | 0.005 | 0.99999 |  |  |  |  | 133 |  | 0 | [742 empty] | 0
P | 2018-07-21T12:13:12+08:00 | 2018-07-21T16:47:46+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | G18FNP551G | CPSM Final Test | FAIL | [13 empty] | 16.3
P | 2018-07-21T14:56:57+08:00 | 2018-07-21T18:07:17+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP789G | CPSM Final Test | PASS | 129 | 162.77 | -0.02 | -0.38 | 56.52 | 48.94 | 1.13 | 0.13 | 0 | -4 | 0 | -0.864 | 0.5252 | 1.6 | -9 |  |  | 0.2214 | 0.1151 | 0.0648 | 0.99992 | [21 empty] | 0 | [15 empty] | 0
P | 2018-07-21T14:52:37+08:00 | 2018-07-21T18:10:06+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP761G | CPSM Final Test | PASS | 129 | 166.73 | 0.27 | -0.61 | 56.22 | 50.04 | 1.59 | 0.13 | 0 | 1 | 0 | -0.8839 | 0.506 | -2.1 | -1 |  |  | 0.0649 | 0.1379 | -0.0124 | 1 | [21 empty] | 0 | [728 empty] | 0
P | 2018-07-21T14:54:38+08:00 | 2018-07-21T18:17:47+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP788G | CPSM Final Test | PASS | 129 | 165.02 | 0.27 | -0.31 | 56.72 | 55.2 | 1.29 | 0.13 | 1 | -4 | 0 | -0.9058 | 0.505 | -3.1 | 2 |  |  | 0.1031 | 0.1381 | 0.0143 | 0.99998 | [21 empty] | 0 | [729 empty] | 0
P | 2018-07-21T14:41:29+08:00 | 2018-07-21T18:15:51+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP786G | CPSM Final Test | PASS | 129 | 157.91 | 0.27 | -0.54 | 56.64 | 49.87 | 0.91 | 0.13 | -1 | 1 | -1 | -0.9 | 0.521 | 2.7 | 6 |  |  | 0.0618 | 0.2337 | -0.0332 | 0.99997 | [21 empty] | 0 | [247 empty] | 0
P | 2018-07-21T14:37:19+08:00 | 2018-07-21T18:25:14+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | G18FNP553G | CPSM Final Test | PASS |  | 164.3 | 0.27 | -0.16 | 56.22 | 49.78 | 1.74 | 0.13 | -1 | -1 | 0 | -0.9158 | 0.509 | 3.3 | -10 |  |  | 0.1522 | 0.076 | -0.0507 | 1 |  |  |  |  | 133 |  | 0 | [745 empty] | 0
P | 2018-07-21T14:33:48+08:00 | 2018-07-21T18:28:13+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | G18FNP552G | CPSM Final Test | PASS |  | 162.05 | 0.59 | -0.61 | 56.23 | 48.85 | 1.44 | 0.13 | 1 | -2 | 1 | -0.9119 | 0.496 | 3.8 | 15 |  |  | 0.1512 | 0.3497 | -0.0034 | 1 |  |  |  |  | 133 |  | 0 | [493 empty] | 0
P | 2018-07-19T18:59:01+08:00 | 2018-07-23T10:30:10+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP762G | CPSM Final Test | FAIL | [767 empty] | 0
P | 2018-07-21T14:44:35+08:00 | 2018-07-23T11:35:50+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP762G | CPSM Final Test | PASS | 129 | 170.24 | 0.27 | -0.69 | 56.19 | 51.05 | 1.21 | 0.13 | 1 | 1 | 0 | -0.9259 | 0.523 | 13 | -1 |  |  | 0.0671 | 0.208 | -0.0454 | 0.99998 | [21 empty] | 0 | [724 empty] | 0
P | 2018-07-23T10:03:26+08:00 | 2018-07-23T13:57:18+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP780G | CPSM Final Test | FAIL
P | 2018-07-23T11:28:48+08:00 | 2018-07-23T15:23:51+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP770G | CPSM Final Test | PASS | 129 | 167.36 | -0.02 | -0.77 | 56.43 | 55.03 | 1.29 | 0.13 | 0 | 1 | 0 | -0.86 | 0.4986 | 9 | 25 |  |  | 0.0814 | 0.2336 | -0.0185 | 0.99999 | [21 empty] | 0 | [227 empty] | 0
P | 2018-07-23T11:41:04+08:00 | 2018-07-23T15:33:49+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP781G | CPSM Final Test | PASS | 129 | 166.28 | -0.02 | -0.54 | 56.32 | 50.8 | 1.13 | 0.13 | -1 | 1 | 0 | -0.888 | 0.5208 | 3.9 | 3 |  |  | 0.1189 | 0.402 | 0.0144 | 0.99998 | [21 empty] | 0 | [731 empty] | 0
P | 2018-07-23T08:00:47+08:00 | 2018-07-23T15:30:02+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP752G | CPSM Final Test | FAIL
P | 2018-07-23T08:08:23+08:00 | 2018-07-23T15:30:55+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP776G | CPSM Final Test | FAIL | [637 empty] | 5271
P | 2018-07-23T12:00:22+08:00 | 2018-07-23T15:41:48+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP780G | CPSM Final Test | PASS | 129 | 158.09 | 0.37 | -0.38 | 56.04 | 48.85 | 0.91 | 0.13 | -1 | 1 | 0 | -0.8659 | 0.5284 | 1.8 | -6 |  |  | 0.1057 | 0.1354 | -0.0146 | 1 | [21 empty] | 0 | [732 empty] | 0
P | 2018-07-23T14:14:06+08:00 | 2018-07-23T15:43:04+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP554G | CPSM Final Test | FAIL
P | 2018-07-23T11:45:24+08:00 | 2018-07-23T16:36:11+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP752G | CPSM Final Test | PASS | 129 | 156.02 | -0.02 | -0.46 | 56.52 | 48.43 | 1.36 | 0.13 | 0 | -3 | 0 | -0.886 | 0.4972 | 4.4 | 47 |  |  | 0.0967 | 0.2116 | -0.0414 | 0.99994 | [21 empty] | 0 | [95 empty] | 0
P | 2018-07-23T11:49:39+08:00 | 2018-07-23T16:36:33+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP776G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.31 | 56.16 | 50.88 | 1.29 | 0.13 | 0 | -1 | 0 | -0.8699 | 0.512 | 8.8 | 3 |  |  | 0.1827 | 0.5441 | 0.0276 | 0.99999 | [21 empty] | 0 | [594 empty] | 0
P | 2018-07-23T12:25:24+08:00 | 2018-07-23T17:31:46+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP554G | CPSM Final Test | FAIL
P | 2018-07-23T15:30:04+08:00 | 2018-07-23T17:33:49+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP551G | CPSM Final Test | FAIL | [13 empty] | -8.8
P | 2018-07-23T12:10:12+08:00 | 2018-07-23T17:46:58+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP554G | CPSM Final Test | FAIL
P | 2018-07-23T15:43:24+08:00 | 2018-07-23T18:26:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP777G | CPSM Final Test | PASS | 129 | 166.28 | 0.37 | -0.61 | 56.41 | 53.42 | 1.29 | 0.13 | -1 | 3 | 0 | -0.9478 | 0.5374 | 7.6 | -12 |  |  | 0.1435 | 0.6258 | -0.0048 | 0.99992 | [21 empty] | 0 | [733 empty] | 0
P | 2018-07-23T04:06:06+08:00 | 2018-07-23T18:59:29+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP787G | CPSM Final Test | FAIL | [14 empty] | 1203
P | 2018-07-23T10:45:46+08:00 | 2018-07-23T19:11:24+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP554G | CPSM Final Test | FAIL
P | 2018-07-23T12:45:02+08:00 | 2018-07-23T20:18:51+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP551G | CPSM Final Test | FAIL
P | 2018-07-23T14:58:35+08:00 | 2018-07-23T20:24:36+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | G18FNP554G | CPSM Final Test | PASS |  | 160.79 | 0.27 | 0.07 | 56.71 | 55.29 | 1.9 | 0.13 | -1 | -2 | -1 | -0.9477 | 0.4978 | 9.2 | 4 |  |  | 0.0832 | 0.2122 | -0.0451 | 0.99999 |  |  |  |  | 133 |  | 0 | [37 empty] | 0
P | 2018-07-23T11:32:48+08:00 | 2018-07-23T20:46:15+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP787G | CPSM Final Test | PASS | 129 | 166.19 | -0.02 | -0.54 | 56.49 | 51.73 | 1.82 | 0.13 | -1 | 1 | 0 | -0.8121 | 0.5066 | 10.2 | 19 |  |  | 0.1356 | 0.2419 | 0.0319 | 0.99992 | [21 empty] | 0.00565 | [734 empty] | 0
P | 2018-07-24T09:02:18+08:00 | 2018-07-24T09:48:41+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP750G | CPSM Final Test | FAIL | [348 empty] | FALSE
P | 2018-07-23T06:55:50+08:00 | 2018-07-24T09:59:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP751G | CPSM Final Test | FAIL | [5 empty] | 43.52
P | 2018-07-23T06:04:26+08:00 | 2018-07-24T10:51:03+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP751G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-07-24T08:01:21+08:00 | 2018-07-24T10:51:19+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP766G | CPSM Final Test | FAIL | [13 empty] | 21.4
P | 2018-07-22T21:32:42+08:00 | 2018-07-24T11:31:11+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP551G | CPSM Final Test | FAIL
P | 2018-07-23T20:13:58+08:00 | 2018-07-24T11:55:39+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP749G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -0.08 | 56.73 | 55.12 | 1.13 | 0.13 | -1 | 2 | 0 | -0.852 | 0.5312 | 1.8 | 15 |  |  | 0.4376 | 0.4665 | 0.0223 | 0.99999 | [21 empty] | 0 | [594 empty] | 0
P | 2018-07-23T20:10:42+08:00 | 2018-07-24T11:57:01+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP759G | CPSM Final Test | PASS | 129 | 167.45 | 0.65 | -0.31 | 56.63 | 48.68 | 1.51 | 0.13 | 0 | 1 | -1 | -0.898 | 0.5202 | 5 | 27 |  |  | 0.1365 | 0.1192 | 0.0005 | 0.99999 | [21 empty] | 0 | [95 empty] | 0
P | 2018-07-24T09:25:30+08:00 | 2018-07-24T12:13:33+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP750G | CPSM Final Test | PASS | 129 | 167.9 | 0.17 | -0.69 | 56.32 | 49.95 | 1.44 | 0.13 | 0 | -2 | 0 | -0.7761 | 0.514 | 5.8 | 10 |  |  | 0.1195 | 0.3064 | 0.051 | 0.99993 | [21 empty] | 0 | [735 empty] | 0
P | 2018-07-23T04:36:57+08:00 | 2018-07-24T12:14:31+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP763G | CPSM Final Test | FAIL
P | 2018-07-23T20:25:44+08:00 | 2018-07-24T12:15:01+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP763G | CPSM Final Test | PASS | 129 | 160.97 | -0.02 | -0.46 | 56.23 | 50.21 | 1.06 | 0.13 | -1 | -2 | 0 | -0.848 | 0.521 | 0 | -85 |  |  | 0.0744 | 0.1146 | -0.0226 | 1 | [21 empty] | 0 | [732 empty] | 0
P | 2018-07-23T16:31:57+08:00 | 2018-07-24T12:34:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | G18FNP551G | CPSM Final Test | PASS |  | 169.7 | 0.37 | 0.38 | 56.29 | 51.14 | 1.21 | 0.13 | -1 | 1 | 0 | -0.9678 | 0.4886 | 9.3 | 4 |  |  | 0.054 | 0.2638 | -0.0683 | 1 |  |  |  |  | 133 |  | 0 | [213 empty] | 0
P | 2018-07-24T09:26:20+08:00 | 2018-07-24T12:55:11+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP766G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -0.46 | 56.45 | 50.97 | 1.13 | 0.13 | 1 | -2 | -1 | -0.8939 | 0.482 | 2.6 | 5 |  |  | 0.08 | 0.2641 | -0.0205 | 1 | [21 empty] | 0 | [22 empty] | 0
P | 2018-07-23T03:20:42+08:00 | 2018-07-24T13:34:47+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP751G | CPSM Final Test | FAIL
P | 2018-07-23T20:27:45+08:00 | 2018-07-24T13:35:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP751G | CPSM Final Test | PASS | 129 | 163.13 | 0.27 | -0.61 | 56.01 | 49.19 | 1.59 | 0.13 | -1 | 1 | 0 | -0.8439 | 0.5048 | 5 | 20 |  |  | 0.1635 | 0.352 | -0.0038 | 0.99998 | [21 empty] | 0 | [736 empty] | 0
P | 2018-07-24T14:37:46+08:00 | 2018-07-24T17:37:51+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP792G | CPSM Final Test | PASS | 129 | 158.72 | 0.37 | -0.54 | 56.41 | 53.68 | 0.98 | 0.13 | 0 | -2 | 0 | -0.9278 | 0.5468 | 8.2 | 15 |  |  | 0.0659 | 0.2404 | 0.0206 | 0.99997 | [21 empty] | 0 | [737 empty] | 0
P | 2018-07-24T14:24:21+08:00 | 2018-07-24T17:33:05+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP794G | CPSM Final Test | PASS | 129 | 151.7 | 0.46 | -0.31 | 56.32 | 51.14 | 1.13 | 0.13 | 1 | 2 | 0 | -0.8021 | 0.5084 | 6 | 12 |  |  | 0.0898 | 0.1679 | -0.0231 | 1 | [21 empty] | 0 | [198 empty] | 0
P | 2018-07-24T14:28:01+08:00 | 2018-07-24T17:50:43+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP791G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.16 | 56.18 | 51.31 | 1.29 | 0.13 | 1 | -2 | 0 | -0.8959 | 0.4828 | 2.5 | 16 |  |  | 0.242 | 0.4716 | -0.0844 | 1 | [21 empty] | 0 | [738 empty] | 0
P | 2018-07-24T13:41:13+08:00 | 2018-07-24T17:56:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP769G | CPSM Final Test | PASS | 129 | 167.72 | 0.27 | -0.46 | 56.4 | 55.88 | 1.36 | 0.13 | -1 | 1 | -1 | -0.872 | 0.505 | 12.8 | -21 |  |  | 0.0283 | 0.0601 | -0.0037 | 1 | [21 empty] | 0 | [103 empty] | 0
P | 2018-07-24T14:45:34+08:00 | 2018-07-24T18:03:35+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP779G | CPSM Final Test | PASS | 129 | 165.56 | -0.02 | -0.16 | 56.71 | 55.03 | 1.29 | 0.13 | 0 | -1 | 0 | -0.8621 | 0.5182 | 5 | 15 |  |  | 0.0889 | 0.1672 | -0.0223 | 1 | [21 empty] | 0 | [739 empty] | 0
P | 2018-07-22T21:55:37+08:00 | 2018-07-24T18:19:01+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP760G | CPSM Final Test | FAIL
P | 2018-07-24T06:44:32+08:00 | 2018-07-24T18:25:05+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP756G | CPSM Final Test | FAIL | [200 empty] | 0
P | 2018-07-24T12:34:48+08:00 | 2018-07-24T19:30:47+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP756G | CPSM Final Test | PASS | 129 | 160.43 | -0.21 | -0.46 | 56.61 | 51.39 | 1.44 | 0.13 | 0 | -3 | 1 | -1.0335 | 0.538 | 7.2 | 13 |  |  | 0.1051 | 0.1026 | 0.0216 | 0.99999 | [21 empty] | 0 | [157 empty] | 0
P | 2018-07-23T20:07:19+08:00 | 2018-07-24T20:17:07+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP760G | CPSM Final Test | PASS | 129 | 160.43 | -0.02 | -0.31 | 56.33 | 49.36 | 1.29 | 0.13 | -1 | -1 | -1 | -0.834 | 0.512 | 8.4 | 3 |  |  | 0.0406 | 0.1655 | 0.0157 | 0.99999 | [21 empty] | 0 | [740 empty] | 0
P | 2018-07-24T17:46:47+08:00 | 2018-07-24T20:24:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP765G | CPSM Final Test | FAIL | [13 empty] | 26.7
P | 2018-07-24T16:58:24+08:00 | 2018-07-24T20:25:31+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP771G | CPSM Final Test | FAIL
P | 2018-07-24T19:03:59+08:00 | 2018-07-24T21:56:15+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP755G | CPSM Final Test | PASS | 129 | 163.76 | -0.02 | -0.69 | 56.41 | 55.03 | 1.44 | 0.13 | 0 | -3 | 1 | -0.924 | 0.511 | 0.2 | -22 |  |  | 0.0637 | 0.1431 | -0.005 | 1 | [21 empty] | 0 | [741 empty] | 0
P | 2018-07-24T18:41:58+08:00 | 2018-07-24T22:10:05+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP771G | CPSM Final Test | PASS | 129 | 168.8 | 0.27 | -0.92 | 56.47 | 51.05 | 0.98 | 0.13 | 1 | -4 | 1 | -0.9379 | 0.5372 | 6.6 | -22 |  |  | 0.0517 | 0.1349 | 0.017 | 1 | [21 empty] | 0 | [742 empty] | 0
P | 2018-07-24T19:05:43+08:00 | 2018-07-24T22:24:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP765G | CPSM Final Test | PASS | 129 | 159.98 | -0.02 | -0.84 | 56.4 | 55.88 | 1.06 | 0.13 | -1 | 1 | 0 | -0.9278 | 0.5384 | 13.3 | 4 |  |  | 0.0727 | 0.3109 | -0.0266 | 1 | [21 empty] | 0 | [381 empty] | 0
P | 2018-07-23T18:58:13+08:00 | 2018-07-25T08:33:27+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP773G | CPSM Final Test | FAIL | [14 empty] | -129
P | 2018-07-24T04:28:39+08:00 | 2018-07-25T09:12:13+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP782G | CPSM Final Test | FAIL | [786 empty] | 0
P | 2018-07-24T04:15:02+08:00 | 2018-07-25T09:19:09+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP785G | CPSM Final Test | FAIL | [41 empty] | 0
P | 2018-07-23T17:10:21+08:00 | 2018-07-25T10:21:19+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP773G | CPSM Final Test | FAIL
P | 2018-07-24T13:45:50+08:00 | 2018-07-25T10:21:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | G18FTP773G | CPSM Final Test | PASS | 129 | 165.29 | -0.02 | -0.77 | 56.42 | 54.19 | 1.13 | 0.13 | 1 | -2 | 1 | -0.874 | 0.534 | 13 | -78 |  |  | 0.1045 | 0.3826 | 0.0287 | 0.99995 | [21 empty] | 0 | [530 empty] | 0
P | 2018-07-24T18:47:05+08:00 | 2018-07-25T10:24:39+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP785G | CPSM Final Test | PASS | 129 | 160.34 | 0.46 | -0.84 | 56.42 | 55.96 | 0.98 | 0.13 | 1 | -2 | 1 | -0.9378 | 0.5184 | 4.3 | -6 |  |  | 0.0785 | 0.1682 | -0.024 | 1 | [20 empty] | 0 | 0
P | 2018-07-25T08:55:18+08:00 | 2018-07-25T11:27:33+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP782G | CPSM Final Test | FAIL | [13 empty] | 16.5
P | 2018-07-25T10:02:07+08:00 | 2018-07-25T12:52:51+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP753G | CPSM Final Test | PASS | 129 | 166.82 | 0.27 | 0.53 | 56.47 | 52.15 | 1.36 | 0.13 | 1 | 3 | 0 | -0.9099 | 0.5116 | 4.2 | -10 |  |  | 0.123 | 0.1846 | 0.0173 | 0.99999 | [21 empty] | 0.00578 | [744 empty] | 0
P | 2018-07-25T09:36:34+08:00 | 2018-07-25T12:53:03+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP757G | CPSM Final Test | PASS | 129 | 161.15 | -0.02 | -0.54 | 56.47 | 51.05 | 1.21 | 0.13 | 0 | -2 | 0 | -0.88 | 0.4708 | -0.3 | 26 |  |  | 0.0777 | 0.6073 | 0.0013 | 0.99999 | [21 empty] | 0 | [157 empty] | 0
P | 2018-07-25T09:47:52+08:00 | 2018-07-25T13:02:09+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP754G | CPSM Final Test | PASS | 129 | 165.65 | 0.27 | -0.08 | 56.53 | 48.51 | 1.9 | 0.13 | 1 | -4 | 0 | -0.8381 | 0.5028 | 3.6 | 14 |  |  | 0.1135 | 0.1128 | -0.0424 | 1 | [21 empty] | 0 | [82 empty] | 0
P | 2018-07-25T09:49:41+08:00 | 2018-07-25T13:07:45+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP764G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.46 | 56.51 | 51.48 | 1.21 | 0.13 | 0 | -1 | 0 | -0.8899 | 0.5272 | 10.6 | 2 |  |  | 0.1284 | 0.2663 | -0.0006 | 0.99999 | [21 empty] | 0 | [745 empty] | 0
P | 2018-07-25T10:11:26+08:00 | 2018-07-25T13:26:03+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP782G | CPSM Final Test | PASS | 129 | 165.47 | -0.21 | -0.54 | 56.23 | 52.83 | 0.98 | 0.13 | 1 | -3 | 0 | -0.9896 | 0.477 | 7.9 | -15 |  |  | 0.156 | 0.3652 | 0.0029 | 0.99999 | [21 empty] | 0 | [743 empty] | 0
P | 2018-07-25T10:43:09+08:00 | 2018-07-25T13:57:11+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP790G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-07-25T14:24:56+08:00 | 2018-07-25T15:22:51+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP758G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-07-25T14:17:01+08:00 | 2018-07-25T15:27:49+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP774G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-07-25T14:12:32+08:00 | 2018-07-25T15:36:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP783G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-07-25T14:09:04+08:00 | 2018-07-25T15:41:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP772G | CPSM Final Test | FAIL | [5 empty] | 46.57
P | 2018-07-25T13:40:42+08:00 | 2018-07-25T16:08:11+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP783G | CPSM Final Test | FAIL | [348 empty] | FALSE
P | 2018-07-25T12:20:10+08:00 | 2018-07-25T17:16:43+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP790G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -0.54 | 56.02 | 50.12 | 1.21 | 0.13 | 1 | 3 | 0 | -0.844 | 0.511 | 6.1 | 5 |  |  | 0.1772 | 0.277 | -0.01 | 0.99999 | [20 empty] | 0 | 0
P | 2018-07-25T09:23:29+08:00 | 2018-07-25T17:22:30+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP793G | CPSM Final Test | FAIL
P | 2018-07-25T11:23:47+08:00 | 2018-07-25T18:21:03+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP774G | CPSM Final Test | FAIL
P | 2018-07-25T14:55:13+08:00 | 2018-07-25T18:37:01+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | G18FTP772G | CPSM Final Test | PASS | 129 | 167.63 | 0.17 | -0.69 | 56.37 | 50.8 | 1.51 | 0.13 | 1 | -2 | -1 | -0.888 | 0.4872 | -0.6 | -9 |  |  | 0.0671 | 0.3616 | 0.032 | 0.99995 | [21 empty] | 0 | [724 empty] | 0
P | 2018-07-25T15:43:20+08:00 | 2018-07-25T18:41:18+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | G18FTP758G | CPSM Final Test | PASS | 129 | 161.42 | 0.37 | -0.38 | 56.5 | 49.95 | 1.21 | 0.13 | 0 | -2 | 0 | -0.892 | 0.4918 | 3.8 | 21 |  |  | 0.2036 | 0.2973 | 0.0018 | 0.99997 | [21 empty] | 0 | [746 empty] | 0
P | 2018-07-25T14:54:27+08:00 | 2018-07-25T18:48:11+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | G18FTP783G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -0.92 | 56.31 | 48.26 | 0.98 | 0.13 | 1 | -3 | 0 | -0.9198 | 0.491 | 4.5 | -5 |  |  | 0.1731 | 0.2052 | -0.0057 | 0.99997 | [21 empty] | 0 | [22 empty] | 0
P | 2018-07-25T14:46:16+08:00 | 2018-07-25T18:44:37+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | G18FTP768G | CPSM Final Test | PASS | 129 | 162.68 | -0.02 | -0.84 | 56.41 | 52.07 | 1.29 | 0.13 | 0 | -4 | 0 | -0.834 | 0.5108 | 7.1 | 18 |  |  | 0.0526 | 0.1375 | -0.0038 | 0.99997 | [21 empty] | 0 | [747 empty] | 0
P | 2018-07-25T19:10:10+08:00 | 2018-07-25T19:58:41+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP767G | CPSM Final Test | FAIL | [5 empty] | 46.57
P | 2018-07-25T14:52:25+08:00 | 2018-07-25T20:05:17+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | G18FTP774G | CPSM Final Test | PASS | 129 | 162.05 | 0.17 | -0.84 | 56.55 | 48.85 | 0.98 | 0.13 | 0 | 1 | 0 | -0.912 | 0.502 | 5 | -10 |  |  | 0.0545 | 0.2319 | -0.0018 | 1 | [21 empty] | 0 | [748 empty] | 0
P | 2018-07-25T06:34:55+08:00 | 2018-07-25T20:11:04+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP793G | CPSM Final Test | FAIL
P | 2018-07-25T13:23:00+08:00 | 2018-07-25T21:14:04+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | G18FTP793G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -0.31 | 56.41 | 51.22 | 1.06 | 0.13 | 0 | -1 | 1 | -0.856 | 0.5226 | 6.6 | -11 |  |  | 0.1518 | 0.2706 | -0.009 | 0.99999 | [21 empty] | 0 | [749 empty] | 0
P | 2018-07-25T06:46:18+08:00 | 2018-07-26T08:22:33+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP767G | CPSM Final Test | FAIL
P | 2018-07-25T19:34:25+08:00 | 2018-07-26T09:28:05+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | G18FTP767G | CPSM Final Test | PASS | 129 | 168.71 | -0.02 | -0.84 | 56.02 | 48.34 | 0.91 | 0.13 | 1 | -2 | -1 | -0.9439 | 0.5198 | 10.7 | 1 |  |  | 0.3158 | 0.4461 | 0.0332 | 1 | [21 empty] | 0 | [204 empty] | 0
P | 2018-07-26T13:05:48+08:00 | 2018-07-26T14:17:45+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | G18USM058G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-07-26T12:10:20+08:00 | 2018-07-26T15:13:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | G18USM058G | CPSM Final Test | FAIL | [13 empty] | -14.4
P | 2018-07-26T12:04:42+08:00 | 2018-07-26T15:14:18+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | G18USM054G | CPSM Final Test | FAIL | [13 empty] | -9.6
P | 2018-07-26T13:30:17+08:00 | 2018-07-26T16:33:35+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | G18USM051G | CPSM Final Test | PASS |  | 158.45 | 0.08 | -0.99 | 56.56 | 48.77 | 0.98 | 0.13 | -1 | 3 | 0 | -0.8122 | 0.5112 | 13 | 1 |  |  | 0.1503 | 0.2271 | 0.0461 | 0.99997 |  |  |  |  | 133 |  | 0 | [765 empty] | 0
P | 2018-07-26T13:35:00+08:00 | 2018-07-26T16:41:10+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | G18USM059G | CPSM Final Test | PASS |  | 162.68 | 0.08 | -0.61 | 56.86 | 49.53 | 1.21 | 0.13 | 0 | -1 | 0 | -0.9638 | 0.5108 | -3.2 | 10 |  |  | 0.4218 | 0.2709 | -0.0403 | 0.99999 |  |  |  |  | 133 |  | 0 | [612 empty] | 0
P | 2018-07-26T13:31:37+08:00 | 2018-07-26T16:43:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | G18USM056G | CPSM Final Test | PASS |  | 148.37 | 0.08 | -0.23 | 56.84 | 50.72 | 1.36 | 0.13 | 0 | 1 | 0 | -0.852 | 0.52 | 0.3 | 6 |  |  | 0.0591 | 0.0932 | 0.0071 | 0.99999 |  |  |  |  | 133 |  | 0 | [700 empty] | 0
P | 2018-07-26T13:36:19+08:00 | 2018-07-26T16:43:27+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | G18USM052G | CPSM Final Test | PASS |  | 156.92 | 0.08 | -0.54 | 56.67 | 51.99 | 1.74 | 0.13 | 0 | -4 | 0 | -0.868 | 0.49 | 4.4 | 21 |  |  | 0.0315 | 0.1721 | -0.0141 | 0.99999 |  |  |  |  | 133 |  | 0 | [219 empty] | 0
P | 2018-07-26T13:38:17+08:00 | 2018-07-26T16:47:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | G18USM053G | CPSM Final Test | PASS |  | 164.12 | 0.08 | -0.77 | 56.84 | 50.55 | 0.75 | 0.13 | 1 | 1 | 0 | -0.9119 | 0.4686 | 0.8 | 3 |  |  | 0.0633 | 0.3027 | 0.0256 | 1 |  |  |  |  | 133 |  | 0 | [764 empty] | 0
P | 2018-07-26T13:41:47+08:00 | 2018-07-26T17:17:41+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | G18USM058G | CPSM Final Test | PASS |  | 164.66 | 0.37 | -0.77 | 56.66 | 48.68 | 1.06 | 0.13 | 0 | 1 | 0 | -0.9438 | 0.529 | 4.3 | 3 |  |  | 0.1667 | 0.3605 | -0.0176 | 1 |  |  |  |  | 133 |  | 0 | [766 empty] | 0
P | 2018-07-26T13:39:30+08:00 | 2018-07-26T17:18:16+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | G18USM054G | CPSM Final Test | PASS |  | 168.98 | -0.02 | -0.92 | 56.79 | 50.97 | 1.44 | 0.13 | 1 | 3 | 1 | -0.9378 | 0.468 | 4.6 | 30 |  |  | 0.1515 | 0.1995 | 0.0499 | 0.99999 |  |  |  |  | 133 |  | 0.02307 | [254 empty] | 0
P | 2018-07-26T08:58:09+08:00 | 2018-07-26T18:18:18+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | G18USM060G | CPSM Final Test | FAIL | [20 empty] | 0.80234
P | 2018-07-26T13:38:14+08:00 | 2018-07-26T19:21:16+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | G18USM060G | CPSM Final Test | PASS |  | 167.9 | 0.08 | -0.46 | 57.04 | 49.53 | 1.51 | 0.13 | 0 | -1 | 0 | -0.9239 | 0.5206 | 8.1 | 17 |  |  | 0.1959 | 0.2334 | 0.0252 | 0.99998 |  |  |  |  | 133 |  | 0 | [767 empty] | 0
P | 2018-07-26T17:03:16+08:00 | 2018-07-26T21:19:48+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | G18USM055G | CPSM Final Test | PASS |  | 160.7 | 0.17 | -0.84 | 56.7 | 54.1 | 0.98 | 0.13 | 1 | -2 | 1 | -0.866 | 0.528 | 2.8 | -7 |  |  | 0.2108 | 0.1671 | 0.0409 | 0.99993 |  |  |  |  | 133 |  | 0 | [139 empty] | 0
P | 2018-07-26T17:19:39+08:00 | 2018-07-26T21:25:04+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI130G | CPSM Final Test | PASS | 129 | 161.78 | 0.68 | -1.07 | 56.34 | 48.6 | 1.13 | 0.13 | 0 | -2 | 0 | -0.9539 | 0.5058 | 2.3 | 23 |  | 0 | 0.1077 | 0.138 | 0.0046 | 1 | [451 empty] | 0
P | 2018-07-26T17:28:34+08:00 | 2018-07-26T21:28:25+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI128G | CPSM Final Test | PASS | 129 | 166.46 | 0.27 | -1.15 | 56.47 | 49.28 | 1.21 | 0.13 | 0 | 2 | 0 | -0.9299 | 0.546 | 3.8 | 4 |  | 0 | 0.0065 | 0.1811 | -0.0398 | 1 | [775 empty] | 0
P | 2018-07-26T18:39:53+08:00 | 2018-07-26T21:37:43+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | F18SDI093G | CPSM Final Test | PASS | 129 | 163.67 | 0.37 | -0.54 | 56.43 | 53.42 | 1.21 | 0.13 | 0 | -2 | 0 | -0.9338 | 0.462 | -2.8 | -12 |  | 0 | 0.0625 | 0.125 | -0.0424 | 0.99997 | [776 empty] | 0
P | 2018-07-26T04:33:16+08:00 | 2018-07-27T08:28:10+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI127G | CPSM Final Test | FAIL | [14 empty] | 655
P | 2018-07-26T00:34:41+08:00 | 2018-07-27T09:13:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | G18USM057G | CPSM Final Test | FAIL | [798 empty] | 0
P | 2018-07-26T18:30:43+08:00 | 2018-07-27T10:13:14+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI127G | CPSM Final Test | PASS | 129 | 158.9 | 0.27 | -0.77 | 56.52 | 48.18 | 1.44 | 0.13 | 0 | 3 | 0 | -0.9578 | 0.502 | 8.9 | 8 |  | 0 | 0.2093 | 0.1297 | 0.0082 | 0.99999 | [778 empty] | 0
P | 2018-07-26T16:54:15+08:00 | 2018-07-27T10:19:23+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | G18USM057G | CPSM Final Test | PASS |  | 161.06 | 0.27 | -1.22 | 56.51 | 52.07 | 0.83 | 0.13 | 1 | -2 | 0 | -0.854 | 0.5002 | -0.7 | -3 |  |  | 0.1126 | 0.3074 | 0.054 | 0.99987 |  |  |  |  | 133 |  | 0 | [770 empty] | 0
P | 2018-07-27T13:01:01+08:00 | 2018-07-27T13:47:46+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI129G | CPSM Final Test | FAIL
P | 2018-07-27T12:46:08+08:00 | 2018-07-27T14:02:26+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI125G | CPSM Final Test | FAIL | [5 empty] | 46.4
P | 2018-07-27T12:46:17+08:00 | 2018-07-27T14:02:30+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI129G | CPSM Final Test | FAIL | [5 empty] | 45.38
P | 2018-07-27T12:14:25+08:00 | 2018-07-27T14:34:22+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI129G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-07-25T19:29:16+08:00 | 2018-07-27T15:52:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI107G | CPSM Final Test | FAIL
P | 2018-07-27T13:16:06+08:00 | 2018-07-27T16:19:45+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI115G | CPSM Final Test | PASS | 129 | -43.87 | 0.17 | 0.15 | 56.51 | 50.29 | 1.97 | 0.18 | 1 | 3 | -1 | -0.864 | 0.5074 | 1.8 | 2 |  | 0 | 0.0714 | 0.2477 | -0.0485 | 1 | [779 empty] | 0
P | 2018-07-27T12:38:48+08:00 | 2018-07-27T16:23:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI103G | CPSM Final Test | PASS | 129 | 159.89 | -0.02 | -0.77 | 56.61 | 49.02 | 1.74 | 0.13 | 1 | -2 | 0 | -0.9199 | 0.5132 | 3.5 | 108 |  | 0 | 0.1394 | 0.4715 | -0.016 | 0.99999 | [780 empty] | 0
P | 2018-07-27T13:20:22+08:00 | 2018-07-27T16:26:44+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI108G | CPSM Final Test | PASS | 129 | 163.04 | -0.21 | -1.15 | 56.25 | 51.31 | 1.36 | 0.13 | 0 | 1 | 0 | -0.8419 | 0.5008 | -0.6 | 0 |  | 0.04202 | 0.1019 | 0.3858 | -0.0057 | 0.99999 | [781 empty] | 0
P | 2018-07-27T09:49:21+08:00 | 2018-07-27T16:51:25+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI124G | CPSM Final Test | FAIL | [14 empty] | 872
P | 2018-07-26T17:41:03+08:00 | 2018-07-27T16:56:15+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI107G | CPSM Final Test | PASS | 129 | 162.41 | 0.27 | -0.77 | 56.33 | 50.04 | 1.29 | 0.18 | 0 | -1 | 0 | -0.9697 | 0.5232 | 3.3 | -35 |  | 0 | 0.1086 | 0.2166 | -0.1692 | 0.99912 | [782 empty] | 0
P | 2018-07-27T09:49:15+08:00 | 2018-07-27T16:59:32+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI129G | CPSM Final Test | FAIL | [13 empty] | -5.3
P | 2018-07-27T09:26:27+08:00 | 2018-07-27T17:22:20+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI129G | CPSM Final Test | FAIL | [13 empty] | -9
P | 2018-07-27T16:34:36+08:00 | 2018-07-27T17:42:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI136G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-07-27T13:24:17+08:00 | 2018-07-27T18:03:20+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI125G | CPSM Final Test | PASS | 129 | 162.05 | 0.37 | -0.69 | 56.72 | 48.68 | 1.59 | 0.13 | -1 | 1 | 0 | -0.9259 | 0.5026 | -2.2 | 14 |  | 0 | 0.0964 | 0.1374 | -0.045 | 0.99998 | [277 empty] | 0
P | 2018-07-27T16:00:04+08:00 | 2018-07-27T18:16:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI136G | CPSM Final Test | FAIL
P | 2018-07-27T16:31:34+08:00 | 2018-07-27T18:26:29+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI109G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-07-27T13:20:23+08:00 | 2018-07-27T18:36:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI124G | CPSM Final Test | PASS | 129 | 166.55 | 0.56 | -0.08 | 57.03 | 51.73 | 1.21 | 0.13 | -1 | 1 | 0 | -0.8759 | 0.508 | 1.6 | 13 |  | 0 | 0.0748 | 0.1943 | -0.037 | 1 | [51 empty] | 0
P | 2018-07-27T15:31:54+08:00 | 2018-07-27T19:18:54+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI121G | CPSM Final Test | FAIL | [13 empty] | -7.5
P | 2018-07-27T15:42:06+08:00 | 2018-07-27T19:15:57+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI109G | CPSM Final Test | FAIL
P | 2018-07-27T13:24:23+08:00 | 2018-07-27T19:22:22+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI129G | CPSM Final Test | PASS | 129 | 154.76 | 0.08 | -1.22 | 56.57 | 48.34 | 1.51 | 0.13 | 0 | -1 | 0 | -0.838 | 0.52 | 12.2 | 6 |  | 0 | 0.0812 | 0.2558 | -0.0173 | 0.99998 | [777 empty] | 0
P | 2018-07-27T16:50:11+08:00 | 2018-07-27T20:34:18+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI131G | CPSM Final Test | PASS | 129 | 162.5 | 0.46 | -1.07 | 56.31 | 50.97 | 1.36 | 0.13 | 0 | -2 | 0 | -0.866 | 0.49 | -2.8 | 5 |  | 0 | 0.0833 | 0.3135 | 0.0172 | 0.99999 | [207 empty] | 0
P | 2018-07-27T17:36:20+08:00 | 2018-07-27T20:57:17+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI126G | CPSM Final Test | PASS | 129 | 160.97 | 0.46 | -0.69 | 56.71 | 55.29 | 1.21 | 0.13 | 0 | -4 | 0 | -0.9517 | 0.5232 | -2.2 | 13 |  | 0 | 0.0461 | 0.2423 | -0.0123 | 0.99999 | [650 empty] | 0
P | 2018-07-27T17:25:24+08:00 | 2018-07-27T21:18:42+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI121G | CPSM Final Test | PASS | 129 | 164.84 | 0.37 | -0.54 | 57.21 | 50.97 | 1.21 | 0.13 | 0 | 1 | 0 | -0.9558 | 0.4808 | 1.9 | -28 |  | 0 | 0.0722 | 0.278 | -0.0046 | 0.99999 | [783 empty] | 0
P | 2018-07-27T17:29:02+08:00 | 2018-07-27T21:21:27+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI109G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.84 | 56.57 | 49.7 | 1.74 | 0.13 | 1 | -2 | 0 | -0.9159 | 0.4984 | 5.7 | 12 |  | 0 | 0.0932 | 0.208 | 0.0079 | 0.99999 | [784 empty] | 0
P | 2018-07-28T08:40:07+08:00 | 2018-07-28T10:11:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI139G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-07-26T22:26:20+08:00 | 2018-07-28T11:50:33+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI136G | CPSM Final Test | FAIL | [806 empty] | 0
P | 2018-07-28T05:39:16+08:00 | 2018-07-28T12:43:56+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI138G | CPSM Final Test | FAIL
P | 2018-07-28T06:22:41+08:00 | 2018-07-28T12:56:06+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI111G | CPSM Final Test | FAIL | [13 empty] | 27.8
P | 2018-07-28T08:54:07+08:00 | 2018-07-28T13:11:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI132G | CPSM Final Test | PASS | 129 | 161.6 | 0.56 | -0.99 | 56.41 | 55.03 | 1.21 | 0.13 | 0 | -2 | 0 | -0.9139 | 0.4866 | 1.3 | -2 |  | 0 | 0.0397 | 0.1924 | -0.0107 | 0.99999 | [164 empty] | 0
P | 2018-07-28T09:18:15+08:00 | 2018-07-28T13:12:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI137G | CPSM Final Test | PASS | 129 | 162.86 | 0.56 | 0.15 | 56.54 | 51.22 | 1.44 | 0.13 | 1 | 2 | 0 | -0.834 | 0.5424 | 11.5 | -27 |  | 0 | 0.1127 | 0.1536 | -0.0062 | 1 | [218 empty] | 0
P | 2018-07-28T09:25:37+08:00 | 2018-07-28T13:13:51+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI139G | CPSM Final Test | PASS | 129 | 160.97 | 0.27 | -0.92 | 56.58 | 48.94 | 1.06 | 0.13 | -1 | 2 | -1 | -0.8141 | 0.5102 | 4.4 | -9 |  | 0.01508 | 0.0961 | 0.3497 | -0.011 | 0.99999 | [786 empty] | 0
P | 2018-07-28T09:43:00+08:00 | 2018-07-28T13:17:12+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI110G | CPSM Final Test | PASS | 129 | 165.2 | 0.17 | -0.61 | 56.64 | 50.12 | 1.59 | 0.13 | 1 | -2 | 0 | -0.834 | 0.5412 | 9.9 | 4 |  | 0 | 0.0395 | 0.0513 | -0.0231 | 1 | [787 empty] | 0
P | 2018-07-28T09:50:25+08:00 | 2018-07-28T13:19:08+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI095G | CPSM Final Test | PASS | 129 | 162.32 | 0.17 | -0.84 | 56.76 | 51.82 | 1.29 | 0.13 | -1 | 1 | 0 | -0.8181 | 0.4512 | -3.5 | -22 |  | 0 | 0.0615 | 0.1964 | 0.0391 | 0.99989 | [788 empty] | 0
P | 2018-07-26T20:53:38+08:00 | 2018-07-28T13:23:15+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI136G | CPSM Final Test | FAIL
P | 2018-07-27T17:08:27+08:00 | 2018-07-28T14:26:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI136G | CPSM Final Test | PASS | 129 | 159.62 | 0.08 | -0.92 | 56.55 | 48.51 | 1.13 | 0.13 | 0 | 0 | 1 | -0.9398 | 0.5 | 4.9 | -89 |  | 0 | 0.0525 | 0.1821 | 0.0036 | 1 | [785 empty] | 0
P | 2018-07-28T09:11:36+08:00 | 2018-07-28T14:28:22+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI138G | CPSM Final Test | PASS | 129 | 161.87 | 0.37 | -0.46 | 56.72 | 55.88 | 1.59 | 0.13 | 1 | 2 | 0 | -0.828 | 0.501 | -4.3 | 10 |  | 0 | 0.0762 | 0.3035 | 0.0205 | 0.99998 | [542 empty] | 0
P | 2018-07-28T09:39:23+08:00 | 2018-07-28T15:30:52+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI111G | CPSM Final Test | PASS | 129 | 160.97 | 0.17 | -0.84 | 56.79 | 50.55 | 1.74 | 0.13 | 1 | -4 | 0 | -0.7921 | 0.5116 | 1.5 | 4 |  | 0 | 0.0885 | 0.0863 | 0.0164 | 0.99998 | [761 empty] | 0
P | 2018-07-28T13:27:30+08:00 | 2018-07-28T16:11:43+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | G18SDI105G | CPSM Final Test | PASS | 129 | 173.39 | 0.56 | -0.92 | 56.5 | 50.97 | 1.06 | 0.13 | 0 | -1 | 0 | -0.8579 | 0.5212 | 1.6 | -14 |  | 0 | 0.0488 | 0.2472 | -0.0272 | 0.99998 | [355 empty] | 0
P | 2018-07-28T10:50:02+08:00 | 2018-07-28T16:22:59+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI134G | CPSM Final Test | FAIL | [16 empty] | 0.19044
P | 2018-07-28T15:58:47+08:00 | 2018-07-28T16:43:14+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI102G | CPSM Final Test | FAIL | [5 empty] | 46.48
P | 2018-07-28T11:49:48+08:00 | 2018-07-28T17:04:54+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI099G | CPSM Final Test | FAIL
P | 2018-07-28T12:25:54+08:00 | 2018-07-28T17:11:26+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI098G | CPSM Final Test | FAIL
P | 2018-07-28T13:46:12+08:00 | 2018-07-28T17:21:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI117G | CPSM Final Test | PASS | 129 | 168.53 | -0.02 | -0.61 | 56.7 | 52.75 | 1.59 | 0.13 | 1 | -2 | 0 | -0.83 | 0.5188 | -2 | -19 |  | 0 | 0.0804 | 0.441 | 0.008 | 0.99999 | [789 empty] | 0
P | 2018-07-28T13:36:31+08:00 | 2018-07-28T17:26:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI134G | CPSM Final Test | PASS | 129 | 167.72 | 0.27 | -0.61 | 56.71 | 50.88 | 1.13 | 0.13 | 1 | 2 | 0 | -0.844 | 0.5304 | 3.9 | -3 |  | 0.01819 | 0.0401 | 0.2301 | -0.0254 | 1 | [673 empty] | 0
P | 2018-07-28T14:17:33+08:00 | 2018-07-28T17:55:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI114G | CPSM Final Test | PASS | 129 | 161.87 | 0.08 | -0.08 | 56.5 | 51.14 | 1.51 | 0.13 | 1 | 3 | 0 | -0.8222 | 0.5048 | 1.3 | -1 |  | 0 | 0.1231 | 0.1277 | 0.007 | 1 | [790 empty] | 0
P | 2018-07-28T16:05:03+08:00 | 2018-07-28T18:47:34+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | G18SDI140G | CPSM Final Test | PASS | 129 | 159.44 | -0.02 | -0.99 | 56.49 | 50.88 | 1.13 | 0.13 | 1 | 1 | 0 | -0.8301 | 0.5282 | -1.3 | -4 |  | 0 | 0.0647 | 0.1419 | -0.0125 | 1 | [785 empty] | 0
P | 2018-07-28T14:27:21+08:00 | 2018-07-28T18:49:30+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI099G | CPSM Final Test | PASS | 129 | 159.44 | -0.02 | -1.15 | 56.44 | 53.09 | 1.13 | 0.13 | 0 | -2 | 0 | -0.864 | 0.5062 | 2.1 | -7 |  | 0 | 0.1058 | 0.2028 | 0.0584 | 0.99985 | [791 empty] | 0
P | 2018-07-28T14:48:40+08:00 | 2018-07-28T18:56:06+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | G18SDI098G | CPSM Final Test | PASS | 129 | 161.33 | 0.56 | -0.99 | 56.61 | 48.34 | 1.59 | 0.13 | 0 | 3 | 0 | -0.864 | 0.4882 | 10.3 | -1 |  | 0 | 0.133 | 0.4184 | 0.0037 | 0.99998 | [175 empty] | 0
P | 2018-07-27T23:51:45+08:00 | 2018-07-29T08:50:16+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI102G | CPSM Final Test | FAIL | [14 empty] | 2799
P | 2018-07-29T08:26:18+08:00 | 2018-07-29T09:28:56+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | G18SDI135G | CPSM Final Test | FAIL | [5 empty] | 46.48
P | 2018-07-29T07:45:14+08:00 | 2018-07-29T10:17:24+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI094G | CPSM Final Test | FAIL | [13 empty] | -7.2
P | 2018-07-27T21:48:13+08:00 | 2018-07-29T10:53:48+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI102G | CPSM Final Test | FAIL
P | 2018-07-28T16:21:01+08:00 | 2018-07-29T10:54:28+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | G18SDI102G | CPSM Final Test | PASS | 129 | 162.05 | 0.46 | -0.99 | 56.25 | 48.18 | 1.36 | 0.13 | -1 | -4 | 0 | -0.8819 | 0.5366 | 4.4 | 15 |  | 0 | 0.0637 | 0.2519 | -0.0171 | 0.99998 | [792 empty] | 0
P | 2018-07-29T06:34:34+08:00 | 2018-07-29T11:35:27+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | G18SDI104G | CPSM Final Test | FAIL | [21 empty] | 8540
P | 2018-07-29T08:53:45+08:00 | 2018-07-29T11:43:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI100G | CPSM Final Test | PASS | 129 | 162.68 | 0.87 | -1.15 | 56.49 | 51.05 | 0.98 | 0.13 | 1 | 2 | 0 | -0.8301 | 0.515 | 5.9 | -2 |  | 0.00239 | 0.1365 | 0.1823 | 0.097 | 0.9999 | [793 empty] | 0
P | 2018-07-29T09:03:01+08:00 | 2018-07-29T11:49:24+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI101G | CPSM Final Test | PASS | 129 | 161.87 | 0.46 | -0.92 | 56.61 | 48.26 | 1.06 | 0.13 | 0 | -3 | 0 | -0.7922 | 0.5058 | 1.3 | 27 |  | 0 | 0.0333 | 0.0884 | -0.0265 | 0.99999 | [240 empty] | 0
P | 2018-07-29T09:08:35+08:00 | 2018-07-29T11:53:27+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | G18SDI097G | CPSM Final Test | PASS | 129 | 167.09 | 0.17 | -0.61 | 56.62 | 50.04 | 1.51 | 0.13 | 0 | -4 | 0 | -0.8002 | 0.5124 | 1.3 | -3 |  | 0.02403 | 0.0541 | 0.2786 | -0.0277 | 1 | [69 empty] | 0
P | 2018-07-29T08:54:42+08:00 | 2018-07-29T11:58:49+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | G18SDI106G | CPSM Final Test | PASS | 129 | 158.45 | 0.08 | -0.84 | 56.32 | 50.04 | 1.51 | 0.13 | -1 | 1 | 0 | -0.8341 | 0.5248 | 10.3 | -4 |  | 0 | 0.0705 | 0.0704 | -0.0123 | 0.99998 | [794 empty] | 0
P | 2018-07-29T08:57:37+08:00 | 2018-07-29T12:09:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | G18SDI135G | CPSM Final Test | PASS | 129 | 158.54 | 0.17 | 0.07 | 56.85 | 49.02 | 1.13 | 0.13 | 0 | 2 | 0 | -0.842 | 0.513 | -4.6 | 68 |  | 0 | 0.1518 | 0.1743 | 0.0086 | 0.99999 | [795 empty] | 0
P | 2018-07-29T09:01:19+08:00 | 2018-07-29T12:17:34+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | F18SDI094G | CPSM Final Test | PASS | 129 | 166.37 | -1 | -0.77 | 56.48 | 49.95 | 1.82 | 0.13 | 0 | 2 | 0 | -0.82 | 0.545 | 6.6 | -8 |  | 0 | 0.048 | 0.1257 | -0.0091 | 1 | [796 empty] | 0
P | 2018-07-29T09:05:01+08:00 | 2018-07-29T12:41:03+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | G18SDI104G | CPSM Final Test | PASS | 129 | 163.31 | 0.17 | -1.22 | 56.31 | 49.45 | 1.13 | 0.13 | -1 | -4 | 0 | -0.7762 | 0.522 | -5 | 32 |  | 0 | 0.0643 | 0.1986 | -0.0165 | 0.99999 | 0
P | 2018-07-29T10:03:12+08:00 | 2018-07-29T14:04:48+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI113G | CPSM Final Test | FAIL
P | 2018-07-29T10:25:52+08:00 | 2018-07-29T14:16:53+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI096G | CPSM Final Test | FAIL | [14 empty] | 661
P | 2018-07-29T09:52:18+08:00 | 2018-07-29T14:49:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI092G | CPSM Final Test | FAIL
P | 2018-07-29T12:00:52+08:00 | 2018-07-29T15:23:48+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | G18SDI116G | CPSM Final Test | PASS | 129 | 157.46 | 0.37 | -0.99 | 56.63 | 48.94 | 1.13 | 0.13 | -1 | 2 | 0 | -0.7861 | 0.4812 | -3.7 | 18 |  | 0 | 0.0775 | 0.2906 | -0.0156 | 0.99998 | [797 empty] | 0
P | 2018-07-29T11:57:44+08:00 | 2018-07-29T15:20:00+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | G18SDI119G | CPSM Final Test | PASS | 129 | 164.93 | -0.02 | -0.61 | 56.73 | 51.73 | 1.67 | 0.13 | 0 | -2 | 0 | -0.8062 | 0.5162 | 5.6 | 41 |  | 0 | 0.1089 | 0.1399 | -0.0168 | 0.99999 | [143 empty] | 0
P | 2018-07-29T08:37:18+08:00 | 2018-07-29T15:30:42+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI113G | CPSM Final Test | FAIL | [24 empty] | 0
P | 2018-07-29T12:21:23+08:00 | 2018-07-29T16:01:09+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | G18SDI096G | CPSM Final Test | PASS | 129 | 151.7 | -0.02 | -0.61 | 56.51 | 49.45 | 1.67 | 0.13 | -1 | -3 | 0 | -0.858 | 0.501 | 12.7 | 0 |  | 0.02396 | 0.0487 | 0.2111 | -0.0456 | 0.99999 | [798 empty] | 0
P | 2018-07-29T12:07:08+08:00 | 2018-07-29T16:30:25+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | G18SDI133G | CPSM Final Test | PASS | 129 | 164.75 | 0.46 | -0.54 | 56.7 | 54.61 | 1.67 | 0.13 | 0 | 1 | 1 | -0.9398 | 0.4928 | 0.3 | -15 |  | 0 | 0.2088 | 0.2438 | -0.0226 | 0.99999 |  |  | 0
P | 2018-07-29T12:21:06+08:00 | 2018-07-29T16:34:12+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | F18SDI092G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -1.15 | 56.73 | 55.37 | 0.91 | 0.13 | 0 | 1 | 0 | -0.8341 | 0.516 | -4.8 | 2 |  | 0 | 0.1202 | 0.1143 | 0.0041 | 1 | [799 empty] | 0
P | 2018-07-29T12:04:00+08:00 | 2018-07-29T16:36:26+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | G18SDI113G | CPSM Final Test | PASS | 129 | 162.86 | 0.27 | -0.99 | 56.48 | 48.43 | 1.51 | 0.13 | 0 | -2 | 0 | -0.7803 | 0.5172 | -1.7 | 6 |  | 0.04804 | 0.1484 | 0.2602 | 0.0412 | 0.99992 |  |  |  | 0
P | 2018-03-13T10:14:11+08:00 | 2018-03-13T12:59:43+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI626G | CPSM Final Test | PASS | 129 | 158.63 | -0.02 | -0.84 | 56.84 | 53.34 | 1.13 | 0.13 | 1 | -2 | 1 | -0.8619 | 0.5002 | 6.6 | 1 |  | 0.01217 | 0.173 | 0.4666 | -0.0288 | 1 | [700 empty] | 0
P | 2018-03-15T06:49:38+08:00 | 2018-03-15T12:16:13+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | C18USM892G | CPSM Final Test | FAIL | [14 empty] | 709
P | 2018-03-15T09:48:57+08:00 | 2018-03-15T12:49:39+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | C18USM893G | CPSM Final Test | PASS |  | 163.4 | -0.02 | -0.23 | 56.79 | 54.02 | 1.74 | 0.13 | 0 | 4 | 1 | -0.7801 | 0.513 | 1.1 | 51 |  |  | 0.1511 | 0.2888 | -0.0192 | 1 |  |  |  |  | 133 |  | 0 | [793 empty] | 0
P | 2018-03-15T09:22:18+08:00 | 2018-03-15T12:52:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | C18USM894G | CPSM Final Test | PASS |  | 166.19 | 0.37 | -0.16 | 57.06 | 54.27 | 1.44 | 0.13 | 1 | -2 | 1 | -0.7263 | 0.5218 | 5.3 | 43 |  |  | 0.139 | 0.1283 | -0.0522 | 1 |  |  |  |  | 133 |  | 0 | [248 empty] | 0
P | 2018-03-15T10:26:30+08:00 | 2018-03-15T13:36:44+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | C18USM896G | CPSM Final Test | PASS |  | 164.57 | 0.37 | -0.61 | 56.78 | 53.85 | 1.44 | 0.13 | 0 | 4 | -1 | -0.8899 | 0.5152 | 3.6 | 77 |  |  | 0.1742 | 0.1677 | -0.003 | 1 |  |  |  |  | 133 |  | 0.03074 | [52 empty] | 0
P | 2018-03-15T06:25:29+08:00 | 2018-03-15T13:57:05+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | C18USM895G | CPSM Final Test | FAIL | [822 empty] | 0
P | 2018-03-15T10:55:43+08:00 | 2018-03-15T14:13:55+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | C18USM897G | CPSM Final Test | PASS |  | 161.78 | 0.37 | -0.61 | 57.03 | 55.2 | 1.67 | 0.13 | 1 | -2 | 1 | -0.804 | 0.4682 | 5.3 | 10 |  |  | 0.0993 | 0.3954 | -0.0143 | 0.99983 |  |  |  |  | 133 |  | 0.00844 | [795 empty] | 0
P | 2018-03-15T09:32:56+08:00 | 2018-03-15T14:16:15+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | C18USM892G | CPSM Final Test | PASS |  | 160.61 | -0.02 | -0.61 | 57.07 | 56.81 | 1.44 | 0.13 | 1 | -12 | 1 | -0.7383 | 0.464 | 1.3 | 3 |  |  | 0.1513 | 0.3286 | -0.0539 | 0.99999 |  |  |  |  | 133 |  | 0 | [695 empty] | 0
P | 2018-03-15T10:11:17+08:00 | 2018-03-15T15:03:07+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | C18USM895G | CPSM Final Test | PASS |  | 165.83 | -0.02 | -0.16 | 57.41 | 54.78 | 1.59 | 0.13 | -1 | -2 | 2 | -0.7901 | 0.4612 | 2.2 | 51 |  |  | 0.0903 | 0.2881 | -0.0862 | 1 |  |  |  |  | 133 |  | 0.00385 | [794 empty] | 0
P | 2018-03-15T12:00:24+08:00 | 2018-03-15T15:21:17+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | C18USM898G | CPSM Final Test | PASS |  | 162.32 | -0.11 | 0 | 57.1 | 55.88 | 1.51 | 0.13 | -1 | 2 | 1 | -0.7862 | 0.5362 | 1.1 | -10 |  |  | 0.083 | 0.3559 | 0.0237 | 0.99994 |  |  |  |  | 133 |  | 0 | [353 empty] | 0
P | 2018-03-15T15:11:51+08:00 | 2018-03-15T16:53:19+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | C18USM901G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-15T14:20:44+08:00 | 2018-03-15T17:26:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | C18USM900G | CPSM Final Test | PASS |  | 163.49 | 0.65 | -0.77 | 57.07 | 53.51 | 1.82 | 0.13 | -1 | -1 | 1 | -0.8141 | 0.5446 | 1.7 | -13 |  |  | 0.0814 | 0.0822 | -0.035 | 1 |  |  |  |  | 133 |  | 0.01013 | [796 empty] | 0
P | 2018-03-15T14:43:48+08:00 | 2018-03-15T18:01:51+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | C18USM899G | CPSM Final Test | PASS |  | 166.28 | 0.27 | -0.54 | 57.01 | 53.51 | 1.82 | 0.13 | -1 | -11 | 1 | -0.8241 | 0.5404 | 8 | -6 |  |  | 0.119 | 0.4202 | -0.0049 | 1 |  |  |  |  | 133 |  | 0 | [76 empty] | 0
P | 2018-03-15T16:15:32+08:00 | 2018-03-15T19:08:31+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | C18USM902G | CPSM Final Test | PASS |  | 163.58 | 0.08 | -0.61 | 56.73 | 59.27 | 1.74 | 0.13 | -1 | -2 | 1 | -0.852 | 0.492 | 3.1 | 29 |  |  | 0.0704 | 0.2383 | 0.0009 | 0.99997 |  |  |  |  | 133 |  | 0.0316 | [80 empty] | 0
P | 2018-03-15T12:09:57+08:00 | 2018-03-15T19:55:13+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | C18USM901G | CPSM Final Test | FAIL | [578 empty] | 2271
P | 2018-03-15T16:42:50+08:00 | 2018-03-15T20:22:47+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | C18USM903G | CPSM Final Test | PASS |  | 165.65 | 0.65 | -1.07 | 56.63 | 55.2 | 1.67 | 0.13 | -1 | -1 | 1 | -0.9038 | 0.499 | 2.2 | 23 |  |  | 0.0704 | 0.2107 | -0.0337 | 0.99998 |  |  |  |  | 133 |  | 0.00561 | [797 empty] | 0
P | 2018-03-15T17:31:19+08:00 | 2018-03-15T20:51:23+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | C18USM904G | CPSM Final Test | PASS |  | 165.11 | -0.11 | -1.15 | 56.49 | 55.2 | 1.74 | 0.13 | 0 | -2 | 1 | -0.7901 | 0.488 | 2 | -9 |  |  | 0.1237 | 0.123 | 0.0441 | 0.99987 |  |  |  |  | 133 |  | 0 | [798 empty] | 0
P | 2018-03-15T16:02:35+08:00 | 2018-03-15T21:00:59+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | C18USM901G | CPSM Final Test | PASS |  | 166.46 | 0.08 | -0.16 | 57.33 | 52.75 | 1.74 | 0.13 | 0 | -8 | 0 | -0.8939 | 0.4968 | 0.8 | 42 |  |  | 0.2298 | 0.1816 | 0.0042 | 0.99998 |  |  |  |  | 133 |  | 0.00472 | [550 empty] | 0
P | 2018-03-16T07:33:19+08:00 | 2018-03-16T11:27:32+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM909G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-16T09:26:06+08:00 | 2018-03-16T12:41:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM908G | CPSM Final Test | PASS |  | 164.21 | 0.17 | -0.54 | 56.69 | 58.76 | 1.67 | 0.13 | 1 | 3 | 3 | -0.7542 | 0.5004 | 1.5 | 49 |  |  | 0.0601 | 0.2262 | 0.0388 | 0.99999 |  |  |  |  | 133 |  | 0.02843 | [58 empty] | 0
P | 2018-03-16T09:19:52+08:00 | 2018-03-16T12:44:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM905G | CPSM Final Test | PASS |  | 165.74 | -0.02 | -0.46 | 56.77 | 54.69 | 1.59 | 0.13 | 1 | -8 | 0 | -0.8559 | 0.4946 | 3 | 68 |  |  | 0.1154 | 0.1202 | -0.0193 | 0.99999 |  |  |  |  | 133 |  | 0 | [799 empty] | 0
P | 2018-03-16T08:36:44+08:00 | 2018-03-16T14:32:17+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM907G | CPSM Final Test | FAIL | [14 empty] | 1041
P | 2018-03-16T11:32:33+08:00 | 2018-03-16T14:48:23+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM906G | CPSM Final Test | PASS |  | 159.8 | -0.02 | -0.46 | 57.02 | 53.68 | 1.67 | 0.13 | 1 | 2 | 1 | -0.832 | 0.493 | 0.9 | -39 |  |  | 0.0244 | 0.3491 | 0.016 | 0.99997 |  |  |  |  | 133 |  | 0.0351 | [166 empty] | 0
P | 2018-03-16T11:52:50+08:00 | 2018-03-16T15:39:34+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM910G | CPSM Final Test | PASS |  | 164.39 | -0.49 | -0.31 | 57.07 | 54.69 | 1.36 | 0.13 | 1 | -4 | 1 | -0.816 | 0.507 | 5.4 | 53 |  |  | 0.3277 | 0.2987 | -0.0069 | 0.99997 |  |  |  |  | 133 |  | 0.02403 | [550 empty] | 0
P | 2018-03-16T03:10:29+08:00 | 2018-03-16T15:50:22+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM909G | CPSM Final Test | FAIL
P | 2018-03-16T07:14:31+08:00 | 2018-03-16T15:54:30+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM907G | CPSM Final Test | FAIL | [828 empty] | 16978
P | 2018-03-16T13:09:57+08:00 | 2018-03-16T17:30:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI632G | CPSM Final Test | FAIL
P | 2018-03-16T09:30:25+08:00 | 2018-03-16T17:40:30+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM909G | CPSM Final Test | PASS |  | 158.81 | -0.02 | -0.84 | 56.79 | 54.61 | 1.06 | 0.13 | 1 | -11 | 1 | -0.8021 | 0.5088 | 2.9 | 9 |  |  | 0.0957 | 0.0793 | 0.0075 | 0.99999 |  |  |  |  | 133 |  | 0.00627 | [794 empty] | 0
P | 2018-03-16T12:40:03+08:00 | 2018-03-16T18:21:15+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI633G | CPSM Final Test | FAIL
P | 2018-03-16T15:24:41+08:00 | 2018-03-16T18:25:27+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI631G | CPSM Final Test | PASS | 129 | 153.77 | -0.02 | -0.99 | 56.83 | 54.69 | 1.21 | 0.13 | 1 | 1 | 1 | -0.87 | 0.491 | 1.9 | -19 |  | 0 | 0.0842 | 0.2574 | -0.0317 | 1 | [384 empty] | 0
P | 2018-03-16T15:39:10+08:00 | 2018-03-16T18:47:15+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI634G | CPSM Final Test | PASS | 129 | 170.87 | -0.02 | -0.61 | 57.08 | 54.61 | 1.29 | 0.13 | 0 | -12 | 1 | -0.864 | 0.5498 | 6.7 | 0 |  | 0.03509 | 0.0482 | 0.0703 | 0.0091 | 0.99999 | [470 empty] | 0
P | 2018-03-16T04:10:51+08:00 | 2018-03-16T18:58:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM907G | CPSM Final Test | FAIL | [828 empty] | 0
P | 2018-03-16T15:20:18+08:00 | 2018-03-16T19:15:27+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI632G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.38 | 57.02 | 53.59 | 1.74 | 0.13 | 1 | -2 | 1 | -0.9379 | 0.5032 | 1.8 | -1 |  | 0 | 0.0549 | 0.1078 | -0.0336 | 1 | [808 empty] | 0
P | 2018-03-16T17:25:49+08:00 | 2018-03-16T20:00:24+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI636G | CPSM Final Test | FAIL
P | 2018-03-16T10:59:29+08:00 | 2018-03-16T20:01:49+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI633G | CPSM Final Test | FAIL | [830 empty] | 0
P | 2018-03-16T10:33:53+08:00 | 2018-03-16T20:27:25+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI633G | CPSM Final Test | FAIL
P | 2018-03-16T11:34:30+08:00 | 2018-03-16T20:32:06+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM907G | CPSM Final Test | PASS |  | 164.03 | -0.02 | -0.92 | 56.79 | 55.29 | 1.21 | 0.13 | 1 | 4 | -1 | -0.7482 | 0.5044 | 1.6 | -3 |  |  | 0.1533 | 0.1794 | -0.0168 | 0.99999 |  |  |  |  | 133 |  | 0 | [800 empty] | 0
P | 2018-03-16T15:30:39+08:00 | 2018-03-16T21:33:09+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI633G | CPSM Final Test | PASS | 129 | 162.14 | 0.17 | -1.22 | 56.76 | 53.42 | 1.21 | 0.13 | -1 | 2 | 1 | -0.9398 | 0.5136 | 0.6 | 12 |  | 0 | 0.1961 | 0.2178 | -0.0511 | 1 | [809 empty] | 0
P | 2018-03-16T18:38:43+08:00 | 2018-03-16T21:35:24+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI635G | CPSM Final Test | PASS | 129 | 165.74 | -0.21 | -0.54 | 56.77 | 53.42 | 1.36 | 0.13 | -1 | -12 | 1 | -0.844 | 0.4938 | 13.2 | 17 |  | 0 | 0.2275 | 0.1261 | -0.0063 | 0.99996 | [374 empty] | 0
P | 2018-03-16T18:43:06+08:00 | 2018-03-16T21:45:50+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI636G | CPSM Final Test | PASS | 129 | 161.51 | 0.17 | -0.54 | 56.73 | 58.08 | 1.29 | 0.13 | 1 | -2 | 1 | -0.872 | 0.5006 | 2.9 | 8 |  | 0 | 0.0886 | 0.0581 | 0.0115 | 0.99999 | [354 empty] | 0
P | 2018-03-17T06:44:30+08:00 | 2018-03-17T11:50:25+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI638G | CPSM Final Test | FAIL | [831 empty] | 1269
P | 2018-03-17T06:19:17+08:00 | 2018-03-17T11:56:21+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI639G | CPSM Final Test | FAIL
P | 2018-03-17T09:08:46+08:00 | 2018-03-17T12:16:56+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI637G | CPSM Final Test | PASS | 129 | 161.87 | -0.02 | -0.92 | 56.94 | 55.2 | 1.44 | 0.13 | -2 | -12 | -1 | -0.9656 | 0.5486 | 2.5 | 7 |  | 0.0155 | 0.2221 | 0.4335 | -0.0046 | 1 | [807 empty] | 0
P | 2018-03-17T09:17:28+08:00 | 2018-03-17T12:55:59+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI638G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | -0.54 | 56.92 | 52.58 | 2.05 | 0.13 | -1 | -2 | 0 | -0.9198 | 0.5058 | 2.3 | -8 |  | 0 | 0.0575 | 0.1658 | -0.0297 | 0.99999 | [810 empty] | 0
P | 2018-03-17T12:14:23+08:00 | 2018-03-17T13:50:33+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI642G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-17T09:07:49+08:00 | 2018-03-17T13:55:37+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI639G | CPSM Final Test | PASS | 129 | 158.36 | -0.02 | -0.99 | 56.79 | 53 | 1.51 | 0.13 | 1 | -2 | 0 | -0.87 | 0.5422 | 5.1 | -7 |  | 0 | 0.0996 | 0.321 | 0.0864 | 0.99995 | [811 empty] | 0
P | 2018-03-17T12:57:45+08:00 | 2018-03-17T15:57:04+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI641G | CPSM Final Test | PASS | 129 | 164.57 | -0.02 | -0.54 | 57.14 | 53.68 | 1.29 | 0.13 | 1 | -13 | -1 | -0.8679 | 0.5138 | 3.7 | 39 |  | 0.00749 | 0.0381 | 0.0643 | 0.0307 | 0.99999 | [354 empty] | 0
P | 2018-03-17T12:53:16+08:00 | 2018-03-17T16:01:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI640G | CPSM Final Test | PASS | 129 | 168.44 | -0.02 | -0.08 | 57.13 | 51.82 | 1.9 | 0.13 | 0 | -11 | -1 | -0.8101 | 0.502 | 4.9 | 15 |  | 0 | 0.0767 | 0.2655 | 0.0137 | 0.99996 | [812 empty] | 0
P | 2018-03-17T10:37:46+08:00 | 2018-03-17T16:15:21+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI644G | CPSM Final Test | FAIL | [830 empty] | 0
P | 2018-03-17T09:32:33+08:00 | 2018-03-17T16:32:23+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI642G | CPSM Final Test | FAIL
P | 2018-03-17T13:26:34+08:00 | 2018-03-17T17:21:13+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI644G | CPSM Final Test | PASS | 129 | 164.84 | 0.37 | -0.31 | 56.94 | 53.93 | 1.51 | 0.13 | -1 | -13 | 1 | -0.9099 | 0.542 | 3.4 | 50 |  | 0 | 0.071 | 0.0841 | 0.0289 | 0.99999 | [809 empty] | 0
P | 2018-03-17T13:02:28+08:00 | 2018-03-17T18:17:21+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI642G | CPSM Final Test | PASS | 129 | 165.02 | 0.17 | -0.54 | 56.72 | 58.42 | 1.9 | 0.13 | 0 | -4 | 1 | -0.8699 | 0.491 | 4.6 | 14 |  | 0.01165 | 0.2117 | 0.1892 | -0.0312 | 1 | [104 empty] | 0
P | 2018-03-15T22:17:58+08:00 | 2018-03-19T10:52:40+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI647G | CPSM Final Test | FAIL
P | 2018-03-19T08:51:09+08:00 | 2018-03-19T11:28:40+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI649G | CPSM Final Test | FAIL
P | 2018-03-17T16:30:46+08:00 | 2018-03-19T11:35:35+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI645G | CPSM Final Test | PASS | 129 | 164.39 | 0.08 | 0 | 57.2 | 54.95 | 1.74 | 0.13 | -1 | 4 | -1 | -0.7702 | 0.52 | 3.3 | 17 |  | 0 | 0.1631 | 0.1344 | 0.0572 | 0.99994 | [148 empty] | 0
P | 2018-03-17T16:35:19+08:00 | 2018-03-19T12:46:36+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI647G | CPSM Final Test | PASS | 129 | 166.37 | -0.02 | -0.38 | 57 | 54.1 | 1.51 | 0.13 | -1 | -2 | -3 | -0.824 | 0.4794 | 3.4 | -28 |  | 0 | 0.1336 | 0.5088 | 0.0026 | 1 | [267 empty] | 0
P | 2018-03-19T06:43:12+08:00 | 2018-03-19T13:16:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI648G | CPSM Final Test | FAIL
P | 2018-03-19T10:20:54+08:00 | 2018-03-19T13:18:38+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI650G | CPSM Final Test | PASS | 129 | 164.84 | 0.17 | -0.46 | 56.61 | 53.59 | 1.74 | 0.13 | 0 | -2 | 2 | -0.9039 | 0.5326 | 7.6 | 2 |  | 0.0157 | 0.0757 | 0.0524 | 0.0112 | 0.99998 | [813 empty] | 0
P | 2018-03-19T05:35:35+08:00 | 2018-03-19T14:44:14+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI649G | CPSM Final Test | FAIL
P | 2018-03-19T09:59:59+08:00 | 2018-03-19T15:02:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI648G | CPSM Final Test | PASS | 129 | 166.01 | -0.02 | -0.23 | 56.82 | 52.15 | 1.74 | 0.13 | 1 | -2 | 0 | -0.8959 | 0.4894 | 0.4 | 6 |  | 0 | 0.0775 | 0.1103 | -0.0018 | 1 | [231 empty] | 0
P | 2018-03-19T11:35:31+08:00 | 2018-03-19T15:20:33+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI651G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.54 | 57.04 | 51.39 | 1.44 | 0.13 | 0 | -2 | 1 | -0.92 | 0.476 | 1.3 | 147 |  | 0.00519 | 0.085 | 0.1465 | 0.0025 | 1 | [809 empty] | 0
P | 2018-03-19T10:09:54+08:00 | 2018-03-19T15:49:54+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI649G | CPSM Final Test | PASS | 129 | 163.76 | 0.08 | -0.16 | 56.82 | 55.46 | 2.05 | 0.13 | -1 | -1 | 1 | -0.8859 | 0.5254 | 6 | 3 |  | 0 | 0.0936 | 0.1407 | -0.0178 | 0.99999 | [814 empty] | 0
P | 2018-03-19T11:35:19+08:00 | 2018-03-19T16:01:50+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI652G | CPSM Final Test | FAIL
P | 2018-03-19T14:02:16+08:00 | 2018-03-19T16:58:07+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI653G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.54 | 57.05 | 57.23 | 1.51 | 0.13 | 0 | -2 | -1 | -0.8321 | 0.5194 | 0.3 | 47 |  | 0.05991 | 0.0646 | 0.244 | 0.0433 | 0.99997 | [782 empty] | 0
P | 2018-03-15T16:11:29+08:00 | 2018-03-19T17:08:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI646G | CPSM Final Test | FAIL | [17 empty] | 0.6759
P | 2018-03-19T14:22:19+08:00 | 2018-03-19T17:33:56+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI654G | CPSM Final Test | PASS | 129 | 166.73 | -0.02 | -0.16 | 57.2 | 53.51 | 1.74 | 0.13 | -1 | -1 | 1 | -0.9199 | 0.4734 | 1.8 | 37 |  | 0.00637 | 0.0394 | 0.0656 | -0.0417 | 0.99998 | [700 empty] | 0
P | 2018-03-17T16:39:55+08:00 | 2018-03-19T18:11:23+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI646G | CPSM Final Test | PASS | 129 | 160.79 | -0.02 | -0.46 | 56.83 | 51.31 | 1.9 | 0.13 | -1 | 2 | -1 | -0.7821 | 0.544 | 1.8 | -1 |  | 0 | 0.1113 | 0.3501 | 0.0178 | 0.99999 | [810 empty] | 0
P | 2018-03-19T16:16:53+08:00 | 2018-03-19T19:10:36+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI655G | CPSM Final Test | PASS | 129 | 165.65 | 0.08 | -0.23 | 57.05 | 53 | 1.59 | 0.13 | -1 | 3 | 1 | -0.8221 | 0.467 | 1.8 | 7 |  | 0 | 0.1269 | 0.4458 | 0.0497 | 0.99964 | [814 empty] | 0
P | 2018-03-19T08:09:33+08:00 | 2018-03-19T19:27:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI652G | CPSM Final Test | FAIL
P | 2018-03-19T16:32:57+08:00 | 2018-03-19T19:30:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI656G | CPSM Final Test | PASS | 129 | 158 | 0.08 | -0.77 | 57.02 | 55.29 | 1.36 | 0.13 | -1 | 2 | 1 | -0.88 | 0.5052 | 12.5 | 36 |  | 0 | 0.0285 | 0.0762 | -0.0086 | 0.99998 | [802 empty] | 0
P | 2018-03-19T16:44:16+08:00 | 2018-03-19T19:36:49+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI657G | CPSM Final Test | PASS | 129 | 158.99 | -0.02 | -0.31 | 57.01 | 53.93 | 1.44 | 0.13 | 0 | -3 | 0 | -0.8181 | 0.4576 | 1.2 | 37 |  | 0 | 0.0742 | 0.1099 | 0.0066 | 0.99999 | [60 empty] | 0
P | 2018-03-19T17:23:43+08:00 | 2018-03-19T19:58:21+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI659G | CPSM Final Test | FAIL
P | 2018-03-19T18:28:53+08:00 | 2018-03-19T21:21:26+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI658G | CPSM Final Test | PASS | 129 | 161.15 | -0.11 | -0.38 | 56.92 | 55.03 | 1.44 | 0.13 | 0 | 11 | 0 | -0.846 | 0.4952 | 3.8 | 29 |  | 0.00639 | 0.1581 | 0.0844 | -0.0174 | 1 | [815 empty] | 0
P | 2018-03-19T18:41:02+08:00 | 2018-03-19T21:42:45+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI659G | CPSM Final Test | PASS | 129 | 153.59 | 0.08 | -0.38 | 56.74 | 59.44 | 1.51 | 0.13 | 0 | -2 | 1 | -0.918 | 0.492 | 0.3 | -6 |  | 0 | 0.1493 | 0.1414 | 0.0195 | 0.99998 | [816 empty] | 0
P | 2018-03-18T18:38:35+08:00 | 2018-03-20T08:58:34+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI652G | CPSM Final Test | FAIL
P | 2018-03-19T13:48:34+08:00 | 2018-03-20T08:59:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI652G | CPSM Final Test | PASS | 129 | 159.08 | -0.02 | -0.77 | 56.5 | 52.41 | 2.05 | 0.13 | -1 | -2 | 0 | -0.842 | 0.476 | 5.3 | -9 |  | 0 | 0.0531 | 0.2658 | 0.0223 | 1 | [817 empty] | 0
P | 2018-03-20T08:40:28+08:00 | 2018-03-20T11:38:17+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI662G | CPSM Final Test | FAIL
P | 2018-03-20T07:17:08+08:00 | 2018-03-20T11:41:13+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI661G | CPSM Final Test | FAIL | [13 empty] | 18.9
P | 2018-03-20T06:29:26+08:00 | 2018-03-20T12:28:55+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI661G | CPSM Final Test | FAIL | [14 empty] | 1127
P | 2018-03-20T09:38:54+08:00 | 2018-03-20T13:14:42+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI660G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -0.61 | 56.32 | 49.45 | 1.51 | 0.13 | 1 | -10 | 1 | -0.8261 | 0.5058 | 4.3 | 6 |  | 0.00558 | 0.1055 | 0.5077 | 0.0074 | 0.99997 | [817 empty] | 0
P | 2018-03-20T10:01:28+08:00 | 2018-03-20T13:22:33+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI663G | CPSM Final Test | PASS | 129 | 161.69 | 0.08 | -0.84 | 56.43 | 54.78 | 1.21 | 0.13 | 0 | 3 | 1 | -0.8461 | 0.492 | 3.4 | 4 |  | 0.02259 | 0.0851 | 0.0955 | 0.0195 | 0.99999 | [45 empty] | 0
P | 2018-03-20T10:09:23+08:00 | 2018-03-20T13:40:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI662G | CPSM Final Test | PASS | 129 | 159.98 | 0.37 | -0.69 | 56.02 | 50.88 | 1.44 | 0.13 | 1 | -11 | 1 | -0.8381 | 0.5262 | 8.4 | 6 |  | 0 | 0.0673 | 0.2339 | 0.0146 | 0.99999 | [816 empty] | 0
P | 2018-03-20T10:49:56+08:00 | 2018-03-20T13:49:01+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI664G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.92 | 56.22 | 50.63 | 1.44 | 0.13 | -1 | -2 | 1 | -0.8042 | 0.5028 | 3.4 | -9 |  | 0 | 0.0466 | 0.0826 | -0.0243 | 0.99999 | [818 empty] | 0
P | 2018-03-20T05:09:02+08:00 | 2018-03-20T13:49:19+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI661G | CPSM Final Test | FAIL | [383 empty] | 6301
P | 2018-03-20T09:29:10+08:00 | 2018-03-20T14:55:25+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI661G | CPSM Final Test | PASS | 129 | 165.29 | 0.37 | -0.46 | 56.4 | 54.61 | 1.44 | 0.13 | 1 | 2 | 0 | -0.914 | 0.4932 | 3.8 | 60 |  | 0 | 0.1462 | 0.1591 | 0.0029 | 0.99997 | [362 empty] | 0
P | 2018-03-20T12:21:01+08:00 | 2018-03-20T15:40:34+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI643G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.16 | 56.53 | 52.75 | 1.82 | 0.13 | -1 | 2 | 0 | -0.8121 | 0.511 | 3.3 | 20 |  | 0 | 0.0572 | 0.3119 | -0.0568 | 1 | [819 empty] | 0
P | 2018-03-20T15:54:32+08:00 | 2018-03-20T18:43:42+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI665G | CPSM Final Test | FAIL | [841 empty] | FALSE
P | 2018-03-20T15:46:38+08:00 | 2018-03-20T18:53:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI666G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.69 | 56.18 | 49.95 | 1.21 | 0.13 | 1 | -3 | 0 | -0.9457 | 0.532 | 3.4 | 23 |  | 0 | 0.1601 | 0.2925 | 0.0117 | 1 | [821 empty] | 0
P | 2018-03-20T15:37:54+08:00 | 2018-03-20T18:55:58+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI667G | CPSM Final Test | PASS | 129 | 163.4 | 0.08 | 0.07 | 56.8 | 50.12 | 1.59 | 0.13 | -1 | -11 | 1 | -0.9575 | 0.5386 | 9.9 | -39 |  | 0 | 0.0726 | 0.0863 | -0.0214 | 0.99999 | [806 empty] | 0
P | 2018-03-20T14:40:39+08:00 | 2018-03-20T18:58:26+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI669G | CPSM Final Test | PASS | 129 | 159.44 | 0.56 | -0.46 | 56.34 | 50.55 | 1.36 | 0.13 | 1 | 2 | -1 | -0.9159 | 0.5054 | 9.2 | 23 |  | 0 | 0.0579 | 0.1932 | -0.0038 | 0.99997 | [822 empty] | 0
P | 2018-03-20T14:51:09+08:00 | 2018-03-20T19:00:20+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI668G | CPSM Final Test | PASS | 129 | 167.72 | 0.46 | 0.61 | 56.44 | 51.31 | 1.21 | 0.13 | 1 | 2 | -3 | -0.9219 | 0.5462 | 4.8 | 7 |  | 0 | 0.0962 | 0.0864 | -0.0306 | 0.99999 | [551 empty] | 0
P | 2018-03-20T14:03:10+08:00 | 2018-03-20T19:09:38+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI670G | CPSM Final Test | FAIL
P | 2018-03-20T13:02:36+08:00 | 2018-03-20T20:10:12+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI670G | CPSM Final Test | FAIL
P | 2018-03-20T16:30:31+08:00 | 2018-03-20T20:20:25+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI671G | CPSM Final Test | PASS | 129 | 168.17 | 0.17 | -0.23 | 56.33 | 50.72 | 1.67 | 0.13 | 0 | -3 | 0 | -0.9458 | 0.5156 | 3.8 | 14 |  | 0 | 0.1715 | 0.1653 | 0.0179 | 0.99998 | [346 empty] | 0
P | 2018-03-20T17:08:09+08:00 | 2018-03-20T20:50:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI672G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | 0.07 | 56.81 | 50.72 | 1.74 | 0.13 | -1 | 3 | 1 | -0.8681 | 0.4938 | 10.4 | 0 |  | 0 | 0.099 | 0.084 | 0.035 | 0.99997 | [362 empty] | 0
P | 2018-03-20T17:19:07+08:00 | 2018-03-20T21:32:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI665G | CPSM Final Test | PASS | 129 | 157.1 | 0.27 | -0.46 | 56.33 | 51.65 | 1.44 | 0.13 | 0 | -2 | 1 | -0.9378 | 0.535 | 6.7 | 0 |  | 0 | 0.0686 | 0.622 | 0.0289 | 1 | [823 empty] | 0
P | 2018-03-19T19:13:48+08:00 | 2018-03-21T13:59:00+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI670G | CPSM Final Test | FAIL
P | 2018-03-20T16:36:24+08:00 | 2018-03-21T15:02:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI670G | CPSM Final Test | PASS | 129 | 169.34 | 0.08 | -0.08 | 56.53 | 49.87 | 1.82 | 0.13 | -1 | 4 | 1 | -0.7602 | 0.5048 | 0.3 | -2 |  | 0.00656 | 0.1931 | 0.394 | -0.0142 | 0.99998 | [824 empty] | 0
P | 2018-03-22T13:32:36+08:00 | 2018-03-22T16:02:56+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM911G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-03-22T18:22:57+08:00 | 2018-03-22T21:18:38+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | C18USM916G | CPSM Final Test | PASS |  | 164.66 | 0.78 | -0.54 | 56.41 | 54.19 | 1.74 | 0.13 | 1 | -4 | 0 | -0.806 | 0.5102 | 1.2 | 15 |  |  | 0.1432 | 0.2003 | 0.038 | 0.99997 |  |  |  |  | 133 |  | 0 | [818 empty] | 0
P | 2018-03-22T18:48:58+08:00 | 2018-03-22T21:35:12+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM914G | CPSM Final Test | PASS |  | 164.12 | 0.17 | -0.61 | 56.23 | 48.85 | 1.44 | 0.13 | -1 | -1 | 0 | -0.9837 | 0.544 | 4.2 | 0 |  |  | 0.0494 | 0.2837 | -0.0329 | 1 |  |  |  |  | 133 |  | 0 | [78 empty] | 0
P | 2018-03-22T18:49:39+08:00 | 2018-03-22T21:36:34+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM911G | CPSM Final Test | PASS |  | 164.21 | 0.37 | -0.31 | 56.53 | 51.9 | 1.36 | 0.13 | 0 | -2 | 0 | -0.9816 | 0.499 | 0.6 | 35 |  |  | 0.0455 | 0.1417 | 0.0233 | 0.99995 |  |  |  |  | 133 |  | 0 | [819 empty] | 0
P | 2018-03-22T17:32:33+08:00 | 2018-03-22T22:03:00+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM917G | CPSM Final Test | PASS |  | 164.3 | -0.02 | -0.54 | 56.32 | 49.36 | 1.59 | 0.13 | 0 | 1 | 1 | -0.8759 | 0.4902 | 2.2 | 19 |  |  | 0.0461 | 0.0912 | 0.0426 | 1 |  |  |  |  | 133 |  | 0 | [817 empty] | 0
P | 2018-03-22T19:04:12+08:00 | 2018-03-22T22:12:54+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM915G | CPSM Final Test | PASS |  | 159.17 | 0.17 | -0.61 | 56.25 | 49.61 | 1.44 | 0.13 | 1 | 4 | 0 | -0.858 | 0.4944 | 0.8 | 18 |  |  | 0.0752 | 0.4344 | 0.0369 | 0.99998 |  |  |  |  | 133 |  | 0.03247 | [141 empty] | 0
P | 2018-03-22T03:57:20+08:00 | 2018-03-23T09:43:15+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM912G | CPSM Final Test | FAIL | [52 empty] | 5235
P | 2018-03-22T03:54:18+08:00 | 2018-03-23T09:46:17+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM912G | CPSM Final Test | FAIL
P | 2018-03-23T09:18:19+08:00 | 2018-03-23T10:39:24+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM919G | CPSM Final Test | FAIL | [5 empty] | 47.33
P | 2018-03-22T18:50:18+08:00 | 2018-03-23T10:51:57+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM912G | CPSM Final Test | PASS |  | 170.78 | -0.02 | -0.84 | 56.49 | 50.88 | 1.36 | 0.13 | 1 | -3 | 0 | -0.9897 | 0.48 | 9.9 | 14 |  |  | 0.1055 | 0.7074 | -0.0456 | 0.99999 |  |  |  |  | 133 |  | 0 | [24 empty] | 0
P | 2018-03-23T09:53:47+08:00 | 2018-03-23T13:18:40+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM918G | CPSM Final Test | PASS |  | 160.88 | -0.02 | -0.54 | 56.12 | 50.55 | 1.44 | 0.13 | 0 | -3 | 1 | -0.808 | 0.501 | 1.1 | 37 |  |  | 0.112 | 0.164 | -0.0183 | 1 |  |  |  |  | 133 |  | 0 | [18 empty] | 0
P | 2018-03-23T10:06:19+08:00 | 2018-03-23T13:24:06+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM920G | CPSM Final Test | PASS |  | 159.8 | -0.02 | -0.77 | 56.17 | 50.04 | 1.36 | 0.13 | -1 | -2 | -2 | -0.7862 | 0.488 | 0.9 | 15 |  |  | 0.1356 | 0.1574 | -0.0074 | 0.99999 |  |  |  |  | 133 |  | 0.03008 | [577 empty] | 0
P | 2018-03-23T06:46:41+08:00 | 2018-03-23T13:22:57+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM913G | CPSM Final Test | FAIL
P | 2018-03-23T10:11:58+08:00 | 2018-03-23T13:24:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM922G | CPSM Final Test | PASS |  | 173.21 | 0.17 | -0.46 | 56.24 | 49.45 | 1.44 | 0.13 | 0 | -2 | -1 | -0.834 | 0.5046 | 5.4 | 1 |  |  | 0.0356 | 0.3092 | -0.0156 | 0.99999 |  |  |  |  | 133 |  | 0.00789 | [793 empty] | 0
P | 2018-03-23T06:29:05+08:00 | 2018-03-23T13:28:38+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM919G | CPSM Final Test | FAIL
P | 2018-03-23T09:58:51+08:00 | 2018-03-23T13:35:20+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM919G | CPSM Final Test | PASS |  | 163.94 | 0.65 | -0.54 | 56.34 | 49.87 | 1.67 | 0.13 | -1 | 2 | 1 | -0.8481 | 0.512 | 7.9 | 10 |  |  | 0.053 | 0.2427 | 0.0145 | 1 |  |  |  |  | 133 |  | 0 | [820 empty] | 0
P | 2018-03-23T10:08:46+08:00 | 2018-03-23T13:47:18+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | C18USM923G | CPSM Final Test | PASS |  | 170.6 | -0.21 | -0.54 | 56.62 | 51.39 | 1.97 | 0.13 | 1 | -2 | 1 | -0.8101 | 0.487 | 5.2 | 22 |  |  | 0.1555 | 0.1993 | 0.0047 | 0.99997 |  |  |  |  | 133 |  | 0.01322 | [525 empty] | 0
P | 2018-03-23T14:31:22+08:00 | 2018-03-23T14:45:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM921G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-23T10:04:49+08:00 | 2018-03-23T15:12:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM913G | CPSM Final Test | PASS |  | 161.6 | 0.17 | -0.61 | 56.37 | 50.88 | 1.36 | 0.13 | 1 | 2 | 0 | -0.9577 | 0.4768 | 1.9 | 42 |  |  | 0.0894 | 0.0722 | -0.0249 | 0.99999 |  |  |  |  | 133 |  | 0 | [821 empty] | 0
P | 2018-03-23T13:57:06+08:00 | 2018-03-23T15:19:55+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM921G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-23T12:12:12+08:00 | 2018-03-23T15:48:39+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM924G | CPSM Final Test | PASS |  | 161.51 | -0.3 | -0.54 | 56.73 | 55.12 | 1.82 | 0.13 | 0 | 3 | 1 | -0.7941 | 0.5312 | 5.9 | 33 |  |  | 0.1067 | 0.4275 | -0.0067 | 0.99996 |  |  |  |  | 133 |  | 0 | [24 empty] | 0
P | 2018-03-23T11:25:44+08:00 | 2018-03-23T16:32:06+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI673G | CPSM Final Test | FAIL
P | 2018-03-23T10:39:20+08:00 | 2018-03-23T17:18:30+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI673G | CPSM Final Test | FAIL
P | 2018-03-23T13:49:06+08:00 | 2018-03-23T17:19:34+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI674G | CPSM Final Test | PASS | 129 | 160.79 | -0.02 | -0.84 | 56.23 | 48.85 | 1.51 | 0.13 | 1 | -2 | -1 | -0.8201 | 0.5392 | 2.3 | 1 |  | 0 | 0.0362 | 0.1008 | -0.0634 | 1 | [381 empty] | 0
P | 2018-03-23T13:53:01+08:00 | 2018-03-23T17:21:02+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI675G | CPSM Final Test | PASS | 129 | 166.73 | -0.02 | -0.92 | 56.17 | 50.88 | 1.67 | 0.13 | -1 | -11 | 0 | -0.8421 | 0.546 | 0.3 | 0 |  | 0 | 0.0705 | 0.2068 | -0.0357 | 1 | [373 empty] | 0
P | 2018-03-23T10:53:20+08:00 | 2018-03-23T18:23:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM921G | CPSM Final Test | FAIL
P | 2018-03-23T13:58:55+08:00 | 2018-03-23T18:31:58+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI673G | CPSM Final Test | PASS | 129 | 157.01 | -0.21 | -1.07 | 56.11 | 48.68 | 1.74 | 0.13 | 0 | 1 | 0 | -0.852 | 0.502 | 4.3 | 34 |  | 0 | 0.1206 | 0.3125 | -0.0131 | 0.99999 | [816 empty] | 0
P | 2018-03-23T16:30:32+08:00 | 2018-03-23T20:29:20+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI677G | CPSM Final Test | PASS | 129 | 163.67 | -0.21 | -0.92 | 56.47 | 48.43 | 1.13 | 0.13 | 0 | -2 | 1 | -0.9139 | 0.5 | 2.1 | -9 |  | 0.01847 | 0.1055 | 0.5733 | 0.0161 | 0.99999 | [829 empty] | 0
P | 2018-03-23T16:27:17+08:00 | 2018-03-23T20:31:38+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI676G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | -0.77 | 56.58 | 48.09 | 1.51 | 0.13 | 0 | -2 | 3 | -0.832 | 0.5086 | 4.8 | 13 |  | 0 | 0.0715 | 0.1973 | -0.0237 | 0.99997 | [830 empty] | 0
P | 2018-03-23T16:35:29+08:00 | 2018-03-23T20:58:47+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI679G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.46 | 56.49 | 50.63 | 1.44 | 0.13 | 0 | -3 | -1 | -0.808 | 0.525 | 2 | 28 |  | 0 | 0.1137 | 0.2717 | -0.021 | 0.99999 | [352 empty] | 0
P | 2018-03-23T16:45:33+08:00 | 2018-03-23T21:09:45+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI678G | CPSM Final Test | PASS | 129 | 152.33 | 0.17 | -0.92 | 56.31 | 50.29 | 1.13 | 0.13 | 0 | -12 | 0 | -0.8979 | 0.487 | 5.3 | 39 |  | 0 | 0.0981 | 0.1537 | -0.0372 | 0.99999 | [296 empty] | 0
P | 2018-03-24T09:01:27+08:00 | 2018-03-24T09:49:08+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI684G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-03-24T08:44:10+08:00 | 2018-03-24T09:50:39+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI680G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-03-24T07:35:46+08:00 | 2018-03-24T10:59:03+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI680G | CPSM Final Test | FAIL | [14 empty] | 1820
P | 2018-03-22T17:33:14+08:00 | 2018-03-24T11:43:47+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM921G | CPSM Final Test | FAIL | [27 empty] | 0.84002
P | 2018-03-24T06:57:43+08:00 | 2018-03-24T11:44:56+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI682G | CPSM Final Test | FAIL
P | 2018-03-24T09:13:33+08:00 | 2018-03-24T12:18:33+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI681G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.54 | 56.47 | 49.19 | 1.59 | 0.13 | -1 | 2 | 0 | -0.7522 | 0.5272 | 3 | 11 |  | 0 | 0.1232 | 0.1391 | -0.0203 | 0.99999 | [831 empty] | 0
P | 2018-03-24T09:25:18+08:00 | 2018-03-24T12:21:50+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI684G | CPSM Final Test | PASS | 129 | 164.21 | -0.21 | -1.22 | 56.02 | 50.12 | 1.36 | 0.13 | 1 | -2 | 1 | -0.7583 | 0.493 | 6.3 | -14 |  | 0 | 0.1197 | 0.1147 | 0.0069 | 1 | [822 empty] | 0
P | 2018-03-23T14:38:30+08:00 | 2018-03-24T12:46:51+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM921G | CPSM Final Test | PASS |  | 162.77 | 0.27 | -0.77 | 56.22 | 52.07 | 1.9 | 0.13 | 1 | -8 | 1 | -0.9577 | 0.494 | 1.8 | 11 |  |  | 0.0505 | 0.2852 | 0.0083 | 1 |  |  |  |  | 133 |  | 0.01396 | [825 empty] | 0
P | 2018-03-24T09:17:25+08:00 | 2018-03-24T12:45:05+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI680G | CPSM Final Test | PASS | 129 | 167.09 | -0.02 | -0.61 | 56.36 | 51.39 | 1.59 | 0.13 | 0 | -3 | 0 | -0.7662 | 0.5374 | 1.3 | 13 |  | 0 | 0.1039 | 0.0722 | -0.0008 | 1 | [833 empty] | 0
P | 2018-03-24T09:21:19+08:00 | 2018-03-24T13:31:00+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI682G | CPSM Final Test | PASS | 129 | 167 | -0.02 | -0.92 | 56.33 | 51.39 | 1.67 | 0.13 | -1 | 1 | 1 | -0.7881 | 0.485 | 3.1 | 15 |  | 0 | 0.0987 | 0.1289 | -0.0117 | 1 | [7 empty] | 0
P | 2018-03-24T07:51:19+08:00 | 2018-03-24T14:30:36+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI685G | CPSM Final Test | FAIL
P | 2018-03-24T11:20:06+08:00 | 2018-03-24T15:34:58+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI683G | CPSM Final Test | PASS | 129 | 161.69 | 0.17 | -0.92 | 56.21 | 50.97 | 1.9 | 0.13 | -1 | -2 | 1 | -0.8859 | 0.492 | 3.8 | 50 |  | 0.00817 | 0.0231 | 0.2453 | 0.0061 | 1 | [490 empty] | 0
P | 2018-03-24T11:16:24+08:00 | 2018-03-24T15:35:22+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI686G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | -0.99 | 56.32 | 50.55 | 1.44 | 0.13 | 1 | -2 | 0 | -0.6704 | 0.503 | 0.4 | 7 |  | 0 | 0.1056 | 0.2251 | 0.0307 | 0.99999 | [210 empty] | 0
P | 2018-03-24T11:10:57+08:00 | 2018-03-24T16:19:52+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI685G | CPSM Final Test | PASS | 129 | 153.59 | -0.02 | -0.84 | 56.33 | 50.21 | 1.21 | 0.13 | -1 | -2 | 0 | -0.9139 | 0.5074 | 0.2 | 10 |  | 0.00552 | 0.1059 | 0.191 | 0.0301 | 0.99999 | [787 empty] | 0
P | 2018-03-24T13:36:03+08:00 | 2018-03-24T17:28:01+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI688G | CPSM Final Test | PASS | 129 | 167.36 | -0.02 | 0.23 | 56.43 | 56.3 | 1.51 | 0.13 | 0 | -11 | 1 | -0.7402 | 0.493 | 0.8 | 12 |  | 0.0289 | 0.0426 | 0.158 | -0.0157 | 0.99999 | [834 empty] | 0
P | 2018-03-24T14:07:48+08:00 | 2018-03-24T17:30:18+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI687G | CPSM Final Test | PASS | 129 | 162.41 | 0.17 | -0.46 | 56.58 | 50.97 | 1.82 | 0.13 | -1 | -1 | 3 | -0.8819 | 0.5456 | 5.4 | 9 |  | 0 | 0.1159 | 0.1019 | 0.0214 | 0.99999 | [833 empty] | 0
P | 2018-03-26T07:56:23+08:00 | 2018-03-26T12:01:56+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI692G | CPSM Final Test | FAIL
P | 2018-03-26T09:49:32+08:00 | 2018-03-26T13:07:20+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI689G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -0.84 | 56.22 | 50.29 | 1.59 | 0.13 | 0 | -3 | -1 | -0.852 | 0.4988 | 0.5 | 10 |  | 0 | 0.0348 | 0.1386 | 0.0296 | 0.99999 | [241 empty] | 0
P | 2018-03-26T09:59:10+08:00 | 2018-03-26T13:46:24+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI692G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -1.07 | 56.12 | 50.46 | 1.21 | 0.13 | -1 | 1 | -1 | -0.85 | 0.5032 | 5.2 | 3 |  | 0.00575 | 0.104 | 0.034 | 0.0218 | 0.99999 | [608 empty] | 0
P | 2018-03-26T11:40:26+08:00 | 2018-03-26T15:36:31+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI694G | CPSM Final Test | FAIL
P | 2018-03-26T12:31:47+08:00 | 2018-03-26T16:30:13+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI690G | CPSM Final Test | PASS | 129 | 165.11 | 0.27 | -0.99 | 56.21 | 49.53 | 1.51 | 0.13 | -1 | -3 | 1 | -0.8899 | 0.499 | 1.3 | 5 |  | 0 | 0.0395 | 0.2648 | -0.012 | 0.99999 | [835 empty] | 0
P | 2018-03-26T13:44:18+08:00 | 2018-03-26T16:40:40+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI695G | CPSM Final Test | PASS | 129 | 161.24 | -0.3 | -0.61 | 56.42 | 54.95 | 1.44 | 0.13 | 1 | 5 | 1 | -0.8759 | 0.512 | 4 | 42 |  | 0.00808 | 0.0937 | 0.2525 | 0.0016 | 1 | [836 empty] | 0
P | 2018-03-26T13:41:10+08:00 | 2018-03-26T16:45:50+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI696G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.61 | 56.48 | 50.04 | 1.67 | 0.13 | 0 | -2 | 1 | -0.806 | 0.512 | 1.1 | 5 |  | 0.00885 | 0.1407 | 0.0807 | 0.0465 | 0.99996 | [719 empty] | 0
P | 2018-03-26T13:38:29+08:00 | 2018-03-26T17:20:29+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI694G | CPSM Final Test | PASS | 129 | 166.82 | -0.02 | -1.15 | 56.14 | 51.99 | 1.51 | 0.13 | 0 | -2 | 0 | -0.87 | 0.4848 | 1.5 | 13 |  | 0.0082 | 0.3136 | 0.6036 | -0.0211 | 0.99999 | [832 empty] | 0
P | 2018-03-26T09:31:31+08:00 | 2018-03-26T17:41:46+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI693G | CPSM Final Test | FAIL
P | 2018-03-26T13:02:56+08:00 | 2018-03-26T18:22:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI697G | CPSM Final Test | FAIL
P | 2018-03-26T12:56:49+08:00 | 2018-03-26T18:48:50+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI698G | CPSM Final Test | FAIL
P | 2018-03-26T15:52:50+08:00 | 2018-03-26T18:49:20+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI698G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -0.46 | 56.5 | 51.82 | 1.44 | -0.14 | -2 | -3 | 1 | -0.8599 | 0.523 | 0.5 | -4 |  | 0 | 0.1328 | 0.0963 | -0.0219 | 0.99999 | [837 empty] | 0
P | 2018-03-26T07:42:37+08:00 | 2018-03-26T19:30:40+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI693G | CPSM Final Test | FAIL
P | 2018-03-26T15:42:57+08:00 | 2018-03-26T20:07:52+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI697G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.69 | 56.6 | 50.46 | 1.74 | 0.13 | 1 | -11 | 2 | -0.8739 | 0.4928 | 4.2 | 17 |  | 0.00797 | 0.0431 | 0.215 | -0.0011 | 0.99999 | [838 empty] | 0
P | 2018-03-26T17:17:10+08:00 | 2018-03-26T20:38:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI702G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.69 | 56.51 | 50.8 | 1.59 | 0.13 | 1 | 3 | -1 | -0.8459 | 0.4946 | 0.6 | 9 |  | 0 | 0.0606 | 0.4698 | -0.0089 | 1 | [808 empty] | 0
P | 2018-03-26T17:06:13+08:00 | 2018-03-26T20:47:36+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI700G | CPSM Final Test | PASS | 129 | 174.38 | -0.02 | -0.77 | 56.47 | 51.73 | 1.59 | 0.13 | 0 | 2 | 0 | -0.862 | 0.4918 | 1.7 | 35 |  | 0 | 0.1286 | 0.2305 | 0.0419 | 0.99998 | [839 empty] | 0
P | 2018-03-26T17:09:25+08:00 | 2018-03-26T20:49:22+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI699G | CPSM Final Test | PASS | 129 | 169.7 | -0.02 | -0.77 | 56.11 | 51.48 | 1.51 | 0.13 | 0 | 2 | 0 | -0.8918 | 0.548 | 3.2 | 2 |  | 0.03814 | 0.0597 | 0.2876 | -0.0205 | 1 | [840 empty] | 0
P | 2018-03-26T13:36:39+08:00 | 2018-03-26T21:17:10+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI693G | CPSM Final Test | PASS | 129 | 162.5 | -0.02 | -0.69 | 56.33 | 49.95 | 1.51 | 0.13 | -1 | -2 | 1 | -0.8839 | 0.501 | 5.5 | 2 |  | 0.02241 | 0.0827 | 0.144 | 0.0146 | 1 | [234 empty] | 0
P | 2018-03-26T19:53:34+08:00 | 2018-03-26T22:33:52+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI691G | CPSM Final Test | PASS | 129 | 164.93 | -0.02 | -0.99 | 56.21 | 50.04 | 1.29 | 0.13 | 0 | -2 | 0 | -0.9517 | 0.4592 | 6.9 | -1 |  | 0 | 0.0257 | 0.0911 | 0.0045 | 1 | [841 empty] | 0
P | 2018-03-26T02:06:54+08:00 | 2018-03-27T08:59:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI701G | CPSM Final Test | FAIL
P | 2018-03-26T17:33:07+08:00 | 2018-03-27T09:06:05+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI701G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -0.54 | 56.34 | 50.55 | 1.36 | 0.13 | -1 | -11 | 1 | -0.878 | 0.506 | 8.7 | 21 |  | 0.01309 | 0.0265 | 0.2817 | -0.03 | 0.99999 | [775 empty] | 0
P | 2018-03-27T09:12:45+08:00 | 2018-03-27T10:48:04+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI708G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-03-27T10:04:32+08:00 | 2018-03-27T13:13:29+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI703G | CPSM Final Test | PASS | 129 | 167.36 | -0.02 | -0.54 | 56.34 | 49.87 | 1.21 | 0.13 | -1 | -2 | 1 | -0.7921 | 0.53 | 1.7 | 7 |  | 0 | 0.0708 | 0.1945 | -0.0197 | 1 | [842 empty] | 0
P | 2018-03-27T07:01:25+08:00 | 2018-03-27T13:24:04+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI704G | CPSM Final Test | FAIL | [864 empty] | 12642
P | 2018-03-27T10:20:57+08:00 | 2018-03-27T13:26:06+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI707G | CPSM Final Test | PASS | 129 | 164.39 | 0.17 | -1.07 | 56.22 | 51.65 | 1.36 | 0.13 | -1 | 2 | 1 | -0.7722 | 0.478 | 5.5 | 14 |  | 0.00677 | 0.0583 | 0.2134 | 0.0137 | 0.99997 | [844 empty] | 0
P | 2018-03-27T10:16:59+08:00 | 2018-03-27T13:32:12+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI706G | CPSM Final Test | PASS | 129 | 171.41 | 0.17 | 0.38 | 56.47 | 49.7 | 1.59 | 0.13 | 0 | -13 | -1 | -0.8681 | 0.494 | 0.7 | 2 |  | 0.0251 | 0.0265 | 0.5196 | -0.03 | 0.99999 | [845 empty] | 0
P | 2018-03-27T06:41:42+08:00 | 2018-03-27T13:42:25+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI705G | CPSM Final Test | FAIL
P | 2018-03-27T10:12:04+08:00 | 2018-03-27T13:51:43+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI705G | CPSM Final Test | PASS | 129 | 166.55 | 0.46 | -0.77 | 56.63 | 51.22 | 1.29 | 0.13 | 0 | 2 | 1 | -0.856 | 0.4958 | 8.8 | 3 |  | 0.01092 | 0.0361 | 0.0258 | 0.0043 | 0.99999 | [846 empty] | 0
P | 2018-03-27T10:59:30+08:00 | 2018-03-27T14:08:15+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI708G | CPSM Final Test | PASS | 129 | 160.97 | -0.21 | -0.61 | 56.96 | 54.61 | 1.36 | 0.13 | -1 | -2 | 1 | -0.8879 | 0.545 | 6 | 13 |  | 0 | 0.0775 | 0.0982 | -0.0291 | 1 | [551 empty] | 0
P | 2018-03-27T10:12:45+08:00 | 2018-03-27T14:29:50+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI704G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.54 | 56.5 | 52.75 | 1.59 | 0.13 | -1 | 2 | 1 | -0.864 | 0.5374 | 1.6 | 79 |  | 0.00699 | 0.0816 | 0.1668 | -0.0238 | 1 | [843 empty] | 0
P | 2018-03-28T16:52:33+08:00 | 2018-03-28T20:24:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | C18USM925G | CPSM Final Test | PASS |  | 166.55 | 0.08 | 0 | 57.08 | 55.03 | 1.44 | 0.13 | 1 | 4 | -1 | -0.9497 | 0.5148 | 2.6 | 9 |  |  | 0.1365 | 0.0825 | -0.0088 | 0.99999 |  |  |  |  | 133 |  | 0 | [840 empty] | 0
P | 2018-03-28T17:03:14+08:00 | 2018-03-28T20:25:13+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | C18USM926G | CPSM Final Test | PASS |  | 165.38 | 0.65 | -0.38 | 56.93 | 54.86 | 1.21 | 0.13 | 0 | -3 | -1 | -0.9497 | 0.5144 | 4.8 | 40 |  |  | 0.0909 | 0.1281 | -0.0356 | 1 |  |  |  |  | 133 |  | 0 | [841 empty] | 0
P | 2018-03-29T07:50:41+08:00 | 2018-03-29T11:28:41+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | C18USM927G | CPSM Final Test | FAIL
P | 2018-03-29T09:59:35+08:00 | 2018-03-29T12:45:48+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | C18USM930G | CPSM Final Test | PASS |  | 159.71 | 0.08 | -0.61 | 56.5 | 52.49 | 1.21 | 0.13 | -1 | -2 | 0 | -0.8441 | 0.5368 | 0.6 | 12 |  |  | 0.0839 | 0.1027 | 0.0059 | 0.99999 |  |  |  |  | 133 |  | 0 | [842 empty] | 0
P | 2018-03-29T05:20:39+08:00 | 2018-03-29T13:16:55+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | C18USM929G | CPSM Final Test | FAIL
P | 2018-03-29T09:18:47+08:00 | 2018-03-29T13:17:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | C18USM929G | CPSM Final Test | PASS |  | 167.18 | 0.37 | -0.31 | 57.08 | 53.34 | 1.51 | 0.13 | -1 | -3 | 0 | -0.8759 | 0.4808 | 4.9 | 10 |  |  | 0.0502 | 0.1113 | -0.0148 | 1 |  |  |  |  | 133 |  | 0 | [843 empty] | 0
P | 2018-03-29T10:31:08+08:00 | 2018-03-29T13:23:36+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | C18USM928G | CPSM Final Test | PASS |  | 164.3 | 0.37 | -0.08 | 57.06 | 51.31 | 1.29 | 0.13 | 1 | -2 | -1 | -0.9178 | 0.4924 | 1.8 | -2 |  |  | 0.0891 | 0.2397 | -0.0273 | 0.99999 |  |  |  |  | 133 |  | 0 | [844 empty] | 0
P | 2018-03-29T05:37:53+08:00 | 2018-03-29T13:41:29+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | C18USM927G | CPSM Final Test | FAIL | [873 empty] | 34672
P | 2018-03-29T11:22:33+08:00 | 2018-03-29T14:19:46+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | C18USM932G | CPSM Final Test | PASS |  | 167.18 | 0.65 | -0.31 | 56.73 | 56.47 | 1.59 | 0.13 | 0 | 2 | 1 | -0.8062 | 0.4932 | 3.2 | 9 |  |  | 0.0638 | 0.4485 | -0.0133 | 1 |  |  |  |  | 133 |  | 0 | [814 empty] | 0
P | 2018-03-29T13:51:16+08:00 | 2018-03-29T17:08:11+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | C18USM931G | CPSM Final Test | PASS |  | 164.03 | 0.17 | -0.46 | 56.98 | 51.14 | 1.36 | 0.13 | 1 | -3 | 0 | -0.9098 | 0.5244 | 4.7 | 32 |  |  | 0.1338 | 0.2337 | 0.0047 | 0.99999 |  |  |  |  | 133 |  | 0 | [227 empty] | 0
P | 2018-03-29T14:26:50+08:00 | 2018-03-29T17:42:28+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | C18USM933G | CPSM Final Test | PASS |  | 163.4 | 0.37 | 0.15 | 56.93 | 53.68 | 1.59 | 0.13 | 0 | -3 | 0 | -0.868 | 0.5424 | 0.3 | 0 |  |  | 0.0715 | 0.1864 | -0.0389 | 1 |  |  |  |  | 133 |  | 0 | [78 empty] | 0
P | 2018-03-29T13:46:08+08:00 | 2018-03-29T18:16:20+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI711G | CPSM Final Test | FAIL
P | 2018-03-29T10:14:34+08:00 | 2018-03-29T18:17:27+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | C18USM934G | CPSM Final Test | FAIL
P | 2018-03-29T14:16:00+08:00 | 2018-03-29T18:18:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | C18USM934G | CPSM Final Test | PASS |  | 167.81 | 0.27 | -0.31 | 57.12 | 53.59 | 1.29 | 0.13 | -1 | 4 | 1 | -0.854 | 0.4622 | 3.4 | 10 |  |  | 0.1166 | 0.1561 | -0.0022 | 1 |  |  |  |  | 133 |  | 0 | [846 empty] | 0
P | 2018-03-29T16:12:23+08:00 | 2018-03-29T19:03:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI710G | CPSM Final Test | PASS | 129 | 170.87 | 0.37 | -0.69 | 57.09 | 52.92 | 1.13 | 0.13 | -1 | 2 | 2 | -0.7722 | 0.5192 | 12.3 | -20 |  | 0 | 0.2514 | 0.4014 | 0.038 | 0.99998 | [196 empty] | 0
P | 2018-03-29T16:22:23+08:00 | 2018-03-29T19:11:32+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI709G | CPSM Final Test | PASS | 129 | 166.46 | 0.37 | -1.22 | 56.72 | 57.49 | 0.83 | 0.13 | 0 | -10 | 0 | -0.9019 | 0.5478 | 3.7 | 46 |  | 0 | 0.0573 | 0.1015 | 0.0056 | 0.99999 | [843 empty] | 0
P | 2018-03-29T16:43:06+08:00 | 2018-03-29T19:37:06+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | C18USM927G | CPSM Final Test | PASS |  | 162.95 | 0.37 | -0.23 | 56.89 | 53.59 | 1.51 | 0.13 | 1 | 3 | 0 | -0.9319 | 0.4968 | 1 | 1 |  |  | 0.1582 | 0.145 | 0.0201 | 0.99999 |  |  |  |  | 133 |  | 0.01383 | [847 empty] | 0
P | 2018-03-29T19:42:50+08:00 | 2018-03-29T20:01:25+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI714G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-29T16:01:14+08:00 | 2018-03-29T20:00:42+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI711G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.84 | 57.06 | 54.36 | 1.44 | 0.13 | 1 | 2 | -1 | -0.84 | 0.543 | 3.8 | 53 |  | 0.01031 | 0.0585 | 0.3098 | 0.0296 | 0.99999 | [855 empty] | 0
P | 2018-03-29T18:40:41+08:00 | 2018-03-29T21:35:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI712G | CPSM Final Test | PASS | 129 | 162.41 | -0.02 | -0.69 | 56.71 | 56.9 | 1.51 | 0.13 | 0 | 2 | 1 | -0.8861 | 0.514 | 5.8 | -2 |  | 0 | 0.1599 | 0.1441 | -0.0043 | 1 | [856 empty] | 0
P | 2018-03-29T18:31:38+08:00 | 2018-03-29T21:38:20+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI713G | CPSM Final Test | PASS | 129 | 164.75 | 0.08 | -0.46 | 57.07 | 51.39 | 1.51 | 0.13 | 1 | -2 | 0 | -0.8979 | 0.5208 | 9.8 | 168 |  | 0 | 0.0325 | 0.2117 | -0.0152 | 1 | [857 empty] | 0
P | 2018-03-28T05:20:38+08:00 | 2018-03-31T10:23:37+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI714G | CPSM Final Test | FAIL
P | 2018-03-29T19:52:08+08:00 | 2018-03-31T11:36:49+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI714G | CPSM Final Test | PASS | 129 | 157.46 | 0.56 | -0.16 | 56.42 | 50.46 | 1.59 | 0.13 | 1 | -2 | 0 | -0.826 | 0.5176 | 8 | 13 |  | 0 | 0.1256 | 0.2957 | 0.0411 | 0.99998 | [251 empty] | 0
P | 2018-03-31T09:28:03+08:00 | 2018-03-31T11:52:14+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI717G | CPSM Final Test | FAIL
P | 2018-03-31T10:16:02+08:00 | 2018-03-31T13:31:22+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI716G | CPSM Final Test | PASS | 129 | 165.74 | 0.37 | -0.38 | 56.84 | 54.53 | 1.67 | 0.13 | 0 | -2 | 1 | -0.8102 | 0.5206 | 10.9 | -5 |  | 0.01909 | 0.0988 | 0.2296 | -0.0098 | 0.99999 | [511 empty] | 0
P | 2018-03-31T07:26:16+08:00 | 2018-03-31T13:31:10+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI715G | CPSM Final Test | FAIL
P | 2018-03-31T10:28:43+08:00 | 2018-03-31T13:37:56+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI715G | CPSM Final Test | PASS | 129 | 166.55 | 0.08 | -0.69 | 56.75 | 54.53 | 1.59 | 0.22 | 0 | -3 | 1 | -0.8998 | 0.5068 | 1 | 11 |  | 0 | 0.0839 | 0.0698 | -0.0161 | 1 | [858 empty] | 0
P | 2018-03-31T13:40:16+08:00 | 2018-03-31T16:51:44+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI718G | CPSM Final Test | PASS | 129 | 163.31 | -0.21 | -0.69 | 56.92 | 54.02 | 1.51 | 0.13 | 0 | -2 | 0 | -0.7702 | 0.513 | 3.5 | -20 |  | 0 | 0.0968 | 0.1844 | 0.0421 | 0.99998 | [859 empty] | 0
P | 2018-03-31T14:13:42+08:00 | 2018-03-31T16:55:14+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI720G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.77 | 56.78 | 53.42 | 1.51 | 0.13 | 1 | 2 | 3 | -0.8839 | 0.538 | 3 | 49 |  | 0 | 0.2188 | 0.5153 | -0.0198 | 0.99999 | [45 empty] | 0
P | 2018-03-31T13:50:30+08:00 | 2018-03-31T16:59:10+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI719G | CPSM Final Test | PASS | 129 | 167.27 | 0.08 | -0.38 | 56.64 | 49.02 | 1.59 | 0.13 | -1 | -11 | 1 | -0.886 | 0.5348 | 2.9 | 13 |  | 0 | 0.0988 | 0.4244 | -0.0109 | 0.99997 | [860 empty] | 0
P | 2018-03-31T15:42:04+08:00 | 2018-03-31T18:21:06+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI723G | CPSM Final Test | PASS | 129 | 172.13 | -0.02 | -0.77 | 56.72 | 55.29 | 1.51 | 0.13 | 1 | -2 | 0 | -0.828 | 0.544 | 3.3 | -3 |  | 0 | 0.0741 | 0.0557 | 0.0023 | 0.99995 | [314 empty] | 0
P | 2018-03-31T15:31:54+08:00 | 2018-03-31T18:23:12+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI721G | CPSM Final Test | PASS | 129 | 165.83 | 0.17 | -0.54 | 57.07 | 56.56 | 1.67 | 0.13 | 0 | -2 | -1 | -0.8381 | 0.547 | 7.3 | 11 |  | 0.00983 | 0.1393 | 0.0911 | -0.0553 | 1 | [404 empty] | 0
P | 2018-03-29T23:01:45+08:00 | 2018-04-02T10:28:15+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI722G | CPSM Final Test | FAIL | [882 empty] | 0
P | 2018-03-29T22:08:45+08:00 | 2018-04-02T11:21:15+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI722G | CPSM Final Test | FAIL | [882 empty] | 0
P | 2018-03-29T15:53:03+08:00 | 2018-04-02T11:44:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI717G | CPSM Final Test | FAIL | [87 empty] | FALSE
P | 2018-03-31T16:45:00+08:00 | 2018-04-02T12:26:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI722G | CPSM Final Test | PASS | 129 | 164.48 | -0.11 | -0.69 | 56.7 | 57.91 | 1.51 | 0.13 | 0 | -13 | 1 | -0.8839 | 0.515 | 3.9 | 34 |  | 0 | 0.1848 | 0.2936 | 0.0392 | 0.99999 | [861 empty] | 0
P | 2018-04-02T09:28:07+08:00 | 2018-04-02T12:29:52+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI725G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -1.07 | 56.7 | 59.52 | 1.21 | 0.13 | 0 | -10 | -1 | -0.81 | 0.528 | 5.3 | 4 |  | 0 | 0.1373 | 0.0844 | 0.0078 | 1 | [862 empty] | 0
P | 2018-04-02T06:45:18+08:00 | 2018-04-02T12:30:42+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI726G | CPSM Final Test | FAIL | [13 empty] | 17.3
P | 2018-04-02T09:19:15+08:00 | 2018-04-02T12:32:36+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI724G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -0.69 | 56.51 | 53.51 | 1.36 | 0.13 | -1 | -2 | -1 | -0.7682 | 0.4966 | 2.9 | -16 |  | 0 | 0.1492 | 0.2413 | 0.0357 | 0.99995 | [863 empty] | 0
P | 2018-04-02T11:34:44+08:00 | 2018-04-02T14:27:24+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI728G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -0.69 | 56.75 | 53.59 | 1.21 | 0.13 | 1 | 3 | 0 | -0.8101 | 0.533 | 2.5 | -7 |  | 0 | 0.1498 | 0.5933 | -0.0156 | 1 | [864 empty] | 0
P | 2018-04-02T11:24:56+08:00 | 2018-04-02T14:30:34+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI727G | CPSM Final Test | PASS | 129 | 158.54 | -0.02 | -0.31 | 57.07 | 52.07 | 1.36 | 0.13 | 1 | -2 | 1 | -0.8002 | 0.498 | 0.9 | -11 |  | 0 | 0.0961 | 0.1694 | 0.0093 | 1 | [865 empty] | 0
P | 2018-03-31T13:48:45+08:00 | 2018-04-02T14:45:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI717G | CPSM Final Test | PASS | 129 | 174.74 | 0.46 | -0.92 | 56.89 | 56.98 | 1.44 | 0.13 | 0 | -3 | 0 | -0.8919 | 0.4638 | 0.6 | 38 |  | 0 | 0.2575 | 0.325 | 0.0425 | 0.99999 | [854 empty] | 0
P | 2018-04-02T08:07:36+08:00 | 2018-04-02T15:41:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI729G | CPSM Final Test | FAIL | [873 empty] | 7280
P | 2018-04-02T13:41:56+08:00 | 2018-04-02T16:44:42+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI731G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -1.15 | 56.63 | 52.83 | 1.29 | 0.13 | 0 | -3 | -1 | -0.9198 | 0.5292 | 4.1 | -16 |  | 0.0244 | 0.1048 | 0.2031 | -0.0009 | 0.99999 | [255 empty] | 0
P | 2018-04-02T11:54:37+08:00 | 2018-04-02T16:47:22+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI729G | CPSM Final Test | PASS | 129 | 166.55 | 0.27 | -0.54 | 56.49 | 49.02 | 1.36 | 0.13 | 0 | -4 | 1 | -0.8201 | 0.5404 | 6.9 | 16 |  | 0.02741 | 0.2799 | 0.1715 | 0.0202 | 1 | [852 empty] | 0
P | 2018-04-02T14:03:08+08:00 | 2018-04-02T16:47:54+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI730G | CPSM Final Test | PASS | 129 | 163.76 | -1 | 1.52 | 57.32 | 49.78 | 2.81 | 0.13 | 0 | -1 | 2 | -0.7801 | 0.5258 | 11.6 | -15 |  | 0 | 0.1127 | 0.244 | 0.043 | 0.99997 | [514 empty] | 0
P | 2018-04-02T10:38:47+08:00 | 2018-04-02T17:15:05+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI732G | CPSM Final Test | FAIL
P | 2018-04-02T13:56:56+08:00 | 2018-04-02T18:59:07+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI732G | CPSM Final Test | PASS | 129 | 168.17 | -0.11 | -0.54 | 57.09 | 54.53 | 1.44 | 0.13 | -1 | -11 | 0 | -0.83 | 0.5356 | 4.4 | 9 |  | 0.00869 | 0.0512 | 0.3902 | -0.0004 | 1 | [861 empty] | 0
P | 2018-04-02T00:14:24+08:00 | 2018-04-02T19:01:36+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI726G | CPSM Final Test | FAIL | [437 empty] | 0
P | 2018-04-02T16:21:45+08:00 | 2018-04-02T19:26:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI734G | CPSM Final Test | PASS | 129 | 164.48 | -0.02 | -0.84 | 57.09 | 51.22 | 1.36 | 0.13 | 1 | -12 | -1 | -0.86 | 0.533 | 6.2 | 83 |  | 0 | 0.1008 | 0.1807 | -0.0305 | 1 | [205 empty] | 0
P | 2018-04-02T16:33:03+08:00 | 2018-04-02T19:29:06+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI735G | CPSM Final Test | PASS | 129 | 169.79 | 0.17 | -0.77 | 56.91 | 56.05 | 1.36 | 0.13 | 1 | 4 | 0 | -0.914 | 0.5188 | 5.6 | 28 |  | 0 | 0.3315 | 0.3202 | -0.0282 | 1 | [866 empty] | 0
P | 2018-04-02T13:40:32+08:00 | 2018-04-02T19:44:05+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI733G | CPSM Final Test | FAIL
P | 2018-04-02T16:42:19+08:00 | 2018-04-02T19:44:35+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI733G | CPSM Final Test | PASS | 129 | 165.47 | 0.17 | -0.99 | 56.6 | 51.39 | 1.74 | 0.13 | 1 | -10 | -1 | -0.9358 | 0.485 | 3 | -3 |  | 0.00715 | 0.1061 | 0.1253 | -0.0124 | 1 | [846 empty] | 0
P | 2018-04-02T09:38:00+08:00 | 2018-04-02T20:07:22+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI726G | CPSM Final Test | PASS | 129 | 170.78 | -0.02 | -0.69 | 57.09 | 52.75 | 1.29 | 0.13 | 1 | 3 | 3 | -0.9178 | 0.5176 | 0.3 | 38 |  | 0 | 0.1725 | 0.1326 | -0.0044 | 0.99999 | [416 empty] | 0
P | 2018-04-02T17:01:31+08:00 | 2018-04-02T20:19:34+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI736G | CPSM Final Test | PASS | 129 | 158 | -0.11 | -0.92 | 56.79 | 51.99 | 1.59 | 0.13 | 0 | 2 | 1 | -0.8959 | 0.5256 | 2.3 | 53 |  | 0 | 0.097 | 0.0826 | 0.0444 | 0.99997 | [852 empty] | 0
P | 2018-04-02T18:23:51+08:00 | 2018-04-02T21:21:20+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI737G | CPSM Final Test | PASS | 129 | 164.03 | 0.27 | -0.84 | 56.8 | 54.27 | 1.13 | 0.13 | 0 | -3 | 3 | -0.826 | 0.518 | 0.3 | 17 |  | 0 | 0.193 | 0.1302 | 0.0231 | 0.99998 | [234 empty] | 0
P | 2018-04-02T18:38:22+08:00 | 2018-04-02T21:22:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI738G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.92 | 56.77 | 53.76 | 1.44 | 0.13 | 0 | 4 | -1 | -0.9138 | 0.5448 | 0.3 | 16 |  | 0 | 0.0847 | 0.1719 | 0.0028 | 0.99999 | [867 empty] | 0
P | 2018-04-03T09:50:24+08:00 | 2018-04-03T12:53:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI739G | CPSM Final Test | PASS | 129 | 162.95 | 0.27 | -0.92 | 56.48 | 53.93 | 1.29 | 0.13 | 0 | -13 | 1 | -0.87 | 0.5268 | 0.9 | -6 |  | 0 | 0.1386 | 0.2558 | -0.0192 | 1 | [868 empty] | 0
P | 2018-04-03T10:10:47+08:00 | 2018-04-03T12:57:56+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI740G | CPSM Final Test | PASS | 129 | -178.6 | -0.02 | -0.77 | 56.72 | 59.44 | 1.13 | 0.13 | 1 | -1 | 0 | -0.8221 | 0.505 | 0.9 | 1 |  | 0 | 0.1973 | 0.1922 | 0.0618 | 0.99995 | [869 empty] | 0
P | 2018-04-03T10:02:25+08:00 | 2018-04-03T13:06:18+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI741G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -1.07 | 56.74 | 58 | 0.91 | 0.13 | -1 | -2 | 1 | -0.834 | 0.5104 | 4.1 | -1 |  | 0.02111 | 0.0523 | 0.0566 | 0.015 | 1 | [271 empty] | 0
P | 2018-04-03T11:11:45+08:00 | 2018-04-03T14:20:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI743G | CPSM Final Test | PASS | 129 | 153.5 | -0.02 | 0.15 | 56.78 | 53.34 | 1.36 | 0.13 | 0 | 2 | 1 | -0.8141 | 0.5362 | 6.3 | 3 |  | 0.00494 | 0.1089 | 0.0849 | -0.0209 | 0.99999 | [870 empty] | 0
P | 2018-04-03T11:57:32+08:00 | 2018-04-03T14:37:17+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI742G | CPSM Final Test | PASS | 129 | 164.66 | 0.46 | -0.61 | 56.76 | 53.85 | 1.51 | 0.13 | 1 | -2 | 1 | -0.8062 | 0.451 | 1.6 | 14 |  | 0.01436 | 0.1925 | 0.1622 | 0.023 | 0.99997 | [862 empty] | 0
P | 2018-04-03T11:44:17+08:00 | 2018-04-03T14:56:48+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI744G | CPSM Final Test | PASS | 129 | 160.79 | -0.02 | -0.84 | 57.02 | 54.19 | 1.06 | 0.13 | 0 | 1 | 1 | -0.884 | 0.526 | 8.6 | 53 |  | 0 | 0.1277 | 0.1722 | -0.0196 | 0.99997 | [416 empty] | 0
P | 2018-04-03T12:09:37+08:00 | 2018-04-03T15:19:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI745G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.77 | 57.04 | 55.29 | 1.51 | 0.13 | -1 | 4 | 1 | -0.9079 | 0.4632 | 4.6 | -22 |  | 0.02363 | 0.1197 | 0.2561 | -0.0523 | 0.99999 | [871 empty] | 0
P | 2018-04-03T14:47:20+08:00 | 2018-04-03T17:54:22+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | C18SDI747G | CPSM Final Test | PASS | 129 | 160.79 | 0.08 | -0.46 | 57.08 | 52.49 | 1.13 | 0.13 | 1 | -8 | 1 | -0.822 | 0.541 | 5 | 22 |  | 0 | 0.1628 | 0.1477 | 0.0226 | 0.99998 | [872 empty] | 0
P | 2018-04-03T10:51:29+08:00 | 2018-04-03T18:10:58+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI748G | CPSM Final Test | FAIL
P | 2018-04-03T14:31:13+08:00 | 2018-04-03T18:11:32+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI748G | CPSM Final Test | PASS | 129 | 165.11 | 0.37 | -0.84 | 56.92 | 51.56 | 1.29 | 0.13 | 1 | -2 | -1 | -0.9218 | 0.536 | 2.8 | 5 |  | 0.01334 | 0.1011 | 0.1875 | 0.0075 | 1 | [296 empty] | 0
P | 2018-04-03T15:26:19+08:00 | 2018-04-03T18:19:29+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | C18SDI746G | CPSM Final Test | PASS | 129 | 165.56 | 0.17 | -0.69 | 56.94 | 52.07 | 1.59 | 0.22 | 1 | 2 | 2 | -0.9358 | 0.528 | 10.1 | -3 |  | 0 | 0.074 | 0.1791 | -0.0167 | 1 | [873 empty] | 0
P | 2018-04-03T16:54:51+08:00 | 2018-04-03T18:28:29+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI751G | CPSM Final Test | FAIL | [83 empty] | FALSE
P | 2018-04-03T14:30:47+08:00 | 2018-04-03T19:58:32+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI749G | CPSM Final Test | FAIL
P | 2018-04-03T17:06:38+08:00 | 2018-04-03T20:51:12+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI750G | CPSM Final Test | PASS | 129 | 160.88 | -0.02 | 0 | 56.73 | 57.91 | 1.29 | 0.13 | 1 | -10 | 1 | -0.9099 | 0.5184 | 5.8 | -8 |  | 0 | 0.1678 | 0.0941 | 0.047 | 0.99994 | [68 empty] | 0
P | 2018-04-03T17:41:40+08:00 | 2018-04-03T21:15:53+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI751G | CPSM Final Test | PASS | 129 | 163.4 | 0.27 | -0.54 | 56.82 | 56.98 | 1.82 | 0.13 | 0 | 3 | -2 | -0.9198 | 0.5426 | 8.6 | 6 |  | 0.00572 | 0.0956 | 0.1101 | 0.0174 | 0.99996 | [874 empty] | 0
P | 2018-04-03T18:29:31+08:00 | 2018-04-03T21:23:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI753G | CPSM Final Test | PASS | 129 | 164.84 | 0.27 | -0.77 | 56.58 | 54.69 | 1.36 | 0.22 | 1 | -2 | 1 | -0.9338 | 0.5342 | 3.7 | 44 |  | 0 | 0.1099 | 0.1597 | 0.0207 | 1 | [875 empty] | 0
P | 2018-04-03T18:53:51+08:00 | 2018-04-03T21:40:34+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI752G | CPSM Final Test | PASS | 129 | 169.97 | 0.17 | -0.84 | 57.01 | 57.49 | 1.36 | 0.13 | -1 | 2 | 0 | -0.9458 | 0.513 | 5.5 | 14 |  | 0 | 0.1175 | 0.1718 | -0.0006 | 1 | [876 empty] | 0
P | 2018-04-03T17:14:40+08:00 | 2018-04-03T21:43:02+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | C18SDI749G | CPSM Final Test | PASS | 129 | 157.28 | 0.56 | -0.92 | 57.01 | 52.75 | 1.44 | 0.13 | 1 | -12 | -1 | -0.9278 | 0.5074 | 7.1 | -4 |  | 0 | 0.1214 | 0.1308 | -0.0186 | 1 | [219 empty] | 0
P | 2018-04-03T04:51:52+08:00 | 2018-04-04T09:59:04+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI754G | CPSM Final Test | FAIL
P | 2018-04-03T19:25:28+08:00 | 2018-04-04T11:04:36+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI754G | CPSM Final Test | PASS | 129 | 165.2 | 0.46 | 0.3 | 56.91 | 52.75 | 1.13 | 0.13 | 1 | -3 | 0 | -0.856 | 0.5306 | 2 | -9 |  | 0 | 0.1265 | 0.1479 | -0.0033 | 0.99999 | [877 empty] | 0
P | 2018-04-04T09:19:33+08:00 | 2018-04-04T12:15:52+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI755G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.84 | 56.51 | 55.29 | 1.36 | 0.13 | 0 | -4 | 1 | -0.9178 | 0.528 | 4.5 | 12 |  | 0.02657 | 0.0374 | 0.2776 | -0.0099 | 1 | [817 empty] | 0
P | 2018-04-04T09:41:52+08:00 | 2018-04-04T12:27:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | C18SDI757G | CPSM Final Test | PASS | 129 | 161.06 | 0.37 | -0.92 | 56.92 | 54.36 | 1.36 | 0.13 | 1 | 2 | 1 | -0.8381 | 0.4984 | 2.5 | 48 |  | 0.0067 | 0.1983 | 0.0877 | 0.0013 | 1 | [878 empty] | 0
P | 2018-04-04T10:01:55+08:00 | 2018-04-04T12:47:52+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | C18SDI758G | CPSM Final Test | PASS | 129 | 163.4 | 0.27 | -0.99 | 56.79 | 52.58 | 1.21 | 0.13 | 1 | 4 | 0 | -0.7942 | 0.5482 | 1.2 | 79 |  | 0 | 0.2216 | 0.0828 | 0.0209 | 0.99999 | [52 empty] | 0
P | 2018-04-04T12:29:27+08:00 | 2018-04-04T15:00:18+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | C18SDI756G | CPSM Final Test | PASS | 129 | 159.98 | -0.59 | -0.54 | 56.79 | 54.44 | 1.29 | 0.13 | 0 | 3 | 0 | -1.0056 | 0.5196 | 0.8 | -2 |  | 0.00845 | 0.0615 | 0.2691 | -0.0077 | 0.99999 | [877 empty] | 0
P | 2018-04-06T18:47:04+08:00 | 2018-04-06T19:54:30+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM938G | CPSM Final Test | FAIL
P | 2018-04-06T17:45:45+08:00 | 2018-04-06T20:40:10+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | C18USM937G | CPSM Final Test | FAIL | [13 empty] | 20.3
P | 2018-04-06T18:56:47+08:00 | 2018-04-06T22:27:34+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM935G | CPSM Final Test | PASS |  | 161.96 | -0.02 | -0.31 | 57.34 | 55.46 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9377 | 0.5466 | 1.9 | 22 |  |  | 0.08 | 0.0811 | -0.017 | 0.99998 |  |  |  |  | 133 |  | 0 | [198 empty] | 0
P | 2018-04-06T19:32:45+08:00 | 2018-04-06T22:44:14+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM941G | CPSM Final Test | PASS |  | 152.87 | 0.27 | -0.54 | 57.04 | 53.42 | 1.06 | 0.13 | 1 | -3 | 0 | -0.9357 | 0.492 | 3.6 | -3 |  |  | 0.0837 | 0.1867 | -0.0151 | 0.99999 |  |  |  |  | 133 |  | 0 | [311 empty] | 0
P | 2018-04-06T06:05:19+08:00 | 2018-04-07T08:52:17+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | C18USM942G | CPSM Final Test | FAIL
P | 2018-04-06T19:04:03+08:00 | 2018-04-07T09:45:34+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM940G | CPSM Final Test | PASS |  | 164.93 | 0.37 | -0.16 | 57.07 | 51.9 | 1.21 | 0.13 | 0 | -8 | 0 | -0.9098 | 0.503 | 1.2 | 4 |  |  | 0.0543 | 0.1214 | -0.0315 | 0.99999 |  |  |  |  | 133 |  | 0 | [307 empty] | 0
P | 2018-04-06T19:09:04+08:00 | 2018-04-07T09:45:40+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM939G | CPSM Final Test | PASS |  | 164.57 | -0.02 | -0.46 | 57.12 | 54.36 | 1.29 | 0.13 | -1 | -2 | 0 | -0.9178 | 0.4932 | 4.7 | 15 |  |  | 0.114 | 0.3263 | -0.0181 | 1 |  |  |  | 0 | 133 |  | 0
P | 2018-04-06T04:05:11+08:00 | 2018-04-07T09:55:29+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM936G | CPSM Final Test | FAIL | [27 empty] | 12.725
P | 2018-04-06T19:28:48+08:00 | 2018-04-07T10:39:57+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | C18USM942G | CPSM Final Test | PASS |  | 156.11 | -0.02 | -0.54 | 56.92 | 50.97 | 1.36 | 0.13 | 1 | -2 | 1 | -0.7901 | 0.4942 | 4.3 | -16 |  |  | 0.1063 | 0.2376 | 0.028 | 0.99995 |  |  |  |  | 133 |  | 0 | [373 empty] | 0
P | 2018-04-06T19:00:20+08:00 | 2018-04-07T10:58:35+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM936G | CPSM Final Test | PASS |  | 164.48 | -0.02 | -0.54 | 57.02 | 54.27 | 1.51 | 0.13 | 1 | -2 | 1 | -0.9178 | 0.496 | 5.3 | 36 |  |  | 0.2905 | 0.2914 | 0.0225 | 0.99996 |  |  |  |  | 133 |  | 0 | [758 empty] | 0
P | 2018-04-07T07:35:29+08:00 | 2018-04-07T11:17:10+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM944G | CPSM Final Test | FAIL
P | 2018-04-07T09:29:35+08:00 | 2018-04-07T12:26:38+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM943G | CPSM Final Test | PASS |  | 164.3 | -0.02 | -0.69 | 56.91 | 53.93 | 1.36 | 0.13 | 0 | 3 | -1 | -0.9277 | 0.4898 | 6.3 | 3 |  |  | 0.04 | 0.2481 | -0.0157 | 1 |  |  |  |  | 133 |  | 0 | [872 empty] | 0
P | 2018-04-07T10:05:31+08:00 | 2018-04-07T12:58:36+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM937G | CPSM Final Test | PASS |  | 160.97 | -0.02 | -0.54 | 57.08 | 55.29 | 1.29 | 0.13 | 0 | -2 | -1 | -0.9418 | 0.4664 | 11.4 | 4 |  |  | 0.0945 | 0.1584 | 0.0058 | 1 |  |  |  |  | 133 |  | 0 | [832 empty] | 0
P | 2018-04-07T09:26:19+08:00 | 2018-04-07T13:05:14+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM944G | CPSM Final Test | PASS |  | 169.7 | -0.02 | -0.08 | 57.34 | 52.66 | 1.82 | 0.13 | 0 | -3 | -1 | -0.844 | 0.496 | 0.2 | 5 |  |  | 0.0762 | 0.0854 | 0.0097 | 0.99999 |  |  |  |  | 133 |  | 0 | [374 empty] | 0
P | 2018-04-07T11:27:42+08:00 | 2018-04-07T14:57:16+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM938G | CPSM Final Test | PASS |  | 158.9 | 0.27 | -0.54 | 56.91 | 50.04 | 1.21 | 0.13 | 1 | -12 | -1 | -0.9738 | 0.4532 | 4 | 42 |  |  | 0.1283 | 0.123 | -0.0132 | 0.99998 |  |  |  |  | 133 |  | 0 | [18 empty] | 0
P | 2018-04-09T10:34:03+08:00 | 2018-04-09T11:47:07+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI773G | CPSM Final Test | FAIL | [201 empty] | FALSE
P | 2018-04-09T08:15:23+08:00 | 2018-04-09T13:48:55+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI776G | CPSM Final Test | FAIL | [392 empty] | 24680
P | 2018-04-09T07:17:25+08:00 | 2018-04-09T13:55:31+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI777G | CPSM Final Test | FAIL
P | 2018-04-09T10:46:48+08:00 | 2018-04-09T14:00:24+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI774G | CPSM Final Test | PASS | 129 | 159.62 | -0.02 | -1.07 | 56.76 | 56.39 | 1.44 | 0.13 | 0 | 2 | 0 | -0.8221 | 0.505 | 6.3 | 54 |  | 0 | 0.1418 | 0.5267 | 0.0032 | 1 | [880 empty] | 0
P | 2018-04-09T11:06:17+08:00 | 2018-04-09T14:21:12+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI772G | CPSM Final Test | PASS | 129 | 159.44 | -0.02 | -0.84 | 57.09 | 55.46 | 1.13 | 0.13 | 0 | -2 | 0 | -0.854 | 0.501 | 2.5 | 26 |  | 0 | 0.0959 | 0.2386 | -0.0233 | 1 | [881 empty] | 0
P | 2018-04-09T07:25:38+08:00 | 2018-04-09T14:33:29+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI782G | CPSM Final Test | FAIL
P | 2018-04-09T06:12:46+08:00 | 2018-04-09T14:51:40+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI783G | CPSM Final Test | FAIL | [14 empty] | 821
P | 2018-04-09T11:02:09+08:00 | 2018-04-09T14:55:27+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI776G | CPSM Final Test | PASS | 129 | 159.17 | -0.02 | -0.99 | 56.79 | 56.56 | 1.59 | 0.13 | -1 | -13 | -1 | -0.7801 | 0.5212 | 1.9 | 80 |  | 0 | 0.1186 | 0.212 | 0.0129 | 0.99999 | [371 empty] | 0
P | 2018-04-09T06:06:15+08:00 | 2018-04-09T15:18:36+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI781G | CPSM Final Test | FAIL | [903 empty] | 0
P | 2018-04-09T06:41:57+08:00 | 2018-04-09T15:39:13+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI773G | CPSM Final Test | FAIL
P | 2018-04-09T11:10:35+08:00 | 2018-04-09T15:39:45+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI773G | CPSM Final Test | PASS | 129 | 159.89 | -0.02 | -0.54 | 56.73 | 58.34 | 1.67 | 0.13 | -1 | 2 | 3 | -0.818 | 0.5038 | 4.4 | 17 |  | 0 | 0.1798 | 0.2555 | 0.0415 | 0.99993 | [883 empty] | 0
P | 2018-04-09T10:36:28+08:00 | 2018-04-09T15:55:49+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI777G | CPSM Final Test | PASS | 129 | 158.99 | -0.02 | -0.99 | 56.69 | 58.93 | 1.21 | 0.13 | -1 | 3 | 0 | -0.8341 | 0.5242 | 4.8 | 7 |  | 0 | 0.3912 | 0.1546 | 0.0108 | 1 | [884 empty] | 0
P | 2018-04-09T05:46:48+08:00 | 2018-04-09T16:12:19+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI782G | CPSM Final Test | FAIL
P | 2018-04-09T10:42:25+08:00 | 2018-04-09T16:24:26+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI781G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.46 | 57.01 | 55.12 | 1.74 | 0.13 | 0 | 2 | -1 | -0.87 | 0.4944 | 1.9 | 1 |  | 0 | 0.0565 | 0.0779 | 0.0122 | 1 | [885 empty] | 0
P | 2018-04-09T10:32:13+08:00 | 2018-04-09T16:36:32+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI783G | CPSM Final Test | PASS | 129 | 163.13 | -0.02 | -0.99 | 56.92 | 51.31 | 1.21 | 0.13 | 1 | -12 | 3 | -0.9218 | 0.5086 | 10.6 | 4 |  | 0 | 0.0602 | 0.1769 | -0.013 | 0.99999 | [413 empty] | 0
P | 2018-04-09T16:54:01+08:00 | 2018-04-09T17:39:25+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI782G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-04-09T15:28:32+08:00 | 2018-04-09T18:16:02+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI762G | CPSM Final Test | PASS | 129 | 160.25 | 0.17 | -1.07 | 56.56 | 53.93 | 1.13 | 0.13 | 0 | -2 | -1 | -0.9159 | 0.5146 | 13.4 | 3 |  | 0 | 0.0314 | 0.274 | 0.0198 | 1 | [886 empty] | 0
P | 2018-04-09T15:31:02+08:00 | 2018-04-09T18:32:47+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI769G | CPSM Final Test | PASS | 129 | 155.57 | -0.02 | -0.99 | 56.75 | 55.12 | 1.36 | 0.13 | -1 | -2 | 1 | -0.9119 | 0.5028 | 8.7 | 13 |  | 0 | 0.0726 | 0.3183 | -0.034 | 0.99997 | [371 empty] | 0
P | 2018-04-09T11:14:32+08:00 | 2018-04-09T18:34:34+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI764G | CPSM Final Test | FAIL
P | 2018-04-09T14:54:33+08:00 | 2018-04-09T18:35:10+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI764G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.99 | 57 | 54.19 | 1.44 | 0.13 | -2 | -2 | 1 | -0.9199 | 0.4934 | 6.8 | -18 |  | 0 | 0.0748 | 0.2916 | -0.0251 | 1 | [810 empty] | 0
P | 2018-04-09T13:59:57+08:00 | 2018-04-09T19:55:44+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI779G | CPSM Final Test | FAIL
P | 2018-04-09T17:09:02+08:00 | 2018-04-09T20:18:51+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI788G | CPSM Final Test | PASS | 129 | 164.84 | 0.27 | -0.99 | 56.91 | 52.49 | 1.59 | 0.13 | 0 | -2 | 0 | -0.8958 | 0.5472 | 0.8 | -24 |  | 0 | 0.0558 | 0.2389 | 0.0353 | 0.99994 | [844 empty] | 0
P | 2018-04-09T17:04:54+08:00 | 2018-04-09T20:23:14+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI786G | CPSM Final Test | PASS | 129 | 165.11 | 0.68 | -0.61 | 57.01 | 52.24 | 1.67 | 0.13 | 1 | 4 | -1 | -0.9219 | 0.536 | 0.9 | -8 |  | 0 | 0.0407 | 0.1258 | -0.0486 | 1 | [885 empty] | 0
P | 2018-04-09T17:01:20+08:00 | 2018-04-09T20:33:31+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI785G | CPSM Final Test | PASS | 129 | 166.28 | 0.37 | -1.07 | 57.31 | 55.12 | 0.83 | 0.13 | 1 | -8 | 1 | -0.9179 | 0.5102 | 0.4 | 7 |  | 0 | 0.1031 | 0.2536 | 0.0356 | 1 | [664 empty] | 0
P | 2018-04-09T17:16:43+08:00 | 2018-04-09T20:34:11+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI782G | CPSM Final Test | PASS | 129 | 157.46 | -0.02 | -0.92 | 56.98 | 55.46 | 1.36 | 0.13 | 1 | 2 | -1 | -0.9418 | 0.479 | 0.2 | 47 |  | 0 | 0.1618 | 0.1351 | -0.0263 | 1 | [887 empty] | 0
P | 2018-04-09T13:21:55+08:00 | 2018-04-09T20:33:46+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI779G | CPSM Final Test | FAIL
P | 2018-04-09T17:27:30+08:00 | 2018-04-09T20:35:48+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI775G | CPSM Final Test | FAIL | [14 empty] | 861
P | 2018-04-09T18:53:19+08:00 | 2018-04-09T21:47:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI789G | CPSM Final Test | PASS | 129 | 171.05 | -0.3 | -1.07 | 56.93 | 54.78 | 1.44 | 0.13 | 1 | 4 | 0 | -0.9259 | 0.4688 | 2.2 | 15 |  | 0.01121 | 0.0415 | 0.2593 | -0.0406 | 0.99999 | [888 empty] | 0
P | 2018-04-09T16:57:50+08:00 | 2018-04-09T21:47:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI779G | CPSM Final Test | PASS | 129 | 156.65 | -0.02 | -0.69 | 57.02 | 53.68 | 1.9 | 0.13 | 1 | -3 | -1 | -0.9418 | 0.511 | 0.1 | 2 |  | 0 | 0.0505 | 0.0444 | 0.0345 | 0.99999 | [866 empty] | 0
P | 2018-04-09T19:01:26+08:00 | 2018-04-09T22:06:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI790G | CPSM Final Test | PASS | 129 | 160.7 | 0.37 | -0.99 | 57.01 | 53.51 | 1.51 | 0.13 | 0 | -2 | 1 | -0.9078 | 0.5118 | 11.4 | 31 |  | 0.00922 | 0.0798 | 0.3816 | -0.0231 | 1 | [889 empty] | 0
P | 2018-04-09T19:01:39+08:00 | 2018-04-09T22:29:18+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI775G | CPSM Final Test | PASS | 129 | 156.38 | -0.02 | -0.92 | 56.71 | 59.01 | 1.21 | 0.13 | 0 | -1 | -1 | -0.9498 | 0.543 | 0.9 | 4 |  | 0 | 0.0284 | 0.2498 | 0.015 | 1 | [105 empty] | 0
P | 2018-04-10T07:08:50+08:00 | 2018-04-10T12:09:52+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI787G | CPSM Final Test | FAIL
P | 2018-04-10T07:00:24+08:00 | 2018-04-10T12:32:41+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI763G | CPSM Final Test | FAIL
P | 2018-04-10T10:06:07+08:00 | 2018-04-10T12:55:27+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI759G | CPSM Final Test | PASS | 129 | 163.13 | -0.02 | -1.07 | 56.77 | 53.26 | 1.13 | 0.13 | 0 | 2 | 1 | -0.8161 | 0.55 | 0.6 | 38 |  | 0 | 0.1075 | 0.2211 | 0.0627 | 0.99996 | [890 empty] | 0
P | 2018-04-10T10:00:17+08:00 | 2018-04-10T12:57:22+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI770G | CPSM Final Test | PASS | 129 | 163.94 | -0.02 | -0.99 | 56.91 | 54.53 | 1.59 | 0.13 | -2 | -2 | -1 | -0.854 | 0.5272 | 1.8 | -21 |  | 0 | 0.1083 | 0.1687 | 0.0088 | 0.99999 | [447 empty] | 0
P | 2018-04-10T09:58:33+08:00 | 2018-04-10T13:06:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI778G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -1.22 | 56.72 | 60.54 | 1.9 | 0.13 | 1 | -2 | 1 | -0.8699 | 0.4976 | 6.9 | 2 |  | 0.01057 | 0.0683 | 0.3259 | 0.0116 | 1 | [872 empty] | 0
P | 2018-04-10T09:50:49+08:00 | 2018-04-10T13:05:03+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI780G | CPSM Final Test | PASS | 129 | 158.99 | -0.02 | -0.69 | 57.02 | 53.09 | 1.74 | 0.13 | 0 | 3 | 0 | -0.906 | 0.5022 | 4.1 | 39 |  | 0 | 0.1281 | 0.188 | -0.0151 | 1 | [350 empty] | 0
P | 2018-04-10T09:43:38+08:00 | 2018-04-10T13:10:22+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI771G | CPSM Final Test | PASS | 129 | 163.13 | 0.27 | -0.54 | 56.79 | 52.41 | 1.06 | 0.13 | 0 | 3 | 3 | -0.82 | 0.5148 | 1.7 | 20 |  | 0.00936 | 0.1533 | 0.2586 | -0.0127 | 0.99997 | [891 empty] | 0
P | 2018-04-10T09:35:17+08:00 | 2018-04-10T13:26:20+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI784G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.61 | 57.07 | 52.49 | 1.44 | 0.13 | 0 | -1 | -1 | -0.8839 | 0.5088 | 2.3 | 18 |  | 0 | 0.282 | 0.2842 | -0.0123 | 1 | [711 empty] | 0
P | 2018-04-10T05:13:46+08:00 | 2018-04-10T14:19:19+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI763G | CPSM Final Test | FAIL | [913 empty] | 0
P | 2018-04-10T09:46:32+08:00 | 2018-04-10T15:24:59+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI763G | CPSM Final Test | PASS | 129 | 164.21 | -0.02 | -0.92 | 57 | 54.69 | 1.36 | 0.13 | -1 | -2 | 0 | -0.8659 | 0.4992 | 0.4 | 6 |  | 0 | 0.2917 | 0.4255 | -0.0265 | 0.99999 | [892 empty] | 0
P | 2018-04-10T14:02:06+08:00 | 2018-04-10T17:16:56+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI800G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -0.54 | 57.21 | 53.17 | 1.59 | 0.13 | 1 | -11 | -1 | -0.9339 | 0.5472 | 0.8 | 14 |  | 0 | 0.1137 | 0.4964 | -0.008 | 0.99988 | [560 empty] | 0
P | 2018-04-10T10:39:22+08:00 | 2018-04-10T17:18:14+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI795G | CPSM Final Test | FAIL
P | 2018-04-10T14:13:44+08:00 | 2018-04-10T17:26:04+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI796G | CPSM Final Test | PASS | 129 | 159.62 | 0.27 | -0.84 | 56.91 | 53.51 | 1.44 | 0.13 | -1 | 10 | -1 | -0.9896 | 0.4994 | 7.3 | 79 |  | 0 | 0.0599 | 0.4059 | -0.0352 | 0.99999 | [893 empty] | 0
P | 2018-04-10T14:17:30+08:00 | 2018-04-10T17:28:09+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI794G | CPSM Final Test | PASS | 129 | 157.64 | -0.02 | -1.07 | 56.94 | 52.41 | 1.44 | 0.13 | -1 | -8 | 1 | -0.874 | 0.55 | 1.8 | -13 |  | 0 | 0.1364 | 0.2932 | 0.0374 | 0.99997 | [894 empty] | 0
P | 2018-04-10T14:13:28+08:00 | 2018-04-10T17:36:31+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI799G | CPSM Final Test | PASS | 129 | 158.36 | -0.02 | -0.92 | 57.02 | 54.86 | 1.59 | 0.13 | -1 | -2 | 0 | -0.9776 | 0.4906 | 5.1 | 3 |  | 0.00973 | 0.0749 | 0.3497 | -0.0044 | 1 | [895 empty] | 0
P | 2018-04-10T01:39:02+08:00 | 2018-04-10T17:39:40+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI787G | CPSM Final Test | FAIL
P | 2018-04-10T09:46:17+08:00 | 2018-04-10T18:24:51+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI802G | CPSM Final Test | FAIL | [17 empty] | 0.644
P | 2018-04-10T09:39:21+08:00 | 2018-04-10T18:45:18+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI787G | CPSM Final Test | PASS | 129 | 160.07 | 0.46 | -0.84 | 56.79 | 53 | 1.44 | 0.13 | 1 | -2 | -1 | -0.8919 | 0.5322 | 3.7 | 7 |  | 0 | 0.0898 | 0.1115 | -0.0426 | 1 | [728 empty] | 0
P | 2018-04-10T13:58:48+08:00 | 2018-04-10T19:03:06+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI795G | CPSM Final Test | PASS | 129 | 158.09 | -0.02 | -0.84 | 56.79 | 55.88 | 1.36 | 0.13 | -1 | 2 | 3 | -0.9598 | 0.521 | 3.5 | -7 |  | 0 | 0.0744 | 0.1464 | -0.0268 | 1 | [896 empty] | 0
P | 2018-04-10T14:05:34+08:00 | 2018-04-10T19:27:55+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI802G | CPSM Final Test | PASS | 129 | 163.04 | 0.17 | -0.92 | 56.8 | 55.63 | 1.21 | 0.13 | 1 | 2 | 1 | -0.9937 | 0.4978 | 0.3 | 42 |  | 0 | 0.1735 | 0.2012 | -0.0411 | 0.99994 | [897 empty] | 0
P | 2018-04-10T17:38:39+08:00 | 2018-04-10T19:53:07+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI797G | CPSM Final Test | FAIL
P | 2018-04-10T17:15:21+08:00 | 2018-04-10T20:16:25+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI797G | CPSM Final Test | FAIL
P | 2018-04-10T18:51:05+08:00 | 2018-04-10T21:34:08+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI808G | CPSM Final Test | PASS | 129 | 159.17 | -0.02 | -0.92 | 57.08 | 54.19 | 1.67 | 0.13 | 0 | 4 | 3 | -0.9937 | 0.5004 | 4.1 | -12 |  | 0 | 0.0884 | 0.5678 | -0.0234 | 0.99999 | [898 empty] | 0
P | 2018-04-10T17:38:10+08:00 | 2018-04-10T21:40:52+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI804G | CPSM Final Test | PASS | 129 | 163.76 | -0.02 | -1.3 | 56.7 | 54.02 | 1.59 | 0.13 | 0 | -11 | 0 | -0.9198 | 0.539 | 0.9 | 32 |  | 0 | 0.0615 | 0.3414 | -0.0156 | 1 | [899 empty] | 0
P | 2018-04-10T17:34:26+08:00 | 2018-04-10T21:53:41+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI760G | CPSM Final Test | PASS | 129 | 159.8 | -0.02 | -0.92 | 56.7 | 55.96 | 1.29 | 0.13 | 0 | -12 | -1 | -0.9338 | 0.479 | 5.6 | 3 |  | 0 | 0.0989 | 0.232 | 0.0202 | 0.99999 | [892 empty] | 0
P | 2018-04-10T18:41:46+08:00 | 2018-04-10T22:00:57+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI761G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -1.53 | 56.76 | 54.78 | 0.91 | 0.13 | 1 | -2 | 0 | -0.9417 | 0.528 | 4.6 | -3 |  | 0.01354 | 0.1037 | 0.1896 | 0.0418 | 1 | [900 empty] | 0
P | 2018-04-10T03:51:25+08:00 | 2018-04-11T09:40:21+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI797G | CPSM Final Test | FAIL
P | 2018-04-10T18:45:53+08:00 | 2018-04-11T10:45:59+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI797G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -1.53 | 56.89 | 55.71 | 0.98 | 0.13 | 1 | -3 | 0 | -0.9059 | 0.5436 | 2.7 | 2 |  | 0 | 0.0656 | 0.2568 | 0.0277 | 0.99995 | [901 empty] | 0
P | 2018-04-11T08:15:48+08:00 | 2018-04-11T12:31:26+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI768G | CPSM Final Test | FAIL
P | 2018-04-11T07:42:50+08:00 | 2018-04-11T13:29:56+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI765G | CPSM Final Test | FAIL
P | 2018-04-11T10:18:53+08:00 | 2018-04-11T13:37:36+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI807G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -1.07 | 56.77 | 52.49 | 1.06 | 0.13 | 1 | -12 | 1 | -0.836 | 0.537 | 7.2 | -7 |  | 0 | 0.0666 | 0.2073 | -0.0527 | 0.99999 | [751 empty] | 0
P | 2018-04-11T10:14:57+08:00 | 2018-04-11T13:42:10+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI767G | CPSM Final Test | PASS | 129 | 159.71 | 0.27 | -0.31 | 57.18 | 54.1 | 1.59 | 0.13 | 0 | -10 | -1 | -0.8499 | 0.4966 | 4.5 | 7 |  | 0 | 0.02 | 0.0547 | -0.0316 | 0.99999 | [728 empty] | 0
P | 2018-04-11T10:36:23+08:00 | 2018-04-11T13:42:30+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI765G | CPSM Final Test | PASS | 129 | 163.31 | -0.02 | -0.61 | 57.32 | 53.93 | 1.29 | 0.13 | 0 | -4 | 3 | -0.8241 | 0.541 | 4.3 | 50 |  | 0 | 0.056 | 0.2907 | 0.0007 | 1 | [251 empty] | 0
P | 2018-04-11T07:41:20+08:00 | 2018-04-11T13:45:31+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI801G | CPSM Final Test | FAIL
P | 2018-04-11T10:36:34+08:00 | 2018-04-11T13:50:45+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI766G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.92 | 56.81 | 55.54 | 1.9 | 0.13 | 1 | -2 | 1 | -0.848 | 0.499 | 2.6 | -11 |  | 0.03288 | 0.06 | 0.4613 | -0.0324 | 0.99999 | [902 empty] | 0
P | 2018-04-11T10:43:25+08:00 | 2018-04-11T13:58:05+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI801G | CPSM Final Test | PASS | 129 | 158.09 | -0.02 | -1.3 | 56.42 | 61.38 | 1.13 | 0.13 | 0 | 3 | 0 | -0.7542 | 0.4984 | 8.1 | 17 |  | 0 | 0.2188 | 0.2781 | -0.0063 | 0.99999 | [903 empty] | 0
P | 2018-04-11T10:23:37+08:00 | 2018-04-11T14:15:50+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI768G | CPSM Final Test | PASS | 129 | 158 | -0.02 | -0.31 | 56.93 | 52.15 | 1.59 | 0.13 | 1 | -3 | -1 | -0.814 | 0.486 | 1.8 | 0 |  | 0 | 0.1081 | 0.1342 | 0.0143 | 0.99998 | [904 empty] | 0
P | 2018-04-11T14:22:12+08:00 | 2018-04-11T15:18:35+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | FAIL | [926 empty] | FALSE
P | 2018-04-11T14:15:06+08:00 | 2018-04-11T15:25:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | FAIL | [926 empty] | FALSE
P | 2018-04-11T14:14:52+08:00 | 2018-04-11T15:25:55+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-04-11T14:13:48+08:00 | 2018-04-11T15:26:59+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-04-11T14:04:18+08:00 | 2018-04-11T15:34:30+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI791G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-04-11T12:48:46+08:00 | 2018-04-11T16:13:46+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI803G | CPSM Final Test | FAIL
P | 2018-04-11T12:33:10+08:00 | 2018-04-11T16:29:22+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI803G | CPSM Final Test | FAIL
P | 2018-04-11T13:10:55+08:00 | 2018-04-11T16:34:35+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI792G | CPSM Final Test | FAIL | [13 empty] | 15.8
P | 2018-04-11T13:51:38+08:00 | 2018-04-11T17:09:25+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | FAIL | [13 empty] | 15.5
P | 2018-04-11T14:58:09+08:00 | 2018-04-11T18:08:17+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI793G | CPSM Final Test | PASS | 129 | 158.9 | -0.02 | -0.99 | 57 | 56.13 | 1.36 | 0.13 | 0 | 3 | -1 | -0.8619 | 0.4958 | 1.8 | 22 |  | 0 | 0.3373 | 0.299 | -0.0043 | 0.99998 | [906 empty] | 0
P | 2018-04-11T14:49:24+08:00 | 2018-04-11T18:10:38+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI791G | CPSM Final Test | PASS | 129 | 161.24 | 0.59 | -0.92 | 56.84 | 54.69 | 1.36 | 0.13 | 0 | -2 | 0 | -0.8859 | 0.539 | 3.8 | 2 |  | 0.00852 | 0.1137 | 0.3198 | 0.0256 | 0.99999 | [907 empty] | 0
P | 2018-04-11T14:31:16+08:00 | 2018-04-11T18:29:04+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI803G | CPSM Final Test | PASS | 129 | 160.88 | -0.02 | -0.69 | 57.21 | 55.12 | 1.21 | 0.13 | 1 | -2 | -1 | -0.9258 | 0.494 | 3.4 | 32 |  | 0 | 0.0496 | 0.219 | -0.0509 | 0.99999 | [908 empty] | 0
P | 2018-04-11T10:37:33+08:00 | 2018-04-11T18:44:31+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI806G | CPSM Final Test | FAIL
P | 2018-04-11T10:14:36+08:00 | 2018-04-11T19:00:52+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL
P | 2018-04-11T14:52:45+08:00 | 2018-04-11T19:25:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI792G | CPSM Final Test | PASS | 129 | 158.09 | -0.02 | -0.69 | 57.11 | 55.03 | 1.36 | 0.13 | 0 | -3 | 1 | -0.7702 | 0.4968 | 0.8 | 1 |  | 0 | 0.0874 | 0.3747 | -0.0114 | 0.99999 | [901 empty] | 0
P | 2018-04-11T11:31:22+08:00 | 2018-04-11T19:29:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | FAIL | [16 empty] | 0.1882
P | 2018-04-11T11:28:04+08:00 | 2018-04-11T19:32:59+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | FAIL | [16 empty] | 0
P | 2018-04-11T15:30:31+08:00 | 2018-04-11T20:35:53+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI798G | CPSM Final Test | PASS | 129 | 155.48 | -0.02 | -0.99 | 57 | 54.86 | 1.51 | 0.13 | 1 | -2 | 1 | -0.9138 | 0.5412 | 3.8 | 36 |  | 0 | 0.0706 | 0.4145 | -0.01 | 1 | [909 empty] | 0
P | 2018-04-11T14:41:02+08:00 | 2018-04-11T20:39:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | C18SDI806G | CPSM Final Test | PASS | 129 | 163.94 | -0.02 | -0.92 | 56.72 | 54.53 | 1.74 | 0.13 | 0 | -2 | 1 | -0.9099 | 0.4986 | 2.1 | 20 |  | 0 | 0.059 | 0.1553 | 0.007 | 0.99999 | [680 empty] | 0
P | 2018-04-10T20:41:02+08:00 | 2018-04-12T08:34:26+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL
P | 2018-04-11T14:37:44+08:00 | 2018-04-12T08:34:56+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -0.84 | 56.99 | 53.68 | 1.59 | 0.13 | -1 | -11 | 0 | -0.9099 | 0.4722 | 2.9 | 12 | 0 | 0 | 0.0395 | 0.0279 | 0.0186 | 1
P | 2018-04-12T09:52:04+08:00 | 2018-04-12T12:35:24+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM945G | CPSM Final Test | PASS |  | 158.09 | 0.17 | -0.23 | 56.9 | 55.03 | 1.21 | 0.13 | 1 | -12 | -1 | -0.826 | 0.5424 | 3.3 | 14 |  |  | 0.1326 | 0.1687 | -0.0213 | 1 |  |  |  |  | 133 |  | 0 | [901 empty] | 0
P | 2018-04-12T15:34:19+08:00 | 2018-04-12T18:29:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | D18USM954G | CPSM Final Test | PASS |  | 164.39 | 0.08 | -0.61 | 56.59 | 56.22 | 1.06 | 0.13 | -1 | -2 | 0 | -0.9217 | 0.5474 | 5 | 62 |  |  | 0.0832 | 0.1319 | 0.0141 | 0.99999 |  |  |  |  | 133 |  | 0.01416 | [108 empty] | 0
P | 2018-04-12T15:33:11+08:00 | 2018-04-12T18:34:02+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | D18USM949G | CPSM Final Test | PASS |  | 164.21 | 0.08 | -0.38 | 57.02 | 54.19 | 1.44 | 0.13 | -1 | -2 | 1 | -0.9398 | 0.498 | 0.3 | 7 |  |  | 0.0381 | 0.0956 | -0.0116 | 1 |  |  |  |  | 133 |  | 0 | [749 empty] | 0
P | 2018-04-12T12:44:13+08:00 | 2018-04-12T18:38:48+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | D18USM952G | CPSM Final Test | FAIL | [931 empty] | 3512
P | 2018-04-12T15:42:47+08:00 | 2018-04-12T19:25:01+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | -0.02 | -0.46 | 57.13 | 55.88 | 1.13 | 0.13 | -1 | 4 | 0 | -0.7941 | 0.542 | 10.5 | 5 |  |  | 0.1008 | 0.2162 | -0.0054 | 0.99999 |  |  |  |  | 133 |  | 0 | [904 empty] | 0
P | 2018-04-12T15:41:30+08:00 | 2018-04-12T19:44:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | D18USM952G | CPSM Final Test | PASS |  | 164.93 | 0.27 | -0.46 | 56.8 | 55.46 | 1.82 | 0.13 | 0 | -1 | -1 | -0.902 | 0.5356 | 1.8 | 5 |  |  | 0.0356 | 0.1463 | -0.0632 | 0.99999 |  |  |  |  | 133 |  | 0.01288 | [903 empty] | 0
P | 2018-04-12T15:19:38+08:00 | 2018-04-12T20:07:16+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM946G | CPSM Final Test | FAIL | [903 empty] | 0
P | 2018-04-12T17:43:27+08:00 | 2018-04-12T21:12:42+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM946G | CPSM Final Test | PASS |  | 162.59 | 0.08 | -0.23 | 56.49 | 53.59 | 1.74 | 0.13 | 1 | -11 | 0 | -0.9837 | 0.513 | 4.4 | 1 |  |  | 0.2954 | 0.3731 | -0.0265 | 1 |  |  |  |  | 133 |  | 0 | [905 empty] | 0
P | 2018-04-12T05:44:09+08:00 | 2018-04-13T09:27:25+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | D18USM951G | CPSM Final Test | FAIL
P | 2018-04-13T10:29:00+08:00 | 2018-04-13T10:58:39+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-04-13T10:36:47+08:00 | 2018-04-13T11:02:03+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-04-13T10:18:58+08:00 | 2018-04-13T11:08:41+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-04-13T10:23:21+08:00 | 2018-04-13T11:15:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-04-13T10:49:25+08:00 | 2018-04-13T11:19:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL
P | 2018-04-13T10:43:49+08:00 | 2018-04-13T11:19:57+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL
P | 2018-04-13T11:19:48+08:00 | 2018-04-13T11:30:19+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL
P | 2018-04-12T19:35:47+08:00 | 2018-04-13T11:30:59+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | D18USM951G | CPSM Final Test | PASS |  | 160.61 | -0.21 | -0.38 | 56.78 | 55.88 | 1.36 | 0.13 | 1 | -2 | -1 | -0.8262 | 0.454 | 6.2 | 26 |  |  | 0.029 | 0.2465 | 0.0198 | 0.99992 |  |  |  |  | 133 |  | 0 | [906 empty] | 0
P | 2018-04-13T10:16:16+08:00 | 2018-04-13T13:25:08+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | D18USM953G | CPSM Final Test | PASS |  | 163.22 | 0.37 | -0.54 | 56.93 | 54.78 | 1.44 | 0.13 | 0 | 3 | 2 | -0.914 | 0.4808 | 1.5 | 7 |  |  | 0.0771 | 0.0693 | -0.0256 | 0.99999 |  |  |  |  | 133 |  | 0 | [848 empty] | 0
P | 2018-04-11T16:42:52+08:00 | 2018-04-13T14:41:25+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | D18USM947G | CPSM Final Test | FAIL | [935 empty] | 0
P | 2018-04-13T14:14:18+08:00 | 2018-04-13T14:38:43+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL
P | 2018-04-13T06:57:50+08:00 | 2018-04-13T15:03:56+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI810G | CPSM Final Test | FAIL
P | 2018-04-13T11:55:41+08:00 | 2018-04-13T15:38:32+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | -0.02 | -1.38 | 56.73 | 54.78 | 1.13 | 0.13 | 1 | 1 | 1 | -0.8299 | 0.501 | 2.6 | 2 |  |  | 0.2551 | 0.3278 | -0.0096 | 1 |  |  |  |  | 133 |  | 0 | [908 empty] | 0
P | 2018-04-12T15:42:08+08:00 | 2018-04-13T15:47:27+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | D18USM947G | CPSM Final Test | PASS |  | 155.75 | 0.17 | -0.23 | 57.11 | 53 | 1.21 | 0.13 | 0 | -11 | -1 | -0.9279 | 0.4712 | 2.8 | -15 |  |  | 0.0584 | 0.2635 | -0.0093 | 1 |  |  |  |  | 133 |  | 0 | [907 empty] | 0
P | 2018-04-13T06:10:22+08:00 | 2018-04-13T15:51:23+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI810G | CPSM Final Test | FAIL
P | 2018-04-13T16:07:37+08:00 | 2018-04-13T16:21:52+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL
P | 2018-04-13T05:47:58+08:00 | 2018-04-13T16:40:05+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI809G | CPSM Final Test | FAIL
P | 2018-04-13T14:34:36+08:00 | 2018-04-13T17:31:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | D18USM001G | CPSM Final Test | PASS |  | 162.41 | -0.02 | -0.92 | 56.64 | 51.48 | 0.91 | 0.13 | 0 | -2 | -1 | -0.8121 | 0.5258 | 12.8 | 26 |  |  | 0.0642 | 0.3103 | 0.038 | 0.99998 |  |  |  |  | 133 |  | 0 | [111 empty] | 0
P | 2018-04-13T11:00:53+08:00 | 2018-04-13T17:35:45+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI810G | CPSM Final Test | PASS | 129 | 161.15 | -0.02 | -1.22 | 56.77 | 53.09 | 1.51 | 0.13 | 0 | -2 | -1 | -0.9518 | 0.5168 | 0 | -12 |  | 0 | 0.1186 | 0.2335 | -0.0019 | 1 | [910 empty] | 0
P | 2018-04-13T14:50:19+08:00 | 2018-04-13T17:39:57+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.72 | 0.97 | -1.3 | 57.19 | 56.3 | 1.13 | 0.13 | 1 | -2 | 0 | -0.9239 | 0.5154 | -0.1 | 22 |  |  | 0.0955 | 0.1611 | 0.0114 | 0.99999 | [21 empty] | 0 | [326 empty] | 0
P | 2018-04-13T14:29:11+08:00 | 2018-04-13T17:45:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | D18USM003G | CPSM Final Test | PASS |  | 158.9 | 0.65 | -0.77 | 56.91 | 53.85 | 1.44 | 0.13 | -1 | -2 | -2 | -0.9557 | 0.492 | 11.8 | 9 |  |  | 0.0514 | 0.0588 | -0.0312 | 0.99999 |  |  |  |  | 133 |  | 0 | [909 empty] | 0
P | 2018-04-13T11:14:02+08:00 | 2018-04-13T18:27:25+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI809G | CPSM Final Test | PASS | 129 | 162.68 | -0.02 | -0.99 | 57.01 | 55.03 | 1.21 | 0.13 | 1 | -4 | -1 | -0.9257 | 0.5418 | 1.3 | -14 |  | 0.0102 | 0.0895 | 0.5083 | -0.0161 | 0.99998 | [917 empty] | 0
P | 2018-04-13T10:07:58+08:00 | 2018-04-13T18:47:02+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM948G | CPSM Final Test | PASS |  | 160.79 | -0.68 | -0.46 | 56.94 | 54.27 | 1.44 | 0.13 | 0 | -2 | 0 | -0.924 | 0.489 | 6.3 | 34 |  |  | 0.0762 | 0.4982 | 0.0483 | 0.99999 |  |  |  |  | 133 |  | 0 | [905 empty] | 0
P | 2018-04-13T15:59:04+08:00 | 2018-04-13T18:53:31+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI002G | CPSM Final Test | PASS | 129 | 161.33 | 0.08 | -0.84 | 56.63 | 52.83 | 1.67 | 0.13 | -1 | -10 | -2 | -0.908 | 0.55 | 0.3 | 12 |  | 0 | 0.2147 | 0.229 | 0.0132 | 0.99999 | [914 empty] | 0
P | 2018-04-13T18:48:56+08:00 | 2018-04-13T21:41:23+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI812G | CPSM Final Test | PASS | 129 | 157.37 | 0.08 | -0.99 | 56.77 | 53.34 | 1.67 | 0.13 | 1 | -4 | -1 | -0.9378 | 0.481 | 4 | 34 |  | 0.00938 | 0.0319 | 0.279 | -0.0183 | 1 | [918 empty] | 0
P | 2018-04-13T18:30:24+08:00 | 2018-04-13T21:42:32+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI813G | CPSM Final Test | PASS | 129 | 161.33 | 0.27 | -1.38 | 56.8 | 51.39 | 1.21 | 0.13 | -1 | -3 | -2 | -0.9678 | 0.509 | 0.3 | 15 |  | 0.00756 | 0.0477 | 0.2425 | 0.0035 | 1 | [919 empty] | 0
P | 2018-04-13T17:38:25+08:00 | 2018-04-13T21:56:57+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | D18SDI003G | CPSM Final Test | PASS | 129 | 158.18 | 0.17 | -0.77 | 56.61 | 52.41 | 1.36 | 0.13 | 0 | -2 | -1 | -0.86 | 0.5378 | 4.5 | 27 |  | 0 | 0.1112 | 0.106 | 0.0411 | 0.99994 | [417 empty] | 0
P | 2018-04-13T06:11:39+08:00 | 2018-04-14T08:32:36+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI815G | CPSM Final Test | FAIL
P | 2018-04-13T19:22:07+08:00 | 2018-04-14T08:33:08+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI815G | CPSM Final Test | PASS | 129 | 164.03 | 0.08 | -0.77 | 56.9 | 53.26 | 1.74 | 0.13 | 1 | -2 | 0 | -0.912 | 0.5484 | 2.8 | 80 |  | 0 | 0.1157 | 0.4273 | -0.0062 | 0.99999 | [43 empty] | 0
P | 2018-04-13T05:00:10+08:00 | 2018-04-14T09:10:45+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI811G | CPSM Final Test | FAIL | [14 empty] | 868
P | 2018-04-13T19:39:55+08:00 | 2018-04-14T10:53:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI814G | CPSM Final Test | PASS | 129 | 166.28 | 0.17 | -0.84 | 56.71 | 54.53 | 1.51 | 0.13 | 0 | 3 | 1 | -0.9578 | 0.4886 | 4.1 | -1 |  | 0 | 0.0584 | 0.3612 | 0.0071 | 0.99999 | [920 empty] | 0
P | 2018-04-13T03:15:12+08:00 | 2018-04-14T10:55:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI811G | CPSM Final Test | FAIL
P | 2018-04-13T19:05:27+08:00 | 2018-04-14T10:56:11+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI811G | CPSM Final Test | PASS | 129 | 153.68 | -0.02 | -0.77 | 57.06 | 53.17 | 1.06 | 0.13 | 0 | 2 | 1 | -0.9219 | 0.533 | 6.3 | 7 |  | 0.02063 | 0.2933 | 0.3609 | -0.0384 | 0.99999 | [785 empty] | 0
P | 2018-04-14T10:41:46+08:00 | 2018-04-14T13:34:24+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI817G | CPSM Final Test | PASS | 129 | 160.25 | 0.27 | -0.92 | 56.9 | 55.63 | 0.91 | 0.13 | 0 | -1 | 0 | -0.8421 | 0.548 | 1 | -9 |  | 0 | 0.1314 | 0.3182 | 0.0095 | 0.99999 | [921 empty] | 0
P | 2018-04-14T11:03:29+08:00 | 2018-04-14T13:49:11+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI818G | CPSM Final Test | PASS | 129 | 159.62 | 0.37 | -0.69 | 57.09 | 56.22 | 1.36 | 0.13 | -1 | 2 | 1 | -0.836 | 0.516 | 0.3 | 10 |  | 0 | 0.116 | 0.5637 | 0.0108 | 0.99999 | [424 empty] | 0
P | 2018-04-14T10:32:16+08:00 | 2018-04-14T13:56:53+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI816G | CPSM Final Test | PASS | 129 | 157.82 | 0.37 | -0.92 | 56.73 | 58.5 | 1.59 | 0.13 | 0 | -2 | -1 | -0.8679 | 0.5376 | 1.7 | 43 |  | 0.01938 | 0.0761 | 0.6799 | -0.0469 | 1 | [922 empty] | 0
P | 2018-04-14T11:29:56+08:00 | 2018-04-14T14:45:15+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI819G | CPSM Final Test | PASS | 129 | 166.19 | 0.27 | -0.69 | 56.83 | 54.44 | 0.98 | 0.13 | -1 | 2 | 0 | -0.826 | 0.468 | 6.9 | 3 |  | 0.01897 | 0.0679 | 0.0473 | 0.0128 | 1 | [923 empty] | 0
P | 2018-04-14T11:39:52+08:00 | 2018-04-14T15:36:05+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.72 | 0.49 | -0.84 | 57.19 | 56.73 | 1.13 | 0.13 | 1 | -2 | 0 | -0.884 | 0.5 | -3.1 | 10 |  |  | 0.2208 | 0.2051 | -0.042 | 0.99996 | [21 empty] | 0 | [575 empty] | 0
P | 2018-04-14T08:06:14+08:00 | 2018-04-14T15:42:23+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | FAIL
P | 2018-04-14T11:54:18+08:00 | 2018-04-14T15:42:51+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | 0.27 | -0.84 | 56.81 | 52.15 | 1.13 | 0.13 | 1 | 1 | -1 | -0.8201 | 0.4682 | 5.9 | 8 |  |  | 0.0932 | 0.0736 | -0.0285 | 0.99999 |  |  |  |  | 133 |  | 0.00405 | [854 empty] | 0
P | 2018-04-14T13:04:12+08:00 | 2018-04-14T15:43:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | D18SDI820G | CPSM Final Test | PASS | 129 | 158.81 | -0.02 | -0.77 | 56.72 | 54.1 | 1.82 | 0.13 | 1 | -2 | 1 | -0.7942 | 0.518 | 6.8 | 25 |  | 0 | 0.0805 | 0.2756 | 0.0735 | 0.99994 | [360 empty] | 0
P | 2018-04-14T11:41:25+08:00 | 2018-04-14T17:12:15+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI822G | CPSM Final Test | FAIL
P | 2018-04-14T14:31:52+08:00 | 2018-04-14T17:16:28+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI821G | CPSM Final Test | PASS | 129 | 160.25 | 0.17 | -0.92 | 56.91 | 52.75 | 1.67 | 0.13 | 0 | 2 | 1 | -0.8839 | 0.519 | 3.7 | -30 |  | 0.00706 | 0.0444 | 0.7822 | -0.0219 | 1 | [924 empty] | 0
P | 2018-04-14T15:46:24+08:00 | 2018-04-14T18:54:18+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI824G | CPSM Final Test | PASS | 129 | 161.87 | 0.08 | -0.54 | 57.02 | 54.78 | 1.36 | 0.13 | 1 | -3 | -1 | -0.8281 | 0.52 | 0.8 | -1 |  | 0 | 0.0752 | 0.3579 | -0.0161 | 0.99999 | [925 empty] | 0
P | 2018-04-14T15:28:48+08:00 | 2018-04-14T19:00:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI825G | CPSM Final Test | PASS | 129 | 166.64 | 0.17 | -0.99 | 56.64 | 53.68 | 1.36 | 0.13 | 1 | -2 | 0 | -0.8379 | 0.534 | 4.1 | 9 |  | 0 | 0.0464 | 0.1211 | -0.0068 | 1 | [926 empty] | 0
P | 2018-04-14T14:26:50+08:00 | 2018-04-14T19:06:21+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI822G | CPSM Final Test | PASS | 129 | 161.42 | 0.17 | -0.92 | 56.93 | 55.88 | 1.29 | 0.13 | 1 | -2 | 1 | -0.8879 | 0.541 | 4.3 | 6 |  | 0.02 | 0.0913 | 0.2154 | 0.0109 | 1 | [927 empty] | 0
P | 2018-04-14T16:16:16+08:00 | 2018-04-14T19:03:57+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | D18SDI826G | CPSM Final Test | PASS | 129 | 159.8 | 0.37 | -0.54 | 56.52 | 51.31 | 0.98 | 0.13 | 0 | -2 | 0 | -0.874 | 0.5344 | 2.3 | 27 |  | 0 | 0.0917 | 0.231 | 0.0731 | 0.99999 | [928 empty] | 0
P | 2018-04-14T16:00:38+08:00 | 2018-04-14T19:06:23+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI823G | CPSM Final Test | PASS | 129 | 163.13 | 0.46 | -0.92 | 56.82 | 53.85 | 1.06 | 0.13 | -1 | -12 | 1 | -0.8659 | 0.4882 | 4.9 | 22 |  | 0.01174 | 0.0727 | 0.2406 | -0.0295 | 1 | [778 empty] | 0
P | 2018-04-14T17:53:12+08:00 | 2018-04-14T20:32:54+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI827G | CPSM Final Test | PASS | 129 | 159.71 | 0.37 | -0.61 | 57.09 | 51.05 | 0.75 | 0.13 | 1 | -2 | 1 | -0.8561 | 0.5272 | 3.3 | -1 |  | 0.00977 | 0.2087 | 0.6054 | -0.0211 | 1 | [280 empty] | 0
P | 2018-04-14T19:01:47+08:00 | 2018-04-14T21:52:16+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI828G | CPSM Final Test | PASS | 129 | 157.55 | 0.27 | -0.69 | 56.61 | 49.53 | 0.91 | 0.13 | 0 | -3 | 0 | -0.8441 | 0.472 | 4.4 | -1 |  | 0 | 0.096 | 0.3955 | 0.0114 | 1 | [929 empty] | 0
P | 2018-04-14T19:13:35+08:00 | 2018-04-14T22:04:05+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI831G | CPSM Final Test | PASS | 129 | 151.97 | 0.37 | -0.92 | 56.7 | 54.53 | 1.29 | 0.13 | 0 | 2 | 0 | -0.9038 | 0.536 | 0.3 | 28 |  | 0.0096 | 0.0474 | 0.1261 | 0.0079 | 0.99999 | [930 empty] | 0
P | 2018-04-16T08:30:05+08:00 | 2018-04-16T11:07:21+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL | [13 empty] | -9.4
P | 2018-04-16T09:51:17+08:00 | 2018-04-16T12:50:47+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.76 | 0.17 | -0.69 | 57.02 | 51.73 | 1.51 | 0.13 | 1 | -3 | 0 | -0.7982 | 0.4716 | 4.9 | 4 |  | 0 | 0.0606 | 0.1306 | 0.0346 | 0.99996 | [8 empty] | 0
P | 2018-04-16T10:11:50+08:00 | 2018-04-16T12:54:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI829G | CPSM Final Test | PASS | 129 | 158.63 | 0.08 | -0.84 | 56.52 | 52.24 | 1.51 | 0.13 | 0 | -11 | 0 | -0.8519 | 0.4832 | 2.6 | 4 |  | 0 | 0.0753 | 0.1328 | -0.0039 | 1 | [931 empty] | 0
P | 2018-04-16T10:46:12+08:00 | 2018-04-16T13:33:16+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI832G | CPSM Final Test | PASS | 129 | 164.3 | 0.08 | -0.77 | 56.48 | 52.92 | 1.21 | 0.13 | 0 | -2 | 2 | -0.836 | 0.5092 | 2.4 | 2 |  | 0 | 0.2096 | 0.2129 | 0.0389 | 0.99998 | [932 empty] | 0
P | 2018-04-16T11:00:03+08:00 | 2018-04-16T13:42:35+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI830G | CPSM Final Test | PASS | 129 | 160.52 | 0.17 | -0.92 | 57.01 | 53.42 | 1.06 | 0.13 | -1 | -2 | -1 | -0.8321 | 0.528 | 1.8 | 17 |  | 0.01589 | 0.1793 | 0.1726 | -0.0241 | 0.99998 | [933 empty] | 0
P | 2018-04-16T11:16:14+08:00 | 2018-04-16T14:14:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI833G | CPSM Final Test | PASS | 129 | 157.28 | 0.56 | -0.77 | 57.03 | 53.34 | 2.05 | 0.13 | -1 | 4 | -1 | -0.9159 | 0.4694 | 6.1 | -6 |  | 0 | 0.0935 | 0.264 | -0.0312 | 1 | [934 empty] | 0
P | 2018-04-16T05:20:01+08:00 | 2018-04-16T14:17:25+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL | [17 empty] | 1.581
P | 2018-04-16T12:03:59+08:00 | 2018-04-16T15:17:08+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI835G | CPSM Final Test | PASS | 129 | 157.82 | 0.46 | -1.15 | 56.6 | 53 | 1.21 | 0.13 | 1 | 5 | -1 | -0.8998 | 0.5386 | 2.8 | 55 |  | 0.00946 | 0.0292 | 0.2875 | 0.0053 | 1 | [935 empty] | 0
P | 2018-04-16T12:12:58+08:00 | 2018-04-16T15:19:29+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | D18SDI834G | CPSM Final Test | PASS | 129 | 161.51 | -0.4 | -0.92 | 56.77 | 52.92 | 1.67 | 0.13 | -1 | 4 | 1 | -0.9498 | 0.5198 | 0.7 | 13 |  | 0 | 0.0772 | 0.1735 | 0.0281 | 0.99991 | [936 empty] | 0
P | 2018-04-16T14:10:33+08:00 | 2018-04-16T16:58:40+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI836G | CPSM Final Test | PASS | 129 | 156.47 | 0.46 | -0.77 | 56.83 | 51.56 | 1.44 | 0.13 | 1 | 2 | 1 | -0.9398 | 0.5392 | 0.4 | -2 |  | 0 | 0.0778 | 0.5629 | 0.0116 | 0.99999 | [937 empty] | 0
P | 2018-04-16T14:03:01+08:00 | 2018-04-16T17:01:01+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI837G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.99 | 56.8 | 54.69 | 1.21 | 0.13 | 0 | -2 | 1 | -0.9497 | 0.4646 | 9.8 | 65 |  | 0.01061 | 0.1426 | 0.138 | 0.0081 | 1 | [19 empty] | 0
P | 2018-04-16T14:06:16+08:00 | 2018-04-16T17:23:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | 0.27 | -0.08 | 56.85 | 53.17 | 1.13 | 0.13 | 1 | 1 | 1 | -0.9857 | 0.462 | 7.5 | 37 |  |  | 0.067 | 0.0687 | -0.0509 | 1 |  |  |  |  | 133 |  | 0.00664 | [931 empty] | 0
P | 2018-04-16T02:11:23+08:00 | 2018-04-16T17:26:03+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL
P | 2018-04-16T01:39:55+08:00 | 2018-04-16T17:57:31+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL
P | 2018-04-16T09:48:43+08:00 | 2018-04-16T19:00:37+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | -0.02 | -0.46 | 57.09 | 56.9 | 1.13 | 0.13 | -1 | -2 | 1 | -0.8141 | 0.534 | 6.6 | 44 |  |  | 0.1082 | 0.0561 | 0.018 | 1 |  |  |  |  | 133 |  | 0 | [932 empty] | 0
P | 2018-04-16T16:38:50+08:00 | 2018-04-16T19:23:15+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI839G | CPSM Final Test | PASS | 129 | 161.06 | 0.17 | -1.07 | 56.61 | 53.17 | 1.44 | 0.13 | 0 | -1 | -1 | -0.9557 | 0.5164 | 12.9 | 23 |  | 0 | 0.0458 | 0.2917 | -0.0161 | 0.99999 | [940 empty] | 0
P | 2018-04-16T16:20:29+08:00 | 2018-04-16T19:27:35+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI838G | CPSM Final Test | PASS | 129 | 161.33 | 0.17 | -0.69 | 56.61 | 53.51 | 1.21 | 0.13 | 1 | -4 | 1 | -0.9458 | 0.5126 | 1.8 | 67 |  | 0.0157 | 0.0418 | 0.1686 | -0.0262 | 0.99999 | [261 empty] | 0
P | 2018-04-16T16:43:47+08:00 | 2018-04-16T19:37:08+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI841G | CPSM Final Test | PASS | 129 | 156.65 | 0.37 | -1.38 | 56.7 | 56.05 | 0.68 | 0.13 | 0 | -2 | -1 | -0.9298 | 0.4736 | 1.3 | 0 |  | 0.02175 | 0.096 | 0.2929 | 0.0176 | 0.99999 | [941 empty] | 0
P | 2018-04-16T17:01:07+08:00 | 2018-04-16T19:42:35+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | D18SDI840G | CPSM Final Test | PASS | 129 | 161.15 | -0.02 | -1.15 | 57.02 | 53.93 | 1.44 | 0.13 | -1 | 3 | 1 | -0.8279 | 0.5164 | 1.2 | 22 |  | 0 | 0.1324 | 0.1613 | 0.0082 | 0.99997 | [256 empty] | 0
P | 2018-04-16T18:12:21+08:00 | 2018-04-16T21:10:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.81 | 0.78 | -0.77 | 57.19 | 57.4 | 1.06 | 0.13 | 1 | 3 | 0 | -0.9378 | 0.498 | 0.6 | 61 |  |  | 0.0544 | 0.1518 | -0.0106 | 0.99998 | [21 empty] | 0.00702 | [326 empty] | 0
P | 2018-04-16T18:35:55+08:00 | 2018-04-16T21:24:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI844G | CPSM Final Test | PASS | 129 | 161.6 | 0.08 | -0.69 | 56.77 | 55.8 | 1.13 | 0.13 | 1 | -2 | 0 | -0.9058 | 0.509 | 5.9 | 2 |  | 0.01231 | 0.0891 | 0.0927 | -0.0168 | 1 | [893 empty] | 0
P | 2018-04-16T18:47:27+08:00 | 2018-04-16T21:28:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI845G | CPSM Final Test | PASS | 129 | 157.19 | 0.17 | -0.54 | 56.84 | 53 | 1.21 | 0.13 | 0 | -2 | 1 | -0.9279 | 0.495 | 1.5 | 4 |  | 0 | 0.0975 | 0.5463 | -0.0067 | 0.99998 | [942 empty] | 0
P | 2018-04-16T05:53:20+08:00 | 2018-04-17T08:31:17+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | FAIL
P | 2018-04-16T19:12:18+08:00 | 2018-04-17T08:31:53+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | -0.21 | -0.99 | 57.01 | 53.09 | 1.13 | 0.13 | 1 | 3 | -1 | -0.9436 | 0.4682 | 4.7 | 40 |  |  | 0.0541 | 0.1772 | -0.0194 | 0.99997 |  |  |  |  | 133 |  | 0 | [765 empty] | 0
P | 2018-04-17T07:50:34+08:00 | 2018-04-17T10:18:08+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | FAIL | [13 empty] | -5.8
P | 2018-04-17T09:07:23+08:00 | 2018-04-17T12:10:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | -0.02 | -0.54 | 57.12 | 56.3 | 1.21 | 0.13 | 0 | -3 | 0 | -0.8261 | 0.544 | -2.5 | 8 |  |  | 0.0601 | 0.1083 | 0.0115 | 1 |  |  |  |  | 133 |  | 0 | [936 empty] | 0
P | 2018-04-17T09:04:21+08:00 | 2018-04-17T12:20:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.81 | 0.78 | -1.3 | 57.2 | 57.57 | 1.13 | 0.13 | 1 | 3 | 1 | -0.9438 | 0.4908 | 9.3 | 7 |  |  | 0.0681 | 0.2079 | 0.0202 | 0.99997 | [21 empty] | 0 | [846 empty] | 0
P | 2018-04-17T07:42:10+08:00 | 2018-04-17T12:35:49+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI847G | CPSM Final Test | FAIL
P | 2018-04-17T09:56:06+08:00 | 2018-04-17T12:38:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI848G | CPSM Final Test | PASS | 129 | 162.14 | 0.27 | -0.69 | 56.72 | 55.96 | 1.21 | 0.13 | 0 | -2 | -1 | -0.7882 | 0.5016 | 5.2 | 6 |  | 0 | 0.0704 | 0.3926 | -0.005 | 1 | [944 empty] | 0
P | 2018-04-17T09:46:48+08:00 | 2018-04-17T12:42:13+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI849G | CPSM Final Test | PASS | 129 | 160.34 | -0.02 | -0.23 | 57.22 | 53.51 | 1.9 | 0.13 | 0 | -3 | 1 | -0.828 | 0.4864 | 3.9 | 17 |  | 0.01114 | 0.1357 | 0.1168 | -0.0103 | 0.99999 | [188 empty] | 0
P | 2018-04-17T10:59:30+08:00 | 2018-04-17T13:50:39+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI846G | CPSM Final Test | PASS | 129 | 161.51 | 0.27 | -0.77 | 56.72 | 58.93 | 1.29 | 0.13 | 0 | 2 | -1 | -0.8061 | 0.5236 | 0.3 | -9 |  | 0 | 0.1814 | 0.3357 | 0.0062 | 1 | [945 empty] | 0
P | 2018-04-17T11:07:51+08:00 | 2018-04-17T13:51:30+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI842G | CPSM Final Test | PASS | 129 | 162.68 | 0.08 | -1.07 | 56.82 | 52.07 | 1.67 | 0.13 | 0 | -1 | 1 | -0.8999 | 0.5056 | 4.3 | -27 |  | 0.01931 | 0.056 | 0.3425 | -0.0051 | 0.99999 | [946 empty] | 0
P | 2018-04-17T10:09:00+08:00 | 2018-04-17T14:20:15+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI847G | CPSM Final Test | PASS | 129 | 162.95 | 0.27 | -0.69 | 56.89 | 55.96 | 1.74 | 0.13 | 0 | -2 | 0 | -0.886 | 0.5032 | 3 | 41 |  | 0.01708 | 0.0843 | 0.1572 | -0.0101 | 1 | [714 empty] | 0
P | 2018-04-17T11:15:08+08:00 | 2018-04-17T14:36:33+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | D18SDI843G | CPSM Final Test | PASS | 129 | 164.12 | 0.08 | -0.61 | 56.83 | 51.56 | 1.59 | 0.13 | -1 | -3 | 0 | -0.8021 | 0.537 | 5.5 | 21 |  | 0 | 0.1299 | 0.1863 | 0.0267 | 0.99992 |  |  |  | 0
P | 2018-04-17T11:45:54+08:00 | 2018-04-17T15:12:25+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI853G | CPSM Final Test | FAIL
P | 2018-04-17T11:30:04+08:00 | 2018-04-17T15:48:44+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI855G | CPSM Final Test | FAIL | [14 empty] | -1
P | 2018-04-17T13:34:45+08:00 | 2018-04-17T16:33:24+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.67 | 0.17 | -1.3 | 57.11 | 54.36 | 1.51 | 0.13 | 1 | -2 | 0 | -0.8441 | 0.465 | 3.6 | 5 |  | 0 | 0.1273 | 0.2794 | 0.0398 | 0.99999 | [947 empty] | 0
P | 2018-04-17T14:05:50+08:00 | 2018-04-17T16:48:25+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI850G | CPSM Final Test | PASS | 129 | 161.15 | 0.08 | -0.46 | 57.03 | 53.76 | 1.59 | 0.13 | 1 | -2 | -2 | -0.9338 | 0.545 | 0.2 | 6 |  | 0.01116 | 0.0852 | 0.2878 | -0.0107 | 0.99999 | [872 empty] | 0
P | 2018-04-17T13:29:09+08:00 | 2018-04-17T16:57:07+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI853G | CPSM Final Test | PASS | 129 | 160.34 | 0.37 | -0.61 | 57.05 | 54.19 | 1.29 | 0.13 | 0 | 4 | 1 | -0.8958 | 0.5118 | 0.8 | 22 |  | 0.01294 | 0.0814 | 0.0385 | -0.0556 | 0.99999 | [948 empty] | 0
P | 2018-04-17T14:27:50+08:00 | 2018-04-17T17:14:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI851G | CPSM Final Test | PASS | 129 | 156.11 | 0.17 | -0.84 | 56.83 | 51.65 | 1.59 | 0.13 | -1 | -11 | 1 | -0.9098 | 0.5284 | 6.9 | 10 |  | 0.03069 | 0.1174 | 0.3156 | -0.0167 | 1 | [949 empty] | 0
P | 2018-04-17T13:39:24+08:00 | 2018-04-17T17:32:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI855G | CPSM Final Test | PASS | 129 | 150.8 | 0.17 | -0.69 | 56.82 | 55.12 | 1.51 | 0.13 | 0 | -11 | 0 | -0.862 | 0.5208 | 6.9 | 6 |  | 0.00706 | 0.1043 | 0.3573 | 0.0173 | 0.99997 | [584 empty] | 0
P | 2018-04-17T14:46:11+08:00 | 2018-04-17T18:00:01+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | D18SDI856G | CPSM Final Test | PASS | 129 | 161.87 | 0.08 | -0.84 | 56.76 | 54.61 | 1.13 | 0.13 | -1 | -3 | 1 | -0.9218 | 0.481 | 1.3 | 39 |  | 0.01738 | 0.2052 | 0.1741 | 0.0081 | 1 | [950 empty] | 0
P | 2018-04-17T15:18:24+08:00 | 2018-04-17T18:13:17+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | D18SDI854G | CPSM Final Test | PASS | 129 | 165.2 | 0.08 | -0.61 | 57.1 | 53.09 | 1.36 | 0.13 | 0 | -3 | 1 | -0.8381 | 0.517 | 11.3 | -27 |  | 0 | 0.0962 | 0.098 | 0.056 | 0.99995 | [785 empty] | 0
P | 2018-04-17T17:07:52+08:00 | 2018-04-17T20:12:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI855G | CPSM Final Test | FAIL | [14 empty] | 6
P | 2018-04-17T17:08:41+08:00 | 2018-04-17T20:34:53+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI858G | CPSM Final Test | PASS | 129 | 157.37 | 0.08 | -0.92 | 56.87 | 56.9 | 0.98 | 0.13 | -1 | -1 | 1 | -0.9219 | 0.4652 | 4.1 | 50 |  | 0 | 0.0319 | 0.0672 | -0.0034 | 1 | [951 empty] | 0
P | 2018-04-17T17:19:51+08:00 | 2018-04-17T20:58:17+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | D18SDI857G | CPSM Final Test | PASS | 129 | 161.6 | 0.37 | -0.54 | 57.11 | 55.12 | 1.29 | 0.13 | -1 | 2 | -1 | -0.8261 | 0.5366 | 4 | 2 |  | 0 | 0.1318 | 0.2704 | 0.0098 | 1 | [97 empty] | 0
P | 2018-04-17T17:35:04+08:00 | 2018-04-17T20:57:44+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI852G | CPSM Final Test | PASS | 129 | 162.14 | -0.02 | -0.77 | 56.84 | 52.15 | 1.29 | 0.13 | 0 | -8 | 0 | -0.8321 | 0.5012 | 2.7 | 7 |  | 0.02102 | 0.0242 | 0.6104 | 0.0142 | 0.99999 | [952 empty] | 0
P | 2018-04-17T18:40:22+08:00 | 2018-04-17T21:56:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI855G | CPSM Final Test | PASS | 129 | 150.8 | 0.17 | -0.69 | 56.83 | 56.05 | 1.51 | 0.13 | 0 | -11 | 0 | -0.9577 | 0.496 | 4.5 | 7 |  | 0.00927 | 0.0589 | 0.3252 | 0.0163 | 0.99998 | [205 empty] | 0
P | 2018-04-18T18:50:33+08:00 | 2018-04-18T21:47:34+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | D18USM955G | CPSM Final Test | PASS |  | 159.89 | 0.56 | -0.31 | 56.46 | 53.17 | 1.51 | 0.13 | -1 | 4 | 0 | -0.886 | 0.5 | 8.1 | 1 |  |  | 0.0519 | 0.3771 | -0.0009 | 0.99995 |  |  |  |  | 133 |  | 0.00614 | [612 empty] | 0
P | 2018-04-18T18:51:55+08:00 | 2018-04-18T22:10:00+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | D18USM958G | CPSM Final Test | PASS |  | 171.95 | -0.02 | -0.54 | 56.48 | 48.94 | 1.13 | 0.13 | 0 | 3 | -1 | -0.9778 | 0.496 | 0.5 | -59 |  |  | 0.0578 | 0.1032 | -0.0192 | 0.99999 |  |  |  |  | 133 |  | 0.01055 | [946 empty] | 0
P | 2018-04-18T04:32:17+08:00 | 2018-04-19T09:31:43+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | D18USM957G | CPSM Final Test | FAIL
P | 2018-04-19T09:59:18+08:00 | 2018-04-19T10:49:46+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | D18USM961G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-04-18T19:02:00+08:00 | 2018-04-19T11:39:07+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | D18USM957G | CPSM Final Test | PASS |  | 163.76 | 0.17 | -0.61 | 56.33 | 51.05 | 0.75 | 0.13 | -1 | 4 | 2 | -0.9577 | 0.511 | 0.1 | -6 |  |  | 0.0737 | 0.3931 | -0.0165 | 0.99999 |  |  |  |  | 133 |  | 0 | [947 empty] | 0
P | 2018-04-19T07:11:26+08:00 | 2018-04-19T11:51:43+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | D18USM959G | CPSM Final Test | FAIL
P | 2018-04-19T09:34:33+08:00 | 2018-04-19T12:33:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | D18USM960G | CPSM Final Test | PASS |  | 164.21 | -0.02 | -0.31 | 56.75 | 49.78 | 0.98 | 0.13 | -1 | -3 | -1 | -0.8281 | 0.5424 | 6.9 | 77 |  |  | 0.022 | 0.2934 | -0.0153 | 1 |  |  |  |  | 133 |  | 0.02533 | [948 empty] | 0
P | 2018-04-19T10:30:28+08:00 | 2018-04-19T13:14:58+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | D18USM962G | CPSM Final Test | PASS |  | 158.54 | -0.49 | 0 | 56.5 | 49.53 | 1.67 | 0.13 | 0 | -1 | -1 | -0.8242 | 0.4854 | 8.3 | 2 |  |  | 0.243 | 0.4843 | -0.0444 | 1 |  |  |  |  | 133 |  | 0 | [949 empty] | 0
P | 2018-04-19T10:24:32+08:00 | 2018-04-19T13:34:40+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | D18USM961G | CPSM Final Test | PASS |  | 158.54 | -0.02 | -0.31 | 56.5 | 51.9 | 1.21 | 0.13 | 0 | -3 | -1 | -0.88 | 0.4934 | 3.8 | 3 |  |  | 0.0562 | 0.1172 | -0.0366 | 1 |  |  |  |  | 133 |  | 0.01264 | [950 empty] | 0
P | 2018-04-19T09:31:34+08:00 | 2018-04-19T13:41:57+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | D18USM959G | CPSM Final Test | PASS |  | 165.29 | 0.17 | -0.46 | 56.48 | 52.49 | 0.91 | 0.13 | 0 | -5 | -1 | -0.7761 | 0.5222 | 2.2 | 4 |  |  | 0.1099 | 0.0864 | 0.017 | 0.99996 |  |  |  |  | 133 |  | 0 | [896 empty] | 0
P | 2018-04-19T12:02:11+08:00 | 2018-04-19T15:45:17+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | D18USM963G | CPSM Final Test | PASS |  | 156.56 | 0.27 | -0.31 | 56.5 | 50.97 | 1.36 | 0.13 | 0 | 2 | 1 | -0.8879 | 0.4868 | 3.7 | -15 |  |  | 0.0478 | 0.2862 | -0.0298 | 0.99999 |  |  |  |  | 133 |  | 0.01188 | [113 empty] | 0
P | 2018-04-19T15:58:26+08:00 | 2018-04-19T16:42:18+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI861G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-04-19T15:24:28+08:00 | 2018-04-19T17:16:16+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI861G | CPSM Final Test | FAIL | [5 empty] | 45.89
P | 2018-04-19T14:57:12+08:00 | 2018-04-19T17:48:03+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI859G | CPSM Final Test | PASS | 129 | 158.36 | 0.65 | -0.99 | 56.13 | 49.53 | 1.44 | 0.13 | 0 | -11 | 1 | -0.7083 | 0.524 | 9.4 | -3 |  | 0 | 0.0707 | 0.1095 | 0.0252 | 0.99996 | [903 empty] | 0
P | 2018-04-19T14:54:19+08:00 | 2018-04-19T17:55:42+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | D18USM956G | CPSM Final Test | PASS |  | 149.45 | -0.21 | -0.54 | 56.46 | 53 | 1.21 | 0.13 | -2 | -3 | 0 | -0.9757 | 0.5014 | 2.9 | 1 |  |  | 0.0536 | 0.3334 | -0.0104 | 0.99999 |  |  |  |  | 133 |  | 0 | [292 empty] | 0
P | 2018-04-19T14:38:33+08:00 | 2018-04-19T18:01:58+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI860G | CPSM Final Test | PASS | 129 | 161.96 | 0.68 | -0.69 | 56.48 | 50.38 | 1.67 | 0.13 | 0 | -11 | 3 | -0.9099 | 0.5096 | 3.8 | 30 |  | 0 | 0.081 | 0.2633 | -0.0018 | 0.99999 | [958 empty] | 0
P | 2018-04-19T13:15:22+08:00 | 2018-04-19T19:25:22+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI861G | CPSM Final Test | FAIL
P | 2018-04-19T16:53:33+08:00 | 2018-04-19T20:25:41+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | D18SDI862G | CPSM Final Test | PASS | 129 | 157.28 | 0.97 | -1.3 | 56.09 | 49.7 | 1.06 | 0.13 | -1 | -2 | 0 | -0.8978 | 0.4938 | 3.9 | 6 |  | 0.02131 | 0.0738 | 0.0964 | -0.0132 | 0.99999 | [74 empty] | 0
P | 2018-04-19T17:29:41+08:00 | 2018-04-19T20:41:03+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI863G | CPSM Final Test | PASS | 129 | 166.28 | 0.17 | -0.92 | 56.15 | 51.56 | 1.06 | 0.13 | 0 | 4 | 1 | -0.9358 | 0.543 | 0 | 15 |  | 0 | 0.1329 | 0.2524 | -0.043 | 0.99998 | [623 empty] | 0
P | 2018-04-19T16:20:22+08:00 | 2018-04-19T21:09:34+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI861G | CPSM Final Test | PASS | 129 | 163.85 | 0.46 | -0.54 | 56.55 | 49.7 | 1.29 | 0.13 | -1 | -2 | 0 | -0.9018 | 0.496 | 3 | 10 |  | 0.0123 | 0.0643 | 0.0722 | -0.0759 | 0.99999 | [959 empty] | 0
P | 2018-04-19T03:39:47+08:00 | 2018-04-20T09:39:23+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI864G | CPSM Final Test | FAIL | [13 empty] | 18.5
P | 2018-04-20T09:19:34+08:00 | 2018-04-20T10:12:26+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI866G | CPSM Final Test | FAIL
P | 2018-04-20T10:30:43+08:00 | 2018-04-20T10:59:18+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | FAIL
P | 2018-04-19T20:04:32+08:00 | 2018-04-20T11:25:36+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI865G | CPSM Final Test | PASS | 129 | 161.87 | 0.17 | -0.99 | 56.31 | 50.12 | 1.21 | 0.13 | 1 | 2 | 2 | -0.8421 | 0.5048 | 3.4 | 33 |  | 0.0451 | 0.0494 | 0.7353 | -0.0238 | 1 | [418 empty] | 0
P | 2018-04-19T18:39:35+08:00 | 2018-04-20T11:39:37+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI864G | CPSM Final Test | PASS | 129 | 160.88 | 0.46 | -0.84 | 56.44 | 54.1 | 1.51 | 0.13 | 0 | -2 | 3 | -0.818 | 0.493 | 6.5 | 20 |  | 0 | 0.1124 | 0.2654 | 0.0462 | 0.99996 | [960 empty] | 0
P | 2018-04-20T09:46:00+08:00 | 2018-04-20T12:52:10+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI866G | CPSM Final Test | PASS | 129 | 158.09 | 0.97 | -0.84 | 56.49 | 50.55 | 1.29 | 0.13 | 1 | 4 | 1 | -0.8839 | 0.5318 | 2.4 | -3 |  | 0.0151 | 0.112 | 0.2268 | -0.0098 | 1 | [407 empty] | 0
P | 2018-04-20T06:55:10+08:00 | 2018-04-20T12:58:28+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI867G | CPSM Final Test | FAIL
P | 2018-04-20T09:56:49+08:00 | 2018-04-20T12:58:58+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI867G | CPSM Final Test | PASS | 129 | 161.06 | -0.02 | -0.92 | 56.32 | 50.72 | 1.29 | 0.13 | 1 | 4 | -1 | -0.834 | 0.505 | 2.6 | 16 |  | 0.01401 | 0.1066 | 0.4501 | -0.0096 | 0.99999 | [961 empty] | 0
P | 2018-04-20T08:40:39+08:00 | 2018-04-20T13:20:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | FAIL | [790 empty] | 1819252
P | 2018-04-20T10:47:32+08:00 | 2018-04-20T13:39:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.72 | 1.25 | -0.99 | 57.2 | 58.08 | 1.13 | 0.13 | 1 | 3 | 1 | -0.8779 | 0.494 | 4.9 | 2 |  |  | 0.0415 | 0.153 | -0.0186 | 1 | [21 empty] | 0 | [722 empty] | 0
P | 2018-04-20T10:38:46+08:00 | 2018-04-20T13:39:37+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI868G | CPSM Final Test | PASS | 129 | 155.66 | -0.3 | -0.69 | 56.62 | 49.95 | 1.13 | 0.13 | 1 | -2 | 0 | -0.8599 | 0.5128 | 6.8 | 15 |  | 0 | 0.0367 | 0.2542 | -0.0192 | 0.99998 | [962 empty] | 0
P | 2018-04-20T10:32:17+08:00 | 2018-04-20T13:48:35+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | D18USM964G | CPSM Final Test | PASS |  | 157.64 | 0.37 | 0.07 | 56.72 | 48.68 | 1.51 | 0.13 | 0 | 3 | 0 | -0.8221 | 0.5074 | 1.5 | 20 |  |  | 0.1212 | 0.2539 | -0.0085 | 0.99996 |  |  |  |  | 133 |  | 0 | [229 empty] | 0
P | 2018-04-20T11:00:20+08:00 | 2018-04-20T14:27:25+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | -0.21 | -1.3 | 56.86 | 51.73 | 1.21 | 0.13 | 1 | 3 | 0 | -0.7861 | 0.4762 | 7.8 | 8 |  |  | 0.1671 | 0.3239 | -0.0092 | 1 |  |  |  |  | 133 |  | 0.01246 | [762 empty] | 0
P | 2018-04-20T13:57:10+08:00 | 2018-04-20T17:19:35+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI874G | CPSM Final Test | PASS | 129 | 159.35 | -0.02 | -0.92 | 56.3 | 50.55 | 1.36 | 0.13 | -1 | 1 | 1 | -0.8101 | 0.5202 | 1.8 | -1 |  | 0 | 0.0386 | 0.1521 | 0.0573 | 0.99997 | [963 empty] | 0
P | 2018-04-20T13:52:52+08:00 | 2018-04-20T17:19:49+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI871G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.31 | 56.71 | 54.44 | 1.9 | 0.13 | -1 | 4 | 1 | -0.898 | 0.5132 | 3.3 | -4 |  | 0 | 0.1589 | 0.2337 | -0.0425 | 1 | [251 empty] | 0
P | 2018-04-20T13:49:02+08:00 | 2018-04-20T17:26:44+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI869G | CPSM Final Test | PASS | 129 | 156.11 | -0.02 | -0.61 | 56.51 | 52.07 | 1.59 | 0.13 | 0 | -12 | 1 | -0.89 | 0.517 | 1.3 | 14 |  | 0.01842 | 0.0444 | 0.0518 | 0.0011 | 1 | [964 empty] | 0
P | 2018-04-20T14:14:39+08:00 | 2018-04-20T17:31:04+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI870G | CPSM Final Test | PASS | 129 | 162.14 | 0.37 | -0.69 | 56.46 | 49.19 | 1.44 | 0.13 | 1 | -3 | 1 | -0.824 | 0.5448 | 14.6 | -10 |  | 0.05506 | 0.047 | 0.3496 | 0.0085 | 1 | [965 empty] | 0
P | 2018-04-20T14:45:44+08:00 | 2018-04-20T17:32:25+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.72 | 0.87 | -1.3 | 57.23 | 58.25 | 1.06 | -0.2 | -1 | 2 | 0 | -1.0017 | 0.4942 | 2.8 | 32 |  |  | 0.0923 | 0.3877 | -0.0252 | 0.99999 | [21 empty] | 0 | [747 empty] | 0
P | 2018-04-20T14:43:23+08:00 | 2018-04-20T17:37:57+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | -0.21 | -1.3 | 56.87 | 52.49 | 1.13 | 0.13 | 1 | 2 | 1 | -0.9817 | 0.4724 | 3.8 | -3 |  |  | 0.1021 | 0.0954 | -0.0331 | 1 |  |  |  |  | 133 |  | 0 | [807 empty] | 0
P | 2018-04-20T14:53:31+08:00 | 2018-04-20T18:51:59+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI873G | CPSM Final Test | PASS | 129 | 160.25 | -0.02 | -0.84 | 56.14 | 49.19 | 1.21 | 0.13 | -2 | -2 | 0 | -0.8978 | 0.531 | 5.6 | 23 |  | 0 | 0.1263 | 0.2248 | 0.0052 | 1 | [966 empty] | 0
P | 2018-04-20T14:48:30+08:00 | 2018-04-20T18:58:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI872G | CPSM Final Test | PASS | 129 | 155.75 | -0.02 | -0.69 | 56.21 | 50.12 | 0.91 | 0.13 | 0 | 1 | 0 | -0.9078 | 0.5102 | 2.7 | 22 |  | 0 | 0.1121 | 0.0659 | 0.0266 | 0.99995 | [825 empty] | 0
P | 2018-04-20T18:44:42+08:00 | 2018-04-20T21:37:43+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI878G | CPSM Final Test | PASS | 129 | 166.28 | 0.27 | -1.07 | 56.31 | 51.65 | 1.36 | 0.13 | 0 | -2 | 0 | -0.854 | 0.4612 | 2.7 | 0 |  | 0 | 0.1359 | 0.1608 | 0.0248 | 0.99995 | [967 empty] | 0
P | 2018-04-20T18:40:10+08:00 | 2018-04-20T21:41:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI877G | CPSM Final Test | PASS | 129 | 157.91 | 0.17 | -0.54 | 56.42 | 55.2 | 1.74 | 0.13 | 1 | 4 | 0 | -0.9019 | 0.505 | 2.2 | 2 |  | 0 | 0.0246 | 0.2915 | 0.0095 | 0.99998 | [307 empty] | 0
P | 2018-04-20T18:33:09+08:00 | 2018-04-20T21:46:38+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI875G | CPSM Final Test | PASS | 129 | 165.29 | -0.02 | -0.92 | 56.2 | 50.8 | 1.36 | 0.13 | 0 | 3 | -1 | -0.9278 | 0.5094 | 1.4 | -5 |  | 0.01289 | 0.0983 | 0.8739 | -0.0124 | 1 | [119 empty] | 0
P | 2018-04-20T18:56:27+08:00 | 2018-04-20T21:51:31+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.67 | 0.17 | -0.84 | 57.14 | 56.9 | 1.51 | 0.13 | 1 | -2 | 0 | -0.9099 | 0.4914 | -3 | 36 |  | 0 | 0.0791 | 0.1896 | 0.0276 | 0.99999 | [606 empty] | 0
P | 2018-04-20T18:52:09+08:00 | 2018-04-20T21:54:33+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | 0.17 | -0.54 | 57.11 | 56.3 | 1.13 | 0.13 | 0 | -3 | 0 | -0.91 | 0.525 | -0.3 | 11 |  |  | 0.0448 | 0.0589 | -0.0169 | 1 |  |  |  |  | 133 |  | 0 | [360 empty] | 0
P | 2018-04-20T03:29:18+08:00 | 2018-04-21T09:52:16+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI876G | CPSM Final Test | FAIL | [699 empty] | 1556
P | 2018-04-21T07:22:20+08:00 | 2018-04-21T10:28:51+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL | [13 empty] | -6.6
P | 2018-04-21T07:34:48+08:00 | 2018-04-21T10:38:57+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI881G | CPSM Final Test | FAIL | [13 empty] | -8.7
P | 2018-04-20T18:40:47+08:00 | 2018-04-21T10:57:48+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI876G | CPSM Final Test | PASS | 129 | 160.34 | -0.02 | -0.84 | 56.32 | 49.95 | 1.29 | 0.13 | -1 | -2 | 1 | -0.9159 | 0.4906 | 6.6 | -6 |  | 0 | 0.0478 | 0.3547 | 0.0191 | 1 | [678 empty] | 0
P | 2018-04-21T06:29:46+08:00 | 2018-04-21T11:43:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI881G | CPSM Final Test | FAIL | [14 empty] | 874
P | 2018-04-21T09:02:29+08:00 | 2018-04-21T12:03:41+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI880G | CPSM Final Test | PASS | 129 | 154.04 | 0.37 | -0.69 | 56.41 | 54.78 | 1.51 | 0.13 | 0 | -3 | 0 | -0.7962 | 0.5324 | 2.7 | 2 |  | 0 | 0.1561 | 0.2341 | -0.0348 | 0.99999 | [430 empty] | 0
P | 2018-04-21T08:58:53+08:00 | 2018-04-21T12:12:55+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | 0.17 | -0.61 | 57.11 | 56.3 | 1.13 | 0.13 | 0 | -3 | 0 | -0.9078 | 0.5122 | 2.1 | 2 |  |  | 0.0538 | 0.2576 | -0.0225 | 1 |  |  |  |  | 133 |  | 0.00305 | [961 empty] | 0
P | 2018-04-21T08:55:35+08:00 | 2018-04-21T12:29:47+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.76 | 0.17 | -0.77 | 57.18 | 55.03 | 1.51 | 0.13 | 1 | -2 | 0 | -0.8959 | 0.4808 | 9.1 | 13 |  | 0 | 0.0718 | 0.1296 | 0.024 | 0.99999 | [969 empty] | 0
P | 2018-04-21T09:56:07+08:00 | 2018-04-21T12:48:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI879G | CPSM Final Test | PASS | 129 | 161.69 | -0.21 | -0.69 | 56.3 | 52.41 | 1.13 | 0.13 | 1 | -2 | 1 | -0.83 | 0.5286 | 1.3 | 12 |  | 0 | 0.0745 | 0.4745 | 0.0102 | 1 | [970 empty] | 0
P | 2018-04-21T10:25:38+08:00 | 2018-04-21T13:23:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI882G | CPSM Final Test | PASS | 129 | 151.97 | 0.87 | -0.84 | 56.56 | 50.55 | 1.74 | 0.13 | -1 | -2 | 0 | -0.8281 | 0.5022 | 4 | -11 |  | 0 | 0.1145 | 0.3036 | 0.058 | 0.99996 | [971 empty] | 0
P | 2018-04-21T09:06:52+08:00 | 2018-04-21T13:28:49+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI881G | CPSM Final Test | PASS | 129 | 161.51 | 0.46 | -0.92 | 56.21 | 48.85 | 1.06 | 0.13 | 0 | -4 | 0 | -0.84 | 0.5048 | 1.7 | 38 |  | 0 | 0.0931 | 0.2076 | 0.0356 | 0.99998 | [590 empty] | 0
P | 2018-04-21T13:49:59+08:00 | 2018-04-21T15:29:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI886G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-04-21T12:59:59+08:00 | 2018-04-21T15:45:37+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI883G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.92 | 56.13 | 50.29 | 1.67 | 0.13 | -1 | -2 | -1 | -0.8658 | 0.528 | 0.2 | 10 |  | 0.01057 | 0.1792 | 0.1267 | -0.0338 | 1 | [972 empty] | 0
P | 2018-04-21T13:27:13+08:00 | 2018-04-21T15:45:53+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI885G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-04-21T13:00:53+08:00 | 2018-04-21T15:53:46+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI884G | CPSM Final Test | PASS | 129 | 159.17 | -0.02 | -0.61 | 56.48 | 51.22 | 0.98 | 0.13 | 1 | 3 | 0 | -0.9058 | 0.5058 | 3 | 18 |  | 0 | 0.0462 | 0.7772 | -0.0088 | 1 | [678 empty] | 0
P | 2018-04-21T13:14:43+08:00 | 2018-04-21T16:05:14+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI886G | CPSM Final Test | FAIL | [5 empty] | 46.06
P | 2018-04-23T09:27:09+08:00 | 2018-04-23T10:13:04+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI890G | CPSM Final Test | FAIL | [5 empty] | 46.82
P | 2018-04-21T15:52:02+08:00 | 2018-04-23T10:55:25+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI885G | CPSM Final Test | PASS | 129 | 163.76 | -0.02 | -1.07 | 56.13 | 50.46 | 1.74 | 0.13 | 0 | -1 | 1 | -0.9277 | 0.5066 | 14 | 42 |  | 0 | 0.1572 | 0.1934 | 0.025 | 1 | [970 empty] | 0
P | 2018-04-21T14:39:59+08:00 | 2018-04-23T11:08:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI886G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -0.54 | 56.52 | 49.02 | 1.59 | 0.13 | -1 | -2 | 0 | -0.9518 | 0.47 | -0.9 | 10 |  | 0 | 0.1881 | 0.1026 | 0.0017 | 0.99998 | [29 empty] | 0
P | 2018-04-21T15:57:00+08:00 | 2018-04-23T11:59:32+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI887G | CPSM Final Test | PASS | 129 | 157.55 | -0.02 | -0.84 | 56.18 | 49.19 | 1.82 | 0.13 | 0 | -2 | 1 | -0.88 | 0.5008 | 3.9 | -6 |  | 0 | 0.0973 | 0.3196 | 0.0492 | 0.99996 | [46 empty] | 0
P | 2018-04-23T09:24:16+08:00 | 2018-04-23T12:12:50+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI888G | CPSM Final Test | PASS | 129 | 161.87 | -0.02 | -0.84 | 56.21 | 51.9 | 1.59 | 0.13 | -1 | 3 | 1 | -0.8221 | 0.51 | 7.4 | 6 |  | 0 | 0.284 | 0.3991 | -0.0048 | 1 | [973 empty] | 0
P | 2018-04-23T09:46:49+08:00 | 2018-04-23T12:53:18+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI889G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.84 | 56.28 | 49.78 | 1.51 | 0.13 | 0 | -11 | 0 | -0.852 | 0.5126 | 1.2 | 20 |  | 0.01205 | 0.0312 | 0.1034 | -0.0069 | 1 | [974 empty] | 0
P | 2018-04-23T09:50:06+08:00 | 2018-04-23T13:08:12+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI890G | CPSM Final Test | PASS | 129 | 157.19 | -0.02 | -0.61 | 56.12 | 48.77 | 1.74 | 0.13 | -1 | -12 | 0 | -0.7742 | 0.511 | 1.3 | 17 |  | 0.03342 | 0.0584 | 0.3445 | -0.0279 | 1 | [975 empty] | 0
P | 2018-04-23T11:52:52+08:00 | 2018-04-23T15:38:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI891G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.54 | 56.48 | 48.6 | 2.12 | 0.13 | -1 | 2 | 0 | -0.7862 | 0.5032 | 2.1 | 16 |  | 0 | 0.1118 | 0.1747 | -0.0067 | 1 | [976 empty] | 0
P | 2018-04-23T11:32:21+08:00 | 2018-04-23T15:41:16+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI892G | CPSM Final Test | PASS | 129 | 162.41 | 0.17 | -0.16 | 56.51 | 51.56 | 2.05 | 0.13 | 0 | -1 | 1 | -0.884 | 0.5014 | -1.5 | 1 |  | 0 | 0.0635 | 0.1292 | 0.0145 | 1 | [977 empty] | 0
P | 2018-04-23T13:52:07+08:00 | 2018-04-23T17:24:56+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI893G | CPSM Final Test | PASS | 129 | 160.79 | -0.02 | -0.77 | 56.44 | 55.03 | 1.51 | 0.13 | 0 | 2 | 1 | -0.9239 | 0.5408 | 3.8 | 17 |  | 0 | 0.0511 | 0.3772 | -0.0012 | 1 | [880 empty] | 0
P | 2018-04-23T13:51:23+08:00 | 2018-04-23T17:35:38+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI894G | CPSM Final Test | PASS | 129 | 161.96 | 0.17 | -0.99 | 56.23 | 48.68 | 1.51 | 0.13 | 0 | -2 | 1 | -0.8939 | 0.5492 | 3 | -21 |  | 0.00842 | 0.0616 | 0.4986 | 0.0039 | 1 | [978 empty] | 0
P | 2018-04-23T14:27:03+08:00 | 2018-04-23T18:02:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI895G | CPSM Final Test | PASS | 129 | 161.6 | -0.02 | -0.31 | 56.31 | 49.36 | 1.44 | 0.13 | -1 | -10 | -4 | -0.8499 | 0.4924 | 1.8 | 19 |  | 0 | 0.0565 | 0.2781 | 0.0066 | 0.99996 | [106 empty] | 0
P | 2018-04-23T13:47:49+08:00 | 2018-04-23T19:17:39+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI898G | CPSM Final Test | FAIL | [14 empty] | 928
P | 2018-04-23T17:26:47+08:00 | 2018-04-23T20:06:30+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI897G | CPSM Final Test | FAIL
P | 2018-04-23T12:46:52+08:00 | 2018-04-23T20:09:58+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI899G | CPSM Final Test | FAIL | [1000 empty] | 0
P | 2018-04-23T16:09:48+08:00 | 2018-04-23T20:22:38+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI896G | CPSM Final Test | PASS | 129 | 160.7 | -0.02 | 0 | 56.3 | 50.46 | 1.51 | 0.13 | 0 | 3 | -1 | -0.9079 | 0.4944 | -1.5 | 18 |  | 0 | 0.0478 | 0.3157 | 0.0178 | 1 | [980 empty] | 0
P | 2018-04-23T16:32:44+08:00 | 2018-04-23T21:02:17+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI898G | CPSM Final Test | PASS | 129 | 159.17 | 0.17 | -0.77 | 56.41 | 51.48 | 1.59 | 0.13 | 0 | 1 | -1 | -0.886 | 0.5048 | -3.3 | -10 |  | 0 | 0.065 | 0.0587 | 0.0237 | 1 | [981 empty] | 0
P | 2018-04-23T16:28:25+08:00 | 2018-04-23T21:15:30+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI899G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.92 | 56.22 | 49.95 | 1.36 | 0.13 | 0 | 4 | 1 | -0.8799 | 0.548 | 1.4 | -10 |  | 0.01152 | 0.0679 | 0.1545 | 0.0213 | 0.99999 | [979 empty] | 0
P | 2018-04-23T18:43:35+08:00 | 2018-04-23T21:58:00+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI900G | CPSM Final Test | PASS | 129 | 158.36 | -0.02 | -0.61 | 56.49 | 48.68 | 1.44 | 0.13 | 0 | -1 | -1 | -0.8818 | 0.544 | 4.9 | 8 |  | 0 | 0.0905 | 0.2772 | -0.0089 | 0.99997 | [439 empty] | 0
P | 2018-04-23T18:46:38+08:00 | 2018-04-23T22:00:24+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI897G | CPSM Final Test | PASS | 129 | 164.93 | -0.02 | -0.61 | 56.49 | 50.88 | 1.51 | 0.13 | 1 | -2 | -1 | -0.9857 | 0.4918 | 3.4 | -3 |  | 0 | 0.0914 | 0.19 | 0.0569 | 0.99999 | [982 empty] | 0
P | 2018-04-24T07:25:58+08:00 | 2018-04-24T11:54:04+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI901G | CPSM Final Test | FAIL
P | 2018-04-24T10:03:23+08:00 | 2018-04-24T13:11:06+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI902G | CPSM Final Test | PASS | 129 | 160.16 | 0.46 | -1.3 | 56.09 | 51.14 | 1.13 | 0.13 | -2 | -11 | -1 | -0.852 | 0.507 | 6.4 | -10 |  | 0.01221 | 0.0513 | 0.3973 | 0.008 | 0.99999 | [983 empty] | 0
P | 2018-04-24T09:40:01+08:00 | 2018-04-24T13:47:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | D18SDI901G | CPSM Final Test | PASS | 129 | 159.71 | 0.37 | -0.92 | 56.34 | 50.04 | 1.36 | 0.13 | -2 | -2 | 1 | -0.8081 | 0.5034 | 1.8 | 2 |  | 0.01122 | 0.0893 | 0.4355 | -0.0114 | 0.99999 | [984 empty] | 0
P | 2018-04-24T05:20:09+08:00 | 2018-04-24T14:41:13+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI903G | CPSM Final Test | FAIL
P | 2018-04-24T11:37:11+08:00 | 2018-04-24T15:22:44+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI904G | CPSM Final Test | PASS | 129 | 164.03 | 0.27 | -0.92 | 56.42 | 54.69 | 1.51 | 0.13 | 1 | -2 | -1 | -0.8361 | 0.543 | 2.3 | 30 |  | 0.00919 | 0.2494 | 0.2644 | -0.027 | 0.99996 | [985 empty] | 0
P | 2018-04-24T11:41:01+08:00 | 2018-04-24T15:26:39+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | D18SDI905G | CPSM Final Test | PASS | 129 | 163.4 | 0.17 | -0.54 | 56.21 | 49.61 | 1.74 | 0.13 | -1 | 2 | 0 | -0.864 | 0.5064 | 0.3 | 18 |  | 0 | 0.0631 | 0.0426 | 0.02 | 0.99999 | [315 empty] | 0
P | 2018-04-24T12:18:48+08:00 | 2018-04-24T15:35:02+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI906G | CPSM Final Test | PASS | 129 | 165.11 | 0.37 | -0.84 | 56.2 | 52.32 | 1.36 | 0.13 | -1 | 3 | -1 | -0.8441 | 0.5096 | 10.3 | 4 |  | 0 | 0.1126 | 0.348 | 0.0084 | 0.99999 | [986 empty] | 0
P | 2018-04-24T16:04:31+08:00 | 2018-04-24T18:57:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | D18SDI908G | CPSM Final Test | PASS | 129 | 163.49 | 0.17 | -0.61 | 56.54 | 49.95 | 1.67 | 0.13 | 0 | 4 | 1 | -0.8919 | 0.518 | 3.7 | 66 |  | 0.01101 | 0.1207 | 0.2609 | -0.0297 | 1 | [987 empty] | 0
P | 2018-04-24T14:28:12+08:00 | 2018-04-24T17:47:58+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | D18SDI907G | CPSM Final Test | PASS | 129 | 159.62 | 0.27 | -0.77 | 56.42 | 52.75 | 1.51 | 0.13 | -2 | -8 | 0 | -0.8978 | 0.546 | 3.1 | 2 |  | 0 | 0.1476 | 0.2328 | 0.0028 | 0.99999 | [319 empty] | 0
P | 2018-04-24T10:00:41+08:00 | 2018-04-24T16:35:55+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | D18SDI903G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -0.84 | 56.18 | 49.02 | 1.51 | 0.13 | -1 | 3 | 0 | -0.8818 | 0.5308 | 8.8 | 2 |  | 0.01034 | 0.0529 | 0.0521 | 0.0139 | 1 | [979 empty] | 0
P | 2018-04-24T17:26:12+08:00 | 2018-04-24T20:50:14+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | D18SDI909G | CPSM Final Test | PASS | 129 | 163.94 | 0.27 | -0.77 | 56.45 | 48.09 | 1.9 | 0.13 | -2 | -2 | -1 | -0.9139 | 0.504 | 0.1 | 7 |  | 0.00813 | 0.2402 | 0.3543 | -0.0049 | 0.99999 | [988 empty] | 0
P | 2018-04-25T07:47:17+08:00 | 2018-04-25T10:54:00+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | FAIL | [13 empty] | 19.8
P | 2018-04-25T09:29:01+08:00 | 2018-04-25T12:30:13+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.72 | 0.87 | -0.92 | 57.21 | 58.17 | 1.13 | 0.13 | 1 | 3 | 1 | -0.8779 | 0.502 | 4.6 | 13 |  |  | 0.1449 | 0.3034 | 0.0018 | 0.99999 | [21 empty] | 0.01838 | [967 empty] | 0
P | 2018-04-25T09:20:38+08:00 | 2018-04-25T12:57:58+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | -0.21 | -0.99 | 56.88 | 50.88 | 1.21 | 0.13 | 1 | 3 | 0 | -0.7942 | 0.479 | 12.8 | 19 |  |  | 0.169 | 0.2032 | 0.0267 | 0.99994 |  |  |  |  | 133 |  | 0 | [983 empty] | 0
P | 2018-04-25T13:41:35+08:00 | 2018-04-25T17:21:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | -0.21 | -0.92 | 56.88 | 52.32 | 1.13 | 0.13 | 1 | 3 | 1 | -0.9277 | 0.46 | 5.7 | -38 |  |  | 0.0779 | 0.4674 | -0.0025 | 0.99999 |  |  |  |  | 133 |  | 0.02235 | [673 empty] | 0
P | 2018-04-25T13:33:23+08:00 | 2018-04-25T17:17:30+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.72 | 0.87 | -0.92 | 57.23 | 58.59 | 1.06 | 0.13 | -1 | -1 | 0 | -0.888 | 0.491 | 0.8 | 74 |  |  | 0.0753 | 0.3011 | 0.0329 | 0.99992 | [21 empty] | 0 | [111 empty] | 0
P | 2018-04-25T15:17:02+08:00 | 2018-04-25T18:10:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | D18USM966G | CPSM Final Test | PASS |  | 163.76 | -0.02 | 0 | 56.51 | 50.21 | 1.82 | 0.13 | 1 | 4 | 0 | -0.9278 | 0.5348 | -3.8 | 19 |  |  | 0.0744 | 0.0763 | 0.0162 | 1 |  |  |  |  | 133 |  | 0.00748 | [984 empty] | 0
P | 2018-04-25T15:27:32+08:00 | 2018-04-25T18:13:22+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | D18USM965G | CPSM Final Test | PASS |  | 163.67 | -0.3 | -0.54 | 56.41 | 51.31 | 1.59 | 0.13 | 0 | -2 | 0 | -0.9179 | 0.523 | 10.6 | 26 |  |  | 0.2261 | 0.1969 | -0.0862 | 0.99991 |  |  |  |  | 133 |  | 0 | [985 empty] | 0
P | 2018-04-25T16:17:22+08:00 | 2018-04-25T19:05:55+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | D18USM967G | CPSM Final Test | PASS |  | 158.36 | 0.37 | -0.69 | 56.31 | 52.24 | 1.67 | 0.13 | 1 | 2 | 0 | -0.9119 | 0.5032 | 1.6 | 5 |  |  | 0.1253 | 0.1448 | -0.047 | 0.99999 |  |  |  |  | 133 |  | 0 | [536 empty] | 0
P | 2018-04-25T14:21:06+08:00 | 2018-04-25T20:06:05+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM968G | CPSM Final Test | FAIL | [911 empty] | 0
P | 2018-04-25T17:13:35+08:00 | 2018-04-25T21:11:55+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM968G | CPSM Final Test | PASS |  | 161.87 | 0.37 | -0.77 | 56.21 | 50.21 | 2.2 | 0.13 | 0 | -2 | 0 | -0.91 | 0.5002 | -3.8 | -4 |  |  | 0.1313 | 0.3668 | -0.0126 | 0.99999 |  |  |  |  | 133 |  | 0 | [883 empty] | 0
P | 2018-04-26T08:52:45+08:00 | 2018-04-26T10:05:07+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | D18USM971G | CPSM Final Test | FAIL | [5 empty] | 46.91
P | 2018-04-26T07:42:28+08:00 | 2018-04-26T10:19:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | D18USM969G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-04-26T07:35:31+08:00 | 2018-04-26T11:22:21+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | D18USM971G | CPSM Final Test | FAIL
P | 2018-04-26T09:13:17+08:00 | 2018-04-26T11:27:01+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM972G | CPSM Final Test | FAIL | [13 empty] | -8.7
P | 2018-04-26T09:00:44+08:00 | 2018-04-26T12:23:08+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | D18USM969G | CPSM Final Test | PASS |  | 163.58 | 0.46 | -0.84 | 56.43 | 49.36 | 1.36 | 0.13 | -1 | 2 | 0 | -0.8101 | 0.4954 | 4.8 | 37 |  |  | 0.2683 | 0.1542 | -0.0562 | 0.99998 |  |  |  |  | 133 |  | 0.00666 | [491 empty] | 0
P | 2018-04-26T09:48:59+08:00 | 2018-04-26T12:35:19+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | D18USM970G | CPSM Final Test | PASS |  | 159.26 | 0.27 | -1.07 | 56.41 | 48.68 | 1.29 | 0.13 | -1 | 2 | 1 | -0.8101 | 0.501 | -3.3 | 17 |  |  | 0.0875 | 0.0547 | -0.0034 | 1 |  |  |  |  | 133 |  | 0 | [544 empty] | 0
P | 2018-04-25T20:12:01+08:00 | 2018-04-26T12:46:55+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | 0.27 | -0.38 | 57.1 | 55.88 | 1.13 | 0.13 | 0 | -2 | 0 | -0.8281 | 0.515 | -2.8 | -20 |  |  | 0.107 | 0.288 | 0.0172 | 1 |  |  |  |  | 133 |  | 0.02473 | [253 empty] | 0
P | 2018-04-25T03:26:39+08:00 | 2018-04-26T12:42:56+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL
P | 2018-04-25T20:04:47+08:00 | 2018-04-26T13:09:30+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.67 | 0.08 | -0.92 | 57.17 | 55.54 | 1.51 | 0.13 | 1 | 1 | 0 | -0.7962 | 0.4762 | 1.4 | 9 |  | 0 | 0.1072 | 0.2647 | 0.0761 | 0.99991 | [312 empty] | 0
P | 2018-04-26T09:28:56+08:00 | 2018-04-26T13:10:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | D18USM971G | CPSM Final Test | PASS |  | 154.04 | 1.03 | -0.99 | 56.19 | 48.26 | 1.44 | 0.13 | 1 | -4 | -1 | -0.864 | 0.527 | 3.8 | 16 |  |  | 0.0837 | 0.1175 | -0.0497 | 0.99999 |  |  |  |  | 133 |  | 0 | [70 empty] | 0
P | 2018-04-26T10:20:09+08:00 | 2018-04-26T13:30:41+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | D18USM972G | CPSM Final Test | PASS |  | 163.31 | 0.08 | -0.61 | 56.61 | 49.53 | 1.51 | 0.13 | -1 | -1 | -1 | -0.8361 | 0.5264 | 5.7 | -4 |  |  | 0.1349 | 0.2737 | -0.0405 | 0.99999 |  |  |  |  | 133 |  | 0 | [883 empty] | 0
P | 2018-04-26T10:48:40+08:00 | 2018-04-26T13:53:22+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | D18USM973G | CPSM Final Test | PASS |  | 163.58 | 0.08 | -1.22 | 56.01 | 49.78 | 1.29 | 0.13 | 0 | -11 | 1 | -0.854 | 0.5414 | 7 | 9 |  |  | 0.2244 | 0.5557 | -0.0332 | 1 |  |  |  |  | 133 |  | 0.02109 | [6 empty] | 0
P | 2018-04-26T13:01:17+08:00 | 2018-04-26T14:04:30+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | D18USM974G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-04-26T12:26:30+08:00 | 2018-04-26T16:04:15+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI911G | CPSM Final Test | FAIL
P | 2018-04-26T13:32:53+08:00 | 2018-04-26T16:42:06+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | D18USM974G | CPSM Final Test | PASS |  | 151.97 | 0.37 | -0.99 | 56.16 | 49.36 | 1.29 | 0.13 | -1 | 1 | 0 | -0.7941 | 0.5124 | 1.3 | 56 |  |  | 0.135 | 0.2166 | -0.0353 | 0.99981 |  |  |  |  | 133 |  | 0.0149 | [752 empty] | 0
P | 2018-04-26T14:02:38+08:00 | 2018-04-26T16:46:29+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI910G | CPSM Final Test | PASS | 129 | 157.37 | 0.27 | -0.77 | 56.23 | 51.39 | 1.59 | 0.13 | 1 | -2 | 1 | -0.906 | 0.521 | 0.6 | 17 |  | 0 | 0.1101 | 0.1687 | 0.0256 | 0.99998 | [993 empty] | 0
P | 2018-04-26T16:09:48+08:00 | 2018-04-26T16:42:54+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI912G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-04-26T14:15:22+08:00 | 2018-04-26T17:48:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI911G | CPSM Final Test | PASS | 129 | 157.19 | 0.27 | -1.07 | 56.4 | 55.29 | 0.91 | 0.13 | 1 | -2 | -1 | -0.874 | 0.4732 | -2.4 | 24 |  | 0 | 0.0892 | 0.3957 | -0.028 | 1 | [994 empty] | 0
P | 2018-04-26T16:08:16+08:00 | 2018-04-26T18:54:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | D18SDI915G | CPSM Final Test | FAIL | [5 empty] | 46.14
P | 2018-04-26T16:14:22+08:00 | 2018-04-26T19:24:43+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.67 | 0.08 | -0.77 | 57.23 | 56.81 | 1.51 | 0.13 | 1 | 1 | 1 | -0.9737 | 0.4784 | 1.4 | 10 |  | 0.01454 | 0.0927 | 0.2435 | 0.0449 | 0.99996 | [372 empty] | 0
P | 2018-04-26T16:13:14+08:00 | 2018-04-26T19:21:01+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI913G | CPSM Final Test | PASS | 129 | 159.08 | 0.37 | -0.84 | 55.99 | 48.34 | 0.68 | 0.13 | 1 | 1 | 0 | -0.914 | 0.5482 | 0.8 | 4 |  | 0 | 0.058 | 0.4161 | -0.0087 | 1 | [995 empty] | 0
P | 2018-04-26T16:26:21+08:00 | 2018-04-26T19:21:38+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI912G | CPSM Final Test | PASS | 129 | 160.07 | 0.46 | -1.38 | 55.87 | 49.45 | 0.75 | 0.13 | -1 | -2 | -1 | -0.9618 | 0.532 | 0.2 | -7 |  | 0.01329 | 0.1062 | 0.3845 | -0.0024 | 1 | [996 empty] | 0
P | 2018-04-26T12:43:40+08:00 | 2018-04-26T19:30:46+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL | [476 empty] | 0
P | 2018-04-26T16:51:29+08:00 | 2018-04-26T19:50:59+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI914G | CPSM Final Test | FAIL
P | 2018-04-26T17:31:27+08:00 | 2018-04-26T21:28:44+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | D18SDI915G | CPSM Final Test | PASS | 129 | 157.01 | 0.56 | -1.3 | 56.15 | 50.29 | 1.13 | 0.13 | 0 | -2 | 0 | -0.9179 | 0.5358 | 2.8 | 3 |  | 0.00952 | 0.1119 | 0.205 | -0.0208 | 0.99999 | [997 empty] | 0
P | 2018-04-26T18:21:14+08:00 | 2018-04-26T21:35:51+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI914G | CPSM Final Test | PASS | 129 | 163.31 | 0.56 | -0.84 | 56.22 | 49.45 | 1.21 | 0.13 | 1 | -2 | 0 | -0.946 | 0.4864 | 0.3 | -5 |  | 0 | 0.1113 | 0.0737 | 0.0012 | 0.99998 | [558 empty] | 0
P | 2018-04-26T18:42:58+08:00 | 2018-04-26T21:41:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI916G | CPSM Final Test | PASS | 129 | 158.72 | -0.02 | -0.69 | 56.2 | 50.8 | 1.74 | 0.13 | 0 | 1 | 0 | -0.8879 | 0.5228 | 2.1 | 1 |  | 0 | 0.0607 | 0.1041 | -0.014 | 1 | [998 empty] | 0
P | 2018-04-25T19:40:14+08:00 | 2018-04-27T12:34:12+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL | [17 empty] | 60.2262
P | 2018-04-27T10:01:17+08:00 | 2018-04-27T12:49:19+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI918G | CPSM Final Test | PASS | 129 | 162.23 | 0.46 | -0.61 | 56.46 | 51.82 | 1.51 | 0.13 | -1 | -3 | 0 | -0.8321 | 0.5356 | -3.4 | 16 |  | 0 | 0.0431 | 0.2576 | -0.0021 | 1 | [784 empty] | 0
P | 2018-04-27T10:15:37+08:00 | 2018-04-27T12:59:25+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI917G | CPSM Final Test | PASS | 129 | 171.86 | 0.46 | -0.99 | 55.79 | 48.26 | 1.59 | 0.13 | 0 | 1 | -1 | -0.7662 | 0.5008 | 0 | 28 |  | 0 | 0.0867 | 0.3504 | -0.0096 | 0.99999 | [123 empty] | 0
P | 2018-04-27T10:32:39+08:00 | 2018-04-27T13:19:28+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | D18SDI919G | CPSM Final Test | PASS | 129 | 166.64 | 0.46 | -0.08 | 56.27 | 50.04 | 1.51 | 0.13 | 1 | 4 | 0 | -0.8281 | 0.4862 | 3.2 | -20 |  | 0.01545 | 0.322 | 0.0899 | -0.0622 | 0.99995 | [999 empty] | 0
P | 2018-04-26T16:07:13+08:00 | 2018-04-27T13:37:12+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | 0.08 | -0.54 | 57.11 | 56.56 | 1.13 | 0.13 | 0 | -3 | 0 | -0.856 | 0.5044 | 10.4 | 2 |  |  | 0.2939 | 0.229 | -0.144 | 0.99989 |  |  |  |  | 133 |  | 0 | [448 empty] | 0
P | 2018-04-27T11:53:15+08:00 | 2018-04-27T14:35:20+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI921G | CPSM Final Test | PASS | 129 | 160.97 | 0.37 | -0.61 | 56.27 | 48.85 | 1.9 | 0.13 | 0 | 4 | 1 | -0.9078 | 0.513 | 4.5 | 19 |  | 0.01361 | 0.0702 | 0.2535 | -0.0132 | 0.99999 | [996 empty] | 0
P | 2018-04-27T09:29:34+08:00 | 2018-04-27T14:36:45+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI920G | CPSM Final Test | FAIL
P | 2018-04-27T12:03:09+08:00 | 2018-04-27T16:21:49+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI920G | CPSM Final Test | PASS | 129 | 165.92 | 0.08 | 0.07 | 56.32 | 49.7 | 1.36 | 0.13 | 1 | -2 | 0 | -0.9418 | 0.5018 | 1.9 | 7 |  | 0 | 0.1329 | 0.1281 | -0.0091 | 1 | [527 empty] | 0
P | 2018-04-27T14:08:07+08:00 | 2018-04-27T16:26:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI924G | CPSM Final Test | FAIL | [13 empty] | -5.6
P | 2018-04-27T14:03:53+08:00 | 2018-04-27T16:51:42+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | D18SDI923G | CPSM Final Test | PASS | 129 | 160.25 | 0.37 | -0.69 | 56.44 | 50.88 | 1.36 | 0.13 | 0 | -4 | 0 | -0.864 | 0.5188 | 1.4 | 18 |  | 0.01424 | 0.1088 | 0.2875 | 0.0425 | 0.99998 | [1000 empty] | 0
P | 2018-04-27T14:24:30+08:00 | 2018-04-27T17:20:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI922G | CPSM Final Test | PASS | 129 | 164.39 | -0.02 | -0.69 | 56.26 | 49.28 | 1.51 | 0.13 | 0 | 1 | 0 | -0.9099 | 0.528 | 2.8 | 26 |  | 0 | 0.081 | 0.1268 | -0.0443 | 1 | [1001 empty] | 0
P | 2018-04-27T15:17:22+08:00 | 2018-04-27T18:26:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI924G | CPSM Final Test | PASS | 129 | 158 | 0.27 | -0.69 | 56.34 | 51.48 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9059 | 0.4968 | 0 | 10 |  | 0 | 0.0504 | 0.3172 | -0.0401 | 0.99999 | [1002 empty] | 0
P | 2018-04-27T16:53:29+08:00 | 2018-04-27T19:38:39+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI925G | CPSM Final Test | PASS | 129 | 157.73 | -0.02 | -1.07 | 55.92 | 50.63 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9418 | 0.4786 | 3.2 | 34 |  | 0 | 0.1189 | 0.1714 | 0.0058 | 1 | [322 empty] | 0
P | 2018-04-27T16:30:04+08:00 | 2018-04-27T19:40:12+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI926G | CPSM Final Test | PASS | 129 | 154.85 | 0.56 | -0.84 | 56.32 | 49.53 | 1.67 | 0.13 | 1 | -12 | 0 | -0.908 | 0.5154 | 4.8 | 80 |  | 0.02211 | 0.0845 | 0.2602 | 0.0031 | 0.99997 | [968 empty] | 0
P | 2018-04-27T17:18:39+08:00 | 2018-04-27T19:51:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI928G | CPSM Final Test | FAIL | [13 empty] | 17.9
P | 2018-04-27T17:39:51+08:00 | 2018-04-27T20:50:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | D18SDI927G | CPSM Final Test | PASS | 129 | 161.87 | 0.37 | -0.31 | 56.09 | 48.18 | 1.44 | 0.13 | -1 | 2 | 0 | -0.8481 | 0.496 | 5.3 | 17 |  | 0.02735 | 0.234 | 0.135 | -0.0094 | 0.99999 | [1003 empty] | 0
P | 2018-04-27T18:34:54+08:00 | 2018-04-27T21:51:31+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI928G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.46 | 56.24 | 50.04 | 1.97 | 0.13 | 0 | -1 | 0 | -0.9238 | 0.5262 | 6.8 | -8 |  | 0.00863 | 0.0966 | 0.1593 | -0.0199 | 0.99999 | [1004 empty] | 0
P | 2018-04-27T19:07:38+08:00 | 2018-04-27T21:51:48+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI929G | CPSM Final Test | PASS | 129 | 162.32 | 0.27 | -1.07 | 55.93 | 49.36 | 0.83 | 0.13 | 0 | 1 | 0 | -0.8659 | 0.5244 | 3.4 | -8 |  | 0 | 0.0687 | 0.2752 | -0.0405 | 1 | [1005 empty] | 0
P | 2018-04-28T10:52:49+08:00 | 2018-04-28T11:27:42+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI931G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-04-28T09:12:59+08:00 | 2018-04-28T12:07:32+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | D18SDI930G | CPSM Final Test | PASS | 129 | 162.59 | 0.17 | -1.3 | 56.37 | 50.29 | 1.21 | 0.13 | -1 | 1 | 0 | -0.7862 | 0.5294 | -4.7 | 10 |  | 0.01131 | 0.1777 | 0.2877 | 0.0273 | 1 | [1006 empty] | 0
P | 2018-04-28T10:59:43+08:00 | 2018-04-28T13:45:47+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI932G | CPSM Final Test | PASS | 129 | 162.77 | 0.08 | -0.46 | 57.04 | 53.26 | 1.97 | 0.13 | 1 | 1 | 0 | -0.8081 | 0.4876 | -3.9 | 15 |  | 0 | 0.1856 | 0.1469 | 0.0146 | 0.99998 | [1007 empty] | 0
P | 2018-04-28T11:38:39+08:00 | 2018-04-28T14:54:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI931G | CPSM Final Test | PASS | 129 | 161.15 | 0.27 | -0.77 | 56.18 | 49.19 | 1.59 | 0.13 | 0 | -1 | 0 | -0.8181 | 0.534 | 5.3 | 61 |  | 0 | 0.1317 | 0.2994 | -0.0245 | 1 | [290 empty] | 0
P | 2018-04-28T13:12:54+08:00 | 2018-04-28T16:04:24+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI935G | CPSM Final Test | PASS | 129 | 161.6 | 0.37 | -0.54 | 56.33 | 48.34 | 1.9 | 0.13 | -1 | -3 | -1 | -0.9118 | 0.457 | 4.6 | -26 |  | 0.02413 | 0.0762 | 0.319 | -0.0413 | 0.99999 | [1008 empty] | 0
P | 2018-04-28T13:26:20+08:00 | 2018-04-28T16:19:31+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | D18SDI934G | CPSM Final Test | PASS | 129 | 155.57 | 0.37 | -0.92 | 56.24 | 51.22 | 1.74 | 0.13 | 1 | 2 | 0 | -0.9338 | 0.4938 | -2.6 | 11 |  | 0 | 0.2336 | 0.2041 | -0.0201 | 0.99997 | [1009 empty] | 0
P | 2018-04-28T14:44:40+08:00 | 2018-04-28T18:03:35+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | D18SDI937G | CPSM Final Test | PASS | 129 | 157.1 | 0.17 | -0.92 | 56.4 | 49.95 | 0.91 | 0.13 | 1 | -2 | -1 | -0.9239 | 0.4974 | 3.4 | 22 |  | 0 | 0.1174 | 0.1506 | -0.0134 | 0.99991 | [495 empty] | 0
P | 2018-04-28T15:34:22+08:00 | 2018-04-28T18:30:57+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | D18SDI936G | CPSM Final Test | PASS | 129 | 159.98 | -0.21 | 0.23 | 56.38 | 49.02 | 1.97 | 0.13 | -1 | 2 | -1 | -0.9598 | 0.4822 | 0.2 | -16 |  | 0.01418 | 0.102 | 0.2234 | -0.0104 | 1 | [1010 empty] | 0
P | 2018-04-28T15:30:00+08:00 | 2018-04-28T18:33:28+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | D18SDI933G | CPSM Final Test | PASS | 129 | 156.2 | 0.97 | -0.54 | 56.51 | 50.12 | 1.29 | 0.13 | -2 | 3 | 0 | -0.9139 | 0.5214 | 0 | -2 |  | 0 | 0.0696 | 0.316 | -0.0046 | 1 | [907 empty] | 0
P | 2018-04-28T16:13:14+08:00 | 2018-04-28T18:53:20+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | D18SDI938G | CPSM Final Test | PASS | 129 | 157.64 | 0.27 | -0.92 | 56.48 | 50.04 | 1.74 | 0.13 | 0 | -2 | 1 | -0.9559 | 0.5338 | 4 | -37 |  | 0.02722 | 0.0213 | 0.2579 | 0.0052 | 0.99998 | [1011 empty] | 0
P | 2018-05-02T11:10:42+08:00 | 2018-05-02T15:55:54+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | FAIL
P | 2018-05-02T13:33:18+08:00 | 2018-05-02T17:44:22+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.62 | -0.21 | -1.3 | 56.87 | 54.02 | 1.13 | 0.13 | 1 | 2 | 1 | -0.9837 | 0.4608 | 1.9 | 11 |  |  | 0.1446 | 0.119 | -0.0184 | 0.99999 |  |  |  |  | 133 |  | 0 | [1005 empty] | 0
P | 2018-05-02T17:51:41+08:00 | 2018-05-02T20:36:06+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.72 | 1.06 | -1.3 | 57.2 | 58.34 | 1.06 | 0.13 | 1 | -2 | 0 | -0.9417 | 0.497 | 5.3 | -10 |  |  | 0.1139 | 0.1758 | -0.027 | 0.99999 | [21 empty] | 0 | [991 empty] | 0
P | 2018-05-03T07:13:59+08:00 | 2018-05-03T10:51:00+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | FAIL | [13 empty] | -6.8
P | 2018-05-03T05:55:35+08:00 | 2018-05-03T12:24:29+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | FAIL
P | 2018-05-03T09:10:02+08:00 | 2018-05-03T12:36:55+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | D18FNP002G | CPSM Final Test | PASS |  | 159.71 | -0.21 | -0.16 | 56.87 | 53 | 1.13 | 0.13 | 1 | 3 | 0 | -0.8181 | 0.486 | 5.9 | -11 |  |  | 0.0828 | 0.4565 | -0.0164 | 1 |  |  |  |  | 133 |  | 0.02903 | [1007 empty] | 0
P | 2018-05-03T09:02:30+08:00 | 2018-05-03T12:52:32+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.67 | 0.17 | -0.69 | 57.22 | 56.81 | 1.51 | 0.13 | 1 | 1 | 0 | -0.8759 | 0.484 | 14.7 | -5 |  | 0 | 0.0628 | 0.0763 | 0.0294 | 0.99999 | [219 empty] | 0
P | 2018-05-03T14:10:31+08:00 | 2018-05-03T17:02:00+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.56 | 0.08 | -0.69 | 57.12 | 55.71 | 1.13 | 0.13 | 0 | -3 | 0 | -0.9718 | 0.5122 | 2.3 | 1 |  |  | 0.062 | 0.2067 | -0.0125 | 1 |  |  |  |  | 133 |  | 0 | [234 empty] | 0
P | 2018-05-03T14:13:10+08:00 | 2018-05-03T17:13:17+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | C18SDI805G | CPSM Final Test | PASS | 129 | 163.76 | 0.17 | -1.07 | 57.25 | 57.23 | 1.51 | 0.13 | 1 | 1 | -1 | -0.9737 | 0.4796 | -2.3 | 9 |  | 0.02283 | 0.104 | 0.4116 | 0.0256 | 0.99999 | [381 empty] | 0
P | 2018-05-03T17:15:38+08:00 | 2018-05-03T19:48:05+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | FAIL | [13 empty] | -8.2
P | 2018-05-03T18:31:52+08:00 | 2018-05-03T21:51:21+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | D18USM950G | CPSM Final Test | PASS |  | 165.65 | 0.37 | -0.61 | 57.12 | 55.46 | 1.06 | 0.13 | 0 | -3 | -1 | -0.9697 | 0.5204 | 7.4 | -17 |  |  | 0.0568 | 0.3354 | -0.0018 | 1 |  |  |  |  | 133 |  | 0.02991 | [466 empty] | 0
P | 2018-05-04T08:51:39+08:00 | 2018-05-04T11:39:19+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | D18FTP001G | CPSM Final Test | PASS | 129 | 158.81 | 1.06 | -1.3 | 57.26 | 58.42 | 1.13 | 0.13 | -1 | 3 | 0 | -0.8959 | 0.496 | -2.7 | 16 |  |  | 0.071 | 0.2875 | -0.0266 | 1 | [21 empty] | 0.02704 | [993 empty] | 0
P | 2018-05-07T12:35:29+08:00 | 2018-05-07T14:51:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP401G | CPSM Final Test | FAIL | [13 empty] | -7.4
P | 2018-05-07T12:25:51+08:00 | 2018-05-07T15:55:55+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP400G | CPSM Final Test | PASS | 129 | 160.07 | 0.27 | -0.77 | 56.18 | 50.12 | 1.44 | 0.13 | -1 | -2 | 0 | -0.9359 | 0.5086 | 4.8 | 6 |  |  | 0.2241 | 0.2064 | -0.0014 | 0.99999 | [21 empty] | 0 | [994 empty] | 0
P | 2018-05-07T13:43:18+08:00 | 2018-05-07T16:58:45+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP401G | CPSM Final Test | PASS | 129 | 160.88 | -0.02 | -0.54 | 56.31 | 49.11 | 1.29 | 0.13 | -1 | -4 | 0 | -0.87 | 0.5072 | 9.3 | 3 |  |  | 0.181 | 0.2796 | -0.0005 | 1 | [21 empty] | 0 | [995 empty] | 0
P | 2018-05-07T12:09:01+08:00 | 2018-05-07T16:59:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP402G | CPSM Final Test | FAIL
P | 2018-05-07T14:40:52+08:00 | 2018-05-07T18:21:29+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP404G | CPSM Final Test | FAIL | [13 empty] | -8.3
P | 2018-05-07T15:44:10+08:00 | 2018-05-07T18:48:02+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP403G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.99 | 56.3 | 48.6 | 1.36 | 0.13 | -1 | 1 | 1 | -0.9019 | 0.5416 | 2.3 | -6 |  |  | 0.2953 | 0.1963 | 0.006 | 0.99997 | [21 empty] | 0.01277 | [996 empty] | 0
P | 2018-05-07T09:12:09+08:00 | 2018-05-07T19:56:49+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP402G | CPSM Final Test | FAIL | [19 empty] | -1.0234
P | 2018-05-07T16:31:11+08:00 | 2018-05-07T20:24:35+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP404G | CPSM Final Test | PASS | 129 | 166.1 | 0.17 | -0.23 | 55.93 | 50.97 | 1.36 | 0.13 | 1 | 3 | 0 | -0.9777 | 0.487 | -1.5 | 109 |  |  | 0.0431 | 0.2718 | -0.0303 | 1 | [21 empty] | 0 | [962 empty] | 0
P | 2018-05-07T14:34:29+08:00 | 2018-05-07T20:59:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP402G | CPSM Final Test | PASS | 129 | 160.7 | -0.02 | -0.23 | 56.43 | 54.27 | 1.59 | 0.13 | -1 | 2 | 0 | -0.9318 | 0.5044 | 3.3 | 10 |  |  | 0.2284 | 0.2028 | 0.0037 | 0.99995 | [21 empty] | 0 | [997 empty] | 0
P | 2018-05-08T15:33:05+08:00 | 2018-05-08T16:55:26+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | E18FNP402G | CPSM Final Test | FAIL | [841 empty] | FALSE
P | 2018-05-08T14:12:34+08:00 | 2018-05-08T17:02:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP400G | CPSM Final Test | PASS |  | 156.92 | 0.27 | -0.38 | 56.17 | 48.09 | 1.29 | 0.13 | 1 | 2 | 0 | -0.9717 | 0.55 | 3.3 | 24 |  |  | 0.0771 | 0.1308 | -0.0443 | 0.99999 |  |  |  |  | 133 |  | 0 | [1013 empty] | 0
P | 2018-05-08T13:18:19+08:00 | 2018-05-08T19:03:40+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP401G | CPSM Final Test | FAIL
P | 2018-05-08T16:11:00+08:00 | 2018-05-08T19:16:36+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP401G | CPSM Final Test | PASS |  | 162.05 | 0.27 | -0.77 | 56.18 | 50.46 | 1.74 | 0.13 | 1 | -2 | 0 | -0.9459 | 0.5492 | 11.4 | 3 |  |  | 0.0783 | 0.1981 | -0.0384 | 0.99998 |  |  |  |  | 133 |  | 0 | [1014 empty] | 0
P | 2018-05-08T16:14:16+08:00 | 2018-05-08T20:31:14+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | E18FNP402G | CPSM Final Test | PASS |  | 164.48 | 0.27 | -0.92 | 56.16 | 50.29 | 1.06 | 0.13 | -1 | -1 | 0 | -1.0116 | 0.5096 | 8.3 | 9 |  |  | 0.0541 | 0.0319 | 0.0242 | 0.9999 |  |  |  |  | 133 |  | 0 | [1015 empty] | 0
P | 2018-05-08T17:27:51+08:00 | 2018-05-08T20:59:19+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP403G | CPSM Final Test | PASS |  | 160.7 | -0.02 | -0.38 | 56.42 | 57.23 | 1.36 | 0.13 | 1 | 3 | 0 | -1.0075 | 0.539 | 0.7 | 17 |  |  | 0.1328 | 0.3931 | -0.0261 | 1 |  |  |  |  | 133 |  | 0 | [1016 empty] | 0
P | 2018-05-08T19:01:38+08:00 | 2018-05-08T21:46:51+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP404G | CPSM Final Test | PASS |  | 160.52 | -0.52 | -0.61 | 55.94 | 49.36 | 1.06 | 0.13 | 0 | 3 | 0 | -0.9039 | 0.5144 | 12.7 | 6 |  |  | 0.0955 | 0.2116 | -0.0315 | 0.99999 |  |  |  |  | 133 |  | 0 | [517 empty] | 0
P | 2018-05-10T08:07:12+08:00 | 2018-05-10T10:21:48+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP405G | CPSM Final Test | FAIL | [13 empty] | -5.4
P | 2018-05-10T08:10:59+08:00 | 2018-05-10T12:05:55+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP406G | CPSM Final Test | FAIL
P | 2018-05-10T09:14:30+08:00 | 2018-05-10T12:21:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP405G | CPSM Final Test | PASS | 129 | 159.98 | 0.46 | -0.61 | 56.13 | 50.04 | 1.13 | 0.13 | 0 | -2 | -1 | -0.8141 | 0.528 | 10.3 | 10 |  |  | 0.1121 | 0.2601 | 0.0253 | 0.99999 | [21 empty] | 0 | [560 empty] | 0
P | 2018-05-10T09:50:53+08:00 | 2018-05-10T12:37:39+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP407G | CPSM Final Test | PASS | 129 | 159.8 | 0.56 | -0.38 | 56.22 | 50.97 | 1.21 | 0.13 | 1 | 1 | 0 | -0.8441 | 0.5304 | 5.8 | -4 |  |  | 0.1153 | 0.1742 | 0.0151 | 1 | [21 empty] | 0 | [1002 empty] | 0
P | 2018-05-10T07:11:18+08:00 | 2018-05-10T13:10:14+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP409G | CPSM Final Test | FAIL
P | 2018-05-10T10:44:24+08:00 | 2018-05-10T13:34:05+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP408G | CPSM Final Test | PASS | 129 | 163.04 | 0.17 | -0.69 | 55.99 | 51.56 | 1.51 | 0.13 | -1 | 3 | 0 | -0.9157 | 0.511 | 3.9 | 3 |  |  | 0.4142 | 0.2436 | -0.0277 | 1 | [21 empty] | 0 | [1003 empty] | 0
P | 2018-05-10T10:08:27+08:00 | 2018-05-10T13:50:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP406G | CPSM Final Test | PASS | 129 | 165.56 | 0.46 | -0.69 | 56.05 | 48.68 | 1.06 | 0.13 | 0 | -3 | 0 | -0.8541 | 0.509 | 0.4 | 0 |  |  | 0.1374 | 0.5135 | -0.047 | 1 | [21 empty] | 0 | [1004 empty] | 0
P | 2018-05-10T08:13:38+08:00 | 2018-05-10T14:36:16+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP410G | CPSM Final Test | FAIL | [875 empty] | 0
P | 2018-05-10T10:10:46+08:00 | 2018-05-10T14:55:18+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP409G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | -0.31 | 56.51 | 48.34 | 1.21 | 0.13 | -1 | -1 | 0 | -0.8121 | 0.4892 | 1.6 | 1 |  |  | 0.0649 | 0.1035 | -0.0113 | 1 | [21 empty] | 0 | [1005 empty] | 0
P | 2018-05-10T11:24:57+08:00 | 2018-05-10T15:41:58+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP410G | CPSM Final Test | PASS | 129 | 163.13 | 0.27 | -0.92 | 56.19 | 49.61 | 1.29 | 0.13 | -1 | -3 | 0 | -0.912 | 0.4994 | 6.5 | 17 |  |  | 0.0549 | 0.2761 | -0.0209 | 1 | [21 empty] | 0.00698 | [832 empty] | 0
P | 2018-05-10T13:54:30+08:00 | 2018-05-10T16:44:33+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP411G | CPSM Final Test | PASS | 129 | 158.9 | 0.46 | -1.15 | 56.32 | 49.45 | 0.75 | 0.13 | -2 | 1 | 0 | -0.9138 | 0.4928 | 0.9 | 4 |  |  | 0.0882 | 0.3981 | -0.0286 | 0.99998 | [21 empty] | 0 | [594 empty] | 0
P | 2018-05-10T14:54:37+08:00 | 2018-05-10T17:53:02+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP412G | CPSM Final Test | PASS | 129 | 165.83 | 0.17 | -0.99 | 56.12 | 49.53 | 1.67 | 0.13 | 1 | 1 | 0 | -0.844 | 0.497 | 1.7 | 3 |  |  | 0.1303 | 0.312 | 0.0336 | 0.99996 | [21 empty] | 0 | [433 empty] | 0
P | 2018-05-10T17:22:18+08:00 | 2018-05-10T20:53:11+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP413G | CPSM Final Test | PASS | 129 | 164.57 | -0.02 | -0.38 | 56.23 | 48.18 | 0.91 | 0.13 | 1 | 1 | 0 | -0.9478 | 0.4898 | 1.6 | 6 |  |  | 0.114 | 0.1353 | -0.0192 | 1 | [21 empty] | 0.01397 | [41 empty] | 0
P | 2018-05-10T18:30:23+08:00 | 2018-05-10T21:16:41+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP415G | CPSM Final Test | PASS | 129 | 161.78 | -0.11 | -0.46 | 56.12 | 49.78 | 1.29 | 0.13 | -1 | -1 | 0 | -0.9199 | 0.5148 | 6.7 | 2 |  |  | 0.1234 | 0.2081 | -0.0041 | 0.99999 | [21 empty] | 0.01377 | [486 empty] | 0
P | 2018-05-10T02:57:47+08:00 | 2018-05-11T10:30:22+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP414G | CPSM Final Test | FAIL | [55 empty] | 0
P | 2018-05-10T18:44:04+08:00 | 2018-05-11T11:37:06+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP414G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.99 | 55.95 | 50.8 | 1.13 | 0.13 | 0 | 1 | 0 | -0.9338 | 0.5244 | 4.5 | 40 |  |  | 0.0542 | 0.2707 | 0.0279 | 1 | [21 empty] | 0 | [962 empty] | 0
P | 2018-05-11T09:46:48+08:00 | 2018-05-11T12:43:06+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP416G | CPSM Final Test | PASS | 129 | 163.13 | 0.17 | -0.31 | 56.35 | 51.05 | 1.44 | 0.13 | 1 | 2 | 0 | -0.856 | 0.4938 | 9.3 | 28 |  |  | 0.1143 | 0.1367 | 0.016 | 0.99999 | [21 empty] | 0 | [1006 empty] | 0
P | 2018-05-11T09:59:11+08:00 | 2018-05-11T12:50:43+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP417G | CPSM Final Test | PASS | 129 | 166.64 | 0.37 | -0.69 | 56.08 | 49.19 | 1.06 | 0.13 | 1 | -2 | 0 | -0.8221 | 0.4968 | 3.7 | -13 |  |  | 0.1069 | 0.3208 | -0.0272 | 0.99999 | [21 empty] | 0 | [129 empty] | 0
P | 2018-05-11T09:10:37+08:00 | 2018-05-11T13:16:52+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP418G | CPSM Final Test | FAIL
P | 2018-05-11T11:13:44+08:00 | 2018-05-11T14:24:08+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP418G | CPSM Final Test | PASS | 129 | 162.68 | -0.02 | -0.84 | 56.01 | 49.28 | 1.21 | 0.13 | 1 | 1 | 1 | -0.9217 | 0.5426 | 7.3 | 12 |  |  | 0.0532 | 0.1284 | -0.0819 | 0.99999 | [21 empty] | 0 | [832 empty] | 0
P | 2018-05-11T12:04:59+08:00 | 2018-05-11T15:19:42+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP419G | CPSM Final Test | PASS | 129 | 157.55 | 0.08 | -0.92 | 55.91 | 50.29 | 0.6 | 0.13 | 0 | 3 | -1 | -0.9078 | 0.55 | 0.4 | 52 |  |  | 0.0904 | 0.3732 | -0.0232 | 1 | [21 empty] | 0 | [1007 empty] | 0
P | 2018-05-11T12:06:50+08:00 | 2018-05-11T15:24:02+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP421G | CPSM Final Test | PASS | 129 | 161.6 | -0.49 | -0.54 | 56.42 | 54.36 | 1.67 | 0.13 | 0 | 2 | 0 | -0.8481 | 0.5208 | 14.6 | 15 |  |  | 0.122 | 0.1209 | -0.0051 | 0.99999 | [21 empty] | 0 | [756 empty] | 0
P | 2018-05-11T13:30:11+08:00 | 2018-05-11T16:19:01+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP420G | CPSM Final Test | PASS | 129 | 157.91 | 0.27 | -0.38 | 55.81 | 49.28 | 1.44 | 0.13 | -1 | 2 | 0 | -0.884 | 0.508 | 4.8 | -8 |  |  | 0.063 | 0.3359 | -0.0369 | 1 | [21 empty] | 0.0189 | [557 empty] | 0
P | 2018-05-11T13:37:10+08:00 | 2018-05-11T16:21:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP422G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.54 | 56.02 | 49.36 | 1.36 | 0.13 | 1 | 3 | 0 | -0.9379 | 0.5422 | 4.8 | -6 |  |  | 0.0878 | 0.1146 | 0.021 | 1 | [21 empty] | 0 | [118 empty] | 0
P | 2018-05-11T16:45:05+08:00 | 2018-05-11T17:16:16+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP426G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-05-11T14:39:03+08:00 | 2018-05-11T17:54:46+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP423G | CPSM Final Test | PASS | 129 | 160.43 | 0.27 | -0.46 | 56.01 | 48.6 | 1.13 | 0.13 | -1 | -1 | 0 | -0.87 | 0.4784 | -3.1 | 38 |  |  | 0.17 | 0.3007 | -0.0431 | 0.99999 | [21 empty] | 0 | [13 empty] | 0
P | 2018-05-11T18:44:07+08:00 | 2018-05-11T19:26:30+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP430G | CPSM Final Test | FAIL | [5 empty] | 47.16
P | 2018-05-11T17:45:16+08:00 | 2018-05-11T20:09:40+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP429G | CPSM Final Test | FAIL | [13 empty] | -9.9
P | 2018-05-11T16:51:55+08:00 | 2018-05-11T20:38:32+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP424G | CPSM Final Test | PASS | 129 | 166.37 | -0.11 | -1.22 | 56.17 | 49.11 | 1.13 | 0.13 | 0 | -2 | 0 | -0.8681 | 0.52 | 5.9 | 7 |  |  | 0.0946 | 0.3664 | 0.0077 | 0.99994 | [21 empty] | 0.02353 | [962 empty] | 0
P | 2018-05-11T18:31:15+08:00 | 2018-05-11T21:18:11+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP428G | CPSM Final Test | PASS | 129 | 163.22 | 0.08 | -0.99 | 56.01 | 48.6 | 1.21 | 0.13 | 1 | 2 | 0 | -0.892 | 0.5346 | 3.3 | -7 |  |  | 0.0751 | 0.1545 | -0.0204 | 1 | [21 empty] | 0 | [748 empty] | 0
P | 2018-05-11T18:25:18+08:00 | 2018-05-11T21:22:11+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP427G | CPSM Final Test | PASS | 129 | 152.87 | 0.17 | -0.61 | 56.1 | 49.36 | 1.36 | 0.13 | -1 | -2 | 0 | -0.872 | 0.4814 | 6.3 | -6 |  |  | 0.1977 | 0.3238 | -0.0228 | 0.99998 | [21 empty] | 0 | [1008 empty] | 0
P | 2018-05-11T17:29:01+08:00 | 2018-05-11T21:24:54+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP425G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -1.07 | 56.18 | 48.94 | 1.06 | 0.13 | -1 | 2 | 0 | -0.9197 | 0.511 | 0.2 | -1 |  |  | 0.2279 | 0.3313 | -0.0213 | 1 | [21 empty] | 0 | [832 empty] | 0
P | 2018-05-11T19:05:18+08:00 | 2018-05-11T21:57:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP430G | CPSM Final Test | PASS | 129 | 159.53 | -0.02 | -0.84 | 56.53 | 48.77 | 1.06 | 0.13 | -1 | -2 | 0 | -0.8719 | 0.5048 | -2.5 | -1 |  |  | 0.0645 | 0.1016 | 0.0238 | 0.99997 | [21 empty] | 0 | [1009 empty] | 0
P | 2018-05-11T18:57:28+08:00 | 2018-05-11T22:10:10+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP429G | CPSM Final Test | PASS | 129 | 165.38 | -0.11 | -0.61 | 56.29 | 50.8 | 1.44 | 0.13 | 1 | -4 | 0 | -0.9059 | 0.5454 | 9.9 | 45 |  |  | 0.1677 | 0.1328 | -0.0103 | 0.99997 | [21 empty] | 0 | [309 empty] | 0
P | 2018-05-12T08:51:32+08:00 | 2018-05-12T13:07:53+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP431G | CPSM Final Test | FAIL | [13 empty] | -7.4
P | 2018-05-12T09:40:52+08:00 | 2018-05-12T13:23:08+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP434G | CPSM Final Test | FAIL | [13 empty] | -8.2
P | 2018-05-12T11:08:15+08:00 | 2018-05-12T13:46:44+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP432G | CPSM Final Test | PASS | 129 | 157.37 | -0.02 | -0.46 | 56.44 | 49.02 | 1.29 | 0.13 | 0 | -1 | 0 | -0.9118 | 0.5304 | 2.8 | 2 |  |  | 0.0519 | 0.2514 | -0.01 | 0.99999 | [21 empty] | 0 | [1010 empty] | 0
P | 2018-05-12T07:52:06+08:00 | 2018-05-12T14:55:46+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP433G | CPSM Final Test | FAIL
P | 2018-05-12T11:23:56+08:00 | 2018-05-12T14:56:26+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP433G | CPSM Final Test | PASS | 129 | 156.56 | -0.4 | -0.92 | 56.21 | 48.85 | 1.21 | 0.13 | 0 | 2 | 0 | -0.888 | 0.4902 | 1.8 | 31 |  |  | 0.054 | 0.2461 | 0.0278 | 0.99999 | [21 empty] | 0 | [1011 empty] | 0
P | 2018-05-12T11:18:24+08:00 | 2018-05-12T14:59:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP436G | CPSM Final Test | FAIL
P | 2018-05-12T10:59:43+08:00 | 2018-05-12T15:08:21+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP431G | CPSM Final Test | PASS | 129 | 159.08 | 0.17 | -0.54 | 56.17 | 48.43 | 1.21 | 0.13 | 0 | -2 | 0 | -0.8261 | 0.4788 | 5.9 | -55 |  |  | 0.1059 | 0.3953 | -0.028 | 1 | [21 empty] | 0 | [100 empty] | 0
P | 2018-05-12T11:43:15+08:00 | 2018-05-12T15:12:10+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP435G | CPSM Final Test | PASS | 129 | 167.09 | -0.02 | -1.38 | 56.09 | 49.61 | 1.67 | 0.13 | -2 | -2 | 0 | -0.8002 | 0.527 | 1.1 | 1 |  |  | 0.1176 | 0.2351 | 0.0736 | 1 | [21 empty] | 0 | [1012 empty] | 0
P | 2018-05-12T11:32:00+08:00 | 2018-05-12T15:23:34+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP434G | CPSM Final Test | PASS | 129 | 161.87 | 0.17 | -0.77 | 56.34 | 50.63 | 1.51 | 0.13 | -1 | -1 | 0 | -0.834 | 0.4874 | 3.4 | -3 |  |  | 0.032 | 0.319 | -0.0364 | 1 | [21 empty] | 0 | [1013 empty] | 0
P | 2018-05-12T13:09:06+08:00 | 2018-05-12T16:44:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP436G | CPSM Final Test | PASS | 129 | 157.1 | -0.02 | -0.38 | 56.1 | 50.88 | 1.82 | 0.13 | 1 | 1 | 0 | -0.8978 | 0.5384 | 1.7 | 4 |  |  | 0.0543 | 0.0447 | -0.0276 | 0.99999 | [21 empty] | 0.01089 | [1014 empty] | 0
P | 2018-05-12T13:55:26+08:00 | 2018-05-12T16:45:46+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP437G | CPSM Final Test | PASS | 129 | 156.65 | -0.71 | -0.99 | 56.18 | 51.31 | 0.91 | 0.13 | -1 | 1 | 0 | -0.8141 | 0.485 | 5 | -49 |  |  | 0.0787 | 0.2532 | -0.0204 | 0.99998 | [21 empty] | 0 | [904 empty] | 0
P | 2018-05-14T10:43:02+08:00 | 2018-05-14T13:49:58+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP442G | CPSM Final Test | PASS | 129 | 158.72 | 0.08 | -0.92 | 56.17 | 51.65 | 1.59 | 0.13 | 1 | -2 | 0 | -0.8261 | 0.5098 | 0.3 | -13 |  |  | 0.3448 | 0.1423 | 0.0029 | 1 | [21 empty] | 0 | [238 empty] | 0
P | 2018-05-14T09:17:06+08:00 | 2018-05-14T14:06:33+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP440G | CPSM Final Test | FAIL
P | 2018-05-14T11:13:24+08:00 | 2018-05-14T14:07:03+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP438G | CPSM Final Test | PASS | 129 | 155.21 | -0.02 | -0.69 | 55.83 | 50.72 | 1.13 | 0.13 | 1 | -2 | 1 | -0.8421 | 0.5386 | 8 | -25 |  |  | 0.0517 | 0.084 | 0.0107 | 0.99999 | [21 empty] | 0 | [1015 empty] | 0
P | 2018-05-14T11:23:16+08:00 | 2018-05-14T14:10:56+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP439G | CPSM Final Test | PASS | 129 | 159.26 | -0.3 | 0.15 | 56.31 | 50.8 | 0.98 | 0.13 | 1 | -3 | 0 | -0.8261 | 0.548 | 1.4 | 6 |  |  | 0.1533 | 0.1712 | 0.0232 | 0.99999 | [21 empty] | 0 | [425 empty] | 0
P | 2018-05-14T11:41:49+08:00 | 2018-05-14T15:12:09+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP440G | CPSM Final Test | PASS | 129 | 161.15 | 0.08 | -0.54 | 56.18 | 49.53 | 1.29 | 0.13 | -1 | 1 | 0 | -0.8301 | 0.5006 | -3.6 | 0 |  |  | 0.0638 | 0.2918 | -0.03 | 0.99999 | [21 empty] | 0 | [1016 empty] | 0
P | 2018-05-14T12:15:08+08:00 | 2018-05-14T15:43:42+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP443G | CPSM Final Test | PASS | 129 | 162.5 | 0.08 | -0.84 | 56.1 | 50.63 | 1.36 | 0.13 | 1 | 3 | 0 | -0.7263 | 0.4934 | -2.9 | 3 |  |  | 0.123 | 0.0945 | 0.1062 | 0.99996 | [21 empty] | 0 | [326 empty] | 0
P | 2018-05-14T12:28:40+08:00 | 2018-05-14T15:46:23+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP444G | CPSM Final Test | PASS | 129 | 171.95 | -0.02 | -1.3 | 56.09 | 50.21 | 1.44 | 0.13 | 0 | -4 | 0 | -0.868 | 0.504 | 4.4 | -17 |  |  | 0.0677 | 0.4455 | 0.0051 | 1 | [21 empty] | 0 | [168 empty] | 0
P | 2018-05-14T13:52:19+08:00 | 2018-05-14T16:35:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP441G | CPSM Final Test | PASS | 129 | 156.92 | 0.17 | -0.92 | 55.99 | 48.26 | 1.67 | 0.13 | 0 | -4 | 1 | -0.9477 | 0.5352 | 7.6 | 44 |  |  | 0.0528 | 0.3619 | -0.0188 | 1 | [21 empty] | 0.00823 | [82 empty] | 0
P | 2018-05-14T16:28:21+08:00 | 2018-05-14T16:35:58+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP426G | CPSM Final Test | FAIL | [87 empty] | FALSE
P | 2018-05-14T16:00:19+08:00 | 2018-05-14T17:04:00+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP426G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-05-14T13:31:37+08:00 | 2018-05-14T19:32:42+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP426G | CPSM Final Test | FAIL
P | 2018-05-14T18:53:23+08:00 | 2018-05-14T21:38:02+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP446G | CPSM Final Test | PASS | 129 | 156.83 | 0.08 | 0.23 | 56.43 | 53.51 | 1.29 | 0.13 | 1 | 3 | 0 | -0.9198 | 0.5498 | 7.2 | -15 |  |  | 0.0922 | 0.1335 | -0.0225 | 1 | [21 empty] | 0 | [70 empty] | 0
P | 2018-05-14T18:35:13+08:00 | 2018-05-14T21:38:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP405G | CPSM Final Test | PASS |  | 164.66 | -0.02 | -0.92 | 55.94 | 51.31 | 1.51 | 0.13 | -1 | -3 | 0 | -0.9218 | 0.526 | 1.5 | -7 |  |  | 0.3064 | 0.2726 | -0.0167 | 0.99999 |  |  |  |  | 133 |  | 0 | [1032 empty] | 0
P | 2018-05-14T18:47:31+08:00 | 2018-05-14T21:43:19+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP447G | CPSM Final Test | PASS | 129 | 160.07 | -0.11 | -0.61 | 56.11 | 49.61 | 1.59 | 0.13 | -1 | -4 | 0 | -0.9178 | 0.5158 | 6.1 | 30 |  |  | 0.0415 | 0.221 | 0.0009 | 1 | [21 empty] | 0 | [1016 empty] | 0
P | 2018-05-14T18:58:57+08:00 | 2018-05-14T21:51:22+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP445G | CPSM Final Test | PASS | 129 | 154.67 | 0.08 | -0.92 | 56.18 | 48.01 | 1.36 | 0.13 | 0 | 2 | 0 | -0.8779 | 0.4878 | -2.4 | 46 |  |  | 0.0772 | 0.2563 | -0.0209 | 0.99999 | [21 empty] | 0 | [1018 empty] | 0
P | 2018-05-14T06:16:54+08:00 | 2018-05-15T08:17:49+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP406G | CPSM Final Test | FAIL
P | 2018-05-14T19:17:21+08:00 | 2018-05-15T08:18:27+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP406G | CPSM Final Test | PASS |  | 161.42 | 0.17 | -0.46 | 56.49 | 50.72 | 1.44 | 0.13 | 0 | -2 | 1 | -0.9079 | 0.4848 | 0.5 | 35 |  |  | 0.1135 | 0.1602 | -0.0074 | 0.99997 |  |  |  |  | 133 |  | 0 | [1034 empty] | 0
P | 2018-05-13T23:33:35+08:00 | 2018-05-15T09:30:44+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP426G | CPSM Final Test | FAIL | [716 empty] | 0
P | 2018-05-13T22:30:25+08:00 | 2018-05-15T10:33:54+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP426G | CPSM Final Test | FAIL
P | 2018-05-14T16:32:09+08:00 | 2018-05-15T10:46:18+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP426G | CPSM Final Test | PASS | 129 | 160.52 | 0.37 | -0.92 | 55.94 | 50.21 | 1.29 | 0.13 | 0 | -1 | 0 | -0.9518 | 0.5224 | 4.7 | 0 |  |  | 0.0718 | 0.1159 | 0.0272 | 1 | [21 empty] | 0 | [673 empty] | 0
P | 2018-05-15T10:01:53+08:00 | 2018-05-15T14:15:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | E18FNP410G | CPSM Final Test | FAIL | [13 empty] | -8.1
P | 2018-05-15T11:23:20+08:00 | 2018-05-15T14:35:11+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP411G | CPSM Final Test | PASS |  | 161.69 | -0.11 | -0.84 | 56.02 | 51.39 | 1.06 | 0.13 | -1 | -2 | 1 | -0.8959 | 0.5466 | 7.9 | -2 |  |  | 0.092 | 0.3538 | -0.0285 | 0.99999 |  |  |  |  | 133 |  | 0 | [559 empty] | 0
P | 2018-05-15T12:21:06+08:00 | 2018-05-15T16:03:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP412G | CPSM Final Test | PASS |  | 163.31 | 0.17 | -1.07 | 55.92 | 50.63 | 1.29 | 0.13 | -1 | -1 | 0 | -0.876 | 0.4842 | -4.4 | 18 |  |  | 0.3293 | 0.1738 | -0.0401 | 0.99999 |  |  |  |  | 133 |  | 0 | [1035 empty] | 0
P | 2018-05-15T13:20:33+08:00 | 2018-05-15T16:16:27+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP408G | CPSM Final Test | PASS |  | 158 | 0.17 | -1.3 | 56.01 | 48.68 | 0.75 | 0.13 | -1 | -2 | 0 | -0.8999 | 0.51 | 3.4 | 23 |  |  | 0.0889 | 0.18 | -0.0447 | 0.99999 |  |  |  |  | 133 |  | 0 | [1036 empty] | 0
P | 2018-05-15T12:08:46+08:00 | 2018-05-15T16:18:10+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | E18FNP410G | CPSM Final Test | FAIL |  | 164.75 | 0.37 | -0.92 | 55.85 | 49.02 | 0.83 | 0.13 | -1 | 1 | 0 | -0.8181 | 0.4812 | -1.3 | 11 | [10 empty] | 133
P | 2018-05-15T14:01:47+08:00 | 2018-05-15T16:53:06+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP407G | CPSM Final Test | PASS |  | 157.55 | -0.59 | -0.69 | 56.16 | 48.09 | 0.98 | 0.13 | -2 | 3 | 0 | -0.9478 | 0.5306 | -3.5 | 30 |  |  | 0.1266 | 0.2075 | -0.0118 | 1 |  |  |  |  | 133 |  | 0 | [349 empty] | 0
P | 2018-05-15T14:22:11+08:00 | 2018-05-15T17:14:27+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | E18FNP409G | CPSM Final Test | PASS |  | 161.42 | 0.17 | -0.69 | 56.19 | 49.36 | 0.83 | 0.13 | 0 | -5 | 1 | -0.898 | 0.4942 | 4.6 | 11 |  |  | 0.0748 | 0.2845 | -0.0058 | 1 |  |  |  |  | 133 |  | 0 | [1037 empty] | 0
P | 2018-05-15T15:19:08+08:00 | 2018-05-15T18:16:26+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | E18FNP414G | CPSM Final Test | PASS |  | 163.13 | -0.02 | 0.23 | 56.16 | 49.61 | 1.13 | 0.13 | 1 | 4 | 0 | -0.844 | 0.5348 | -3.7 | 8 |  |  | 0.0806 | 0.1077 | 0.0032 | 0.99997 |  |  |  |  | 133 |  | 0 | [688 empty] | 0
P | 2018-05-15T13:54:41+08:00 | 2018-05-15T18:34:12+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | E18FNP416G | CPSM Final Test | FAIL | [14 empty] | 638
P | 2018-05-15T16:08:49+08:00 | 2018-05-15T19:47:11+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP413G | CPSM Final Test | PASS |  | 161.51 | 0.08 | -0.38 | 56.5 | 49.28 | 1.51 | 0.13 | -1 | -2 | 1 | -0.8321 | 0.534 | 2.8 | -4 |  |  | 0.0808 | 0.5019 | 0.0122 | 1 |  |  |  |  | 133 |  | 0 | [1038 empty] | 0
P | 2018-05-15T18:47:55+08:00 | 2018-05-15T21:37:04+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP419G | CPSM Final Test | PASS |  | 161.69 | 0.46 | -0.23 | 56.62 | 49.87 | 1.51 | 0.13 | 1 | -2 | 0 | -0.9198 | 0.476 | 2 | 21 |  |  | 0.0862 | 0.6511 | 0.0372 | 0.99997 |  |  |  |  | 133 |  | 0 | [1039 empty] | 0
P | 2018-05-15T18:45:38+08:00 | 2018-05-15T21:44:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP417G | CPSM Final Test | PASS |  | 160.61 | 0.37 | -0.54 | 56.49 | 49.45 | 1.06 | 0.13 | -2 | -3 | 0 | -0.8979 | 0.4892 | -2.3 | 63 |  |  | 0.0575 | 0.2911 | -0.0187 | 0.99999 |  |  |  |  | 133 |  | 0.00772 | [1040 empty] | 0
P | 2018-05-15T18:51:53+08:00 | 2018-05-15T21:44:30+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | E18FNP418G | CPSM Final Test | PASS |  | 155.21 | 0.37 | -0.38 | 56.29 | 48.18 | 1.36 | 0.13 | 1 | -2 | 0 | -0.9238 | 0.476 | 4.4 | -7 |  |  | 0.0636 | 0.0892 | -0.0279 | 1 |  |  |  |  | 133 |  | 0 | [1041 empty] | 0
P | 2018-05-16T09:54:29+08:00 | 2018-05-16T10:37:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP410G | CPSM Final Test | FAIL | [5 empty] | 47.5
P | 2018-05-15T02:37:51+08:00 | 2018-05-16T10:51:35+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP415G | CPSM Final Test | FAIL | [14 empty] | 764
P | 2018-05-14T21:05:23+08:00 | 2018-05-16T11:23:30+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | E18FNP416G | CPSM Final Test | FAIL
P | 2018-05-16T08:16:49+08:00 | 2018-05-16T12:15:27+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP410G | CPSM Final Test | FAIL
P | 2018-05-15T16:14:27+08:00 | 2018-05-16T12:26:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | E18FNP416G | CPSM Final Test | PASS |  | 158.99 | -0.02 | -1.07 | 56.02 | 51.14 | 1.59 | 0.13 | 1 | -3 | 0 | -0.8061 | 0.488 | 4.1 | -4 |  |  | 0.1689 | 0.1477 | 0.0151 | 0.99996 |  |  |  |  | 133 |  | 0 | [1042 empty] | 0
P | 2018-05-15T18:44:43+08:00 | 2018-05-16T12:39:13+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP415G | CPSM Final Test | PASS |  | 165.56 | 0.27 | -0.99 | 55.73 | 49.45 | 1.29 | 0.13 | -1 | 3 | 0 | -0.9099 | 0.506 | 10.3 | 4 |  |  | 0.0876 | 0.0511 | -0.0622 | 1 |  |  |  |  | 133 |  | 0 | [1043 empty] | 0
P | 2018-05-16T04:33:47+08:00 | 2018-05-16T15:58:29+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP410G | CPSM Final Test | FAIL | [1072 empty] | 6390
P | 2018-05-16T03:42:29+08:00 | 2018-05-16T16:49:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP410G | CPSM Final Test | FAIL | [1072 empty] | 0
P | 2018-05-16T10:16:08+08:00 | 2018-05-16T17:56:17+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP410G | CPSM Final Test | PASS |  | 164.75 | 0.27 | -0.54 | 55.97 | 48.09 | 0.83 | 0.13 | -1 | 1 | 0 | -0.9199 | 0.4658 | -3.2 | 6 |  |  | 0.0248 | 0.1982 | -0.0114 | 1 |  |  |  |  | 133 |  | 0 | [1044 empty] | 0
P | 2018-05-17T14:54:04+08:00 | 2018-05-17T16:02:55+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP449G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-05-17T14:46:07+08:00 | 2018-05-17T17:39:49+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP450G | CPSM Final Test | FAIL
P | 2018-05-17T12:55:37+08:00 | 2018-05-17T17:42:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP448G | CPSM Final Test | FAIL
P | 2018-05-17T15:11:13+08:00 | 2018-05-17T18:36:07+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP452G | CPSM Final Test | FAIL | [14 empty] | 3551
P | 2018-05-17T15:28:30+08:00 | 2018-05-17T18:44:49+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP449G | CPSM Final Test | PASS | 129 | 169.43 | 0.27 | -0.61 | 56 | 50.29 | 0.91 | 0.13 | 1 | -1 | 1 | -0.9119 | 0.474 | 5.2 | -9 |  |  | 0.3008 | 0.4048 | 0.027 | 0.99994 | [21 empty] | 0 | [451 empty] | 0
P | 2018-05-17T15:19:07+08:00 | 2018-05-17T19:27:02+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP448G | CPSM Final Test | PASS | 129 | 165.56 | -0.21 | -0.84 | 56.03 | 52.07 | 0.83 | 0.13 | 1 | -2 | 0 | -0.9517 | 0.505 | 5.8 | 5 |  |  | 0.0824 | 0.182 | 0.0255 | 0.99998 | [21 empty] | 0 | [1030 empty] | 0
P | 2018-05-17T16:12:58+08:00 | 2018-05-17T19:40:07+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP450G | CPSM Final Test | PASS | 129 | 163.13 | 0.08 | -1.22 | 55.94 | 50.29 | 1.21 | 0.13 | -1 | -1 | 1 | -0.866 | 0.502 | 11 | 19 |  |  | 0.1168 | 0.3369 | 0.0337 | 0.99999 | [21 empty] | 0 | [1031 empty] | 0
P | 2018-05-17T16:53:40+08:00 | 2018-05-17T20:20:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP452G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.69 | 56.02 | 52.24 | 0.83 | 0.13 | 1 | -2 | 0 | -0.9159 | 0.471 | 3.9 | 4 |  |  | 0.0555 | 0.2961 | 0.0074 | 0.99998 | [21 empty] | 0 | [1029 empty] | 0
P | 2018-05-17T17:19:26+08:00 | 2018-05-17T20:41:51+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP453G | CPSM Final Test | PASS | 129 | 171.86 | -0.02 | -1.3 | 56.2 | 48.68 | 1.13 | 0.13 | 1 | -2 | 0 | -0.8281 | 0.53 | -4.1 | 1 |  |  | 0.0462 | 0.2449 | -0.0077 | 0.99999 | [21 empty] | 0 | [1032 empty] | 0
P | 2018-05-18T08:58:09+08:00 | 2018-05-18T12:06:30+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP456G | CPSM Final Test | PASS | 129 | 176.27 | -0.02 | -0.54 | 56.12 | 53.09 | 1.06 | 0.13 | 1 | -3 | 0 | -0.872 | 0.54 | 5.7 | 16 |  |  | 0.167 | 0.1134 | 0.0222 | 0.99999 | [21 empty] | 0.00868 | [19 empty] | 0
P | 2018-05-18T09:13:30+08:00 | 2018-05-18T12:09:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP454G | CPSM Final Test | PASS | 129 | 171.32 | 0.17 | -0.77 | 56.18 | 52.24 | 0.83 | 0.13 | 1 | -4 | 0 | -0.8659 | 0.505 | 3.7 | 47 |  |  | 0.2187 | 0.1823 | -0.0381 | 1 | [21 empty] | 0 | [1033 empty] | 0
P | 2018-05-18T09:08:30+08:00 | 2018-05-18T12:14:18+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP455G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -0.61 | 56.15 | 50.21 | 0.91 | 0.13 | 1 | -3 | 0 | -0.862 | 0.4864 | 12.4 | 11 |  |  | 0.0876 | 0.1392 | -0.004 | 0.99994 | [21 empty] | 0 | [1034 empty] | 0
P | 2018-05-18T11:41:55+08:00 | 2018-05-18T12:25:12+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP459G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-05-18T09:47:20+08:00 | 2018-05-18T12:45:31+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP457G | CPSM Final Test | PASS | 129 | 160.43 | 0.17 | -1.3 | 56.31 | 51.22 | 1.21 | 0.13 | -1 | -2 | 0 | -0.8439 | 0.496 | -2.5 | 7 |  |  | 0.0913 | 0.2877 | -0.0367 | 1 | [21 empty] | 0 | [1035 empty] | 0
P | 2018-05-18T09:59:21+08:00 | 2018-05-18T12:55:07+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP451G | CPSM Final Test | PASS | 129 | 166.19 | -0.02 | -0.77 | 56.22 | 49.87 | 1.06 | 0.13 | -1 | 1 | 0 | -0.8819 | 0.5306 | -1.2 | 59 |  |  | 0.1393 | 0.2614 | -0.0162 | 1 | [21 empty] | 0 | [1036 empty] | 0
P | 2018-05-18T12:07:05+08:00 | 2018-05-18T15:42:09+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP458G | CPSM Final Test | PASS | 129 | 160.79 | -0.02 | -0.77 | 56.17 | 48.68 | 0.83 | 0.13 | 1 | -2 | 0 | -0.89 | 0.509 | -0.3 | -5 |  |  | 0.3669 | 0.5415 | 0.0001 | 0.99999 | [21 empty] | 0 | [1037 empty] | 0
P | 2018-05-18T12:03:34+08:00 | 2018-05-18T15:56:54+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP459G | CPSM Final Test | PASS | 129 | 174.92 | -0.02 | -1.3 | 55.94 | 48.43 | 0.91 | 0.13 | 0 | 1 | -1 | -0.9378 | 0.5446 | 5.8 | -12 |  |  | 0.1244 | 0.1407 | -0.0335 | 0.99999 | [21 empty] | 0 | [1038 empty] | 0
P | 2018-05-18T13:11:51+08:00 | 2018-05-18T17:42:45+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP460G | CPSM Final Test | FAIL | [14 empty] | 1614
P | 2018-05-18T15:27:02+08:00 | 2018-05-18T18:22:41+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP461G | CPSM Final Test | PASS | 129 | 160.97 | -0.02 | -0.77 | 56.11 | 50.04 | 1.21 | 0.13 | 1 | 1 | 0 | -0.9119 | 0.512 | 2.9 | 3 |  |  | 0.0742 | 0.1118 | -0.0047 | 1 | [21 empty] | 0 | [360 empty] | 0
P | 2018-05-18T15:20:16+08:00 | 2018-05-18T18:29:00+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP462G | CPSM Final Test | PASS | 129 | 171.14 | -0.02 | -0.77 | 56.03 | 49.78 | 1.29 | 0.13 | 1 | 2 | 1 | -0.866 | 0.515 | -1.4 | -3 |  |  | 0.0912 | 0.1636 | 0.0665 | 0.99992 | [21 empty] | 0 | [1027 empty] | 0
P | 2018-05-18T15:27:18+08:00 | 2018-05-18T19:29:31+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP460G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -0.61 | 56.32 | 49.11 | 1.29 | 0.13 | 1 | 1 | 0 | -0.9339 | 0.4702 | -2.1 | 8 |  |  | 0.066 | 0.3787 | -0.0345 | 1 | [21 empty] | 0 | [1039 empty] | 0
P | 2018-05-18T17:33:23+08:00 | 2018-05-18T20:57:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP457G | CPSM Final Test | PASS | 129 | 160.43 | -0.11 | -0.38 | 56.34 | 51.99 | 1.13 | 0.13 | -2 | -2 | 0 | -0.9077 | 0.4698 | 3.8 | -16 |  |  | 0.1124 | 0.2717 | -0.0411 | 1 | [21 empty] | 0 | [1040 empty] | 0
P | 2018-05-18T17:26:29+08:00 | 2018-05-18T20:59:08+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP464G | CPSM Final Test | PASS | 129 | 174.2 | -0.02 | -1.22 | 56.11 | 51.31 | 1.06 | 0.13 | 1 | -2 | -1 | -0.8919 | 0.5322 | 7.6 | 31 |  |  | 0.0825 | 0.1303 | 0.0593 | 0.99998 | [21 empty] | 0.01987 | [1041 empty] | 0
P | 2018-05-17T23:40:42+08:00 | 2018-05-19T11:17:58+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP463G | CPSM Final Test | FAIL | [42 empty] | 0.84703
P | 2018-05-19T07:30:58+08:00 | 2018-05-19T11:21:17+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP466G | CPSM Final Test | FAIL
P | 2018-05-17T23:09:26+08:00 | 2018-05-19T11:49:14+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP463G | CPSM Final Test | FAIL
P | 2018-05-19T06:57:06+08:00 | 2018-05-19T11:55:09+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP466G | CPSM Final Test | FAIL
P | 2018-05-19T09:22:52+08:00 | 2018-05-19T12:24:39+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP465G | CPSM Final Test | PASS | 129 | 164.39 | 0.27 | -1.07 | 56.12 | 50.04 | 1.06 | 0.13 | 1 | -2 | 0 | -0.8021 | 0.482 | 0 | 95 |  |  | 0.0643 | 0.3272 | -0.001 | 0.99999 | [21 empty] | 0 | [1042 empty] | 0
P | 2018-05-18T17:29:20+08:00 | 2018-05-19T12:52:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP463G | CPSM Final Test | PASS | 129 | 163.67 | -0.21 | -0.31 | 55.89 | 48.94 | 1.06 | 0.13 | 1 | 3 | 0 | -0.8579 | 0.5098 | 7.9 | -1 |  |  | 0.077 | 0.1107 | 0.0199 | 0.99999 | [21 empty] | 0.03985 | [1043 empty] | 0
P | 2018-05-19T05:51:38+08:00 | 2018-05-19T13:00:37+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP466G | CPSM Final Test | FAIL
P | 2018-05-19T12:37:32+08:00 | 2018-05-19T13:54:03+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP470G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-05-19T12:40:46+08:00 | 2018-05-19T13:52:24+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP469G | CPSM Final Test | FAIL | [5 empty] | 45.97
P | 2018-05-19T12:18:34+08:00 | 2018-05-19T14:13:01+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP470G | CPSM Final Test | FAIL | [5 empty] | 45.89
P | 2018-05-19T09:26:08+08:00 | 2018-05-19T14:15:33+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP466G | CPSM Final Test | PASS | 129 | 172.22 | -0.02 | -0.46 | 56.23 | 48.77 | 1.44 | 0.13 | -1 | -1 | 0 | -0.854 | 0.548 | -4.3 | -1 |  |  | 0.0624 | 0.1313 | 0.0011 | 0.99999 | [21 empty] | 0 | [1044 empty] | 0
P | 2018-05-19T12:44:46+08:00 | 2018-05-19T16:06:51+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP472G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-05-19T12:49:02+08:00 | 2018-05-19T16:08:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP471G | CPSM Final Test | FAIL | [5 empty] | 47.67
P | 2018-05-19T13:06:47+08:00 | 2018-05-19T16:10:21+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP468G | CPSM Final Test | PASS | 129 | 159.17 | -0.02 | -0.77 | 56.02 | 51.39 | 0.6 | 0.13 | 1 | 1 | 0 | -0.9219 | 0.5478 | 1.8 | -9 |  |  | 0.1489 | 0.4238 | -0.0324 | 0.99993 | [21 empty] | 0 | [869 empty] | 0
P | 2018-05-19T13:08:34+08:00 | 2018-05-19T16:12:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP467G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.99 | 56.18 | 50.12 | 1.13 | 0.13 | 1 | -3 | -1 | -0.8281 | 0.5112 | 3.4 | 10 |  |  | 0.084 | 0.238 | 0.0599 | 0.99998 | [21 empty] | 0 | [1045 empty] | 0
P | 2018-05-19T12:32:08+08:00 | 2018-05-19T16:25:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP471G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-05-19T09:24:58+08:00 | 2018-05-19T17:08:12+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP469G | CPSM Final Test | FAIL | [14 empty] | 627
P | 2018-05-19T14:25:49+08:00 | 2018-05-19T18:53:17+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP472G | CPSM Final Test | PASS | 129 | 169.16 | -0.02 | -0.61 | 56.1 | 49.53 | 1.21 | 0.13 | -1 | -1 | 0 | -0.9258 | 0.4982 | -2.4 | 42 |  |  | 0.118 | 0.236 | -0.0059 | 1 | [21 empty] | 0 | [1046 empty] | 0
P | 2018-05-19T13:16:35+08:00 | 2018-05-19T18:53:20+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP469G | CPSM Final Test | PASS | 129 | 172.4 | -0.02 | -0.92 | 56.01 | 49.61 | 0.98 | 0.13 | -2 | -2 | -1 | -0.8959 | 0.516 | 3.1 | -7 |  |  | 0.1298 | 0.109 | 0.0024 | 1 | [21 empty] | 0 | [1047 empty] | 0
P | 2018-05-19T06:43:22+08:00 | 2018-05-19T19:48:13+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP470G | CPSM Final Test | FAIL | [1091 empty] | 0
P | 2018-05-19T14:28:53+08:00 | 2018-05-19T19:49:35+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP471G | CPSM Final Test | PASS | 129 | 165.29 | 0.17 | -0.31 | 56.39 | 49.11 | 1.21 | 0.13 | -2 | 4 | 0 | -0.9437 | 0.512 | 3.9 | 41 |  |  | 0.0499 | 0.1955 | -0.018 | 1 | [21 empty] | 0 | [151 empty] | 0
P | 2018-05-19T16:34:32+08:00 | 2018-05-19T19:54:16+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP474G | CPSM Final Test | PASS | 129 | 175.82 | -0.02 | -0.84 | 56.14 | 50.38 | 0.98 | 0.13 | 0 | 2 | 0 | -0.8779 | 0.4908 | -0.4 | 9 |  |  | 0.1716 | 0.2236 | -0.0083 | 0.99999 | [21 empty] | 0 | [1049 empty] | 0
P | 2018-05-19T16:38:59+08:00 | 2018-05-19T20:16:08+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP473G | CPSM Final Test | PASS | 129 | 157.1 | -0.02 | -0.54 | 56.24 | 48.68 | 0.83 | 0.13 | 1 | 1 | 0 | -0.9378 | 0.494 | 8.8 | 18 |  |  | 0.082 | 0.1119 | 0.0441 | 0.99999 | [21 empty] | 0.00679 | [1050 empty] | 0
P | 2018-05-19T13:15:48+08:00 | 2018-05-19T20:53:59+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP470G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -0.54 | 55.97 | 49.28 | 0.91 | 0.13 | 0 | 2 | 0 | -0.9458 | 0.502 | 5.3 | 24 |  |  | 0.05 | 0.3028 | -0.0387 | 0.99999 | [21 empty] | 0.00788 | [1048 empty] | 0
P | 2018-05-21T09:25:09+08:00 | 2018-05-21T10:25:49+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP477G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-05-21T09:55:29+08:00 | 2018-05-21T13:18:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP477G | CPSM Final Test | PASS | 129 | 164.93 | -0.02 | -0.84 | 56.34 | 49.87 | 1.36 | 0.13 | 0 | 2 | 0 | -0.8081 | 0.5048 | 8.9 | 7 |  |  | 0.0717 | 0.4454 | -0.0364 | 1 | [21 empty] | 0 | [1048 empty] | 0
P | 2018-05-21T09:55:41+08:00 | 2018-05-21T13:14:43+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP475G | CPSM Final Test | PASS | 129 | 166.37 | -0.02 | -0.99 | 55.92 | 50.55 | 0.91 | 0.13 | -1 | -3 | 0 | -0.7722 | 0.543 | 14.4 | 10 |  |  | 0.1072 | 0.2705 | 0.051 | 0.99997 | [21 empty] | 0.02718 | [1051 empty] | 0
P | 2018-05-21T09:54:51+08:00 | 2018-05-21T13:15:00+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP476G | CPSM Final Test | PASS | 129 | 174.83 | -0.3 | -0.54 | 56.12 | 50.29 | 0.98 | 0.13 | -2 | 1 | 0 | -0.854 | 0.502 | -2.8 | 84 |  |  | 0.0247 | 0.1402 | 0.0156 | 0.99997 | [21 empty] | 0 | [1052 empty] | 0
P | 2018-05-21T09:44:38+08:00 | 2018-05-21T13:14:41+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP478G | CPSM Final Test | PASS | 129 | 165.11 | -0.02 | -0.69 | 55.86 | 49.11 | 1.13 | 0.13 | 1 | 1 | 1 | -0.8181 | 0.539 | -1.5 | -2 |  |  | 0.0832 | 0.541 | -0.0045 | 1 | [21 empty] | 0.02599 | [1053 empty] | 0
P | 2018-05-21T11:49:37+08:00 | 2018-05-21T16:53:16+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP481G | CPSM Final Test | FAIL
P | 2018-05-21T13:44:37+08:00 | 2018-05-21T17:20:40+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP482G | CPSM Final Test | PASS | 129 | 170.78 | -0.02 | -0.31 | 56.52 | 49.53 | 1.36 | 0.13 | -1 | 3 | 0 | -0.8919 | 0.5124 | 5.3 | 47 |  |  | 0.1256 | 0.3526 | -0.0584 | 0.99998 | [21 empty] | 0 | [1054 empty] | 0
P | 2018-05-21T12:40:19+08:00 | 2018-05-21T17:23:15+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP479G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.38 | 56.48 | 51.48 | 1.59 | 0.13 | 1 | 1 | -1 | -0.9059 | 0.456 | 7.2 | 10 |  |  | 0.1324 | 0.174 | -0.0308 | 0.99998 | [21 empty] | 0.0275 | [260 empty] | 0
P | 2018-05-21T13:40:46+08:00 | 2018-05-21T17:25:20+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP480G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -0.61 | 56.44 | 52.75 | 0.83 | 0.13 | -1 | 1 | 0 | -0.868 | 0.5158 | -0.1 | 4 |  |  | 0.171 | 0.2337 | -0.0199 | 1 | [21 empty] | 0.01233 | [722 empty] | 0
P | 2018-05-21T14:45:55+08:00 | 2018-05-21T18:18:27+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP486G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.84 | 56.22 | 49.87 | 0.75 | 0.13 | 0 | 1 | 1 | -0.9119 | 0.547 | -3.9 | 9 |  |  | 0.1174 | 0.3135 | -0.0359 | 0.99999 | [21 empty] | 0.04179 | [1055 empty] | 0
P | 2018-05-21T14:50:26+08:00 | 2018-05-21T18:20:39+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP485G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -0.84 | 56.14 | 50.46 | 1.59 | 0.13 | 1 | -2 | 0 | -0.8899 | 0.544 | -2.9 | 4 |  |  | 0.1252 | 0.1714 | 0.0206 | 0.99995 | [21 empty] | 0 | [1056 empty] | 0
P | 2018-05-21T09:14:50+08:00 | 2018-05-21T18:26:15+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP483G | CPSM Final Test | FAIL | [1100 empty] | 0
P | 2018-05-21T14:21:27+08:00 | 2018-05-21T18:37:34+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP481G | CPSM Final Test | PASS | 129 | 175.28 | -0.3 | -1.3 | 56 | 48.01 | 0.91 | 0.13 | -1 | 2 | 0 | -0.906 | 0.507 | -0.8 | 13 |  |  | 0.025 | 0.1458 | -0.0206 | 0.99999 | [21 empty] | 0 | [1058 empty] | 0
P | 2018-05-21T13:50:33+08:00 | 2018-05-21T19:33:31+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP483G | CPSM Final Test | PASS | 129 | 164.57 | -0.02 | -1.3 | 55.83 | 50.88 | 0.68 | 0.13 | 0 | 1 | 1 | -0.8579 | 0.4892 | 8.3 | 65 |  |  | 0.0689 | 0.4774 | -0.0212 | 0.99998 | [21 empty] | 0 | [1057 empty] | 0
P | 2018-05-21T18:12:38+08:00 | 2018-05-21T20:10:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP489G | CPSM Final Test | FAIL | [201 empty] | FALSE
P | 2018-05-21T18:48:58+08:00 | 2018-05-21T21:46:05+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP490G | CPSM Final Test | PASS | 129 | 171.32 | -0.02 | -0.38 | 56.3 | 50.21 | 1.29 | 0.13 | 1 | 3 | 0 | -0.9319 | 0.5042 | 2.2 | 10 |  |  | 0.2153 | 0.1815 | -0.0108 | 1 | [21 empty] | 0.05251 | [1059 empty] | 0
P | 2018-05-21T18:39:03+08:00 | 2018-05-21T21:51:25+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP488G | CPSM Final Test | PASS | 129 | 161.06 | -0.02 | -0.46 | 56.22 | 48.94 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9119 | 0.4994 | 12.3 | -11 |  |  | 0.0225 | 0.1179 | -0.0264 | 0.99999 | [21 empty] | 0 | [1060 empty] | 0
P | 2018-05-21T04:00:55+08:00 | 2018-05-22T09:10:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP487G | CPSM Final Test | FAIL
P | 2018-05-21T18:35:46+08:00 | 2018-05-22T10:58:06+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP487G | CPSM Final Test | PASS | 129 | 178.16 | -0.02 | -0.54 | 56.13 | 48.94 | 1.21 | 0.13 | 0 | 1 | 1 | -0.9318 | 0.5348 | 4.3 | 65 |  |  | 0.1469 | 0.4095 | -0.0271 | 0.99996 | [21 empty] | 0 | [1061 empty] | 0
P | 2018-05-22T06:53:51+08:00 | 2018-05-22T12:09:51+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP493G | CPSM Final Test | FAIL
P | 2018-05-22T07:11:32+08:00 | 2018-05-22T12:10:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP494G | CPSM Final Test | FAIL | [1105 empty] | 0
P | 2018-05-21T02:00:34+08:00 | 2018-05-22T12:22:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP489G | CPSM Final Test | FAIL
P | 2018-05-22T09:35:03+08:00 | 2018-05-22T12:32:42+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015085 | E18FTP492G | CPSM Final Test | PASS | 129 | 166.01 | 0.17 | -0.92 | 56.17 | 50.63 | 0.75 | 0.13 | -2 | 1 | 0 | -0.892 | 0.5 | -0.1 | 14 |  |  | 0.1008 | 0.2209 | -0.0148 | 1 | [21 empty] | 0 | [295 empty] | 0
P | 2018-05-22T09:43:54+08:00 | 2018-05-22T12:37:57+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015085 | E18FTP491G | CPSM Final Test | PASS | 129 | 169.88 | -0.02 | -1.07 | 55.9 | 49.45 | 1.13 | 0.13 | -1 | 2 | 0 | -0.7922 | 0.4896 | 1.1 | 36 |  |  | 0.0978 | 0.2712 | 0.008 | 1 | [21 empty] | 0 | [727 empty] | 0
P | 2018-05-22T09:31:51+08:00 | 2018-05-22T13:15:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP493G | CPSM Final Test | PASS | 129 | 169.07 | 0.27 | -0.54 | 56.17 | 49.95 | 0.98 | 0.13 | 1 | -3 | 0 | -0.8638 | 0.4972 | 4.3 | 20 |  |  | 0.0611 | 0.3123 | 0.0265 | 0.99999 | [21 empty] | 0 | [769 empty] | 0
P | 2018-05-22T09:41:14+08:00 | 2018-05-22T13:16:21+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP494G | CPSM Final Test | PASS | 129 | 167.9 | -0.02 | -0.31 | 56.62 | 49.87 | 1.13 | 0.13 | 1 | 2 | 0 | -0.7563 | 0.492 | 14.5 | 27 |  |  | 0.0577 | 0.1757 | 0.0564 | 0.99995 | [21 empty] | 0.05062 | [1062 empty] | 0
P | 2018-05-21T19:11:32+08:00 | 2018-05-22T14:08:51+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP489G | CPSM Final Test | PASS | 129 | 168.17 | -0.02 | -0.84 | 56.29 | 51.73 | 0.68 | 0.13 | 1 | -2 | 0 | -0.9138 | 0.5096 | -2.7 | 4 |  |  | 0.0364 | 0.0473 | 0.023 | 0.99998 | [21 empty] | 0.00672 | [1063 empty] | 0
P | 2018-05-20T11:40:01+08:00 | 2018-05-22T17:04:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP484G | CPSM Final Test | FAIL | [19 empty] | -1.1448
P | 2018-05-22T11:01:59+08:00 | 2018-05-22T16:56:33+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP422G | CPSM Final Test | FAIL | [1105 empty] | 34630
P | 2018-05-22T13:46:21+08:00 | 2018-05-22T17:33:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | E18FNP420G | CPSM Final Test | PASS |  | 164.3 | -0.02 | -0.69 | 56.19 | 50.63 | 0.83 | 0.13 | -1 | -3 | -1 | -0.862 | 0.5402 | 8.3 | 10 |  |  | 0.0575 | 0.4024 | -0.0639 | 1 |  |  |  |  | 133 |  | 0.02126 | [784 empty] | 0
P | 2018-05-22T13:49:33+08:00 | 2018-05-22T17:36:22+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP421G | CPSM Final Test | PASS |  | 164.3 | -0.21 | -0.61 | 56.42 | 49.19 | 1.44 | 0.13 | 1 | 4 | 0 | -0.9079 | 0.502 | 3.7 | 4 |  |  | 0.0659 | 0.0789 | -0.0259 | 1 |  |  |  |  | 133 |  | 0 | [522 empty] | 0
P | 2018-05-22T13:54:28+08:00 | 2018-05-22T17:40:47+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP423G | CPSM Final Test | PASS |  | 163.85 | -0.02 | -0.54 | 56.48 | 49.45 | 0.68 | 0.13 | -2 | -3 | 1 | -0.9178 | 0.5012 | 6.5 | -14 |  |  | 0.0439 | 0.277 | -0.0063 | 0.99999 |  |  |  |  | 133 |  | 0 | [1079 empty] | 0
P | 2018-05-21T14:22:20+08:00 | 2018-05-22T18:07:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP484G | CPSM Final Test | PASS | 129 | 169.61 | -0.02 | -0.38 | 56.34 | 49.02 | 1.29 | 0.13 | 1 | 3 | 0 | -0.882 | 0.5326 | -5 | 2 |  |  | 0.0753 | 0.3445 | -0.0212 | 0.99999 | [21 empty] | 0 | [913 empty] | 0
P | 2018-05-22T15:19:34+08:00 | 2018-05-22T18:22:34+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | E18FNP425G | CPSM Final Test | PASS |  | 159.71 | -0.02 | -0.16 | 56.44 | 52.24 | 1.13 | 0.13 | 0 | -1 | 0 | -0.9098 | 0.463 | -0.7 | 35 |  |  | 0.033 | 0.1361 | -0.0344 | 1 |  |  |  |  | 133 |  | 0 | [1080 empty] | 0
P | 2018-05-22T08:54:25+08:00 | 2018-05-22T19:04:07+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP422G | CPSM Final Test | FAIL | [1105 empty] | 11521
P | 2018-05-22T11:03:23+08:00 | 2018-05-22T19:42:10+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | E18FNP424G | CPSM Final Test | FAIL
P | 2018-05-22T18:26:55+08:00 | 2018-05-22T19:53:39+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | E18FNP428G | CPSM Final Test | FAIL | [5 empty] | 47.07
P | 2018-05-22T13:59:16+08:00 | 2018-05-22T20:09:29+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP422G | CPSM Final Test | PASS |  | 167.27 | -0.02 | -0.54 | 56.12 | 49.78 | 0.98 | 0.13 | -1 | 1 | 1 | -0.838 | 0.4998 | 2.8 | 6 |  |  | 0.0236 | 0.2762 | 0.0378 | 0.99992 |  |  |  |  | 133 |  | 0 | [1077 empty] | 0
P | 2018-05-22T15:22:46+08:00 | 2018-05-22T20:45:06+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | E18FNP424G | CPSM Final Test | PASS |  | 169.07 | -0.4 | -0.38 | 56.71 | 54.02 | 1.29 | 0.13 | 0 | 1 | 0 | -0.8701 | 0.5494 | 8 | 50 |  |  | 0.0499 | 0.2667 | 0.0497 | 0.99998 |  |  |  |  | 133 |  | 0 | [1081 empty] | 0
P | 2018-05-22T17:08:29+08:00 | 2018-05-22T21:24:44+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | E18FNP432G | CPSM Final Test | FAIL | [14 empty] | 783
P | 2018-05-22T19:13:17+08:00 | 2018-05-22T22:28:50+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP430G | CPSM Final Test | PASS |  | 165.29 | -0.02 | -0.84 | 56.32 | 51.05 | 0.98 | 0.13 | -1 | 1 | 0 | -0.91 | 0.5402 | -1.3 | 14 |  |  | 0.0766 | 0.092 | -0.0467 | 0.99999 |  |  |  |  | 133 |  | 0 | [1082 empty] | 0
P | 2018-05-22T19:01:23+08:00 | 2018-05-22T22:33:49+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP426G | CPSM Final Test | PASS |  | 162.05 | -0.02 | -0.69 | 56.19 | 51.65 | 0.83 | 0.13 | 1 | -2 | 0 | -0.9259 | 0.533 | 0.7 | -1 |  |  | 0.1029 | 0.3311 | -0.034 | 0.99999 |  |  |  |  | 133 |  | 0 | [1083 empty] | 0
P | 2018-05-22T19:10:08+08:00 | 2018-05-22T22:38:51+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | E18FNP427G | CPSM Final Test | PASS |  | 159.08 | -0.02 | -0.54 | 56.11 | 48.94 | 1.21 | 0.13 | 0 | -2 | -1 | -0.9658 | 0.515 | 2.5 | 31 |  |  | 0.1824 | 0.3039 | -0.0188 | 0.99998 |  |  |  |  | 133 |  | 0 | [660 empty] | 0
P | 2018-05-22T19:16:37+08:00 | 2018-05-22T23:12:48+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | E18FNP432G | CPSM Final Test | PASS |  | 161.06 | -0.02 | -0.46 | 56.31 | 50.8 | 1.13 | 0.13 | 0 | -1 | 0 | -0.8999 | 0.5084 | 5.7 | 15 |  |  | 0.2187 | 0.852 | -0.0444 | 0.99997 |  |  |  |  | 133 |  | 0 | [1084 empty] | 0
P | 2018-05-23T10:37:38+08:00 | 2018-05-23T11:14:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | E18FNP429G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-05-23T09:34:10+08:00 | 2018-05-23T12:17:41+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | E18FNP429G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-05-23T10:33:18+08:00 | 2018-05-23T13:50:40+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP434G | CPSM Final Test | PASS |  | 157.82 | -0.02 | -0.61 | 56.43 | 53 | 1.36 | 0.13 | 1 | 1 | 0 | -0.8879 | 0.547 | 7.3 | 8 |  |  | 0.0638 | 0.1784 | -0.0013 | 1 |  |  |  |  | 133 |  | 0 | [1085 empty] | 0
P | 2018-05-23T11:27:46+08:00 | 2018-05-23T13:50:36+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP437G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-05-23T10:42:19+08:00 | 2018-05-23T13:59:39+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | E18FNP435G | CPSM Final Test | PASS |  | 166.19 | -0.02 | -0.77 | 56.42 | 55.8 | 1.36 | 0.13 | 1 | -2 | 0 | -0.8859 | 0.487 | -0.6 | 11 |  |  | 0.0589 | 0.3284 | -0.0371 | 0.99998 |  |  |  |  | 133 |  | 0 | [1086 empty] | 0
P | 2018-05-23T06:41:07+08:00 | 2018-05-23T14:18:02+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | E18FNP433G | CPSM Final Test | FAIL | [104 empty] | 0
P | 2018-05-23T10:18:42+08:00 | 2018-05-23T14:59:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP437G | CPSM Final Test | FAIL | [11 empty] | -0.5845
P | 2018-05-23T10:03:30+08:00 | 2018-05-23T15:14:52+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP437G | CPSM Final Test | FAIL | [11 empty] | -0.728
P | 2018-05-23T07:16:20+08:00 | 2018-05-23T15:17:14+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | E18FNP436G | CPSM Final Test | FAIL | [1109 empty] | 0
P | 2018-05-23T10:29:35+08:00 | 2018-05-23T15:23:30+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | E18FNP433G | CPSM Final Test | PASS |  | 170.87 | -0.02 | -0.61 | 56.23 | 48.6 | 0.98 | 0.13 | 1 | 3 | 0 | -0.876 | 0.5028 | 2.7 | 13 |  |  | 0.0683 | 0.0795 | -0.0125 | 1 |  |  |  |  | 133 |  | 0 | [76 empty] | 0
P | 2018-05-23T10:55:55+08:00 | 2018-05-23T15:42:17+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | E18FNP429G | CPSM Final Test | PASS |  | 165.83 | -0.02 | -0.69 | 56.4 | 49.36 | 0.83 | 0.13 | -1 | 2 | 0 | -0.8261 | 0.504 | 6.4 | 2 |  |  | 0.0684 | 0.1545 | -0.006 | 1 |  |  |  |  | 133 |  | 0.00747 | [1087 empty] | 0
P | 2018-05-23T05:30:16+08:00 | 2018-05-23T17:03:18+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | E18FNP436G | CPSM Final Test | FAIL
P | 2018-05-23T11:16:47+08:00 | 2018-05-23T17:09:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | E18FNP436G | CPSM Final Test | PASS |  | 163.13 | -0.02 | -0.92 | 55.9 | 49.53 | 0.98 | 0.13 | 0 | -2 | 1 | -0.878 | 0.5058 | 12.9 | 33 |  |  | 0.1288 | 0.1136 | -0.0059 | 1 |  |  |  |  | 133 |  | 0 | [1081 empty] | 0
P | 2018-05-21T18:57:15+08:00 | 2018-05-23T19:23:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | E18FNP428G | CPSM Final Test | FAIL | [27 empty] | 1.0505
P | 2018-05-23T00:28:37+08:00 | 2018-05-23T20:54:17+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP431G | CPSM Final Test | FAIL | [13 empty] | -6.4
P | 2018-05-23T17:31:54+08:00 | 2018-05-23T20:56:23+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP437G | CPSM Final Test | PASS |  | 164.93 | 0.17 | -0.77 | 55.84 | 50.55 | 1.29 | 0.13 | -2 | -4 | 1 | -0.9139 | 0.5018 | 2.7 | 14 |  |  | 0.1295 | 0.1239 | 0.0036 | 0.99991 |  |  |  |  | 133 |  | 0 | [1077 empty] | 0
P | 2018-05-21T16:45:21+08:00 | 2018-05-23T21:35:13+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | E18FNP428G | CPSM Final Test | FAIL | [17 empty] | 1.1859
P | 2018-05-22T19:10:17+08:00 | 2018-05-23T22:38:23+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | E18FNP428G | CPSM Final Test | PASS |  | 166.19 | 0.17 | -0.99 | 56.23 | 52.92 | 0.98 | 0.13 | 1 | 1 | 0 | -0.8838 | 0.542 | 2 | -38 |  |  | 0.2819 | 0.2015 | 0.0189 | 0.99998 |  |  |  |  | 133 |  | 0.01202 | [208 empty] | 0
P | 2018-05-23T10:41:27+08:00 | 2018-05-23T22:58:03+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP431G | CPSM Final Test | PASS |  | 167.54 | -0.02 | -0.61 | 56.16 | 52.49 | 1.06 | 0.13 | -1 | 2 | 0 | -0.8441 | 0.5046 | 3.8 | 25 |  |  | 0.0836 | 0.3255 | -0.0148 | 1 |  |  |  |  | 133 |  | 0 | [896 empty] | 0
P | 2018-05-24T09:23:42+08:00 | 2018-05-24T10:17:48+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP439G | CPSM Final Test | FAIL | [5 empty] | 47.41
P | 2018-05-24T07:57:28+08:00 | 2018-05-24T11:14:42+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP438G | CPSM Final Test | FAIL | [13 empty] | -5.6
P | 2018-05-24T09:53:34+08:00 | 2018-05-24T12:08:57+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP443G | CPSM Final Test | FAIL | [13 empty] | -7.8
P | 2018-05-24T09:50:45+08:00 | 2018-05-24T13:16:36+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP439G | CPSM Final Test | PASS |  | 162.77 | -0.02 | -0.69 | 56.31 | 49.78 | 1.06 | 0.13 | 1 | 1 | 0 | -0.8061 | 0.516 | 3.6 | 5 |  |  | 0.071 | 0.0634 | -0.0535 | 0.99999 |  |  |  |  | 133 |  | 0 | [76 empty] | 0
P | 2018-05-24T09:42:02+08:00 | 2018-05-24T13:17:09+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | E18FNP444G | CPSM Final Test | FAIL
P | 2018-05-24T09:36:05+08:00 | 2018-05-24T13:20:04+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP438G | CPSM Final Test | PASS |  | 158.9 | 0.08 | -0.61 | 56.2 | 50.46 | 0.91 | 0.13 | 1 | 3 | 0 | -0.8161 | 0.541 | 8.1 | 6 |  |  | 0.1652 | 0.3068 | -0.0036 | 0.99998 |  |  |  |  | 133 |  | 0 | [1081 empty] | 0
P | 2018-05-24T10:21:52+08:00 | 2018-05-24T13:20:27+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP440G | CPSM Final Test | PASS |  | 158.99 | -0.02 | -0.38 | 56.61 | 49.87 | 0.6 | 0.13 | 1 | 3 | -1 | -0.7842 | 0.5384 | 7.3 | 6 |  |  | 0.1154 | 0.3171 | -0.0677 | 1 |  |  |  |  | 133 |  | 0 | [1088 empty] | 0
P | 2018-05-24T10:07:44+08:00 | 2018-05-24T13:23:42+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | E18FNP441G | CPSM Final Test | PASS |  | 156.47 | -0.02 | -0.46 | 56.5 | 49.19 | 1.29 | 0.13 | -1 | 1 | 0 | -0.8001 | 0.4714 | 0.4 | 0 |  |  | 0.0766 | 0.063 | -0.0522 | 1 |  |  |  |  | 133 |  | 0 | [1089 empty] | 0
P | 2018-05-24T10:34:53+08:00 | 2018-05-24T13:28:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | E18FNP442G | CPSM Final Test | PASS |  | 162.5 | -0.02 | -0.69 | 56.5 | 48.94 | 0.98 | 0.13 | 1 | -2 | 0 | -0.8421 | 0.5374 | 7.3 | 37 |  |  | 0.0558 | 0.034 | -0.0062 | 0.99999 |  |  |  |  | 133 |  | 0 | [616 empty] | 0
P | 2018-05-24T11:01:16+08:00 | 2018-05-24T14:13:01+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP443G | CPSM Final Test | PASS |  | 161.15 | -0.4 | -0.38 | 56.5 | 50.46 | 1.06 | 0.13 | 0 | -2 | 0 | -0.8081 | 0.542 | 2.7 | 17 |  |  | 0.0814 | 0.3329 | -0.005 | 0.99999 |  |  |  |  | 133 |  | 0 | [1090 empty] | 0
P | 2018-05-24T13:27:20+08:00 | 2018-05-24T16:23:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | E18FNP445G | CPSM Final Test | PASS |  | 163.49 | -0.11 | -0.08 | 56.11 | 50.72 | 1.21 | 0.13 | 0 | 2 | 0 | -0.852 | 0.5004 | -2.2 | 25 |  |  | 0.085 | 0.2632 | 0.0335 | 0.99991 |  |  |  |  | 133 |  | 0 | [1091 empty] | 0
P | 2018-05-24T15:00:43+08:00 | 2018-05-24T16:56:05+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP449G | CPSM Final Test | FAIL | [5 empty] | 45.13
P | 2018-05-24T13:36:28+08:00 | 2018-05-24T17:33:18+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP447G | CPSM Final Test | FAIL | [14 empty] | 813
P | 2018-05-24T14:50:29+08:00 | 2018-05-24T17:55:35+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | E18FNP446G | CPSM Final Test | PASS |  | 173.21 | -0.02 | -1.07 | 56.23 | 50.04 | 0.75 | 0.15 | -1 | 1 | -1 | -0.85 | 0.528 | 6.8 | 21 |  |  | 0.0788 | 0.1736 | -0.0249 | 0.99997 |  |  |  |  | 133 |  | 0 | [768 empty] | 0
P | 2018-05-24T14:23:57+08:00 | 2018-05-24T18:19:00+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | E18FNP450G | CPSM Final Test | FAIL | [14 empty] | 883
P | 2018-05-24T15:20:38+08:00 | 2018-05-24T18:19:20+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP448G | CPSM Final Test | PASS |  | 168.98 | -0.3 | -0.46 | 56.45 | 49.19 | 1.21 | 0.13 | -1 | 1 | 0 | -0.872 | 0.519 | -0.8 | -6 |  |  | 0.0848 | 0.1866 | -0.0428 | 1 |  |  |  |  | 133 |  | 0 | [1092 empty] | 0
P | 2018-05-24T04:21:36+08:00 | 2018-05-24T18:37:35+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | E18FNP444G | CPSM Final Test | FAIL | [27 empty] | 1.1712
P | 2018-05-24T15:34:53+08:00 | 2018-05-24T19:21:30+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP447G | CPSM Final Test | PASS |  | 164.84 | 0.37 | -0.92 | 56.19 | 48.77 | 0.6 | 0.13 | 0 | 1 | 0 | -0.9457 | 0.5158 | 0.8 | -7 |  |  | 0.0507 | 0.1298 | -0.0255 | 1 |  |  |  |  | 133 |  | 0 | [1093 empty] | 0
P | 2018-05-24T15:58:24+08:00 | 2018-05-24T19:39:31+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP449G | CPSM Final Test | PASS |  | 164.57 | -0.02 | -0.92 | 55.92 | 50.12 | 0.91 | 0.13 | -2 | -1 | 1 | -0.8779 | 0.4716 | -3.1 | 7 |  |  | 0.0898 | 0.3608 | -0.053 | 0.99999 |  |  |  |  | 133 |  | 0 | [1094 empty] | 0
P | 2018-05-24T14:49:42+08:00 | 2018-05-24T19:39:52+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | E18FNP451G | CPSM Final Test | FAIL | [14 empty] | 662
P | 2018-05-24T11:29:35+08:00 | 2018-05-24T19:41:27+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015088 | E18FNP444G | CPSM Final Test | PASS |  | 164.21 | 0.17 | -1.45 | 55.91 | 50.63 | 0.98 | 0.13 | 0 | -2 | 1 | -0.846 | 0.4816 | -3.6 | 155 |  |  | 0.0517 | 0.2543 | -0.0298 | 1 |  |  |  |  | 133 |  | 0 | [1095 empty] | 0
P | 2018-05-24T16:21:29+08:00 | 2018-05-24T20:06:40+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015088 | E18FNP450G | CPSM Final Test | PASS |  | 159.08 | -0.02 | -0.99 | 56.33 | 48.18 | 0.91 | 0.13 | -1 | 1 | 0 | -0.9138 | 0.539 | -3.6 | 11 |  |  | 0.1939 | 0.2387 | 0.0101 | 0.99999 |  |  |  |  | 133 |  | 0.03027 | [1096 empty] | 0
P | 2018-05-24T17:14:47+08:00 | 2018-05-24T21:27:52+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015088 | E18FNP451G | CPSM Final Test | PASS |  | 167.18 | 0.08 | -1.07 | 56.32 | 50.21 | 0.6 | 0.13 | 1 | 2 | 0 | -0.9139 | 0.5352 | -1.9 | 1 |  |  | 0.0845 | 0.1463 | 0.0542 | 0.99992 |  |  |  |  | 133 |  | 0 | [579 empty] | 0
P | 2018-05-25T10:04:15+08:00 | 2018-05-25T12:10:56+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP455G | CPSM Final Test | FAIL | [5 empty] | 46.74
P | 2018-05-25T10:58:37+08:00 | 2018-05-25T12:38:23+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP456G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-05-25T09:52:25+08:00 | 2018-05-25T13:27:59+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015088 | E18FNP453G | CPSM Final Test | PASS |  | 167.63 | 0.27 | -0.99 | 56.24 | 48.68 | 0.91 | 0.13 | -1 | 2 | 0 | -0.816 | 0.539 | 1.1 | 39 |  |  | 0.038 | 0.0969 | -0.0132 | 0.99999 |  |  |  |  | 133 |  | 0 | [752 empty] | 0
P | 2018-05-25T10:48:37+08:00 | 2018-05-25T13:39:58+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015088 | E18FNP454G | CPSM Final Test | PASS |  | 161.78 | -0.4 | -0.99 | 56.13 | 50.46 | 1.06 | 0.13 | -1 | 1 | 0 | -0.7882 | 0.5496 | 4.8 | -9 |  |  | 0.0969 | 0.3357 | -0.0124 | 0.99998 |  |  |  |  | 133 |  | 0 | [1097 empty] | 0
P | 2018-05-25T10:35:50+08:00 | 2018-05-25T14:11:32+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP495G | CPSM Final Test | FAIL | [5 empty] | 45.55
P | 2018-05-25T11:07:36+08:00 | 2018-05-25T15:24:38+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015088 | E18FNP455G | CPSM Final Test | PASS |  | 161.51 | -0.02 | -0.77 | 56.36 | 48.43 | 1.29 | 0.13 | -1 | -2 | 0 | -0.8959 | 0.541 | 4.5 | -3 |  |  | 0.0926 | 0.1435 | -0.033 | 1 |  |  |  |  | 133 |  | 0 | [931 empty] | 0
P | 2018-05-23T21:13:20+08:00 | 2018-05-25T15:43:35+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP452G | CPSM Final Test | FAIL
P | 2018-05-25T11:48:30+08:00 | 2018-05-25T16:09:13+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015088 | E18FNP456G | CPSM Final Test | PASS |  | 162.77 | -0.02 | -0.77 | 56.42 | 48.51 | 1.59 | 0.13 | -1 | 1 | 0 | -0.8978 | 0.499 | -1.5 | 33 |  |  | 0.0585 | 0.1749 | -0.0544 | 0.99999 |  |  |  |  | 133 |  | 0 | [1098 empty] | 0
P | 2018-05-25T12:23:41+08:00 | 2018-05-25T16:55:52+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP495G | CPSM Final Test | PASS | 129 | 157.46 | -0.02 | -1.3 | 56.26 | 49.19 | 1.59 | 0.13 | -1 | 1 | -1 | -0.9778 | 0.549 | 2.5 | 18 |  |  | 0.0915 | 0.1033 | -0.0013 | 1 | [21 empty] | 0 | [57 empty] | 0
P | 2018-05-24T18:28:28+08:00 | 2018-05-25T17:31:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015088 | E18FNP452G | CPSM Final Test | PASS |  | 159.26 | -0.02 | -1.68 | 55.92 | 50.46 | 0.98 | 0.13 | 1 | 0 | 1 | -1.0017 | 0.4816 | 0.8 | 12 |  |  | 0.1483 | 0.348 | -0.0336 | 0.99989 |  |  |  |  | 133 |  | 0 | [1099 empty] | 0
P | 2018-05-25T18:25:09+08:00 | 2018-05-25T19:02:01+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP500G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-05-25T18:37:04+08:00 | 2018-05-25T19:14:44+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP501G | CPSM Final Test | FAIL | [5 empty] | 47.58
P | 2018-05-25T18:31:01+08:00 | 2018-05-25T21:39:52+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP498G | CPSM Final Test | PASS | 129 | 156.92 | -0.02 | -0.69 | 56.2 | 52.32 | 0.83 | 0.13 | 0 | -2 | 0 | -0.9598 | 0.4936 | -3 | 4 |  |  | 0.155 | 0.2046 | 0.004 | 0.99999 | [21 empty] | 0 | [1085 empty] | 0
P | 2018-05-25T18:22:11+08:00 | 2018-05-25T21:40:51+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP499G | CPSM Final Test | PASS | 129 | 167.36 | -0.02 | -0.69 | 56.36 | 50.97 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9737 | 0.52 | -2.3 | 12 |  |  | 0.0672 | 0.279 | -0.0357 | 1 | [21 empty] | 0.03718 | [260 empty] | 0
P | 2018-05-25T17:44:19+08:00 | 2018-05-25T21:54:46+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP497G | CPSM Final Test | PASS | 129 | 163.4 | -0.11 | -0.54 | 56.3 | 49.78 | 1.06 | 0.13 | 1 | -3 | 0 | -0.9559 | 0.5288 | 4.3 | -4 |  |  | 0.0777 | 0.624 | -0.0279 | 0.99999 | [21 empty] | 0 | [324 empty] | 0
P | 2018-05-25T18:55:54+08:00 | 2018-05-25T22:03:18+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP501G | CPSM Final Test | PASS | 129 | 165.47 | -0.02 | -0.31 | 56.17 | 48.68 | 0.91 | 0.13 | 1 | 1 | 0 | -0.8701 | 0.5486 | 1.2 | 0 |  |  | 0.1309 | 0.0869 | 0.0439 | 0.99997 | [21 empty] | 0 | [757 empty] | 0
P | 2018-05-25T03:45:01+08:00 | 2018-05-26T09:08:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP496G | CPSM Final Test | FAIL
P | 2018-05-25T03:13:11+08:00 | 2018-05-26T10:13:59+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP500G | CPSM Final Test | FAIL | [1129 empty] | 3155
P | 2018-05-25T18:26:44+08:00 | 2018-05-26T10:53:39+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP496G | CPSM Final Test | PASS | 129 | 160.16 | -0.02 | -0.46 | 56.5 | 48.43 | 1.21 | 0.13 | 1 | -2 | 1 | -0.896 | 0.548 | -0.2 | 7 |  |  | 0.1421 | 0.0726 | 0.0096 | 0.99997 | [21 empty] | 0 | [1087 empty] | 0
P | 2018-05-26T07:01:04+08:00 | 2018-05-26T11:14:32+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP502G | CPSM Final Test | FAIL
P | 2018-05-25T18:43:35+08:00 | 2018-05-26T11:19:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP500G | CPSM Final Test | PASS | 129 | 168.26 | -0.21 | -0.61 | 56.05 | 49.19 | 1.13 | 0.13 | -2 | 1 | 0 | -0.9458 | 0.5032 | 2.8 | 13 |  |  | 0.0869 | 0.3396 | -0.0019 | 1 | [21 empty] | 0 | [1086 empty] | 0
P | 2018-05-26T08:55:39+08:00 | 2018-05-26T11:56:31+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP495G | CPSM Final Test | PASS | 129 | 157.46 | -0.21 | -1.07 | 56.01 | 49.45 | 1.13 | 0.13 | -2 | -3 | -1 | -0.8202 | 0.532 | 3.5 | -16 |  |  | 0.0781 | 0.3078 | -0.0273 | 0.99999 | [21 empty] | 0 | [125 empty] | 0
P | 2018-05-26T09:07:48+08:00 | 2018-05-26T12:59:04+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP502G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.16 | 56.79 | 48.43 | 1.06 | 0.13 | -1 | 3 | 1 | -0.8401 | 0.5372 | 3.9 | -6 |  |  | 0.0274 | 0.1977 | -0.0027 | 0.99999 | [21 empty] | 0.008 | [465 empty] | 0
P | 2018-05-26T12:27:21+08:00 | 2018-05-26T13:22:49+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP503G | CPSM Final Test | FAIL | [5 empty] | 47.92
P | 2018-05-26T10:52:51+08:00 | 2018-05-26T14:57:19+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP503G | CPSM Final Test | FAIL
P | 2018-05-26T13:09:09+08:00 | 2018-05-26T16:00:08+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP504G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -0.92 | 56.32 | 49.28 | 0.75 | 0.13 | -1 | -3 | 0 | -0.8839 | 0.5142 | -3.6 | 4 |  |  | 0.0688 | 0.1335 | -0.0702 | 0.99999 | [21 empty] | 0 | [243 empty] | 0
P | 2018-05-26T13:28:06+08:00 | 2018-05-26T16:16:27+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP506G | CPSM Final Test | PASS | 129 | 165.47 | 0.08 | -1.22 | 56.1 | 49.87 | 0.98 | 0.13 | -1 | 1 | 1 | -0.9139 | 0.497 | -2.9 | 9 |  |  | 0.0825 | 0.2698 | -0.0445 | 0.99999 | [21 empty] | 0 | [1088 empty] | 0
P | 2018-05-26T13:30:04+08:00 | 2018-05-26T16:22:44+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP508G | CPSM Final Test | FAIL
P | 2018-05-26T12:55:05+08:00 | 2018-05-26T17:10:17+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP503G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -0.92 | 56.49 | 48.6 | 1.06 | 0.13 | 1 | 2 | 1 | -0.878 | 0.5088 | 5.6 | -30 |  |  | 0.0872 | 0.0861 | -0.0065 | 0.99995 | [21 empty] | 0 | [1086 empty] | 0
P | 2018-05-26T14:38:09+08:00 | 2018-05-26T17:47:09+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP507G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -0.84 | 56.33 | 50.29 | 1.29 | 0.13 | -1 | 1 | 0 | -0.9318 | 0.5096 | 3.3 | -11 |  |  | 0.0437 | 0.3185 | -0.0011 | 1 | [21 empty] | 0 | [1028 empty] | 0
P | 2018-05-26T14:56:24+08:00 | 2018-05-26T18:07:46+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP508G | CPSM Final Test | PASS | 129 | 162.5 | -0.02 | -0.46 | 56.48 | 48.26 | 1.06 | 0.13 | 1 | 2 | 0 | -0.9059 | 0.5442 | 2.5 | 3 |  |  | 0.0732 | 0.0858 | 0.0069 | 1 | [21 empty] | 0.02649 | [1089 empty] | 0
P | 2018-05-25T22:16:51+08:00 | 2018-05-27T08:32:08+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP509G | CPSM Final Test | FAIL
P | 2018-05-25T22:42:58+08:00 | 2018-05-27T08:33:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP510G | CPSM Final Test | FAIL
P | 2018-05-26T15:37:59+08:00 | 2018-05-27T08:33:33+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP510G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -0.61 | 56.33 | 48.94 | 1.36 | 0.13 | -1 | -3 | -1 | -0.9517 | 0.495 | -0.5 | 13 |  |  | 0.328 | 0.3328 | 0.0312 | 0.99995 | [21 empty] | 0 | [1090 empty] | 0
P | 2018-05-25T17:57:36+08:00 | 2018-05-27T09:32:24+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP505G | CPSM Final Test | FAIL | [19 empty] | -0.7582
P | 2018-05-26T15:24:29+08:00 | 2018-05-27T10:17:22+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP509G | CPSM Final Test | PASS | 129 | 165.2 | -1.1 | -0.61 | 56.71 | 54.02 | 1.29 | 0.13 | -1 | -2 | 0 | -0.9458 | 0.539 | 3.6 | 9 |  |  | 0.145 | 0.1954 | 0.0053 | 0.99964 | [21 empty] | 0 | [662 empty] | 0
P | 2018-05-26T13:45:00+08:00 | 2018-05-27T11:03:42+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP505G | CPSM Final Test | PASS | 129 | 159.89 | -0.3 | -0.61 | 56.41 | 54.27 | 1.44 | 0.13 | -1 | 1 | 0 | -0.8202 | 0.5192 | -2.3 | -2 |  |  | 0.2046 | 0.2173 | -0.0551 | 0.99991 | [21 empty] | 0 | [1091 empty] | 0
P | 2018-05-27T08:45:15+08:00 | 2018-05-27T11:47:26+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP512G | CPSM Final Test | PASS | 129 | 160.07 | -0.02 | -0.92 | 56.23 | 50.72 | 1.21 | 0.13 | 1 | 1 | -1 | -0.8659 | 0.528 | 4.2 | -11 |  |  | 0.1344 | 0.188 | 0.0141 | 0.99999 | [21 empty] | 0 | [1092 empty] | 0
P | 2018-05-27T04:38:24+08:00 | 2018-05-27T13:18:03+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP511G | CPSM Final Test | FAIL
P | 2018-05-27T10:06:42+08:00 | 2018-05-27T13:46:08+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP515G | CPSM Final Test | FAIL
P | 2018-05-27T10:27:48+08:00 | 2018-05-27T14:31:12+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP514G | CPSM Final Test | FAIL | [13 empty] | -5.5
P | 2018-05-27T09:46:02+08:00 | 2018-05-27T14:43:29+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP516G | CPSM Final Test | FAIL
P | 2018-05-27T07:04:06+08:00 | 2018-05-27T15:13:47+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP513G | CPSM Final Test | FAIL | [1136 empty] | 0
P | 2018-05-27T08:36:28+08:00 | 2018-05-27T15:16:22+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP515G | CPSM Final Test | FAIL | [78 empty] | 0
P | 2018-05-27T09:00:26+08:00 | 2018-05-27T15:29:05+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP516G | CPSM Final Test | FAIL
P | 2018-05-27T11:08:56+08:00 | 2018-05-27T16:19:21+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP513G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.92 | 56.12 | 49.95 | 0.75 | 0.13 | 1 | 1 | 1 | -0.9218 | 0.5128 | 2 | 12 |  |  | 0.0435 | 0.3134 | -0.0544 | 0.99999 | [21 empty] | 0 | [1093 empty] | 0
P | 2018-05-27T11:56:25+08:00 | 2018-05-27T16:21:52+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP515G | CPSM Final Test | PASS | 129 | 162.05 | 0.46 | -1.15 | 55.93 | 50.04 | 0.75 | 0.13 | 1 | -4 | 0 | -0.87 | 0.495 | -3.1 | -11 |  |  | 0.1083 | 0.2231 | -0.0434 | 0.99998 | [21 empty] | 0 | [35 empty] | 0
P | 2018-05-27T13:39:25+08:00 | 2018-05-27T16:30:41+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP517G | CPSM Final Test | PASS | 129 | 160.25 | 0.08 | -1.15 | 56.17 | 50.88 | 0.75 | 0.13 | 1 | 1 | -1 | -0.8341 | 0.5308 | -3.8 | 3 |  |  | 0.0871 | 0.301 | -0.0418 | 0.99999 | [21 empty] | 0 | [1094 empty] | 0
P | 2018-05-27T12:29:30+08:00 | 2018-05-27T16:33:56+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP514G | CPSM Final Test | PASS | 129 | 165.47 | -0.02 | -0.77 | 56.11 | 49.61 | 0.68 | 0.13 | 1 | 3 | 0 | -0.9318 | 0.4682 | 1.5 | 0 |  |  | 0.2021 | 0.1442 | -0.0367 | 0.99997 | [21 empty] | 0 | [1095 empty] | 0
P | 2018-05-27T01:14:12+08:00 | 2018-05-27T16:42:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP511G | CPSM Final Test | FAIL
P | 2018-05-27T13:59:09+08:00 | 2018-05-27T16:47:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP518G | CPSM Final Test | PASS | 129 | 162.32 | -0.11 | -0.99 | 56.12 | 50.12 | 0.68 | 0.13 | 1 | 1 | -1 | -0.8081 | 0.5122 | 4.7 | 29 |  |  | 0.0832 | 0.2329 | 0.0398 | 0.99994 | [21 empty] | 0 | [1096 empty] | 0
P | 2018-05-27T16:06:51+08:00 | 2018-05-27T17:32:57+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP522G | CPSM Final Test | FAIL | [5 empty] | 45.72
P | 2018-05-27T15:19:50+08:00 | 2018-05-27T18:04:46+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP519G | CPSM Final Test | PASS | 129 | 162.77 | -0.02 | -0.99 | 56.03 | 51.05 | 1.13 | 0.13 | 1 | -2 | 0 | -0.9379 | 0.5368 | 0.8 | 11 |  |  | 0.0707 | 0.231 | -0.0417 | 1 | [21 empty] | 0 | [858 empty] | 0
P | 2018-05-27T08:58:13+08:00 | 2018-05-27T18:27:53+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP511G | CPSM Final Test | PASS | 129 | 174.02 | -0.21 | -0.92 | 56.17 | 49.45 | 0.91 | 0.13 | -2 | 3 | 0 | -0.9477 | 0.524 | 1.6 | 2 |  |  | 0.0413 | 0.3505 | -0.0452 | 0.99999 | [21 empty] | 0 | [1097 empty] | 0
P | 2018-05-26T21:54:31+08:00 | 2018-05-28T09:38:54+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP516G | CPSM Final Test | FAIL
P | 2018-05-26T22:59:17+08:00 | 2018-05-28T10:08:44+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP521G | CPSM Final Test | FAIL
P | 2018-05-27T15:46:43+08:00 | 2018-05-28T11:24:02+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP516G | CPSM Final Test | PASS | 129 | 159.8 | 0.08 | -0.99 | 55.96 | 50.29 | 0.68 | 0.13 | -2 | -3 | 0 | -0.9638 | 0.482 | 6 | 18 |  |  | 0.4547 | 0.846 | -0.0151 | 0.99999 | [21 empty] | 0.01185 | [1098 empty] | 0
P | 2018-05-26T21:24:03+08:00 | 2018-05-28T11:43:58+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP521G | CPSM Final Test | FAIL | [78 empty] | 0
P | 2018-05-27T16:49:54+08:00 | 2018-05-28T11:47:35+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP522G | CPSM Final Test | PASS | 129 | 167.72 | -0.02 | -1.07 | 55.95 | 49.78 | 0.45 | 0.13 | 0 | -2 | 0 | -0.8021 | 0.4734 | 6.7 | 7 |  |  | 0.1333 | 0.6176 | -0.0324 | 1 | [21 empty] | 0 | [1099 empty] | 0
P | 2018-05-27T17:28:40+08:00 | 2018-05-28T11:57:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP520G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.99 | 55.79 | 50.12 | 0.68 | 0.13 | -2 | -2 | -1 | -0.8681 | 0.5382 | 5.3 | -2 |  |  | 0.1214 | 0.3826 | -0.0254 | 0.99998 | [21 empty] | 0 | [1093 empty] | 0
P | 2018-05-27T16:34:00+08:00 | 2018-05-28T12:49:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP521G | CPSM Final Test | PASS | 129 | 165.02 | -0.02 | -0.54 | 56.32 | 48.43 | 0.83 | 0.13 | 1 | 2 | 0 | -0.9179 | 0.5346 | 4.8 | 20 |  |  | 0.2007 | 0.5335 | 0.0017 | 0.99998 | [21 empty] | 0 | [35 empty] | 0
P | 2018-05-28T09:08:56+08:00 | 2018-05-28T13:59:18+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP523G | CPSM Final Test | FAIL
P | 2018-05-28T10:38:47+08:00 | 2018-05-28T14:14:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP527G | CPSM Final Test | FAIL | [13 empty] | -9
P | 2018-05-28T13:07:47+08:00 | 2018-05-28T14:21:10+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP528G | CPSM Final Test | FAIL | [5 empty] | 46.99
P | 2018-05-28T11:46:49+08:00 | 2018-05-28T14:36:25+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP524G | CPSM Final Test | PASS | 129 | 156.38 | -0.11 | -0.84 | 55.9 | 50.88 | 1.13 | 0.13 | 0 | 3 | 0 | -0.8799 | 0.5278 | 5.5 | -1 |  |  | 0.3168 | 0.805 | -0.0394 | 0.99998 | [21 empty] | 0 | [992 empty] | 0
P | 2018-05-28T11:58:17+08:00 | 2018-05-28T15:23:38+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP525G | CPSM Final Test | PASS | 129 | 165.02 | 0.08 | -0.77 | 56.11 | 49.78 | 1.06 | 0.13 | 1 | -2 | -1 | -0.896 | 0.5144 | -2 | 18 |  |  | 0.135 | 0.1226 | -0.0354 | 1 | [21 empty] | 0 | [1100 empty] | 0
P | 2018-05-28T12:14:23+08:00 | 2018-05-28T15:36:31+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP526G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -1.22 | 56.17 | 48.26 | 0.83 | 0.13 | 0 | -1 | 0 | -0.84 | 0.481 | -2.8 | -3 |  |  | 0.064 | 0.2593 | 0.0519 | 0.99998 | [21 empty] | 0 | [628 empty] | 0
P | 2018-05-28T12:26:24+08:00 | 2018-05-28T16:14:42+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP527G | CPSM Final Test | PASS | 129 | 158.09 | -0.3 | -1.22 | 55.92 | 51.05 | 0.98 | 0.13 | -1 | -3 | 0 | -0.8241 | 0.46 | 9.3 | 3 |  |  | 0.1365 | 0.2936 | 0.064 | 0.99999 | [21 empty] | 0 | [1101 empty] | 0
P | 2018-05-28T11:11:43+08:00 | 2018-05-28T16:50:48+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP529G | CPSM Final Test | FAIL
P | 2018-05-28T13:44:28+08:00 | 2018-05-28T16:51:36+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP528G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -1.38 | 56.1 | 48.01 | 0.68 | 0.13 | 1 | -3 | 0 | -0.8839 | 0.514 | 4.8 | 11 |  |  | 0.1182 | 0.3048 | -0.0247 | 0.99998 | [21 empty] | 0 | [35 empty] | 0
P | 2018-05-28T06:10:26+08:00 | 2018-05-28T16:57:48+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP523G | CPSM Final Test | FAIL | [1145 empty] | 0
P | 2018-05-28T14:55:52+08:00 | 2018-05-28T17:55:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP531G | CPSM Final Test | PASS | 129 | 168.08 | -0.02 | -1.3 | 55.91 | 49.02 | 0.91 | 0.13 | 1 | 2 | 1 | -0.854 | 0.481 | 4.6 | 0 |  |  | 0.044 | 0.2949 | -0.033 | 0.99999 | [21 empty] | 0 | [1103 empty] | 0
P | 2018-05-28T11:34:07+08:00 | 2018-05-28T18:03:26+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP523G | CPSM Final Test | PASS | 129 | 169.88 | -0.02 | -0.99 | 56.04 | 50.72 | 1.13 | 0.13 | -2 | 1 | -1 | -0.902 | 0.539 | 5.9 | 1 |  |  | 0.0448 | 0.0794 | -0.0067 | 0.99998 | [21 empty] | 0.01145 | [1102 empty] | 0
P | 2018-05-28T14:01:16+08:00 | 2018-05-28T18:35:38+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP529G | CPSM Final Test | PASS | 129 | 163.13 | -0.02 | -0.92 | 56.18 | 49.7 | 0.91 | 0.13 | 1 | 1 | 0 | -0.904 | 0.4836 | 0.7 | -1 |  |  | 0.0564 | 0.0846 | -0.0451 | 0.99999 | [21 empty] | 0.0403 | [1104 empty] | 0
P | 2018-05-28T18:14:11+08:00 | 2018-05-28T18:46:04+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP535G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-05-28T16:26:38+08:00 | 2018-05-28T19:22:31+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015085 | E18FTP530G | CPSM Final Test | PASS | 129 | 161.87 | 0.17 | -0.84 | 56.24 | 49.61 | 1.13 | 0.13 | -2 | 1 | 0 | -0.9618 | 0.512 | 5.7 | 24 |  |  | 0.1235 | 0.1741 | 0.0058 | 1 | [21 empty] | 0 | [1105 empty] | 0
P | 2018-05-28T16:20:44+08:00 | 2018-05-28T19:21:16+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015085 | E18FTP532G | CPSM Final Test | PASS | 129 | 161.78 | 0.37 | -0.84 | 56.6 | 48.34 | 0.91 | 0.13 | -2 | 4 | 1 | -0.9239 | 0.5144 | 4.3 | -7 |  |  | 0.0903 | 0.1963 | -0.0242 | 0.99999 | [21 empty] | 0 | [896 empty] | 0
P | 2018-05-28T14:01:29+08:00 | 2018-05-28T19:44:13+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP534G | CPSM Final Test | FAIL
P | 2018-05-28T16:52:51+08:00 | 2018-05-28T19:50:51+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015085 | E18FTP534G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.77 | 56.42 | 54.61 | 1.44 | 0.13 | 1 | -2 | 1 | -0.9537 | 0.511 | 0.8 | 29 |  |  | 0.0578 | 0.3658 | 0.0007 | 1 | [21 empty] | 0 | [1106 empty] | 0
P | 2018-05-28T17:08:08+08:00 | 2018-05-28T20:21:34+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015085 | E18FTP533G | CPSM Final Test | PASS | 129 | 160.07 | -0.02 | -0.77 | 56.2 | 50.97 | 1.06 | 0.13 | -2 | 2 | 0 | -0.832 | 0.5066 | -2.2 | -9 |  |  | 0.1267 | 0.4118 | 0.0021 | 0.99997 | [21 empty] | 0.00796 | [906 empty] | 0
P | 2018-05-28T17:29:57+08:00 | 2018-05-28T20:46:24+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015085 | E18FTP536G | CPSM Final Test | PASS | 129 | 157.82 | 0.08 | -1.3 | 56.11 | 51.14 | 0.75 | 0.13 | -1 | 1 | 0 | -0.8661 | 0.544 | 1.8 | 3 |  |  | 0.2096 | 0.1049 | -0.0239 | 0.99996 | [21 empty] | 0.00909 | [1107 empty] | 0
P | 2018-05-28T18:30:08+08:00 | 2018-05-28T21:23:10+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015085 | E18FTP535G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -1.22 | 56.15 | 49.02 | 0.83 | 0.13 | -1 | -1 | 0 | -0.9278 | 0.4566 | 3.9 | -7 |  |  | 0.2328 | 0.2463 | 0.0072 | 0.99999 | [21 empty] | 0.02098 | [1102 empty] | 0
P | 2018-05-28T18:57:27+08:00 | 2018-05-28T21:38:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015085 | E18FTP537G | CPSM Final Test | PASS | 129 | 163.13 | -0.11 | -1.15 | 56.01 | 48.85 | 0.6 | 0.13 | -1 | 2 | 0 | -0.8699 | 0.5002 | -0.5 | -12 |  |  | 0.1116 | 0.3711 | -0.0432 | 0.99998 | [21 empty] | 0 | [949 empty] | 0
P | 2018-05-28T19:13:34+08:00 | 2018-05-28T21:55:12+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015085 | E18FTP538G | CPSM Final Test | PASS | 129 | 163.94 | -0.02 | 0.38 | 56.47 | 51.73 | 1.44 | 0.13 | -1 | -2 | -1 | -0.888 | 0.5358 | 0.5 | -2 |  |  | 0.0647 | 0.1634 | -0.0167 | 1 | [21 empty] | 0.01244 | [584 empty] | 0
P | 2018-05-30T16:27:03+08:00 | 2018-05-30T17:27:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP457G | CPSM Final Test | FAIL | [5 empty] | 46.82
P | 2018-05-30T14:40:45+08:00 | 2018-05-30T18:55:03+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP461G | CPSM Final Test | FAIL | [5 empty] | 45.8
P | 2018-05-30T13:54:26+08:00 | 2018-05-30T19:59:32+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | E18FNP459G | CPSM Final Test | FAIL
P | 2018-05-30T14:05:43+08:00 | 2018-05-30T19:57:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | E18FNP460G | CPSM Final Test | FAIL | [318 empty] | 1807954
P | 2018-05-30T13:14:45+08:00 | 2018-05-30T20:14:31+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP458G | CPSM Final Test | FAIL | [1151 empty] | 1824018
P | 2018-05-30T13:33:59+08:00 | 2018-05-30T20:20:39+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP457G | CPSM Final Test | FAIL | [1152 empty] | 1843153
P | 2018-05-30T12:55:15+08:00 | 2018-05-30T20:40:33+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP461G | CPSM Final Test | FAIL
P | 2018-05-30T17:01:22+08:00 | 2018-05-30T21:03:17+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015088 | E18FNP460G | CPSM Final Test | PASS |  | 164.03 | -0.02 | -0.46 | 56.46 | 49.28 | 1.21 | 0.13 | -1 | 2 | 1 | -0.9219 | 0.492 | -2 | 5 |  |  | 0.0883 | 0.4305 | -0.0762 | 1 |  |  |  |  | 133 |  | 0.01947 | [290 empty] | 0
P | 2018-05-30T16:44:38+08:00 | 2018-05-30T21:20:03+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015088 | E18FNP458G | CPSM Final Test | PASS |  | 167.72 | -0.02 | -0.92 | 56.04 | 50.55 | 0.91 | 0.13 | 1 | -2 | 0 | -0.7942 | 0.4928 | 6.6 | -10 |  |  | 0.0906 | 0.2204 | 0.0255 | 0.99994 |  |  |  |  | 133 |  | 0 | [1123 empty] | 0
P | 2018-05-30T16:57:19+08:00 | 2018-05-30T21:26:13+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP457G | CPSM Final Test | PASS |  | 165.38 | 0.17 | -0.77 | 56.32 | 51.14 | 0.75 | 0.13 | 1 | -1 | 0 | -0.9039 | 0.4932 | 3.2 | 11 |  |  | 0.057 | 0.1488 | -0.073 | 1 |  |  |  |  | 133 |  | 0 | [1124 empty] | 0
P | 2018-05-30T16:56:59+08:00 | 2018-05-30T21:48:52+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | E18FNP459G | CPSM Final Test | PASS |  | 159.17 | -0.49 | -0.61 | 56.32 | 50.29 | 0.68 | 0.13 | 1 | 1 | 0 | -0.886 | 0.5344 | 7.8 | 4 |  |  | 0.0574 | 0.3433 | 0.0119 | 0.99999 |  |  |  |  | 133 |  | 0 | [1125 empty] | 0
P | 2018-05-30T19:25:07+08:00 | 2018-05-30T22:15:59+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | E18FNP462G | CPSM Final Test | PASS |  | 165.11 | -0.02 | -0.77 | 56.21 | 48.94 | 1.29 | 0.13 | -1 | -4 | 0 | -0.9157 | 0.5 | 1.6 | -10 |  |  | 0.0504 | 0.0891 | -0.0325 | 1 |  |  |  |  | 133 |  | 0.02125 | [1126 empty] | 0
P | 2018-05-29T23:56:51+08:00 | 2018-05-31T09:38:57+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP461G | CPSM Final Test | FAIL
P | 2018-05-30T16:47:54+08:00 | 2018-05-31T11:26:23+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP461G | CPSM Final Test | PASS |  | 162.05 | 0.46 | -1.38 | 55.84 | 51.82 | 1.13 | 0.13 | 1 | -2 | 0 | -0.9158 | 0.49 | -0.3 | -26 |  |  | 0.1039 | 0.2299 | -0.0071 | 0.99999 |  |  |  |  | 133 |  | 0 | [223 empty] | 0
P | 2018-05-31T06:52:04+08:00 | 2018-05-31T13:20:58+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | E18FNP466G | CPSM Final Test | FAIL
P | 2018-05-31T10:18:13+08:00 | 2018-05-31T13:29:51+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015088 | E18FNP464G | CPSM Final Test | PASS |  | 161.51 | -0.02 | -0.61 | 56.18 | 50.29 | 1.51 | 0.13 | 1 | 2 | 0 | -0.8281 | 0.514 | 0.6 | 35 |  |  | 0.1236 | 0.1859 | -0.0278 | 0.99999 |  |  |  |  | 133 |  | 0 | [273 empty] | 0
P | 2018-05-31T10:14:59+08:00 | 2018-05-31T13:31:38+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015088 | E18FNP463G | CPSM Final Test | PASS |  | 165.83 | -0.02 | -0.54 | 56.33 | 49.36 | 1.06 | 0.13 | 1 | -2 | 0 | -0.8021 | 0.4876 | -2 | 16 |  |  | 0.1149 | 0.2673 | -0.0177 | 0.99996 |  |  |  |  | 133 |  | 0 | [933 empty] | 0
P | 2018-05-31T11:48:02+08:00 | 2018-05-31T13:35:57+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | E18FNP468G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-05-31T11:38:12+08:00 | 2018-05-31T13:45:47+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | E18FNP468G | CPSM Final Test | FAIL | [400 empty] | FALSE
P | 2018-05-31T12:18:26+08:00 | 2018-05-31T14:30:03+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP539G | CPSM Final Test | FAIL | [5 empty] | 47.75
P | 2018-05-31T04:42:12+08:00 | 2018-05-31T15:30:50+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | E18FNP466G | CPSM Final Test | FAIL | [27 empty] | 0.0869
P | 2018-05-31T10:59:06+08:00 | 2018-05-31T16:17:23+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP467G | CPSM Final Test | PASS |  | 160.97 | -0.02 | -1.15 | 56.11 | 49.95 | 1.06 | 0.13 | -1 | 1 | 0 | -0.8959 | 0.523 | -2.4 | 4 |  |  | 0.3709 | 0.5197 | 0.014 | 0.99999 |  |  |  |  | 133 |  | 0.00666 | [1124 empty] | 0
P | 2018-05-31T10:06:31+08:00 | 2018-05-31T16:33:50+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015088 | E18FNP466G | CPSM Final Test | PASS |  | 160.97 | -0.02 | -0.77 | 56.71 | 54.19 | 1.21 | 0.13 | 1 | -4 | 0 | -0.856 | 0.4974 | -1.8 | 1 |  |  | 0.1453 | 0.2362 | -0.0179 | 1 |  |  |  |  | 133 |  | 0 | [1127 empty] | 0
P | 2018-05-31T12:41:59+08:00 | 2018-05-31T17:25:57+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015088 | E18FNP468G | CPSM Final Test | PASS |  | 163.67 | -0.02 | -0.46 | 56.49 | 49.95 | 1.13 | 0.13 | 1 | 1 | 0 | -0.882 | 0.5026 | 0.5 | 7 |  |  | 0.0403 | 0.3195 | -0.0045 | 0.99997 |  |  |  |  | 133 |  | 0.0175 | [1128 empty] | 0
P | 2018-05-31T13:24:14+08:00 | 2018-05-31T17:23:47+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015085 | E18FTP539G | CPSM Final Test | PASS | 129 | 165.38 | 0.17 | -0.92 | 56.26 | 51.22 | 1.21 | 0.13 | 1 | 1 | 0 | -0.7862 | 0.49 | 5.8 | 30 |  |  | 0.1073 | 0.1694 | 0.0292 | 0.99997 | [21 empty] | 0 | [1108 empty] | 0
P | 2018-05-31T15:38:03+08:00 | 2018-05-31T18:36:57+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015085 | E18FTP540G | CPSM Final Test | PASS | 129 | 166.73 | -0.02 | -0.77 | 56.33 | 49.02 | 0.91 | 0.13 | 0 | -1 | 0 | -0.8878 | 0.4996 | 0.5 | -17 |  |  | 0.0511 | 0.1682 | -0.0448 | 1 | [21 empty] | 0.03326 | [690 empty] | 0
P | 2018-05-31T12:05:18+08:00 | 2018-05-31T19:24:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP542G | CPSM Final Test | FAIL | [55 empty] | 0
P | 2018-05-31T15:34:05+08:00 | 2018-05-31T19:25:38+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP541G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | -0.77 | 56.19 | 49.02 | 1.51 | 0.13 | 1 | -4 | 0 | -0.8639 | 0.495 | 6.3 | -9 |  |  | 0.24 | 0.2174 | -0.0597 | 0.99998 | [21 empty] | 0.02704 | [1114 empty] | 0
P | 2018-05-31T19:44:24+08:00 | 2018-05-31T20:26:57+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP465G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-05-31T15:57:04+08:00 | 2018-05-31T20:29:11+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015088 | E18FNP469G | CPSM Final Test | PASS |  | 165.92 | -0.02 | -1.07 | 56 | 50.55 | 0.83 | 0.13 | 1 | -2 | -1 | -0.9257 | 0.5304 | -2 | 26 |  |  | 0.0594 | 0.3576 | -0.0307 | 0.99999 |  |  |  |  | 133 |  | 0 | [557 empty] | 0
P | 2018-05-31T15:44:40+08:00 | 2018-05-31T20:29:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015085 | E18FTP542G | CPSM Final Test | PASS | 129 | 161.6 | -0.02 | -0.77 | 56.2 | 50.04 | 1.29 | 0.13 | 1 | 1 | 1 | -0.9019 | 0.5016 | 10.7 | -7 |  |  | 0.0674 | 0.0579 | -0.0108 | 1 | [21 empty] | 0 | [275 empty] | 0
P | 2018-05-31T19:47:22+08:00 | 2018-05-31T20:36:23+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP543G | CPSM Final Test | FAIL | [5 empty] | 47.84
P | 2018-05-31T04:38:24+08:00 | 2018-06-01T11:45:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP543G | CPSM Final Test | FAIL | [42 empty] | 1.5489
P | 2018-05-31T20:05:41+08:00 | 2018-06-01T11:44:35+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015088 | E18FNP465G | CPSM Final Test | PASS |  | 169.43 | 0.27 | -0.69 | 56.63 | 49.78 | 0.91 | 0.13 | 0 | 2 | -1 | -0.8879 | 0.545 | -1.2 | 23 |  |  | 0.22 | 0.1953 | -0.0171 | 0.99999 |  |  |  |  | 133 |  | 0.01284 | [1130 empty] | 0
P | 2018-05-31T02:57:02+08:00 | 2018-06-01T13:26:43+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP543G | CPSM Final Test | FAIL
P | 2018-05-31T20:11:52+08:00 | 2018-06-01T13:27:11+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015085 | E18FTP543G | CPSM Final Test | PASS | 129 | 166.82 | -0.02 | -1.07 | 56.23 | 48.34 | 1.29 | 0.13 | 1 | 1 | -1 | -0.91 | 0.5252 | 0.5 | 6 |  |  | 0.0545 | 0.4189 | -0.0302 | 0.99996 | [21 empty] | 0.00739 | [1113 empty] | 0
P | 2018-06-01T02:50:57+08:00 | 2018-06-01T16:58:26+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP547G | CPSM Final Test | FAIL
P | 2018-05-31T10:33:32+08:00 | 2018-06-02T09:23:22+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP546G | CPSM Final Test | FAIL
P | 2018-06-01T01:03:00+08:00 | 2018-06-02T13:58:00+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP555G | CPSM Final Test | FAIL | [452 empty] | 1614
P | 2018-05-31T23:06:20+08:00 | 2018-06-02T15:54:40+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015085 | E18FTP555G | CPSM Final Test | FAIL | [452 empty] | 0
P | 2018-05-30T05:47:58+08:00 | 2018-06-03T14:08:56+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015085 | E18FTP546G | CPSM Final Test | FAIL
P | 2018-01-03T07:53:27+08:00 | 2018-01-03T15:43:23+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI200G | CPSM Final Test | FAIL
P | 2018-01-03T11:48:25+08:00 | 2018-01-03T15:44:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI200G | CPSM Final Test | PASS | 129 | 159.71 | 0.08 | -0.54 | 56.47 | 54.27 | 1.13 | 0.13 | 0 | -11 | -2 | -0.8818 | 0.5132 | 1.6 | -15 |  | 0 | 0.1085 | 0.1047 | 0.01 | 1 | [144 empty] | 0
P | 2018-01-03T13:16:04+08:00 | 2018-01-03T16:03:19+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | M17SDI196G | CPSM Final Test | PASS | 129 | 167.9 | -0.3 | -0.46 | 56.78 | 54.69 | 1.36 | 0.13 | 0 | 4 | -1 | -0.862 | 0.4946 | 9.5 | -1 |  | 0 | 0.1799 | 0.1942 | 0.0198 | 0.99999 | [219 empty] | 0
P | 2018-01-03T15:32:37+08:00 | 2018-01-03T16:07:17+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI168G | CPSM Final Test | FAIL | [5 empty] | 40.3
P | 2018-01-03T13:23:08+08:00 | 2018-01-03T16:10:26+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | M17SDI198G | CPSM Final Test | PASS | 129 | 163.49 | 0.27 | -0.38 | 56.79 | 52.41 | 0.98 | 0.13 | 1 | 2 | -1 | -0.9079 | 0.5436 | 1.5 | 11 |  | 0 | 0.029 | 0.0419 | 0.0105 | 1 | [608 empty] | 0
P | 2018-01-03T15:49:57+08:00 | 2018-01-03T18:35:09+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI168G | CPSM Final Test | PASS | 129 | 169.97 | 0.27 | -0.46 | 56.99 | 55.12 | 1.21 | 0.13 | -1 | -4 | 0 | -0.8779 | 0.5302 | 6.1 | 3 |  | 0 | 0.093 | 0.0771 | 0.0099 | 0.99997 | [1138 empty] | 0
P | 2018-01-03T16:10:16+08:00 | 2018-01-03T18:48:01+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | M17SDI191G | CPSM Final Test | PASS | 129 | 166.37 | -0.02 | 0 | 57.19 | 52.32 | 0.91 | 0.16 | 1 | -2 | 1 | -0.9378 | 0.533 | 0.8 | 5 |  | 0 | 0.0857 | 0.1552 | 0.0118 | 0.99998 | [1139 empty] | 0
P | 2018-01-03T16:17:22+08:00 | 2018-01-03T19:12:46+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | M17SDI197G | CPSM Final Test | PASS | 129 | 164.66 | 0.17 | -0.77 | 56.54 | 52.75 | 0.98 | 0.13 | 0 | -4 | 1 | -0.9378 | 0.4814 | 5.8 | 27 |  | 0 | 0.0454 | 0.0503 | -0.0057 | 1 | [1140 empty] | 0
P | 2018-01-04T07:18:54+08:00 | 2018-01-04T13:17:19+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | M17SDI173G | CPSM Final Test | FAIL
P | 2018-01-04T10:18:06+08:00 | 2018-01-04T13:17:49+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | M17SDI173G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -0.46 | 56.72 | 58.93 | 1.21 | 0.13 | 0 | 4 | 1 | -0.8201 | 0.4748 | 1.4 | 10 |  | 0.02833 | 0.0891 | 0.0606 | 0.0245 | 0.99999 | [286 empty] | 0
P | 2018-01-04T07:48:34+08:00 | 2018-01-04T13:20:12+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI167G | CPSM Final Test | FAIL
P | 2018-01-04T10:34:23+08:00 | 2018-01-04T13:20:54+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI167G | CPSM Final Test | PASS | 129 | 166.01 | -0.02 | -0.38 | 56.47 | 55.03 | 1.44 | -0.2 | 0 | -2 | 0 | -0.8419 | 0.4988 | 3 | 7 |  | 0 | 0.1115 | 0.1585 | 0.026 | 0.99998 | [1141 empty] | 0
P | 2018-01-04T10:32:28+08:00 | 2018-01-04T13:24:54+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | M17SDI192G | CPSM Final Test | PASS | 129 | 166.64 | -0.02 | 0.3 | 57.3 | 52.75 | 1.67 | 0.13 | 0 | 4 | -1 | -0.9478 | 0.4918 | 2.3 | 8 |  | 0 | 0.0369 | 0.0716 | -0.029 | 0.99999 | [228 empty] | 0
P | 2018-01-04T13:39:25+08:00 | 2018-01-04T16:32:06+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI193G | CPSM Final Test | PASS | 129 | 165.65 | -0.11 | -0.38 | 56.63 | 53.93 | 1.06 | 0.13 | 0 | -12 | 1 | -0.7941 | 0.5084 | 0.8 | 8 |  | 0 | 0.0588 | 0.1903 | 0.001 | 0.99999 | [1142 empty] | 0
P | 2018-01-04T13:51:44+08:00 | 2018-01-04T16:35:59+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | M17SDI199G | CPSM Final Test | PASS | 129 | 164.57 | 0.17 | -0.61 | 56.51 | 53.17 | 1.44 | -0.22 | 1 | -10 | 0 | -0.7781 | 0.539 | 1.2 | 17 |  | 0.01477 | 0.0771 | 0.1585 | 0.0025 | 0.99999 | [1143 empty] | 0
P | 2018-01-04T14:01:39+08:00 | 2018-01-04T16:56:52+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | M17SDI195G | CPSM Final Test | PASS | 129 | 158.09 | -0.02 | 0 | 56.78 | 51.73 | 1.21 | 0.13 | 1 | 2 | 0 | -0.8021 | 0.528 | 0.2 | 8 |  | 0 | 0.0518 | 0.2974 | -0.0169 | 0.99999 | [1144 empty] | 0
P | 2018-01-04T13:52:22+08:00 | 2018-01-04T20:05:06+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI194G | CPSM Final Test | FAIL
P | 2018-01-04T17:31:31+08:00 | 2018-01-04T20:47:26+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | M17SDI190G | CPSM Final Test | PASS | 129 | 160.43 | -0.11 | 0 | 56.8 | 52.83 | 1.59 | 0.21 | 1 | 4 | 1 | -0.9657 | 0.5022 | 5.7 | 33 |  | 0 | 0.0665 | 0.0543 | -0.0306 | 1 | [320 empty] | 0
P | 2018-01-04T18:23:39+08:00 | 2018-01-04T21:15:44+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | M17SDI187G | CPSM Final Test | PASS | 129 | 165.11 | 0.08 | -0.69 | 56.71 | 54.69 | 1.13 | 0.13 | 0 | 1 | 0 | -0.9637 | 0.5078 | 3.6 | 5 |  | 0.03665 | 0.086 | 0.063 | 0.0165 | 0.99999 | [718 empty] | 0
P | 2018-01-04T16:58:44+08:00 | 2018-01-04T21:50:16+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI194G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -0.38 | 56.6 | 53.34 | 1.13 | 0.13 | 0 | 4 | 1 | -0.882 | 0.4828 | 0 | 6 |  | 0 | 0.1444 | 0.2186 | 0.0027 | 0.99998 | [1144 empty] | 0
P | 2018-01-05T08:41:43+08:00 | 2018-01-05T12:10:36+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM740G | CPSM Final Test | FAIL | [13 empty] | 16.8
P | 2018-01-05T09:49:05+08:00 | 2018-01-05T12:41:50+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI189G | CPSM Final Test | PASS | 129 | 162.41 | 0.08 | -0.16 | 56.92 | 55.46 | 0.98 | 0.13 | 0 | 2 | 0 | -0.864 | 0.4892 | 0.5 | 2 |  | 0 | 0.219 | 0.109 | -0.0045 | 0.99999 | [826 empty] | 0
P | 2018-01-05T09:57:48+08:00 | 2018-01-05T12:59:16+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM741G | CPSM Final Test | PASS |  | 168.98 | 0.37 | -0.69 | 56.83 | 52.15 | 1.29 | 0.13 | -1 | -2 | 1 | -0.7662 | 0.5448 | 2.5 | 16 |  |  | 0.2271 | 0.0902 | -0.0186 | 0.99999 |  |  |  |  | 133 |  | 0 | [691 empty] | 0
P | 2018-01-05T10:26:10+08:00 | 2018-01-05T14:24:34+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM740G | CPSM Final Test | PASS |  | 164.84 | 0.08 | 0.68 | 56.61 | 52.15 | 0.98 | 0.13 | 0 | 6 | 0 | -0.822 | 0.538 | 2.7 | 25 |  |  | 0.129 | 0.1081 | -0.0433 | 1 |  |  |  |  | 133 |  | 0 | [778 empty] | 0
P | 2018-01-05T13:38:14+08:00 | 2018-01-05T16:51:58+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM746G | CPSM Final Test | PASS |  | 167.36 | -0.4 | -0.77 | 56.54 | 52.24 | 0.75 | 0.13 | 0 | -2 | -2 | -0.828 | 0.5004 | 3.2 | 3 |  |  | 0.0887 | 0.0587 | -0.001 | 1 |  |  |  |  | 133 |  | 0 | [733 empty] | 0
P | 2018-01-05T14:03:14+08:00 | 2018-01-05T17:46:53+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM747G | CPSM Final Test | PASS |  | 160.79 | -0.11 | -0.84 | 56.32 | 51.99 | 0.91 | 0.13 | 0 | -3 | -1 | -0.8301 | 0.5192 | 14.5 | 23 |  |  | 0.2717 | 0.2115 | 0.0157 | 0.99999 |  |  |  |  | 133 |  | 0 | [667 empty] | 0
P | 2018-01-05T14:46:44+08:00 | 2018-01-05T17:49:56+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM748G | CPSM Final Test | PASS |  | 167.99 | -0.11 | -0.38 | 56.62 | 55.29 | 1.36 | 0.13 | 0 | -2 | 1 | -0.9079 | 0.457 | 3.8 | 2 |  |  | 0.1087 | 0.1251 | -0.02 | 0.99999 |  |  |  |  | 133 |  | 0 | [1138 empty] | 0
P | 2018-01-05T18:22:51+08:00 | 2018-01-05T22:00:45+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM749G | CPSM Final Test | PASS |  | 163.49 | -0.21 | 0 | 57.21 | 52.07 | 1.74 | 0.13 | 0 | -2 | 1 | -0.9498 | 0.5106 | 1 | 39 |  |  | 0.1016 | 0.1661 | 0.0208 | 0.99998 |  |  |  |  | 133 |  | 0 | [937 empty] | 0
P | 2018-01-05T19:06:44+08:00 | 2018-01-05T22:16:22+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM742G | CPSM Final Test | PASS |  | 163.49 | -0.43 | 1.06 | 56.83 | 55.88 | 2.58 | 0.13 | -1 | 1 | 1 | -0.9018 | 0.5056 | 2.9 | 5 |  |  | 0.0526 | 0.1744 | 0.015 | 0.99996 |  |  |  |  | 133 |  | 0 | [1139 empty] | 0
P | 2018-01-05T19:17:25+08:00 | 2018-01-05T22:26:57+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM750G | CPSM Final Test | PASS |  | 166.01 | -0.11 | -0.46 | 56.8 | 53.34 | 1.36 | 0.13 | 1 | -2 | 1 | -0.9298 | 0.5 | 7.1 | 11 |  |  | 0.0663 | 0.0764 | 0.002 | 1 |  |  |  |  | 133 |  | 0 | [1140 empty] | 0
P | 2018-01-05T19:32:12+08:00 | 2018-01-06T10:41:42+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM751G | CPSM Final Test | PASS |  | 167.18 | -0.02 | -0.38 | 56.61 | 53.09 | 1.59 | 0.13 | 0 | 1 | -1 | -0.8979 | 0.536 | 1.1 | 14 |  |  | 0.1357 | 0.1569 | -0.0189 | 1 |  |  |  |  | 133 |  | 0 | [67 empty] | 0
P | 2018-01-06T11:30:14+08:00 | 2018-01-06T14:54:38+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM744G | CPSM Final Test | PASS |  | 161.33 | 0.37 | 0.38 | 56.53 | 51.65 | 1.36 | 0.13 | 1 | 2 | -1 | -0.818 | 0.477 | 4.3 | 12 |  |  | 0.0921 | 0.0861 | 0.0021 | 0.9999 |  |  |  |  | 133 |  | 0 | [849 empty] | 0
P | 2018-01-06T11:16:21+08:00 | 2018-01-06T14:57:08+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM745G | CPSM Final Test | PASS |  | 163.94 | -0.02 | -0.31 | 56.74 | 57.74 | 1.51 | 0.13 | 0 | -12 | 1 | -0.8041 | 0.537 | 3.3 | -3 |  |  | 0.0436 | 0.0581 | -0.0032 | 1 |  |  |  |  | 133 |  | 0 | [742 empty] | 0
P | 2018-01-06T11:55:28+08:00 | 2018-01-06T15:49:03+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM743G | CPSM Final Test | PASS |  | 164.21 | -0.3 | 0.07 | 56.78 | 54.53 | 1.82 | 0.13 | 0 | -11 | 3 | -0.9219 | 0.4784 | 5.2 | -4 |  |  | 0.104 | 0.1119 | 0.0281 | 0.99999 |  |  |  |  | 133 |  | 0 | [381 empty] | 0
P | 2018-01-06T15:22:12+08:00 | 2018-01-06T16:01:42+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI201G | CPSM Final Test | FAIL
P | 2018-01-06T07:14:46+08:00 | 2018-01-06T16:09:55+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM739G | CPSM Final Test | FAIL
P | 2018-01-06T15:28:37+08:00 | 2018-01-06T18:17:38+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM752G | CPSM Final Test | PASS |  | 166.01 | -0.02 | -0.38 | 56.94 | 55.03 | 1.36 | 0.13 | 0 | -13 | 1 | -0.8979 | 0.5302 | 2.2 | 13 |  |  | 0.1883 | 0.1891 | -0.036 | 0.99999 |  |  |  |  | 133 |  | 0 | [307 empty] | 0
P | 2018-01-05T21:57:18+08:00 | 2018-01-07T09:26:36+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI201G | CPSM Final Test | FAIL
P | 2018-01-06T15:41:57+08:00 | 2018-01-07T11:12:18+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | M17SDI201G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.61 | 57.23 | 54.95 | 0.98 | 0.13 | 0 | 2 | 0 | -0.8819 | 0.4864 | 0.1 | 7 |  | 0 | 0.1414 | 0.1312 | 0.0058 | 1 | [1148 empty] | 0
P | 2018-01-07T10:36:23+08:00 | 2018-01-07T13:28:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM739G | CPSM Final Test | PASS |  | 162.59 | -0.52 | -0.08 | 56.66 | 53.51 | 1.51 | 0.16 | 0 | -2 | 0 | -0.9598 | 0.4914 | 6.6 | -5 |  |  | 0.0691 | 0.1212 | -0.0447 | 0.99999 |  |  |  |  | 133 |  | 0 | [775 empty] | 0
P | 2018-01-06T22:12:13+08:00 | 2018-01-08T09:02:26+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI202G | CPSM Final Test | FAIL
P | 2018-01-07T15:37:20+08:00 | 2018-01-08T09:09:26+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI202G | CPSM Final Test | PASS | 129 | 163.4 | 0.17 | -0.31 | 56.84 | 54.53 | 1.82 | 0.13 | -1 | 4 | -1 | -0.9058 | 0.499 | 2.1 | 46 |  | 0 | 0.0169 | 0.0991 | -0.031 | 1 | [286 empty] | 0
P | 2018-01-08T06:30:08+08:00 | 2018-01-08T11:44:40+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI205G | CPSM Final Test | FAIL | [903 empty] | 0
P | 2018-01-08T09:15:05+08:00 | 2018-01-08T12:19:37+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI206G | CPSM Final Test | PASS | 129 | 165.83 | -0.02 | -0.54 | 56.41 | 58.59 | 1.13 | 0.13 | -1 | 3 | 1 | -0.8979 | 0.5016 | 0.2 | 6 |  | 0 | 0.0587 | 0.104 | -0.0154 | 1 | [191 empty] | 0
P | 2018-01-08T09:07:24+08:00 | 2018-01-08T12:50:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI205G | CPSM Final Test | PASS | 129 | 162.14 | -0.02 | -0.31 | 56.79 | 56.39 | 1.51 | 0.13 | 0 | -2 | -1 | -0.876 | 0.503 | 1 | 3 |  | 0.01966 | 0.1144 | 0.1412 | 0.006 | 1 | [1149 empty] | 0
P | 2018-01-08T10:10:35+08:00 | 2018-01-08T13:14:30+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI207G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -0.08 | 56.79 | 54.53 | 1.67 | 0.13 | 1 | 4 | 0 | -0.846 | 0.496 | 3.3 | 4 |  | 0.01498 | 0.0798 | 0.1603 | -0.0167 | 1 | [313 empty] | 0
P | 2018-01-06T18:05:54+08:00 | 2018-01-08T13:26:01+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI204G | CPSM Final Test | FAIL | [17 empty] | 244.9507
P | 2018-01-06T17:49:39+08:00 | 2018-01-08T13:28:38+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI203G | CPSM Final Test | FAIL | [17 empty] | 84.2537
P | 2018-01-07T15:45:57+08:00 | 2018-01-08T14:29:21+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI204G | CPSM Final Test | PASS | 129 | 165.02 | 0.17 | -0.69 | 56.64 | 53.34 | 1.13 | 0.13 | -1 | -2 | 0 | -0.8959 | 0.4972 | 0.8 | 0 |  | 0 | 0.0519 | 0.0495 | -0.0132 | 1 | [72 empty] | 0
P | 2018-01-07T15:39:08+08:00 | 2018-01-08T14:31:52+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI203G | CPSM Final Test | PASS | 129 | 164.75 | 0.37 | -0.54 | 56.54 | 54.36 | 1.06 | 0.13 | -1 | 3 | 1 | -0.9158 | 0.452 | 4.9 | 3 |  | 0 | 0.049 | 0.2335 | 0.0223 | 0.99996 | [1150 empty] | 0
P | 2018-01-08T12:02:57+08:00 | 2018-01-08T15:33:38+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI211G | CPSM Final Test | PASS | 129 | 168.71 | -0.02 | -0.31 | 57 | 52.66 | 1.82 | 0.13 | 0 | 2 | 0 | -0.7522 | 0.5088 | 12.6 | 20 |  | 0 | 0.1213 | 0.1093 | 0.0511 | 0.99995 | [203 empty] | 0
P | 2018-01-08T13:50:50+08:00 | 2018-01-08T17:16:56+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI212G | CPSM Final Test | PASS | 129 | 160.43 | -0.02 | 0 | 56.77 | 54.44 | 1.21 | 0.13 | 0 | -12 | 0 | -0.894 | 0.5106 | 2.8 | 11 |  | 0 | 0.0618 | 0.1245 | 0.0337 | 0.99999 | [855 empty] | 0
P | 2018-01-08T13:53:23+08:00 | 2018-01-08T17:20:32+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI208G | CPSM Final Test | PASS | 129 | 165.02 | -0.02 | -0.46 | 56.77 | 54.61 | 1.29 | 0.13 | 1 | 4 | 1 | -0.8201 | 0.4984 | 7.3 | 12 |  | 0 | 0.0976 | 0.1275 | -0.0199 | 0.99999 | [1149 empty] | 0
P | 2018-01-08T13:59:09+08:00 | 2018-01-08T17:27:29+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI209G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.46 | 56.81 | 55.54 | 1.74 | 0.13 | 1 | -8 | -1 | -0.8161 | 0.4892 | 0.8 | 22 |  | 0 | 0.0611 | 0.2413 | -0.0129 | 0.99998 | [532 empty] | 0
P | 2018-01-08T13:35:44+08:00 | 2018-01-08T17:31:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI213G | CPSM Final Test | FAIL
P | 2018-01-08T15:36:26+08:00 | 2018-01-08T18:28:37+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI210G | CPSM Final Test | PASS | 129 | 163.31 | -0.02 | 0 | 57.07 | 55.71 | 1.67 | 0.13 | 1 | -10 | 2 | -0.9019 | 0.5052 | 1.5 | 81 |  | 0 | 0.0363 | 0.08 | -0.0153 | 1 | [134 empty] | 0
P | 2018-01-08T15:33:33+08:00 | 2018-01-08T19:20:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI213G | CPSM Final Test | PASS | 129 | 166.01 | -0.02 | 0.07 | 57.08 | 56.13 | 1.51 | 0.13 | 0 | 3 | 0 | -0.858 | 0.5412 | 6.2 | 6 |  | 0 | 0.094 | 0.195 | -0.0123 | 0.99999 | [1151 empty] | 0
P | 2018-01-08T19:01:12+08:00 | 2018-01-08T21:48:33+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI214G | CPSM Final Test | PASS | 129 | 165.11 | -0.02 | -0.08 | 56.91 | 53.93 | 1.74 | 0.13 | -1 | -11 | 1 | -0.9657 | 0.5382 | 0.6 | 0 |  | 0 | 0.026 | 0.0335 | 0.0348 | 1 | [1152 empty] | 0
P | 2018-01-08T18:56:05+08:00 | 2018-01-08T21:48:10+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI215G | CPSM Final Test | PASS | 129 | 158.63 | -0.02 | -0.31 | 57.07 | 52.92 | 1.51 | 0.13 | 0 | -2 | -1 | -0.8879 | 0.4958 | 2.9 | -15 |  | 0 | 0.114 | 0.0859 | 0.0067 | 0.99999 | [1153 empty] | 0
P | 2018-01-08T18:53:07+08:00 | 2018-01-08T21:52:44+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI218G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -0.31 | 57.02 | 54.44 | 1.36 | 0.13 | 1 | -2 | 0 | -0.9538 | 0.49 | 5.4 | 0 |  | 0 | 0.0875 | 0.1045 | 0.0073 | 1 | [1151 empty] | 0
P | 2018-01-08T19:04:28+08:00 | 2018-01-08T22:18:18+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI216G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | 0 | 57.12 | 57.57 | 1.44 | 0.13 | 1 | 3 | -1 | -0.8978 | 0.542 | 3.1 | 40 |  | 0 | 0.0613 | 0.0767 | 0.0743 | 0.99993 | [1154 empty] | 0
P | 2018-01-09T09:38:30+08:00 | 2018-01-09T10:18:02+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI217G | CPSM Final Test | FAIL
P | 2018-01-09T09:51:58+08:00 | 2018-01-09T12:56:54+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI221G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | 0.45 | 57.08 | 53.34 | 2.05 | 0.13 | 0 | 3 | -1 | -0.8001 | 0.4972 | 0.7 | 50 |  | 0 | 0.0426 | 0.0914 | -0.0005 | 0.99995 | [364 empty] | 0
P | 2018-01-09T09:55:29+08:00 | 2018-01-09T13:03:48+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI220G | CPSM Final Test | PASS | 129 | 168.08 | -0.02 | 0.23 | 57.23 | 53.26 | 1.9 | 0.13 | 0 | 3 | 0 | -0.7482 | 0.502 | 6.9 | -7 |  | 0 | 0.3959 | 0.1755 | -0.0571 | 0.99999 | [1155 empty] | 0
P | 2018-01-09T09:58:16+08:00 | 2018-01-09T13:05:58+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI217G | CPSM Final Test | PASS | 129 | 166.1 | 0.37 | -0.54 | 56.8 | 53.42 | 1.13 | 0.13 | 0 | -12 | -2 | -0.7781 | 0.5054 | 0.8 | 16 |  | 0 | 0.1208 | 0.3476 | 0.004 | 0.99999 | [881 empty] | 0
P | 2018-01-09T10:02:29+08:00 | 2018-01-09T13:12:01+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI219G | CPSM Final Test | PASS | 129 | 166.28 | 0.17 | -0.54 | 56.62 | 53.85 | 1.13 | 0.13 | 0 | -11 | 1 | -0.7882 | 0.5108 | 0.4 | -4 |  | 0 | 0.0475 | 0.1199 | -0.024 | 1 | [1156 empty] | 0
P | 2018-01-09T11:23:53+08:00 | 2018-01-09T14:18:57+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI223G | CPSM Final Test | PASS | 129 | 166.82 | -0.02 | -0.23 | 56.77 | 54.44 | 1.29 | 0.22 | -1 | -3 | 0 | -0.872 | 0.4744 | 0.6 | 0 |  | 0.01275 | 0.0629 | 0.0638 | 0.0079 | 0.99998 | [234 empty] | 0
P | 2018-01-09T13:20:45+08:00 | 2018-01-09T14:19:06+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI225G | CPSM Final Test | FAIL | [5 empty] | 47.24
P | 2018-01-09T12:44:30+08:00 | 2018-01-09T14:48:40+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI224G | CPSM Final Test | FAIL
P | 2018-01-09T07:51:08+08:00 | 2018-01-09T14:56:49+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI222G | CPSM Final Test | FAIL | [811 empty] | 4449
P | 2018-01-09T11:23:59+08:00 | 2018-01-09T16:02:25+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI222G | CPSM Final Test | PASS | 129 | 166.64 | -0.3 | -0.31 | 57.06 | 55.71 | 1.97 | 0.13 | 0 | -10 | 1 | -0.8959 | 0.497 | 4.3 | 0 |  | 0 | 0.1281 | 0.095 | 0.0657 | 0.99996 | [790 empty] | 0
P | 2018-01-09T13:46:35+08:00 | 2018-01-09T16:49:22+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI224G | CPSM Final Test | PASS | 129 | 160.52 | -0.02 | -0.46 | 56.7 | 57.66 | 1.67 | 0.13 | 1 | -2 | 1 | -0.9159 | 0.505 | 0.3 | 17 |  | 0 | 0.1129 | 0.1331 | -0.0101 | 0.99999 | [1157 empty] | 0
P | 2018-01-09T13:49:56+08:00 | 2018-01-09T16:51:42+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI225G | CPSM Final Test | PASS | 129 | 161.15 | -0.02 | -0.54 | 56.7 | 48.6 | 1.13 | 0.13 | 0 | -2 | -1 | -0.9039 | 0.4984 | 5.1 | 15 |  | 0 | 0.1402 | 0.236 | -0.0124 | 0.99998 | [771 empty] | 0
P | 2018-01-09T16:18:32+08:00 | 2018-01-09T19:12:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI226G | CPSM Final Test | PASS | 129 | 162.05 | -0.02 | -0.38 | 56.82 | 51.82 | 1.13 | 0.13 | 0 | -2 | 1 | -0.9517 | 0.504 | 0.1 | 17 |  | 0 | 0.0664 | 0.0302 | -0.0305 | 1 | [205 empty] | 0
P | 2018-01-09T16:22:29+08:00 | 2018-01-09T19:17:58+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI227G | CPSM Final Test | PASS | 129 | 164.39 | -0.02 | -0.84 | 56.47 | 54.78 | 0.98 | 0.13 | 0 | -3 | -1 | -0.9118 | 0.4952 | 0.3 | -5 |  | 0 | 0.0804 | 0.0815 | -0.0011 | 0.99998 | [334 empty] | 0
P | 2018-01-09T18:32:37+08:00 | 2018-01-09T21:25:29+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI228G | CPSM Final Test | PASS | 129 | 160.43 | -0.02 | -0.31 | 57.12 | 55.88 | 1.44 | 0.13 | 1 | -3 | 1 | -0.9119 | 0.5486 | 3.5 | -1 |  | 0 | 0.0789 | 0.0412 | 0.0073 | 0.99999 | [817 empty] | 0
P | 2018-01-09T18:34:14+08:00 | 2018-01-09T21:46:39+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI229G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | 0.45 | 56.76 | 53.34 | 1.44 | 0.13 | 1 | 12 | -1 | -0.884 | 0.5348 | 4.3 | 100 |  | 0 | 0.1745 | 0.1081 | 0.0774 | 0.9999 | [790 empty] | 0
P | 2018-01-10T08:33:48+08:00 | 2018-01-10T09:17:52+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI232G | CPSM Final Test | FAIL
P | 2018-01-10T09:02:58+08:00 | 2018-01-10T12:09:18+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI231G | CPSM Final Test | PASS | 129 | 171.05 | -0.02 | -0.31 | 56.9 | 55.03 | 1.59 | 0.13 | 1 | 4 | -1 | -0.8401 | 0.4998 | 0.7 | -17 |  | 0 | 0.1312 | 0.1675 | 0.0058 | 0.99999 | [1158 empty] | 0
P | 2018-01-10T08:55:50+08:00 | 2018-01-10T12:12:44+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI232G | CPSM Final Test | PASS | 129 | 169.52 | 0.17 | -0.16 | 56.82 | 50.21 | 1.67 | 0.13 | 0 | -3 | 1 | -0.842 | 0.5378 | 8.2 | 2 |  | 0.07108 | 0.1382 | 0.2918 | -0.0063 | 0.99997 | [862 empty] | 0
P | 2018-01-10T08:50:14+08:00 | 2018-01-10T12:17:20+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI230G | CPSM Final Test | PASS | 129 | 164.12 | -0.02 | -0.08 | 56.72 | 57.15 | 2.05 | 0.13 | 0 | -2 | -1 | -0.8819 | 0.513 | 0.9 | 24 |  | 0 | 0.132 | 0.0927 | 0.0124 | 0.99999 | [139 empty] | 0
P | 2018-01-10T10:53:38+08:00 | 2018-01-10T13:41:35+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI233G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -1.15 | 57 | 50.88 | 1.13 | 0.13 | -1 | 5 | -2 | -0.8839 | 0.496 | 0.4 | 12 |  | 0 | 0.0761 | 0.1081 | 0.022 | 0.99999 | [228 empty] | 0
P | 2018-01-10T10:47:50+08:00 | 2018-01-10T14:04:37+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI234G | CPSM Final Test | PASS | 129 | 164.03 | -0.02 | -0.46 | 56.73 | 57.57 | 1.21 | 0.13 | 1 | -2 | -1 | -0.8002 | 0.494 | 1.3 | -6 |  | 0 | 0.1229 | 0.2392 | 0.0042 | 0.99996 | [1159 empty] | 0
P | 2018-01-10T14:19:34+08:00 | 2018-01-10T17:18:04+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI239G | CPSM Final Test | PASS | 129 | 162.95 | 0.37 | -0.46 | 56.82 | 55.8 | 1.29 | 0.13 | 0 | -2 | -1 | -0.86 | 0.462 | 1.6 | 18 |  | 0 | 0.0918 | 0.053 | 0.023 | 0.99998 | [619 empty] | 0
P | 2018-01-10T13:56:44+08:00 | 2018-01-10T17:20:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI238G | CPSM Final Test | PASS | 129 | 166.64 | -0.02 | -0.46 | 56.78 | 52.49 | 1.21 | 0.13 | 1 | 2 | -1 | -0.872 | 0.4976 | 3.7 | 14 |  | 0 | 0.0648 | 0.1721 | -0.0353 | 1 | [854 empty] | 0
P | 2018-01-10T11:43:52+08:00 | 2018-01-10T17:21:18+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI235G | CPSM Final Test | FAIL
P | 2018-01-10T13:52:57+08:00 | 2018-01-10T17:24:15+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI237G | CPSM Final Test | PASS | 129 | 164.48 | -0.02 | -0.31 | 56.83 | 55.03 | 1.51 | 0.13 | 1 | -2 | 0 | -0.8539 | 0.5468 | 2.2 | -3 |  | 0 | 0.0438 | 0.0542 | -0.0324 | 1 | [289 empty] | 0
P | 2018-01-10T14:32:35+08:00 | 2018-01-10T17:34:14+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI235G | CPSM Final Test | PASS | 129 | 159.71 | -0.02 | -0.54 | 57.05 | 54.86 | 1.44 | 0.13 | 1 | -11 | 0 | -0.8559 | 0.544 | 5.1 | 1 |  | 0 | 0.3993 | 0.2897 | -0.012 | 1 | [1160 empty] | 0
P | 2018-01-10T13:50:12+08:00 | 2018-01-10T17:42:44+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI236G | CPSM Final Test | PASS | 129 | 164.48 | 0.27 | -0.61 | 56.76 | 53 | 1.21 | 0.13 | 0 | -4 | -1 | -0.85 | 0.5054 | 1.4 | 0 |  | 0.0228 | 0.1494 | 0.0851 | -0.0113 | 0.99999 | [862 empty] | 0
P | 2018-01-10T18:48:06+08:00 | 2018-01-10T21:44:56+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI240G | CPSM Final Test | PASS | 129 | 168.62 | 0.17 | -0.69 | 56.32 | 51.65 | 0.98 | 0.13 | 0 | -2 | 3 | -0.8959 | 0.5494 | 6.9 | -1 |  | 0 | 0.1426 | 0.1125 | -0.0021 | 1 | [258 empty] | 0
P | 2018-01-10T18:52:14+08:00 | 2018-01-10T21:47:02+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI241G | CPSM Final Test | PASS | 129 | 159.08 | -0.02 | -0.54 | 56.77 | 59.18 | 1.67 | 0.13 | 0 | -3 | 3 | -0.8421 | 0.4986 | 1.9 | 6 |  | 0 | 0.1119 | 0.2646 | 0.0248 | 0.99999 | [1161 empty] | 0
P | 2018-01-10T18:55:22+08:00 | 2018-01-10T21:48:43+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI242G | CPSM Final Test | PASS | 129 | 164.84 | 0.27 | -0.46 | 57.1 | 55.96 | 1.36 | 0.13 | -1 | 5 | 0 | -0.874 | 0.526 | 1.2 | 5 |  | 0.00786 | 0.0928 | 0.1046 | -0.0465 | 1 | [43 empty] | 0
P | 2018-01-11T08:12:42+08:00 | 2018-01-11T11:04:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI243G | CPSM Final Test | FAIL
P | 2018-01-11T09:50:15+08:00 | 2018-01-11T12:36:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI245G | CPSM Final Test | PASS | 129 | 161.33 | 0.37 | -0.31 | 56.71 | 58.76 | 1.36 | 0.13 | 0 | 4 | -1 | -0.8719 | 0.542 | 6.2 | 88 |  | 0 | 0.0643 | 0.1321 | -0.0453 | 1 | [942 empty] | 0
P | 2018-01-11T09:48:53+08:00 | 2018-01-11T12:40:04+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI247G | CPSM Final Test | PASS | 129 | 167 | -0.52 | -0.54 | 56.79 | 55.12 | 1.29 | 0.13 | 1 | -12 | 1 | -0.8879 | 0.535 | 2.9 | -6 |  | 0 | 0.1129 | 0.0704 | -0.0018 | 1 | [711 empty] | 0
P | 2018-01-11T09:42:53+08:00 | 2018-01-11T12:46:04+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI244G | CPSM Final Test | PASS | 129 | 164.12 | 0.37 | -0.69 | 56.71 | 58.5 | 1.29 | 0.13 | 1 | 2 | 1 | -0.9498 | 0.538 | 1.3 | 5 |  | 0 | 0.0616 | 0.1912 | 0.0176 | 0.99999 | [803 empty] | 0
P | 2018-01-11T10:07:54+08:00 | 2018-01-11T12:53:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI246G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.46 | 56.78 | 54.1 | 0.98 | 0.13 | 1 | -3 | 1 | -0.8979 | 0.4982 | 2.3 | 74 |  | 0 | 0.0412 | 0.0898 | 0.0343 | 0.99998 | [92 empty] | 0
P | 2018-01-11T09:38:39+08:00 | 2018-01-11T13:14:18+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI243G | CPSM Final Test | PASS | 129 | 165.11 | 0.27 | -0.54 | 56.5 | 52.66 | 1.44 | 0.13 | 0 | -3 | 1 | -0.9099 | 0.5442 | 3 | 1 |  | 0 | 0.0698 | 0.1794 | -0.0524 | 1 | [686 empty] | 0
P | 2018-01-12T08:45:21+08:00 | 2018-01-12T14:03:10+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM754G | CPSM Final Test | FAIL | [1183 empty] | 0
P | 2018-01-12T07:32:58+08:00 | 2018-01-12T14:06:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM753G | CPSM Final Test | FAIL | [14 empty] | 1203
P | 2018-01-12T10:32:05+08:00 | 2018-01-12T14:33:19+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | A18USM756G | CPSM Final Test | FAIL
P | 2018-01-12T11:41:11+08:00 | 2018-01-12T14:35:35+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM755G | CPSM Final Test | PASS |  | 164.12 | -0.02 | 0 | 57.04 | 52.92 | 1.36 | 0.13 | 0 | 4 | 0 | -0.8599 | 0.5076 | 6.2 | -7 |  |  | 0.052 | 0.1002 | -0.0219 | 0.99999 |  |  |  |  | 133 |  | 0 | [245 empty] | 0
P | 2018-01-12T10:49:36+08:00 | 2018-01-12T15:54:41+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM753G | CPSM Final Test | PASS |  | 162.05 | -0.02 | -0.54 | 56.84 | 50.97 | 0.98 | 0.13 | 1 | -3 | -1 | -0.8421 | 0.5218 | 0.5 | 86 |  |  | 0.1527 | 0.1371 | -0.0196 | 0.99997 |  |  |  |  | 133 |  | 0 | [289 empty] | 0
P | 2018-01-12T12:32:42+08:00 | 2018-01-12T16:20:51+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | A18USM756G | CPSM Final Test | PASS |  | 160.16 | 0.37 | -0.31 | 56.74 | 60.45 | 1.44 | 0.13 | 0 | -2 | 1 | -0.8101 | 0.4902 | 0.3 | 11 |  |  | 0.1306 | 0.0759 | -0.0483 | 1 |  |  |  |  | 133 |  | 0 | [601 empty] | 0
P | 2018-01-12T06:26:41+08:00 | 2018-01-12T16:21:50+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM754G | CPSM Final Test | FAIL
P | 2018-01-12T06:21:55+08:00 | 2018-01-12T16:26:36+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM754G | CPSM Final Test | FAIL | [27 empty] | 0
P | 2018-01-12T13:39:59+08:00 | 2018-01-12T16:35:49+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM757G | CPSM Final Test | PASS |  | 167.54 | -0.21 | -0.54 | 56.64 | 53.26 | 1.51 | 0.13 | 0 | 3 | 1 | -0.832 | 0.53 | 0.8 | 15 |  |  | 0.0513 | 0.1357 | 0.0042 | 0.99999 |  |  |  |  | 133 |  | 0 | [872 empty] | 0
P | 2018-01-12T11:24:15+08:00 | 2018-01-12T17:29:32+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM754G | CPSM Final Test | PASS |  | 163.04 | -0.02 | -0.23 | 57.1 | 53.34 | 1.21 | 0.13 | 0 | -12 | 2 | -0.8859 | 0.54 | 6 | 23 |  |  | 0.1587 | 0.1559 | 0.01 | 0.99997 |  |  |  |  | 133 |  | 0 | [1155 empty] | 0
P | 2018-01-12T15:34:05+08:00 | 2018-01-12T18:23:57+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM759G | CPSM Final Test | PASS |  | 167.09 | -0.02 | -0.46 | 56.76 | 53.68 | 1.67 | 0.13 | 0 | 4 | -1 | -0.786 | 0.5048 | 6 | 5 |  |  | 0.088 | 0.1139 | 0.0049 | 0.99999 |  |  |  |  | 133 |  | 0 | [857 empty] | 0
P | 2018-01-12T10:20:31+08:00 | 2018-01-12T18:25:01+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | A18USM758G | CPSM Final Test | FAIL
P | 2018-01-12T13:39:49+08:00 | 2018-01-12T19:16:39+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM761G | CPSM Final Test | FAIL
P | 2018-01-12T14:22:46+08:00 | 2018-01-12T20:14:27+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | A18USM758G | CPSM Final Test | PASS |  | 161.24 | 0.08 | -0.84 | 56.54 | 51.48 | 1.82 | 0.13 | 0 | 3 | -1 | -0.7962 | 0.529 | 6.1 | -117 |  |  | 0.1748 | 0.0942 | 0.0338 | 0.99994 |  |  |  |  | 133 |  | 0 | [558 empty] | 0
P | 2018-01-12T17:00:19+08:00 | 2018-01-12T20:32:31+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM760G | CPSM Final Test | PASS |  | 158.63 | -0.02 | -0.46 | 57.01 | 55.03 | 0.83 | 0.13 | 0 | 2 | 1 | -0.9138 | 0.512 | 3.2 | -11 |  |  | 0.1556 | 0.2963 | -0.0033 | 0.99999 |  |  |  |  | 133 |  | 0 | [1156 empty] | 0
P | 2018-01-12T17:12:26+08:00 | 2018-01-12T20:37:15+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | A18USM762G | CPSM Final Test | PASS |  | 172.67 | 0.08 | -0.38 | 56.47 | 49.28 | 1.74 | 0.13 | 1 | 3 | 1 | -0.8501 | 0.5134 | 5.6 | 6 |  |  | 0.085 | 0.0572 | -0.0132 | 0.99996 |  |  |  |  | 133 |  | 0 | [857 empty] | 0
P | 2018-01-12T16:28:14+08:00 | 2018-01-12T21:05:07+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM761G | CPSM Final Test | PASS |  | 164.39 | 0.17 | -0.31 | 56.9 | 54.95 | 1.67 | 0.13 | 1 | 4 | 1 | -0.862 | 0.542 | 0.8 | 11 |  |  | 0.092 | 0.1314 | 0.0083 | 1 |  |  |  |  | 133 |  | 0 | [695 empty] | 0
P | 2018-01-12T18:52:10+08:00 | 2018-01-12T21:49:45+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM763G | CPSM Final Test | PASS |  | 166.19 | -0.02 | -0.31 | 56.78 | 52.83 | 1.44 | 0.13 | 1 | 11 | 0 | -0.8679 | 0.5022 | 8.3 | -8 |  |  | 0.0384 | 0.0708 | -0.0051 | 1 |  |  |  |  | 133 |  | 0 | [198 empty] | 0
P | 2018-01-12T19:02:40+08:00 | 2018-01-12T21:57:10+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM764G | CPSM Final Test | PASS |  | 164.93 | -0.11 | -0.23 | 56.98 | 54.27 | 1.59 | 0.13 | 0 | -2 | 1 | -0.9179 | 0.4754 | 8.2 | 24 |  |  | 0.2986 | 0.1335 | 0.0002 | 1 |  |  |  |  | 133 |  | 0 | [1157 empty] | 0
P | 2018-01-13T09:23:02+08:00 | 2018-01-13T14:08:07+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI248G | CPSM Final Test | FAIL
P | 2018-01-13T11:05:58+08:00 | 2018-01-13T14:46:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM766G | CPSM Final Test | PASS |  | 169.79 | -0.02 | -0.31 | 56.91 | 53.09 | 1.36 | 0.13 | 1 | 4 | -1 | -0.9139 | 0.5428 | 6.8 | 0 |  |  | 0.1144 | 0.1372 | -0.0245 | 0.99999 |  |  |  |  | 133 |  | 0 | [273 empty] | 0
P | 2018-01-13T11:18:23+08:00 | 2018-01-13T14:51:43+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM765G | CPSM Final Test | PASS |  | 163.76 | 0.08 | -0.16 | 57.02 | 55.03 | 1.59 | 0.13 | -1 | 4 | 3 | -0.9297 | 0.523 | 3.1 | 42 |  |  | 0.0576 | 0.1493 | 0.0337 | 0.99999 |  |  |  |  | 133 |  | 0 | [407 empty] | 0
P | 2018-01-13T11:42:56+08:00 | 2018-01-13T14:49:34+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI251G | CPSM Final Test | FAIL
P | 2018-01-13T07:58:57+08:00 | 2018-01-13T15:01:08+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI249G | CPSM Final Test | FAIL | [937 empty] | 9916
P | 2018-01-13T11:45:34+08:00 | 2018-01-13T15:53:01+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI248G | CPSM Final Test | PASS | 129 | 164.03 | -0.02 | -0.23 | 56.72 | 57.49 | 1.51 | 0.13 | 1 | 3 | -1 | -0.9139 | 0.5396 | 2.3 | 6 |  | 0 | 0.1538 | 0.1688 | -0.0539 | 1 | [119 empty] | 0
P | 2018-01-13T13:16:15+08:00 | 2018-01-13T16:33:38+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI251G | CPSM Final Test | PASS | 129 | 164.21 | -0.02 | -0.16 | 57.1 | 54.1 | 1.51 | 0.13 | 1 | -3 | 4 | -0.8779 | 0.5216 | 6 | -4 |  | 0.01464 | 0.113 | 0.1755 | -0.0074 | 0.99999 | [782 empty] | 0
P | 2018-01-13T13:30:13+08:00 | 2018-01-13T17:02:37+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI250G | CPSM Final Test | PASS | 129 | 163.67 | 0.27 | 0.15 | 56.82 | 53.26 | 1.36 | 0.13 | 0 | 2 | 1 | -0.7782 | 0.5344 | 5.5 | 27 |  | 0 | 0.1633 | 0.126 | 0.0388 | 0.99993 | [1133 empty] | 0
P | 2018-01-13T14:52:07+08:00 | 2018-01-13T17:50:31+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI255G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.77 | 56.82 | 54.69 | 1.36 | 0.13 | 1 | -12 | -1 | -0.85 | 0.4866 | 5.9 | 19 |  | 0 | 0.1102 | 0.1996 | -0.0305 | 0.99999 | [203 empty] | 0
P | 2018-01-13T15:21:07+08:00 | 2018-01-13T18:04:01+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI254G | CPSM Final Test | PASS | 129 | 166.37 | -0.11 | -0.69 | 57 | 56.47 | 1.21 | 0.13 | 0 | 2 | 1 | -0.9139 | 0.5144 | 1.3 | -6 |  | 0 | 0.1117 | 0.1703 | -0.0188 | 0.99998 | [1165 empty] | 0
P | 2018-01-13T16:06:34+08:00 | 2018-01-13T18:51:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI256G | CPSM Final Test | PASS | 129 | 160.52 | 0.17 | -0.54 | 56.5 | 49.45 | 1.67 | 0.13 | 1 | 2 | -1 | -0.836 | 0.4962 | 4.3 | -10 |  | 0 | 0.0349 | 0.1379 | -0.0192 | 0.99998 | [685 empty] | 0
P | 2018-01-12T12:16:01+08:00 | 2018-01-14T10:44:04+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI249G | CPSM Final Test | FAIL | [937 empty] | 14201
P | 2018-01-13T11:30:02+08:00 | 2018-01-14T11:49:46+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI249G | CPSM Final Test | PASS | 129 | 166.91 | 0.17 | -0.84 | 56.44 | 58.42 | 1.51 | 0.13 | 0 | -2 | 3 | -0.87 | 0.5128 | 9.9 | 5 |  | 0.01159 | 0.1522 | 0.2619 | 0.0168 | 0.99999 | [916 empty] | 0
P | 2018-01-14T10:25:14+08:00 | 2018-01-14T13:16:43+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI257G | CPSM Final Test | PASS | 129 | 168.17 | -0.11 | -0.46 | 56.73 | 59.1 | 1.06 | 0.13 | 0 | 4 | -1 | -0.8081 | 0.524 | 0.3 | 8 |  | 0 | 0.0997 | 0.1362 | 0.0065 | 0.99999 | [811 empty] | 0
P | 2018-01-14T10:36:47+08:00 | 2018-01-14T13:20:46+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI252G | CPSM Final Test | PASS | 129 | 164.39 | 0.17 | -0.08 | 57.22 | 55.37 | 1.59 | 0.13 | 1 | -3 | 1 | -0.8161 | 0.536 | 1.7 | 16 |  | 0 | 0.0649 | 0.1267 | -0.024 | 0.99998 | [1166 empty] | 0
P | 2018-01-14T10:46:06+08:00 | 2018-01-14T13:29:30+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI259G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | -0.31 | 56.64 | 51.73 | 1.51 | 0.13 | 0 | 4 | -1 | -0.7924 | 0.5024 | 0 | 34 |  | 0 | 0.0542 | 0.0976 | -0.013 | 1 | [625 empty] | 0
P | 2018-01-14T10:19:21+08:00 | 2018-01-14T13:32:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI253G | CPSM Final Test | PASS | 129 | 171.05 | -0.02 | -0.23 | 56.92 | 53.42 | 1.51 | 0.13 | 0 | -12 | 1 | -0.8321 | 0.49 | 2 | 23 |  | 0.01843 | 0.1155 | 0.1049 | 0.009 | 0.99994 | [74 empty] | 0
P | 2018-01-14T10:56:34+08:00 | 2018-01-14T13:46:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI258G | CPSM Final Test | PASS | 129 | 167.09 | 0.78 | -0.54 | 56.7 | 60.79 | 1.06 | 0.13 | 1 | 4 | 1 | -0.8141 | 0.4824 | 2.2 | 22 |  | 0 | 0.0856 | 0.1031 | -0.0063 | 0.99999 | [521 empty] | 0
P | 2018-01-14T12:00:21+08:00 | 2018-01-14T15:20:02+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI262G | CPSM Final Test | PASS | 129 | 160.07 | -0.02 | -0.46 | 57.06 | 53.68 | 1.51 | 0.13 | 0 | 2 | 2 | -0.8461 | 0.5118 | 4.4 | 6 |  | 0.01193 | 0.1008 | 0.1197 | -0.051 | 1 | [916 empty] | 0
P | 2018-01-14T15:15:14+08:00 | 2018-01-14T18:07:07+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI266G | CPSM Final Test | PASS | 129 | 169.16 | 0.27 | -0.46 | 57.01 | 54.86 | 1.06 | 0.13 | -1 | -2 | -1 | -0.8479 | 0.5108 | 0.1 | -3 |  | 0 | 0.0782 | 0.062 | -0.0153 | 0.99999 | [1167 empty] | 0
P | 2018-01-14T15:23:38+08:00 | 2018-01-14T18:13:06+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI268G | CPSM Final Test | PASS | 129 | 164.3 | 0.17 | -0.54 | 56.49 | 57.66 | 1.44 | 0.13 | 0 | -3 | 1 | -0.888 | 0.5412 | 1.4 | 39 |  | 0 | 0.0583 | 0.1627 | -0.0302 | 1 | [1168 empty] | 0
P | 2018-01-14T15:38:56+08:00 | 2018-01-14T18:28:19+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI261G | CPSM Final Test | PASS | 129 | 169.97 | -0.02 | -0.46 | 57.08 | 53.09 | 1.29 | 0.13 | 0 | -1 | 3 | -0.8481 | 0.4628 | 2.7 | 10 |  | 0 | 0.0352 | 0.1345 | -0.0432 | 1 | [1015 empty] | 0
P | 2018-01-14T15:52:22+08:00 | 2018-01-14T18:42:37+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI267G | CPSM Final Test | PASS | 129 | 164.66 | 0.27 | -0.61 | 56.7 | 57.49 | 1.9 | 0.13 | 0 | -12 | 1 | -0.8401 | 0.504 | 9.8 | 23 |  | 0.00448 | 0.1286 | 0.146 | -0.0224 | 0.99999 | [134 empty] | 0
P | 2018-01-14T16:24:06+08:00 | 2018-01-14T19:06:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI269G | CPSM Final Test | PASS | 129 | 168.62 | -0.11 | -0.54 | 57.13 | 54.78 | 1.21 | 0.13 | 0 | -2 | -1 | -0.8998 | 0.52 | 5.7 | 2 |  | 0.02569 | 0.0509 | 0.1123 | 0.0004 | 1 | [391 empty] | 0
P | 2018-01-13T20:09:04+08:00 | 2018-01-15T09:31:09+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI260G | CPSM Final Test | FAIL
P | 2018-01-14T14:50:07+08:00 | 2018-01-15T11:16:39+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI260G | CPSM Final Test | PASS | 129 | 157.19 | 0.27 | -0.69 | 56.73 | 58.34 | 1.59 | 0.13 | 0 | -3 | 0 | -0.9558 | 0.5202 | 6.3 | 1 |  | 0 | 0.1324 | 0.0983 | 0.0769 | 0.99997 | [1169 empty] | 0
P | 2018-01-15T09:34:57+08:00 | 2018-01-15T12:12:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI274G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-01-15T07:45:57+08:00 | 2018-01-15T13:26:12+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI273G | CPSM Final Test | FAIL
P | 2018-01-15T10:27:02+08:00 | 2018-01-15T13:58:17+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI271G | CPSM Final Test | PASS | 129 | 162.77 | -0.02 | -0.08 | 57.1 | 57.32 | 1.97 | 0.13 | -1 | 2 | 1 | -0.852 | 0.4916 | 5.7 | 58 |  | 0 | 0.1253 | 0.1075 | -0.0216 | 1 | [746 empty] | 0
P | 2018-01-15T10:41:58+08:00 | 2018-01-15T13:59:42+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI272G | CPSM Final Test | PASS | 129 | 166.73 | -0.02 | 0.15 | 56.9 | 54.78 | 1.51 | 0.13 | 0 | -1 | 0 | -0.9359 | 0.4758 | 2.8 | 36 |  | 0 | 0.103 | 0.0726 | -0.0556 | 1 | [849 empty] | 0
P | 2018-01-15T10:40:05+08:00 | 2018-01-15T14:07:23+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI277G | CPSM Final Test | FAIL
P | 2018-01-15T10:36:04+08:00 | 2018-01-15T14:39:50+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI273G | CPSM Final Test | PASS | 129 | 161.78 | -0.21 | -0.38 | 56.84 | 53.26 | 1.44 | 0.13 | 0 | -2 | -1 | -0.8521 | 0.5382 | 3.6 | 8 |  | 0.01401 | 0.0863 | 0.3233 | -0.0061 | 0.99999 | [1170 empty] | 0
P | 2018-01-15T10:53:56+08:00 | 2018-01-15T15:33:16+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI274G | CPSM Final Test | PASS | 129 | 162.32 | 0.08 | -0.23 | 56.49 | 51.9 | 0.91 | 0.13 | 0 | -2 | 0 | -0.8759 | 0.4992 | 0.3 | -10 |  | 0 | 0.061 | 0.1094 | 0.0018 | 1 | [212 empty] | 0
P | 2018-01-15T13:45:36+08:00 | 2018-01-15T16:32:49+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI279G | CPSM Final Test | PASS | 129 | 168.89 | -0.02 | -0.38 | 56.82 | 54.61 | 1.13 | 0.13 | -1 | 3 | 0 | -0.9178 | 0.481 | 4.3 | 34 |  | 0.02185 | 0.0853 | 0.3191 | 0.0259 | 0.99998 | [140 empty] | 0
P | 2018-01-15T14:33:22+08:00 | 2018-01-15T17:57:32+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI265G | CPSM Final Test | PASS | 129 | 164.12 | -0.11 | -0.99 | 56.52 | 53.42 | 1.51 | 0.13 | 0 | -2 | -1 | -0.9019 | 0.478 | 4.1 | 89 |  | 0 | 0.2 | 0.3324 | 0.0027 | 1 | [1171 empty] | 0
P | 2018-01-15T14:43:29+08:00 | 2018-01-15T18:01:45+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI264G | CPSM Final Test | PASS | 129 | 166.91 | 0.56 | -0.92 | 56.41 | 57.07 | 1.44 | 0.13 | 0 | -12 | 1 | -0.9039 | 0.525 | 0.1 | 5 |  | 0.00534 | 0.1138 | 0.0741 | 0.036 | 0.99999 | [455 empty] | 0
P | 2018-01-15T16:29:10+08:00 | 2018-01-15T19:21:10+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI276G | CPSM Final Test | PASS | 129 | 168.08 | -0.11 | -0.61 | 56.76 | 56.13 | 1.06 | 0.13 | 1 | -2 | 0 | -0.9197 | 0.504 | 0.6 | 3 |  | 0.01486 | 0.0966 | 0.3549 | -0.0047 | 0.99999 | [8 empty] | 0
P | 2018-01-15T16:38:00+08:00 | 2018-01-15T19:25:24+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI280G | CPSM Final Test | PASS | 129 | 165.2 | 0.08 | -0.08 | 57.05 | 54.36 | 1.97 | 0.13 | 0 | -3 | 3 | -0.9399 | 0.468 | 5.7 | 5 |  | 0 | 0.0749 | 0.3902 | -0.0493 | 0.99998 | [802 empty] | 0
P | 2018-01-15T17:06:54+08:00 | 2018-01-15T20:18:53+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI281G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.46 | 57.07 | 56.05 | 2.12 | 0.13 | 0 | 2 | 0 | -0.9198 | 0.5492 | 13.6 | 38 |  | 0 | 0.1507 | 0.1796 | -0.0304 | 0.99999 | [1172 empty] | 0
P | 2018-01-15T18:29:05+08:00 | 2018-01-15T21:11:05+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI278G | CPSM Final Test | PASS | 129 | 171.5 | -0.02 | -0.69 | 56.82 | 56.47 | 1.36 | 0.13 | 1 | -2 | -1 | -0.9339 | 0.4924 | 2.7 | 0 |  | 0.04032 | 0.0813 | 0.2592 | -0.0203 | 0.99999 | [1173 empty] | 0
P | 2018-01-15T18:39:15+08:00 | 2018-01-15T21:21:46+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI275G | CPSM Final Test | PASS | 129 | 159.8 | -0.24 | -0.38 | 56.62 | 52.15 | 2.05 | 0.13 | 0 | 3 | 0 | -0.9259 | 0.547 | 5.2 | 40 |  | 0 | 0.0491 | 0.0736 | -0.0256 | 1 | [1174 empty] | 0
P | 2018-01-16T11:04:57+08:00 | 2018-01-16T12:29:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI270G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-01-16T11:01:21+08:00 | 2018-01-16T12:33:15+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI270G | CPSM Final Test | FAIL | [114 empty] | FALSE
P | 2018-01-16T10:52:36+08:00 | 2018-01-16T14:18:53+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI289G | CPSM Final Test | PASS | 129 | 160.7 | -0.02 | -0.16 | 57.14 | 54.27 | 1.51 | 0.13 | -1 | 3 | -1 | -0.9199 | 0.5358 | 4.4 | -2 |  | 0.0056 | 0.1607 | 0.4833 | -0.0107 | 0.99998 | [1054 empty] | 0
P | 2018-01-16T11:24:45+08:00 | 2018-01-16T14:25:04+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI286G | CPSM Final Test | PASS | 129 | 167.72 | 0.17 | -0.31 | 57.01 | 51.65 | 1.21 | 0.13 | 1 | -2 | 0 | -0.9238 | 0.4872 | 3.4 | 1 |  | 0.00821 | 0.0689 | 0.3871 | -0.0454 | 1 | [798 empty] | 0
P | 2018-01-14T10:17:39+08:00 | 2018-01-16T14:29:49+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI277G | CPSM Final Test | FAIL
P | 2018-01-16T11:53:14+08:00 | 2018-01-16T15:25:26+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI263G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.38 | 56.93 | 54.02 | 1.74 | 0.13 | 1 | 2 | 1 | -0.8401 | 0.5356 | 0.9 | 6 |  | 0 | 0.1683 | 0.2756 | -0.0001 | 0.99999 | [827 empty] | 0
P | 2018-01-16T12:07:30+08:00 | 2018-01-16T15:34:49+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI282G | CPSM Final Test | PASS | 129 | 177.53 | 0.17 | -0.23 | 57.01 | 54.27 | 1.44 | 0.13 | 1 | -11 | -1 | -0.854 | 0.5134 | 4.8 | 15 |  | 0 | 0.4121 | 0.6465 | -0.0024 | 0.99999 | [1005 empty] | 0
P | 2018-01-16T12:35:44+08:00 | 2018-01-16T15:56:53+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI270G | CPSM Final Test | PASS | 129 | 165.11 | -0.02 | -0.38 | 57.1 | 53.68 | 1.67 | 0.13 | 1 | -12 | 2 | -0.9358 | 0.4774 | 1.8 | 8 |  | 0.02416 | 0.1501 | 0.1787 | -0.0424 | 1 | [119 empty] | 0
P | 2018-01-14T07:36:35+08:00 | 2018-01-16T17:10:53+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI277G | CPSM Final Test | FAIL | [74 empty] | 0
P | 2018-01-16T13:15:30+08:00 | 2018-01-16T17:18:45+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI288G | CPSM Final Test | FAIL | [12 empty] | 0.001
P | 2018-01-16T14:45:14+08:00 | 2018-01-16T17:53:03+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI285G | CPSM Final Test | PASS | 129 | 167.27 | 0.17 | -0.92 | 56.78 | 53.68 | 1.13 | 0.13 | 0 | -1 | -1 | -0.846 | 0.491 | 5.8 | -17 |  | 0.00518 | 0.0787 | 0.0946 | -0.0081 | 1 | [236 empty] | 0
P | 2018-01-15T12:23:44+08:00 | 2018-01-16T18:16:19+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI277G | CPSM Final Test | PASS | 129 | 160.97 | 0.37 | -1.22 | 56.52 | 53.68 | 1.36 | 0.13 | -1 | -3 | 1 | -0.854 | 0.512 | 1.2 | 11 |  | 0 | 0.134 | 0.1656 | 0.0011 | 0.99996 | [53 empty] | 0
P | 2018-01-16T16:36:01+08:00 | 2018-01-16T19:30:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI284G | CPSM Final Test | PASS | 129 | 166.01 | 0.08 | -0.38 | 56.93 | 54.78 | 1.67 | 0.13 | 0 | -2 | 1 | -0.9398 | 0.458 | 2.5 | 4 |  | 0 | 0.0424 | 0.0147 | -0.0253 | 0.99999 | [1175 empty] | 0
P | 2018-01-16T15:17:07+08:00 | 2018-01-16T19:33:51+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI288G | CPSM Final Test | PASS | 129 | 167.09 | 0.17 | -0.92 | 56.47 | 52.75 | 0.91 | 0.13 | -1 | -1 | 0 | -0.8361 | 0.536 | 1.1 | 13 |  | 0.02314 | 0.1014 | 0.2285 | 0.0121 | 0.99999 | [323 empty] | 0
P | 2018-01-16T16:45:15+08:00 | 2018-01-16T19:33:08+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI290G | CPSM Final Test | PASS | 129 | 159.98 | -0.02 | -0.31 | 56.74 | 57.4 | 1.51 | 0.13 | 0 | 2 | 3 | -0.8679 | 0.541 | 7 | 22 |  | 0 | 0.1298 | 0.0926 | -0.0129 | 0.99999 | [1176 empty] | 0
P | 2018-01-16T16:56:00+08:00 | 2018-01-16T19:38:09+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI283G | CPSM Final Test | PASS | 129 | 159.08 | -0.4 | -0.92 | 56.82 | 53.76 | 1.21 | 0.13 | 0 | -10 | -1 | -0.8639 | 0.539 | 0.6 | -3 |  | 0 | 0.0556 | 0.3329 | -0.0108 | 1 | [1011 empty] | 0
P | 2018-01-16T18:22:24+08:00 | 2018-01-16T21:04:41+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI287G | CPSM Final Test | PASS | 129 | 166.64 | 0.27 | -0.46 | 57.07 | 53.17 | 1.21 | 0.13 | 1 | -2 | -1 | -0.926 | 0.459 | 5.2 | 12 |  | 0.00444 | 0.0543 | 0.0994 | 0.0028 | 0.99999 | [367 empty] | 0
P | 2018-01-17T19:16:36+08:00 | 2018-01-17T22:16:44+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | A18USM769G | CPSM Final Test | PASS |  | 162.68 | 0.27 | -0.31 | 56.81 | 53.85 | 1.59 | 0.13 | -1 | -2 | -1 | -0.9398 | 0.4944 | 2.5 | 49 |  |  | 0.0833 | 0.1694 | -0.0395 | 0.99999 |  |  |  |  | 133 |  | 0 | [843 empty] | 0
P | 2018-01-17T19:12:51+08:00 | 2018-01-17T22:23:17+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | A18USM768G | CPSM Final Test | PASS |  | 164.57 | -0.02 | -0.77 | 57 | 56.56 | 1.21 | 0.13 | 1 | -3 | 0 | -0.9217 | 0.493 | 3.9 | 8 |  |  | 0.062 | 0.1305 | -0.0269 | 1 |  |  |  |  | 133 |  | 0 | [950 empty] | 0
P | 2018-01-17T19:21:05+08:00 | 2018-01-17T22:31:11+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | A18USM770G | CPSM Final Test | PASS |  | 159.53 | -0.02 | -0.61 | 56.82 | 53.68 | 1.29 | 0.13 | 0 | -10 | 0 | -0.9677 | 0.5054 | 2.3 | -37 |  |  | 0.2705 | 0.2273 | -0.001 | 0.99999 |  |  |  |  | 133 |  | 0 | [1170 empty] | 0
P | 2018-01-17T04:10:45+08:00 | 2018-01-18T10:00:27+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | A18USM771G | CPSM Final Test | FAIL | [312 empty] | 0
P | 2018-01-17T04:48:25+08:00 | 2018-01-18T10:01:12+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | A18USM767G | CPSM Final Test | FAIL | [12 empty] | 0.0018
P | 2018-01-17T19:05:36+08:00 | 2018-01-18T11:06:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | A18USM771G | CPSM Final Test | PASS |  | 156.11 | 0.37 | -0.31 | 56.72 | 57.66 | 1.82 | 0.13 | -1 | -1 | -1 | -0.912 | 0.5402 | 3.6 | 14 |  |  | 0.1857 | 0.3567 | -0.0121 | 1 |  |  |  |  | 133 |  | 0.00665 | [284 empty] | 0
P | 2018-01-18T07:52:12+08:00 | 2018-01-18T11:21:31+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | A18USM772G | CPSM Final Test | FAIL
P | 2018-01-17T02:24:15+08:00 | 2018-01-18T12:25:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | A18USM767G | CPSM Final Test | FAIL | [17 empty] | 4.9471
P | 2018-01-18T06:58:55+08:00 | 2018-01-18T12:35:55+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | A18USM773G | CPSM Final Test | FAIL | [27 empty] | 0.10964
P | 2018-01-18T09:40:45+08:00 | 2018-01-18T12:39:06+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | A18USM774G | CPSM Final Test | PASS |  | 160.61 | -0.02 | -0.46 | 56.69 | 57.15 | 1.21 | 0.13 | 0 | -9 | 1 | -0.854 | 0.498 | 2.8 | 19 |  |  | 0.1088 | 0.0721 | 0.0219 | 0.99999 |  |  |  |  | 133 |  | 0 | [1171 empty] | 0
P | 2018-01-18T10:03:54+08:00 | 2018-01-18T13:02:33+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | A18USM775G | CPSM Final Test | PASS |  | 161.69 | -0.02 | -0.46 | 56.61 | 53.09 | 1.06 | 0.13 | 0 | -12 | -1 | -0.848 | 0.539 | 0.4 | -7 |  |  | 0.0398 | 0.0833 | -0.0194 | 1 |  |  |  | 0 | 133 |  | 0
P | 2018-01-18T09:36:52+08:00 | 2018-01-18T13:27:59+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | A18USM772G | CPSM Final Test | PASS |  | 164.84 | -0.02 | -1.15 | 56.64 | 54.95 | 0.83 | 0.13 | 1 | 2 | 2 | -0.844 | 0.499 | 2.4 | 13 |  |  | 0.0598 | 0.293 | 0.0159 | 0.99999 |  |  |  |  | 133 |  | 0 | [46 empty] | 0
P | 2018-01-17T19:24:48+08:00 | 2018-01-18T13:32:48+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | A18USM767G | CPSM Final Test | PASS |  | 164.84 | 0.17 | -0.61 | 56.71 | 54.95 | 1.21 | 0.13 | -1 | -1 | 1 | -0.8879 | 0.4574 | 2.5 | 13 |  |  | 0.1593 | 0.2185 | -0.0585 | 0.99993 |  |  |  |  | 133 |  | 0.01499 | [798 empty] | 0
P | 2018-01-18T09:47:25+08:00 | 2018-01-18T13:39:05+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | A18USM773G | CPSM Final Test | PASS |  | 166.55 | 0.37 | -0.92 | 56.69 | 59.77 | 1.36 | 0.21 | 1 | 3 | 1 | -0.8499 | 0.5452 | 2.6 | 5 |  |  | 0.2361 | 0.4749 | -0.0096 | 1 |  |  |  |  | 133 |  | 0 | [830 empty] | 0
P | 2018-01-18T14:03:05+08:00 | 2018-01-18T16:49:50+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | A18USM779G | CPSM Final Test | PASS |  | 163.49 | -0.02 | -0.61 | 56.46 | 53 | 1.13 | 0.13 | 0 | -10 | -2 | -0.8799 | 0.5122 | 5.2 | 23 |  |  | 0.0747 | 0.1777 | -0.009 | 1 |  |  |  |  | 133 |  | 0 | [307 empty] | 0
P | 2018-01-18T13:58:52+08:00 | 2018-01-18T16:53:55+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | A18USM778G | CPSM Final Test | PASS |  | 164.3 | -0.02 | -0.46 | 57.01 | 53.59 | 1.36 | 0.13 | 1 | -2 | -1 | -0.8858 | 0.4944 | 5.2 | 15 |  |  | 0.0933 | 0.1715 | -0.0475 | 0.99998 |  |  |  |  | 133 |  | 0 | [1172 empty] | 0
P | 2018-01-18T10:26:07+08:00 | 2018-01-18T17:25:26+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | A18USM776G | CPSM Final Test | FAIL
P | 2018-01-18T13:52:26+08:00 | 2018-01-18T17:39:57+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | A18USM777G | CPSM Final Test | PASS |  | 165.11 | -0.02 | -0.46 | 56.79 | 54.19 | 1.06 | 0.13 | 1 | 2 | -1 | -0.872 | 0.4984 | 6.7 | -1 |  |  | 0.1147 | 0.2605 | -0.0153 | 0.99998 |  |  |  |  | 133 |  | 0.00736 | [284 empty] | 0
P | 2018-01-18T13:55:46+08:00 | 2018-01-18T18:31:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | A18USM776G | CPSM Final Test | PASS |  | 158.63 | -0.02 | -0.77 | 56.83 | 56.22 | 1.21 | 0.13 | 0 | -2 | 1 | -0.8899 | 0.496 | 0.2 | 16 |  |  | 0.0973 | 0.4277 | -0.0195 | 1 |  |  |  |  | 133 |  | 0 | [364 empty] | 0
P | 2018-01-18T09:05:59+08:00 | 2018-01-18T19:12:07+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | A18USM780G | CPSM Final Test | FAIL
P | 2018-01-18T18:33:56+08:00 | 2018-01-18T19:18:41+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI294G | CPSM Final Test | FAIL
P | 2018-01-18T14:09:03+08:00 | 2018-01-18T20:15:19+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | A18USM780G | CPSM Final Test | PASS |  | 163.58 | 0.27 | -0.92 | 56.73 | 56.9 | 1.36 | 0.13 | -1 | -2 | 0 | -0.9578 | 0.5018 | 3.1 | 23 |  |  | 0.1467 | 0.3837 | -0.0013 | 1 |  |  |  |  | 133 |  | 0.01453 | [1173 empty] | 0
P | 2018-01-18T17:33:40+08:00 | 2018-01-18T21:26:28+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI291G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.46 | 56.8 | 54.78 | 1.51 | 0.13 | 0 | 2 | 0 | -0.9197 | 0.539 | 6.8 | 11 |  | 0 | 0.1392 | 0.1446 | 0.0424 | 0.99999 | [382 empty] | 0
P | 2018-01-18T19:00:12+08:00 | 2018-01-18T21:44:14+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI296G | CPSM Final Test | PASS | 129 | 171.5 | 0.17 | -0.61 | 57.11 | 55.2 | 1.29 | 0.13 | -1 | 4 | 1 | -0.9198 | 0.4868 | 5.1 | 7 |  | 0 | 0.0648 | 0.213 | 0.0215 | 1 | [769 empty] | 0
P | 2018-01-18T18:56:18+08:00 | 2018-01-18T21:46:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI294G | CPSM Final Test | PASS | 129 | 165.38 | -0.21 | -0.46 | 57.41 | 53.93 | 1.29 | 0.13 | 0 | -2 | 1 | -0.9099 | 0.5428 | 4.2 | -6 |  | 0 | 0.0639 | 0.0857 | -0.0197 | 0.99999 | [1181 empty] | 0
P | 2018-01-18T18:53:21+08:00 | 2018-01-18T21:58:18+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI293G | CPSM Final Test | PASS | 129 | 161.69 | -0.3 | -0.99 | 56.84 | 54.27 | 1.29 | 0.13 | 0 | -2 | 1 | -0.9238 | 0.5112 | 4.4 | 5 |  | 0.00256 | 0.1049 | 0.2288 | -0.0221 | 0.99999 | [371 empty] | 0
P | 2018-01-18T03:28:29+08:00 | 2018-01-19T10:04:23+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI292G | CPSM Final Test | FAIL
P | 2018-01-18T18:46:26+08:00 | 2018-01-19T11:07:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI292G | CPSM Final Test | PASS | 129 | 171.23 | -0.02 | -0.61 | 56.8 | 54.36 | 1.36 | 0.13 | 1 | -2 | 1 | -0.9279 | 0.4972 | 1.7 | 11 |  | 0.00956 | 0.0594 | 0.1403 | -0.0073 | 1 | [1182 empty] | 0
P | 2018-01-19T09:40:38+08:00 | 2018-01-19T12:28:11+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI297G | CPSM Final Test | PASS | 129 | 160.34 | -0.02 | -0.69 | 57.03 | 57.23 | 1.21 | 0.13 | -1 | -2 | 1 | -0.888 | 0.5242 | 4.8 | 3 |  | 0.01529 | 0.0545 | 0.392 | 0.0317 | 0.99999 | [556 empty] | 0
P | 2018-01-19T09:42:00+08:00 | 2018-01-19T12:28:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI298G | CPSM Final Test | PASS | 129 | 158.27 | -0.02 | -0.61 | 57.01 | 53.42 | 1.36 | 0.13 | 0 | -10 | -1 | -0.8404 | 0.5028 | 9.5 | -23 |  | 0 | 0.1135 | 0.0969 | 0.0059 | 0.99997 | [1183 empty] | 0
P | 2018-01-19T09:35:02+08:00 | 2018-01-19T12:32:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI295G | CPSM Final Test | PASS | 129 | 157.64 | 0.27 | -1.3 | 56.72 | 59.44 | 1.29 | 0.13 | 0 | -2 | 0 | -0.8081 | 0.5042 | 5.3 | 2 |  | 0 | 0.2487 | 0.2052 | 0.0323 | 0.99997 | [754 empty] | 0
P | 2018-01-19T09:11:14+08:00 | 2018-01-19T13:50:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI299G | CPSM Final Test | FAIL
P | 2018-01-19T07:48:36+08:00 | 2018-01-19T15:13:32+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI299G | CPSM Final Test | FAIL
P | 2018-01-19T11:35:22+08:00 | 2018-01-19T15:20:05+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI300G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | -0.46 | 56.91 | 51.05 | 1.36 | 0.13 | 1 | -2 | 0 | -0.8401 | 0.487 | 6.8 | 10 |  | 0 | 0.0756 | 0.1248 | -0.0423 | 1 | [744 empty] | 0
P | 2018-01-19T13:51:50+08:00 | 2018-01-19T16:36:50+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI303G | CPSM Final Test | PASS | 129 | 167.63 | -0.02 | -0.77 | 57 | 56.73 | 1.44 | 0.13 | 0 | -3 | 1 | -0.8081 | 0.55 | 0.8 | 7 |  | 0.00919 | 0.1032 | 0.3059 | 0.0133 | 0.99996 | [662 empty] | 0
P | 2018-01-19T13:58:53+08:00 | 2018-01-19T16:44:59+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI302G | CPSM Final Test | PASS | 129 | 156.38 | 0.27 | -0.99 | 56.92 | 55.88 | 1.82 | 0.13 | -1 | 2 | 1 | -0.8819 | 0.5204 | 7 | 59 |  | 0.01106 | 0.1495 | 0.8498 | 0.0107 | 0.99999 | [388 empty] | 0
P | 2018-01-19T13:47:18+08:00 | 2018-01-19T16:47:11+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI301G | CPSM Final Test | PASS | 129 | 163.85 | -0.02 | -0.61 | 56.9 | 55.54 | 1.44 | 0.13 | 0 | 2 | 1 | -0.906 | 0.489 | 2 | 2 |  | 0 | 0.0915 | 0.1022 | -0.0217 | 0.99998 | [1184 empty] | 0
P | 2018-01-19T11:31:04+08:00 | 2018-01-19T17:16:44+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI299G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.77 | 56.81 | 53.42 | 1.21 | 0.13 | 1 | -12 | -1 | -0.8579 | 0.5306 | 3.3 | 21 |  | 0 | 0.0403 | 0.2899 | -0.0115 | 0.99999 | [1185 empty] | 0
P | 2018-01-19T17:26:37+08:00 | 2018-01-19T20:54:23+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI305G | CPSM Final Test | PASS | 129 | 158.63 | 0.27 | -0.61 | 56.83 | 55.2 | 1.51 | 0.13 | 0 | 6 | 0 | -0.9298 | 0.549 | 7.9 | -15 |  | 0.02003 | 0.08 | 0.1185 | -0.0268 | 0.99999 | [836 empty] | 0
P | 2018-01-19T17:30:08+08:00 | 2018-01-19T20:57:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI304G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.84 | 56.77 | 55.29 | 1.44 | 0.13 | 0 | -12 | -1 | -0.9279 | 0.538 | 2.6 | 3 |  | 0 | 0.082 | 0.0349 | -0.0074 | 1 | [790 empty] | 0
P | 2018-01-19T14:06:32+08:00 | 2018-01-19T21:09:19+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI308G | CPSM Final Test | FAIL
P | 2018-01-19T17:37:55+08:00 | 2018-01-20T08:56:17+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI308G | CPSM Final Test | PASS | 129 | 158.54 | -0.02 | -0.69 | 56.92 | 51.99 | 1.44 | 0.13 | 0 | -11 | 1 | -0.9836 | 0.5052 | 0.3 | 16 |  | 0.01149 | 0.268 | 0.43 | -0.0266 | 0.99999 | [1186 empty] | 0
P | 2018-01-19T14:18:40+08:00 | 2018-01-19T20:26:49+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI306G | CPSM Final Test | FAIL
P | 2018-01-19T00:49:16+08:00 | 2018-01-20T09:56:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI306G | CPSM Final Test | FAIL
P | 2018-01-19T17:22:45+08:00 | 2018-01-20T10:08:45+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI306G | CPSM Final Test | PASS | 129 | 166.01 | -0.02 | -0.54 | 56.73 | 55.54 | 1.36 | 0.13 | 0 | 2 | 3 | -0.9279 | 0.501 | 3.6 | 9 |  | 0.00838 | 0.0723 | 0.0555 | -0.0065 | 0.99999 | [818 empty] | 0
P | 2018-01-20T06:58:34+08:00 | 2018-01-20T11:36:26+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI310G | CPSM Final Test | FAIL
P | 2018-01-20T09:20:29+08:00 | 2018-01-20T12:20:21+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI309G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.54 | 57.22 | 56.73 | 1.21 | 0.13 | 1 | -2 | 1 | -0.864 | 0.5102 | 1.5 | 11 |  | 0 | 0.0621 | 0.0604 | -0.0297 | 1 | [1142 empty] | 0
P | 2018-01-20T09:24:02+08:00 | 2018-01-20T12:21:31+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI307G | CPSM Final Test | PASS | 129 | 167.63 | -0.02 | -0.46 | 56.92 | 50.72 | 1.36 | 0.13 | 0 | 3 | -1 | -0.848 | 0.5138 | 10.3 | 2 |  | 0 | 0.0562 | 0.1362 | -0.0115 | 0.99999 | [718 empty] | 0
P | 2018-01-20T08:46:50+08:00 | 2018-01-20T13:01:47+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI312G | CPSM Final Test | FAIL
P | 2018-01-20T09:17:30+08:00 | 2018-01-20T13:47:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI310G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.54 | 56.94 | 51.14 | 1.29 | 0.13 | 1 | 2 | 0 | -0.8579 | 0.5288 | 6.8 | -3 |  | 0 | 0.0693 | 0.2717 | 0.0297 | 0.99999 | [1185 empty] | 0
P | 2018-01-20T10:47:53+08:00 | 2018-01-20T14:52:44+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI311G | CPSM Final Test | PASS | 129 | 163.4 | -0.3 | -0.31 | 57.11 | 56.39 | 1.82 | 0.13 | 0 | -2 | 3 | -0.7302 | 0.492 | 4.2 | 6 |  | 0 | 0.1679 | 0.183 | 0.0282 | 0.99999 | [157 empty] | 0
P | 2018-01-20T07:44:12+08:00 | 2018-01-20T15:13:19+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI314G | CPSM Final Test | FAIL
P | 2018-01-20T07:14:36+08:00 | 2018-01-20T15:42:55+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI314G | CPSM Final Test | FAIL
P | 2018-01-20T05:06:16+08:00 | 2018-01-20T16:42:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI312G | CPSM Final Test | FAIL | [1208 empty] | 0
P | 2018-01-20T11:28:45+08:00 | 2018-01-20T16:46:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI314G | CPSM Final Test | PASS | 129 | 163.49 | -0.3 | 0.07 | 56.94 | 53.26 | 1.29 | 0.13 | 0 | 1 | 0 | -0.846 | 0.53 | 5.7 | 52 |  | 0.01198 | 0.0373 | 0.1124 | 0.0162 | 0.99997 | [818 empty] | 0
P | 2018-01-20T10:54:18+08:00 | 2018-01-20T17:48:33+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI312G | CPSM Final Test | PASS | 129 | 162.41 | -0.02 | -0.69 | 56.76 | 54.19 | 1.21 | 0.13 | 0 | -11 | -1 | -0.9058 | 0.49 | 9.8 | 3 |  | 0.01636 | 0.1201 | 0.1928 | 0.0299 | 0.99995 | [1187 empty] | 0
P | 2018-01-20T15:17:30+08:00 | 2018-01-20T18:09:28+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI318G | CPSM Final Test | PASS | 129 | 161.87 | -0.02 | -0.84 | 56.8 | 53.59 | 1.06 | 0.13 | -1 | 2 | -1 | -0.9218 | 0.5054 | 1.2 | 64 |  | 0 | 0.0687 | 0.2253 | 0.0045 | 1 | [750 empty] | 0
P | 2018-01-20T15:20:09+08:00 | 2018-01-20T18:09:21+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI317G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.23 | 57.04 | 53.09 | 1.44 | 0.13 | -1 | -12 | 1 | -0.9158 | 0.5006 | 4 | 77 |  | 0 | 0.1324 | 0.1325 | -0.0082 | 0.99999 | [230 empty] | 0
P | 2018-01-20T15:23:56+08:00 | 2018-01-20T18:13:29+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI316G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -0.54 | 57.41 | 56.47 | 1.59 | 0.13 | 0 | 6 | 3 | -0.8779 | 0.4894 | 4.4 | 90 |  | 0.02738 | 0.1042 | 0.1656 | -0.0121 | 0.99999 | [289 empty] | 0
P | 2018-01-20T16:24:47+08:00 | 2018-01-20T19:02:36+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI320G | CPSM Final Test | PASS | 129 | 162.86 | 0.17 | -0.31 | 56.83 | 53.09 | 1.59 | 0.13 | 0 | 2 | 2 | -0.8819 | 0.5204 | 8.1 | 15 |  | 0 | 0.1867 | 0.1027 | -0.0138 | 0.99996 | [1188 empty] | 0
P | 2018-01-21T07:01:53+08:00 | 2018-01-21T12:07:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI321G | CPSM Final Test | FAIL | [1210 empty] | 3359
P | 2018-01-21T06:50:53+08:00 | 2018-01-21T12:11:15+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI322G | CPSM Final Test | FAIL
P | 2018-01-21T06:47:11+08:00 | 2018-01-21T12:14:57+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI322G | CPSM Final Test | FAIL
P | 2018-01-21T09:28:32+08:00 | 2018-01-21T12:36:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI323G | CPSM Final Test | PASS | 129 | 158.45 | -0.02 | -0.54 | 56.79 | 55.8 | 1.21 | 0.13 | 0 | -3 | -2 | -0.796 | 0.5124 | 2.8 | 69 |  | 0 | 0.2069 | 0.1739 | -0.0097 | 0.99999 | [1190 empty] | 0
P | 2018-01-21T09:38:29+08:00 | 2018-01-21T12:41:43+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI315G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -0.61 | 57.23 | 53.93 | 1.36 | 0.13 | 1 | -1 | 1 | -0.84 | 0.5078 | 1.3 | 1 |  | 0 | 0.0878 | 0.2572 | 0.0337 | 0.99999 | [38 empty] | 0
P | 2018-01-21T09:44:50+08:00 | 2018-01-21T12:51:29+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI313G | CPSM Final Test | PASS | 129 | 161.24 | 0.17 | -1.07 | 56.8 | 53.17 | 1.44 | 0.13 | 0 | -2 | -1 | -0.8619 | 0.5024 | 1.7 | -6 |  | 0.012 | 0.1857 | 0.2822 | 0.0332 | 0.99999 | [1187 empty] | 0
P | 2018-01-21T09:34:49+08:00 | 2018-01-21T13:13:37+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI321G | CPSM Final Test | PASS | 129 | 162.77 | -0.59 | -0.31 | 56.8 | 55.29 | 1.13 | 0.13 | 0 | -10 | 3 | -0.846 | 0.495 | 9.1 | 16 |  | 0.00355 | 0.0985 | 0.1049 | -0.0327 | 1 | [1189 empty] | 0
P | 2018-01-21T09:31:04+08:00 | 2018-01-21T13:20:33+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI322G | CPSM Final Test | PASS | 129 | 168.44 | 0.46 | 0.15 | 57.52 | 56.05 | 1.74 | 0.13 | 0 | 2 | -1 | -0.884 | 0.5296 | 5.1 | 0 |  | 0 | 0.0384 | 0.1255 | -0.0481 | 1 | [200 empty] | 0
P | 2018-01-21T06:02:07+08:00 | 2018-01-21T13:29:54+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI319G | CPSM Final Test | FAIL
P | 2018-01-21T09:46:01+08:00 | 2018-01-21T13:30:26+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI319G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.46 | 56.82 | 53.85 | 1.51 | 0.13 | 0 | -3 | -1 | -0.7662 | 0.5484 | 2.4 | 5 |  | 0 | 0.1466 | 0.1291 | 0.0363 | 0.99998 | [89 empty] | 0
P | 2018-01-21T14:22:52+08:00 | 2018-01-21T17:17:44+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI327G | CPSM Final Test | PASS | 129 | 166.01 | 0.17 | -0.31 | 57.2 | 54.61 | 1.36 | 0.13 | 0 | 5 | 1 | -0.8101 | 0.5022 | 11.9 | 1 |  | 0 | 0.1258 | 0.2304 | 0.0129 | 0.99998 | [280 empty] | 0
P | 2018-01-21T14:21:46+08:00 | 2018-01-21T17:24:01+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI326G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.31 | 56.9 | 55.63 | 1.51 | 0.13 | 0 | -1 | 0 | -0.83 | 0.5102 | 4 | -9 |  | 0.01132 | 0.1184 | 0.1184 | -0.0162 | 1 | [124 empty] | 0
P | 2018-01-21T14:15:24+08:00 | 2018-01-21T17:23:31+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI329G | CPSM Final Test | PASS | 129 | 167.54 | -0.52 | -0.38 | 56.91 | 51.22 | 1.44 | 0.13 | 1 | -2 | 1 | -0.7762 | 0.5486 | 4.7 | 32 |  | 0 | 0.0397 | 0.1999 | 0.0075 | 0.99999 | [212 empty] | 0
P | 2018-01-21T14:11:53+08:00 | 2018-01-21T17:26:23+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI328G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | 0 | 57.01 | 53.59 | 1.21 | 0.13 | 0 | 2 | 1 | -0.8919 | 0.507 | 5.5 | 6 |  | 0 | 0.0884 | 0.0668 | -0.0091 | 1 | [1189 empty] | 0
P | 2018-01-21T14:08:37+08:00 | 2018-01-21T17:35:01+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI324G | CPSM Final Test | PASS | 129 | 160.61 | 0.78 | -0.31 | 57.21 | 55.12 | 1.29 | 0.13 | 1 | -2 | -1 | -0.862 | 0.4974 | 4.8 | 17 |  | 0 | 0.0986 | 0.165 | 0.0162 | 1 | [200 empty] | 0
P | 2018-01-21T14:05:47+08:00 | 2018-01-21T17:43:04+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI325G | CPSM Final Test | PASS | 129 | 161.15 | -0.02 | -0.84 | 56.21 | 51.9 | 1.21 | 0.13 | -1 | 2 | 1 | -0.7942 | 0.5048 | 3 | 5 |  | 0 | 0.1694 | 0.1664 | -0.015 | 1 | [443 empty] | 0
P | 2018-01-22T10:02:19+08:00 | 2018-01-22T12:54:45+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI334G | CPSM Final Test | PASS | 129 | 159.71 | -0.02 | -0.31 | 57.31 | 54.61 | 1.36 | 0.21 | -1 | 4 | 1 | -0.874 | 0.5014 | 14 | 11 |  | 0 | 0.0579 | 0.1913 | 0.0201 | 0.99999 | [25 empty] | 0
P | 2018-01-22T09:49:19+08:00 | 2018-01-22T12:58:40+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI332G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -0.69 | 56.93 | 53.68 | 1.44 | 0.13 | 0 | -11 | -1 | -0.8081 | 0.5304 | 2.9 | 38 |  | 0 | 0.2006 | 0.1531 | 0.0252 | 0.99998 | [1191 empty] | 0
P | 2018-01-22T10:08:54+08:00 | 2018-01-22T13:04:19+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI335G | CPSM Final Test | PASS | 129 | 168.17 | 0.37 | -0.54 | 56.92 | 54.95 | 1.36 | 0.13 | 1 | 2 | 0 | -0.9258 | 0.4916 | 12.6 | 26 |  | 0.0115 | 0.1594 | 0.1617 | -0.0033 | 0.99998 | [363 empty] | 0
P | 2018-01-22T10:17:29+08:00 | 2018-01-22T13:04:15+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI330G | CPSM Final Test | PASS | 129 | 161.51 | -0.02 | 0 | 57.12 | 55.54 | 1.51 | 0.13 | 1 | 4 | 1 | -0.8979 | 0.4994 | 5.9 | 2 |  | 0 | 0.042 | 0.2197 | -0.022 | 0.99999 | [856 empty] | 0
P | 2018-01-22T05:59:03+08:00 | 2018-01-22T13:21:40+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI333G | CPSM Final Test | FAIL
P | 2018-01-22T03:36:25+08:00 | 2018-01-22T15:44:18+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI333G | CPSM Final Test | FAIL
P | 2018-01-22T03:16:13+08:00 | 2018-01-22T16:04:30+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI333G | CPSM Final Test | FAIL | [16 empty] | 0
P | 2018-01-22T13:31:41+08:00 | 2018-01-22T16:24:21+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI336G | CPSM Final Test | PASS | 129 | 171.77 | 0.46 | -0.61 | 56.69 | 56.64 | 1.44 | 0.13 | 1 | 3 | -1 | -0.8818 | 0.4978 | 4.4 | -1 |  | 0.00556 | 0.1151 | 0.3201 | 0.015 | 1 | [141 empty] | 0
P | 2018-01-22T09:40:22+08:00 | 2018-01-22T17:07:24+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI333G | CPSM Final Test | PASS | 129 | 168.44 | -0.21 | -0.54 | 57.05 | 53.93 | 1.9 | 0.13 | -1 | -11 | 0 | -0.7682 | 0.539 | 0.8 | 12 |  | 0 | 0.1253 | 0.3237 | 0.0664 | 0.99996 | [1192 empty] | 0
P | 2018-01-22T16:51:41+08:00 | 2018-01-22T20:07:17+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI331G | CPSM Final Test | FAIL | [14 empty] | 1263
P | 2018-01-22T14:18:18+08:00 | 2018-01-23T00:36:45+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI336G | CPSM Final Test | FAIL | [903 empty] | 0
P | 2018-01-22T05:36:20+08:00 | 2018-01-23T09:18:43+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI336G | CPSM Final Test | FAIL
P | 2018-01-22T02:27:13+08:00 | 2018-01-23T10:31:45+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI331G | CPSM Final Test | FAIL
P | 2018-01-22T04:19:50+08:00 | 2018-01-23T10:35:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI336G | CPSM Final Test | FAIL
P | 2018-01-22T19:27:31+08:00 | 2018-01-23T11:38:15+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI336G | CPSM Final Test | PASS | 129 | 171.77 | 0.46 | -0.54 | 56.75 | 53.34 | 1.36 | 0.13 | 1 | -2 | 0 | -0.9417 | 0.48 | 2.8 | -9 |  | 0.01714 | 0.0524 | 0.0766 | -0.0084 | 0.99995 | [850 empty] | 0
P | 2018-01-22T18:29:29+08:00 | 2018-01-23T12:16:23+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI331G | CPSM Final Test | PASS | 129 | 157.73 | 0.27 | -0.46 | 56.8 | 55.71 | 1.51 | 0.13 | -1 | -2 | 0 | -1.0216 | 0.4502 | 1 | 7 |  | 0 | 0.0289 | 0.1691 | -0.0071 | 0.99999 | [827 empty] | 0
P | 2018-01-24T09:26:15+08:00 | 2018-01-24T10:21:19+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM785G | CPSM Final Test | FAIL | [5 empty] | 34.46
P | 2018-01-24T09:14:44+08:00 | 2018-01-24T12:10:41+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | A18USM782G | CPSM Final Test | PASS |  | 164.21 | 0.08 | -0.38 | 56.72 | 57.07 | 1.29 | 0.13 | 0 | -2 | 1 | -0.826 | 0.5066 | 0.5 | 29 |  |  | 0.1039 | 0.0549 | -0.0368 | 1 |  |  |  |  | 133 |  | 0 | [820 empty] | 0
P | 2018-01-24T09:06:56+08:00 | 2018-01-24T12:16:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM783G | CPSM Final Test | PASS |  | 162.14 | 0.46 | -0.38 | 56.75 | 55.03 | 1.59 | 0.13 | 0 | -3 | -1 | -0.8878 | 0.5166 | 6.8 | 9 |  |  | 0.2741 | 0.2235 | -0.0686 | 1 |  |  |  |  | 133 |  | 0 | [834 empty] | 0
P | 2018-01-24T06:29:53+08:00 | 2018-01-24T12:22:51+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM781G | CPSM Final Test | FAIL
P | 2018-01-24T09:26:22+08:00 | 2018-01-24T12:23:33+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM781G | CPSM Final Test | PASS |  | 174.38 | 0.56 | -0.61 | 56.47 | 50.38 | 1.06 | 0.13 | 1 | 4 | 1 | -0.8461 | 0.4838 | 2.6 | 13 |  |  | 0.0499 | 0.4618 | -0.0447 | 0.99999 |  |  |  |  | 133 |  | 0.01158 | [1155 empty] | 0
P | 2018-01-24T08:59:55+08:00 | 2018-01-24T12:33:58+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM784G | CPSM Final Test | PASS |  | 161.87 | 0.17 | 0.15 | 57.2 | 55.8 | 1.74 | 0.13 | 0 | -1 | 1 | -0.852 | 0.4706 | 5 | 21 |  |  | 0.3781 | 0.1706 | -0.0152 | 1 |  |  |  |  | 133 |  | 0 | [121 empty] | 0
P | 2018-01-24T10:25:52+08:00 | 2018-01-24T13:23:03+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | A18USM786G | CPSM Final Test | PASS |  | 165.11 | 0.59 | -1.07 | 56.83 | 54.19 | 0.98 | 0.13 | 0 | -3 | 3 | -0.8201 | 0.5442 | 2.7 | 22 |  |  | 0.2873 | 0.1338 | 0.0315 | 0.99999 |  |  |  |  | 133 |  | 0 | [865 empty] | 0
P | 2018-01-24T09:53:47+08:00 | 2018-01-24T13:26:05+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM785G | CPSM Final Test | PASS |  | 162.05 | -0.11 | -0.54 | 57.06 | 52.66 | 1.06 | 0.13 | 0 | 3 | 0 | -0.8839 | 0.5266 | 0.6 | 3 |  |  | 0.1582 | 0.2393 | -0.0039 | 0.99999 |  |  |  |  | 133 |  | 0 | [1058 empty] | 0
P | 2018-01-24T15:14:55+08:00 | 2018-01-24T18:18:17+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | A18USM787G | CPSM Final Test | PASS |  | 166.01 | -0.81 | -0.38 | 57 | 54.44 | 1.44 | 0.13 | -1 | 3 | -2 | -0.8281 | 0.5332 | 6.5 | 5 |  |  | 0.0334 | 0.1444 | -0.0565 | 1 |  |  |  |  | 133 |  | 0 | [601 empty] | 0
P | 2018-01-24T15:21:51+08:00 | 2018-01-24T18:23:05+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM789G | CPSM Final Test | PASS |  | 162.14 | 0.17 | -0.38 | 56.81 | 55.03 | 1.82 | 0.13 | -1 | -2 | 1 | -0.8939 | 0.5036 | 2.7 | -1 |  |  | 0.0655 | 0.1538 | -0.0644 | 0.99999 |  |  |  |  | 133 |  | 0 | [220 empty] | 0
P | 2018-01-24T15:31:22+08:00 | 2018-01-24T18:38:02+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM790G | CPSM Final Test | PASS |  | 171.86 | -0.02 | -0.46 | 56.6 | 53.93 | 1.21 | 0.13 | 1 | -2 | 3 | -0.9019 | 0.547 | 11.8 | 11 |  |  | 0.0686 | 0.1895 | -0.0188 | 0.99999 |  |  |  |  | 133 |  | 0 | [592 empty] | 0
P | 2018-01-24T15:45:40+08:00 | 2018-01-24T18:59:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | A18USM791G | CPSM Final Test | PASS |  | 166.73 | 0.17 | -0.46 | 56.83 | 52.75 | 1.67 | 0.13 | 0 | -2 | 1 | -0.9339 | 0.509 | 6.6 | -16 |  |  | 0.1087 | 0.3712 | -0.0043 | 0.99999 |  |  |  |  | 133 |  | 0 | [1186 empty] | 0
P | 2018-01-24T16:23:51+08:00 | 2018-01-24T19:19:59+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | A18USM792G | CPSM Final Test | PASS |  | 166.64 | 0.37 | -0.69 | 56.62 | 51.31 | 1.51 | 0.13 | 1 | -2 | -1 | -0.9358 | 0.5208 | 1.3 | -17 |  |  | 0.2429 | 0.1664 | 0.0353 | 0.9999 |  |  |  |  | 133 |  | 0.02655 | [692 empty] | 0
P | 2018-01-24T15:55:50+08:00 | 2018-01-24T19:32:57+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | A18USM788G | CPSM Final Test | PASS |  | 161.15 | 0.27 | 0.84 | 57.31 | 56.13 | 2.58 | 0.13 | 0 | 2 | 3 | -0.9558 | 0.5284 | 1.7 | 10 |  |  | 0.0974 | 0.8719 | -0.0417 | 0.99999 |  |  |  |  | 133 |  | 0.01079 | [1187 empty] | 0
P | 2018-01-24T18:50:28+08:00 | 2018-01-24T19:52:46+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM794G | CPSM Final Test | FAIL | [83 empty] | FALSE
P | 2018-01-24T02:48:21+08:00 | 2018-01-25T10:37:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM793G | CPSM Final Test | FAIL | [27 empty] | 0.06002
P | 2018-01-24T19:27:02+08:00 | 2018-01-25T11:31:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI337G | CPSM Final Test | PASS | 129 | 165.38 | 0.37 | -1.07 | 56.7 | 60.37 | 1.29 | 0.13 | 1 | -2 | -1 | -0.8399 | 0.494 | 4.2 | -46 |  | 0 | 0.1791 | 0.1077 | 0.0179 | 1 | [89 empty] | 0
P | 2018-01-24T19:21:37+08:00 | 2018-01-25T11:37:28+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | A18USM794G | CPSM Final Test | PASS |  | 163.22 | -0.02 | -0.46 | 57.09 | 51.9 | 1.9 | 0.17 | 1 | 4 | 1 | -0.82 | 0.544 | 2 | 49 |  |  | 0.1341 | 0.1355 | 0.017 | 0.99999 |  |  |  |  | 133 |  | 0 | [697 empty] | 0
P | 2018-01-24T18:43:06+08:00 | 2018-01-25T11:47:25+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | A18USM793G | CPSM Final Test | PASS |  | 166.82 | 0.17 | -0.54 | 56.9 | 53.59 | 1.51 | 0.13 | -1 | 2 | -1 | -0.9338 | 0.474 | 0.8 | 6 |  |  | 0.1199 | 0.0842 | -0.0294 | 1 |  |  |  |  | 133 |  | 0.01075 | [1188 empty] | 0
P | 2018-01-25T11:10:45+08:00 | 2018-01-25T14:15:45+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI339G | CPSM Final Test | PASS | 129 | 149.81 | 0.17 | -0.54 | 56.83 | 54.69 | 1.74 | 0.13 | 0 | 6 | 1 | -0.9098 | 0.535 | 7.2 | 6 |  | 0.01031 | 0.0714 | 0.1441 | -0.0321 | 0.99998 | [858 empty] | 0
P | 2018-01-25T11:23:46+08:00 | 2018-01-25T14:19:39+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI340G | CPSM Final Test | PASS | 129 | 168.98 | 0.27 | -0.84 | 57.23 | 55.71 | 1.29 | 0.13 | -1 | 3 | 1 | -0.9299 | 0.5302 | 1.8 | 15 |  | 0 | 0.0916 | 0.2234 | -0.0374 | 0.99999 | [447 empty] | 0
P | 2018-01-25T12:08:50+08:00 | 2018-01-25T15:41:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI342G | CPSM Final Test | PASS | 129 | 165.74 | 0.08 | -0.46 | 56.92 | 52.15 | 1.59 | 0.13 | 1 | 3 | -1 | -0.818 | 0.5062 | 8.2 | 12 |  | 0 | 0.1921 | 0.1385 | 0.0436 | 0.99997 | [1196 empty] | 0
P | 2018-01-25T12:23:42+08:00 | 2018-01-25T15:52:25+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI341G | CPSM Final Test | PASS | 129 | 165.92 | 0.17 | -0.84 | 56.83 | 53.85 | 1.44 | 0.13 | 1 | 6 | 0 | -0.854 | 0.489 | 1 | 69 |  | 0 | 0.0453 | 0.2032 | 0.0056 | 1 | [916 empty] | 0
P | 2018-01-25T13:46:45+08:00 | 2018-01-25T16:39:57+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI338G | CPSM Final Test | PASS | 129 | 166.01 | -0.3 | -0.99 | 56.4 | 54.69 | 2.05 | 0.13 | 0 | -2 | 1 | -0.8321 | 0.5422 | 1.8 | -27 |  | 0 | 0.1281 | 0.3089 | -0.0008 | 0.99999 | [1197 empty] | 0
P | 2018-01-25T13:31:29+08:00 | 2018-01-25T18:20:48+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI344G | CPSM Final Test | FAIL | [14 empty] | 626
P | 2018-01-25T15:35:30+08:00 | 2018-01-25T18:32:35+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI343G | CPSM Final Test | PASS | 129 | 164.48 | -0.02 | -0.69 | 57.05 | 54.53 | 1.21 | 0.13 | 0 | -2 | 1 | -0.898 | 0.5188 | 2.8 | 3 |  | 0 | 0.3351 | 0.3752 | -0.1427 | 1 | [89 empty] | 0
P | 2018-01-25T15:48:19+08:00 | 2018-01-25T18:34:11+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI345G | CPSM Final Test | PASS | 129 | 164.66 | 0.17 | -0.54 | 56.84 | 51.99 | 1.67 | 0.13 | -1 | -2 | 0 | -0.9238 | 0.5402 | 5.1 | 53 |  | 0.05005 | 0.1084 | 0.0706 | -0.0206 | 0.99999 | [555 empty] | 0
P | 2018-01-25T13:07:40+08:00 | 2018-01-25T19:21:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI346G | CPSM Final Test | FAIL
P | 2018-01-25T16:14:39+08:00 | 2018-01-25T19:22:10+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI346G | CPSM Final Test | PASS | 129 | 152.15 | -0.11 | -0.99 | 56.7 | 57.23 | 1.82 | 0.13 | -1 | 3 | 1 | -0.9258 | 0.4846 | 1.8 | -2 |  | 0 | 0.0921 | 0.2072 | 0.0334 | 0.99999 | [1198 empty] | 0
P | 2018-01-25T15:56:09+08:00 | 2018-01-25T20:05:12+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI344G | CPSM Final Test | PASS | 129 | 166.64 | 0.08 | -0.77 | 56.83 | 52.41 | 1.44 | 0.13 | 0 | 2 | 1 | -0.9478 | 0.5358 | 6.8 | 138 |  | 0 | 0.0494 | 0.2898 | -0.0083 | 0.99999 | [1199 empty] | 0
P | 2018-01-25T16:49:51+08:00 | 2018-01-25T20:22:21+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI348G | CPSM Final Test | PASS | 129 | 162.23 | -0.21 | -0.23 | 56.78 | 55.37 | 1.82 | 0.13 | 0 | -2 | 3 | -0.9238 | 0.5426 | 2.6 | 62 |  | 0 | 0.0484 | 0.1083 | -0.0183 | 0.99999 | [1161 empty] | 0
P | 2018-01-25T18:36:48+08:00 | 2018-01-25T21:21:53+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | A18SDI347G | CPSM Final Test | PASS | 129 | 164.12 | 0.08 | -0.99 | 56.8 | 55.54 | 1.36 | 0.13 | 1 | 4 | 1 | -0.9697 | 0.5146 | 2.7 | -37 |  | 0 | 0.1378 | 0.2046 | 0.0277 | 1 | [289 empty] | 0
P | 2018-01-25T18:17:28+08:00 | 2018-01-25T21:35:08+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI349G | CPSM Final Test | PASS | 129 | 165.02 | 0.49 | -1.07 | 56.76 | 55.71 | 1.36 | 0.13 | 1 | 3 | -1 | -0.9279 | 0.4746 | 1 | 5 |  | 0 | 0.0912 | 0.1106 | -0.0093 | 0.99998 | [351 empty] | 0
P | 2018-01-26T08:29:43+08:00 | 2018-01-26T14:15:17+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI350G | CPSM Final Test | FAIL | [558 empty] | FALSE
P | 2018-01-26T09:47:51+08:00 | 2018-01-26T15:32:14+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI356G | CPSM Final Test | FAIL | [14 empty] | 891
P | 2018-01-26T11:42:33+08:00 | 2018-01-26T15:31:13+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI355G | CPSM Final Test | PASS | 129 | 164.03 | 0.27 | -0.31 | 56.82 | 56.22 | 1.82 | 0.13 | 1 | 3 | -1 | -0.8879 | 0.5442 | 4 | 4 |  | 0 | 0.1135 | 0.0907 | -0.0318 | 1 | [51 empty] | 0
P | 2018-01-26T08:32:02+08:00 | 2018-01-26T15:47:55+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | A18SDI354G | CPSM Final Test | FAIL
P | 2018-01-26T11:32:37+08:00 | 2018-01-26T15:49:26+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI351G | CPSM Final Test | PASS | 129 | 166.73 | 0.27 | -0.99 | 56.83 | 53.09 | 0.91 | 0.13 | 1 | 2 | 1 | -0.852 | 0.455 | 4.5 | -4 |  | 0 | 0.4453 | 0.224 | -0.0501 | 1 | [1200 empty] | 0
P | 2018-01-26T11:56:47+08:00 | 2018-01-26T16:02:18+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI352G | CPSM Final Test | PASS | 129 | 164.75 | 0.49 | -0.31 | 56.92 | 53.76 | 1.74 | 0.13 | 0 | -2 | -1 | -0.834 | 0.4834 | 0.8 | 84 |  | 0 | 0.0772 | 0.3917 | 0.0111 | 1 | [814 empty] | 0
P | 2018-01-26T11:22:30+08:00 | 2018-01-26T16:17:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI350G | CPSM Final Test | PASS | 129 | 159.53 | 0.27 | -0.84 | 56.73 | 58.76 | 1.51 | 0.13 | -1 | -11 | -1 | -0.8581 | 0.4932 | 4 | 1 |  | 0.01147 | 0.0484 | 0.1114 | 0.0126 | 0.99999 | [146 empty] | 0
P | 2018-01-26T08:02:21+08:00 | 2018-01-26T17:17:44+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI356G | CPSM Final Test | FAIL
P | 2018-01-26T12:40:03+08:00 | 2018-01-26T17:18:14+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI356G | CPSM Final Test | PASS | 129 | 165.2 | 0.17 | -0.54 | 57.03 | 55.29 | 1.29 | 0.13 | -1 | -2 | 1 | -0.8301 | 0.477 | 1.1 | -1 |  | 0 | 0.0713 | 0.3022 | 0.0311 | 0.99998 | [993 empty] | 0
P | 2018-01-26T12:09:59+08:00 | 2018-01-26T17:32:29+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | A18SDI354G | CPSM Final Test | PASS | 129 | 159.44 | -0.02 | -0.38 | 56.8 | 53 | 1.59 | 0.13 | 1 | -4 | -1 | -0.9079 | 0.5162 | 4.4 | 18 |  | 0 | 0.0233 | 0.5055 | -0.0223 | 1 | [1157 empty] | 0
P | 2018-01-26T07:06:42+08:00 | 2018-01-26T17:39:24+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI353G | CPSM Final Test | FAIL | [242 empty] | 0
P | 2018-01-26T15:41:56+08:00 | 2018-01-26T18:29:01+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI357G | CPSM Final Test | PASS | 129 | 158.27 | 0.4 | -0.54 | 56.52 | 52.58 | 1.29 | 0.13 | 1 | -2 | 0 | -0.8939 | 0.536 | 2.1 | 17 |  | 0 | 0.134 | 0.2188 | 0.011 | 1 | [230 empty] | 0
P | 2018-01-26T12:23:03+08:00 | 2018-01-26T18:45:10+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI353G | CPSM Final Test | PASS | 129 | 167.72 | 0.46 | -0.99 | 56.75 | 54.44 | 1.21 | 0.17 | 1 | -11 | 1 | -0.8441 | 0.5166 | 3.5 | 1 |  | 0 | 0.1048 | 0.1464 | 0.0143 | 0.99999 | [221 empty] | 0
P | 2018-01-26T16:04:09+08:00 | 2018-01-26T19:10:06+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI358G | CPSM Final Test | PASS | 129 | 165.11 | 0.27 | 0.38 | 57.02 | 54.1 | 1.29 | 0.13 | 0 | 3 | 1 | -0.858 | 0.4936 | 0 | 5 |  | 0 | 0.1928 | 0.3946 | 0.0334 | 0.99995 | [416 empty] | 0
P | 2018-01-26T16:38:45+08:00 | 2018-01-26T20:27:33+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI359G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | -0.77 | 56.53 | 51.56 | 1.36 | 0.13 | -1 | -3 | 0 | -0.9179 | 0.502 | 1.8 | 26 |  | 0.01197 | 0.1432 | 0.1169 | -0.0224 | 0.99999 | [1201 empty] | 0
P | 2018-01-26T16:52:09+08:00 | 2018-01-26T20:36:30+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI360G | CPSM Final Test | PASS | 129 | 164.12 | 0.08 | -0.92 | 56.52 | 52.32 | 1.21 | 0.13 | 1 | 4 | 0 | -0.9018 | 0.515 | 0.3 | -7 |  | 0 | 0.0735 | 0.2338 | -0.0083 | 0.99999 | [1202 empty] | 0
P | 2018-01-26T18:18:14+08:00 | 2018-01-26T21:05:31+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | A18SDI362G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -1.3 | 56.21 | 54.1 | 1.36 | 0.13 | 0 | -2 | -1 | -0.9418 | 0.5272 | 0.8 | -51 |  | 0 | 0.0347 | 0.3326 | -0.0155 | 0.99999 | [187 empty] | 0
P | 2018-01-26T17:29:39+08:00 | 2018-01-26T21:18:36+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI361G | CPSM Final Test | PASS | 129 | 167.09 | 0.37 | -0.92 | 56.74 | 51.39 | 1.9 | 0.13 | 1 | -5 | -1 | -0.9418 | 0.466 | 3 | 15 |  | 0.02458 | 0.102 | 0.1823 | -0.0448 | 1 | [1145 empty] | 0
P | 2018-01-27T10:38:41+08:00 | 2018-01-27T13:34:35+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI363G | CPSM Final Test | PASS | 129 | 164.75 | -0.49 | -0.92 | 56.44 | 54.53 | 1.51 | 0.13 | 0 | -12 | -1 | -0.7323 | 0.4522 | 9.5 | 26 |  | 0.01426 | 0.1037 | 0.206 | 0.0118 | 1 | [1203 empty] | 0
P | 2018-01-27T10:48:29+08:00 | 2018-01-27T13:47:42+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI364G | CPSM Final Test | PASS | 129 | 160.16 | 0.08 | -0.77 | 56.51 | 53.42 | 1.59 | 0.16 | 0 | 2 | 0 | -0.7621 | 0.535 | 1.1 | -1 |  | 0 | 0.3114 | 0.105 | -0.0647 | 1 | [1204 empty] | 0
P | 2018-01-27T09:03:34+08:00 | 2018-01-27T14:37:59+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | A18SDI368G | CPSM Final Test | FAIL | [14 empty] | 741
P | 2018-01-27T07:51:58+08:00 | 2018-01-27T14:39:32+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI366G | CPSM Final Test | FAIL | [770 empty] | 0
P | 2018-01-27T07:16:15+08:00 | 2018-01-27T14:41:01+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI365G | CPSM Final Test | FAIL | [17 empty] | 1.1674
P | 2018-01-27T11:35:29+08:00 | 2018-01-27T15:03:54+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI367G | CPSM Final Test | PASS | 129 | 165.65 | 0.08 | -0.77 | 56.34 | 50.63 | 1.44 | 0.13 | 0 | -11 | 1 | -0.7942 | 0.523 | 4.2 | 26 |  | 0 | 0.0577 | 0.2447 | 0.0303 | 1 | [907 empty] | 0
P | 2018-01-27T10:58:38+08:00 | 2018-01-27T15:44:03+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI365G | CPSM Final Test | PASS | 129 | 157.82 | -0.02 | -0.77 | 56.53 | 53.42 | 1.67 | 0.13 | -2 | -2 | -1 | -0.7542 | 0.5148 | 1.6 | 9 |  | 0 | 0.051 | 0.059 | -0.0005 | 1 | [223 empty] | 0
P | 2018-01-27T06:45:38+08:00 | 2018-01-27T15:45:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI366G | CPSM Final Test | FAIL
P | 2018-01-27T11:15:45+08:00 | 2018-01-27T15:46:20+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI366G | CPSM Final Test | PASS | 129 | 166.91 | 0.17 | -0.61 | 56.31 | 55.12 | 1.51 | 0.13 | -1 | 2 | 4 | -0.7762 | 0.52 | 5.8 | 13 |  | 0.00851 | 0.0665 | 0.2631 | -0.0156 | 1 | [749 empty] | 0
P | 2018-01-27T11:50:46+08:00 | 2018-01-27T16:22:09+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | A18SDI368G | CPSM Final Test | PASS | 129 | 165.02 | 0.37 | -0.69 | 56.51 | 53.68 | 0.98 | 0.13 | 1 | -11 | 1 | -1.0077 | 0.5428 | 1.4 | 15 |  | 0 | 0.0641 | 0.2864 | -0.0192 | 0.99998 | [334 empty] | 0
P | 2018-01-27T13:47:53+08:00 | 2018-01-27T16:43:41+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI370G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.54 | 56.9 | 50.8 | 1.82 | 0.13 | -1 | 3 | -1 | -0.9877 | 0.52 | 0.3 | 13 |  | 0.01188 | 0.0481 | 0.1041 | 0.0082 | 1 | [1205 empty] | 0
P | 2018-01-27T09:43:01+08:00 | 2018-01-27T16:44:42+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI369G | CPSM Final Test | FAIL
P | 2018-01-27T14:16:00+08:00 | 2018-01-27T16:56:38+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI371G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.54 | 56.87 | 55.29 | 1.44 | 0.13 | 1 | 10 | 2 | -0.9538 | 0.525 | 2.5 | -6 |  | 0 | 0.0752 | 0.1703 | -0.0064 | 0.99999 | [302 empty] | 0
P | 2018-01-27T15:24:57+08:00 | 2018-01-27T18:22:32+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI372G | CPSM Final Test | PASS | 129 | 171.05 | 0.08 | -0.77 | 56.61 | 49.45 | 1.21 | 0.13 | -1 | -12 | 1 | -0.9877 | 0.46 | 1.8 | 16 |  | 0 | 0.1219 | 0.2143 | -0.0507 | 0.99999 | [1206 empty] | 0
P | 2018-01-27T16:06:44+08:00 | 2018-01-27T18:49:29+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI373G | CPSM Final Test | PASS | 129 | 164.3 | 0.27 | -0.69 | 56.8 | 52.32 | 1.44 | 0.13 | 1 | -2 | 0 | -0.9179 | 0.467 | 1.1 | 7 |  | 0 | 0.0567 | 0.0855 | -0.0225 | 1 | [401 empty] | 0
P | 2018-01-26T17:55:03+08:00 | 2018-01-28T08:32:40+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI369G | CPSM Final Test | FAIL
P | 2018-01-27T13:13:52+08:00 | 2018-01-28T08:33:16+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI369G | CPSM Final Test | PASS | 129 | 165.11 | 0.27 | -0.61 | 57 | 55.2 | 1.29 | 0.13 | 0 | -9 | -1 | -0.9837 | 0.522 | 0.1 | -6 |  | 0 | 0.0707 | 0.1834 | 0.036 | 0.99997 | [221 empty] | 0
P | 2018-01-28T11:10:36+08:00 | 2018-01-28T13:51:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | A18SDI376G | CPSM Final Test | PASS | 129 | 166.91 | 0.56 | -0.77 | 56.78 | 58.76 | 1.36 | 0.13 | -1 | -3 | -1 | -0.8261 | 0.5498 | 0.7 | 43 |  | 0 | 0.0674 | 0.055 | -0.0451 | 1 | [856 empty] | 0
P | 2018-01-28T11:30:01+08:00 | 2018-01-28T14:12:02+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | A18SDI377G | CPSM Final Test | PASS | 129 | 166.19 | 0.17 | -0.54 | 56.8 | 56.22 | 1.13 | 0.13 | 0 | 2 | -1 | -0.836 | 0.524 | 4.8 | -10 |  | 0 | 0.0696 | 0.2871 | -0.0099 | 1 | [749 empty] | 0
P | 2018-01-28T10:47:59+08:00 | 2018-01-28T14:20:49+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | A18SDI374G | CPSM Final Test | PASS | 129 | 171.68 | 0.37 | -0.77 | 56.47 | 55.46 | 1.29 | 0.13 | -1 | -1 | 1 | -0.8301 | 0.4914 | 0.6 | 25 |  | 0.01263 | 0.2114 | 0.1892 | 0.0111 | 0.99996 | [1207 empty] | 0
P | 2018-01-28T10:56:24+08:00 | 2018-01-28T14:21:50+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | A18SDI375G | CPSM Final Test | PASS | 129 | 156.65 | -0.3 | -0.46 | 56.77 | 54.1 | 0.98 | 0.13 | 1 | 2 | -1 | -0.8042 | 0.5194 | 1.5 | -1 |  | 0 | 0.0844 | 0.0913 | 0.01 | 0.99999 | [200 empty] | 0
P | 2018-01-28T12:22:22+08:00 | 2018-01-28T15:51:46+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | A18SDI379G | CPSM Final Test | PASS | 129 | 163.85 | 0.49 | -0.61 | 56.78 | 54.53 | 1.29 | 0.13 | -1 | -4 | -1 | -0.8381 | 0.519 | 0.4 | 13 |  | 0 | 0.0638 | 0.2976 | -0.0106 | 1 | [1208 empty] | 0
P | 2018-01-28T09:10:51+08:00 | 2018-01-28T15:51:22+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI378G | CPSM Final Test | FAIL
P | 2018-01-28T12:31:06+08:00 | 2018-01-28T15:58:00+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | A18SDI378G | CPSM Final Test | PASS | 129 | 165.74 | 0.08 | -0.77 | 56.54 | 51.65 | 1.51 | 0.13 | -1 | -3 | -1 | -0.8461 | 0.48 | 2.8 | -5 |  | 0 | 0.1721 | 0.1171 | -0.0012 | 0.99999 | 0
P | 2018-01-30T14:43:10+08:00 | 2018-01-30T17:36:24+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | A18USM797G | CPSM Final Test | PASS |  | 166.01 | -0.02 | -0.61 | 56.81 | 54.1 | 1.29 | 0.13 | 1 | -4 | 0 | -0.87 | 0.4858 | 3.3 | 10 |  |  | 0.1423 | 0.087 | -0.0065 | 0.99999 |  |  |  |  | 133 |  | 0 | [220 empty] | 0
P | 2018-01-30T14:34:24+08:00 | 2018-01-30T17:38:45+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | A18USM798G | CPSM Final Test | PASS |  | 167.63 | 0.27 | -0.69 | 56.61 | 54.44 | 1.06 | 0.13 | 1 | -3 | -1 | -0.886 | 0.5046 | 6.6 | -1 |  |  | 0.1346 | 0.1683 | -0.0454 | 0.99998 |  |  |  |  | 133 |  | 0.01482 | [1202 empty] | 0
P | 2018-01-30T14:47:45+08:00 | 2018-01-30T17:41:43+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | A18USM796G | CPSM Final Test | PASS |  | 167.36 | -0.02 | 0 | 56.78 | 51.73 | 1.44 | 0.13 | 0 | 2 | 1 | -0.8999 | 0.506 | 5.8 | 2 |  |  | 0.0697 | 0.1089 | -0.0193 | 0.99999 |  |  |  |  | 133 |  | 0 | [576 empty] | 0
P | 2018-01-30T14:52:20+08:00 | 2018-01-30T17:44:00+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | A18USM795G | CPSM Final Test | PASS |  | 168.53 | 0.46 | -0.31 | 56.75 | 55.88 | 1.44 | 0.13 | 1 | -2 | -1 | -0.854 | 0.503 | 0.1 | 0 |  |  | 0.0505 | 0.2639 | 0.0024 | 0.99999 |  |  |  |  | 133 |  | 0 | [395 empty] | 0
P | 2018-01-30T16:30:44+08:00 | 2018-01-30T19:38:00+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | A18USM799G | CPSM Final Test | PASS |  | 171.68 | -0.02 | -0.08 | 57.09 | 55.71 | 1.82 | 0.13 | 0 | -12 | 0 | -0.9918 | 0.5036 | 5.1 | 36 |  |  | 0.0328 | 0.2225 | 0.0198 | 0.99999 |  |  |  |  | 133 |  | 0 | [772 empty] | 0
P | 2018-01-30T16:32:14+08:00 | 2018-01-30T20:38:23+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | A18USM800G | CPSM Final Test | PASS |  | 158.72 | 0.59 | -0.54 | 56.64 | 54.86 | 1.59 | 0.13 | 1 | 2 | -1 | -1.0077 | 0.4962 | 1.2 | 7 |  |  | 0.0852 | 0.0829 | -0.0005 | 1 |  |  |  |  | 133 |  | 0 | [1203 empty] | 0
P | 2018-01-29T21:18:47+08:00 | 2018-01-31T11:55:50+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | A18USM801G | CPSM Final Test | FAIL
P | 2018-01-31T09:42:01+08:00 | 2018-01-31T12:48:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | A18USM804G | CPSM Final Test | PASS |  | 159.71 | 0.17 | -0.31 | 57.13 | 54.44 | 1.51 | 0.13 | 0 | -5 | 1 | -0.85 | 0.4948 | 3.8 | 13 |  |  | 0.315 | 0.1413 | -0.0144 | 1 |  |  |  |  | 133 |  | 0 | [1204 empty] | 0
P | 2018-01-31T09:47:20+08:00 | 2018-01-31T12:50:57+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | A18USM807G | CPSM Final Test | PASS |  | 174.11 | 0.37 | -0.46 | 56.93 | 52.32 | 1.9 | 0.13 | 0 | -12 | -1 | -0.9139 | 0.4974 | 1 | 2 |  |  | 0.0434 | 0.1918 | -0.0015 | 1 |  |  |  |  | 133 |  | 0 | [1205 empty] | 0
P | 2018-01-31T09:29:55+08:00 | 2018-01-31T13:00:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | A18USM802G | CPSM Final Test | PASS |  | 166.01 | 0.37 | -0.77 | 56.91 | 52.66 | 0.91 | 0.13 | -1 | -1 | 1 | -0.872 | 0.5428 | 1.5 | 6 |  |  | 0.1011 | 0.1279 | -0.0299 | 0.99999 |  |  |  |  | 133 |  | 0.03956 | [206 empty] | 0
P | 2018-01-31T09:48:49+08:00 | 2018-01-31T13:07:00+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | A18USM805G | CPSM Final Test | PASS |  | 166.01 | -0.02 | -0.69 | 56.78 | 54.78 | 1.44 | 0.13 | 1 | -3 | 1 | -0.8501 | 0.501 | 6.1 | 38 |  |  | 0.1803 | 0.1737 | -0.0046 | 0.99998 |  |  |  |  | 133 |  | 0 | [716 empty] | 0
P | 2018-01-31T09:38:22+08:00 | 2018-01-31T13:13:14+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | A18USM803G | CPSM Final Test | PASS |  | 167.9 | 0.27 | -0.69 | 56.63 | 55.8 | 1.59 | 0.13 | 0 | -2 | 3 | -0.84 | 0.5124 | 3.3 | 15 |  |  | 0.2435 | 0.1488 | -0.0012 | 0.99998 |  |  |  |  | 133 |  | 0 | [78 empty] | 0
P | 2018-01-31T09:33:56+08:00 | 2018-01-31T13:14:29+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | A18USM806G | CPSM Final Test | PASS |  | 166.19 | 0.27 | -0.23 | 56.83 | 53.85 | 1.36 | 0.13 | 1 | 4 | 1 | -0.8858 | 0.5358 | 7.7 | 8 |  |  | 0.103 | 0.1559 | -0.0419 | 1 |  |  |  |  | 133 |  | 0 | [191 empty] | 0
P | 2018-01-29T19:08:59+08:00 | 2018-01-31T14:05:38+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | A18USM801G | CPSM Final Test | FAIL
P | 2018-01-30T16:37:19+08:00 | 2018-01-31T15:54:42+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | A18USM801G | CPSM Final Test | PASS |  | 165.65 | 0.97 | -0.61 | 56.5 | 51.99 | 1.44 | 0.13 | -1 | -3 | 0 | -0.8759 | 0.4932 | 7.3 | -8 |  |  | 0.107 | 0.2498 | 0.0057 | 1 |  |  |  |  | 133 |  | 0 | [1206 empty] | 0
P | 2018-01-31T11:39:46+08:00 | 2018-01-31T16:51:56+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI383G | CPSM Final Test | FAIL | [13 empty] | 17.4
P | 2018-01-31T14:04:09+08:00 | 2018-01-31T17:24:19+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI381G | CPSM Final Test | PASS | 129 | 172.85 | -0.02 | -0.77 | 56.72 | 59.18 | 1.36 | 0.13 | 0 | -1 | 1 | -0.9278 | 0.493 | 5.9 | -8 |  | 0 | 0.1049 | 0.1019 | -0.0374 | 0.99999 | [1015 empty] | 0
P | 2018-01-31T14:09:52+08:00 | 2018-01-31T17:27:20+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI382G | CPSM Final Test | PASS | 129 | 167.18 | -0.21 | -1.15 | 56.58 | 53.34 | 1.29 | 0.13 | 1 | 4 | 1 | -0.852 | 0.4928 | 1.7 | -1 |  | 0 | 0.0889 | 0.1617 | -0.0171 | 0.99999 | [508 empty] | 0
P | 2018-01-30T19:09:46+08:00 | 2018-02-01T09:21:56+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI383G | CPSM Final Test | FAIL
P | 2018-01-31T14:15:51+08:00 | 2018-02-01T09:28:36+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI383G | CPSM Final Test | PASS | 129 | 160.88 | -0.02 | -0.99 | 56.71 | 59.35 | 1.51 | 0.13 | 0 | -2 | -1 | -0.9439 | 0.549 | 3.9 | 31 |  | 0.0101 | 0.0451 | 0.2369 | 0.007 | 1 | [706 empty] | 0
P | 2018-01-30T17:51:04+08:00 | 2018-02-01T10:02:47+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | A18USM808G | CPSM Final Test | FAIL
P | 2018-01-31T13:56:56+08:00 | 2018-02-01T11:06:03+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | A18USM808G | CPSM Final Test | PASS |  | 161.78 | -0.02 | -0.38 | 57.02 | 52.24 | 1.9 | 0.13 | 0 | -11 | 1 | -0.8759 | 0.4872 | 0.8 | 14 |  |  | 0.0834 | 0.1325 | -0.0101 | 0.99993 |  |  |  |  | 133 |  | 0.02073 | [1070 empty] | 0
P | 2018-01-31T17:02:12+08:00 | 2018-02-01T11:25:59+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | A18SDI385G | CPSM Final Test | PASS | 129 | 167.36 | 0.27 | -1.3 | 56.49 | 54.53 | 1.29 | 0.13 | 1 | -10 | -2 | -0.9737 | 0.483 | 1.4 | -6 |  | 0 | 0.4393 | 0.4128 | -0.0308 | 0.99997 | [1214 empty] | 0
P | 2018-01-30T22:39:19+08:00 | 2018-02-01T11:37:22+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI387G | CPSM Final Test | FAIL
P | 2018-01-31T17:08:20+08:00 | 2018-02-01T11:49:50+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI387G | CPSM Final Test | PASS | 129 | 171.05 | -0.02 | -1.07 | 56.53 | 54.19 | 1.51 | 0.13 | 1 | -2 | -2 | -0.9079 | 0.514 | 5.5 | -4 |  | 0 | 0.0814 | 0.401 | -0.0243 | 0.99999 | [1213 empty] | 0
P | 2018-02-01T09:50:29+08:00 | 2018-02-01T12:39:26+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI388G | CPSM Final Test | PASS | 129 | 153.05 | 0.27 | -0.99 | 56.3 | 52.83 | 1.06 | 0.13 | -1 | 2 | -1 | -0.892 | 0.5048 | 4.8 | 18 |  | 0.01173 | 0.0768 | 0.2112 | 0.0077 | 1 | [1215 empty] | 0
P | 2018-02-01T09:54:40+08:00 | 2018-02-01T12:40:49+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI386G | CPSM Final Test | PASS | 129 | 165.92 | -0.21 | -0.99 | 56.73 | 59.44 | 1.36 | 0.13 | 0 | -11 | 0 | -0.8939 | 0.5038 | 10.4 | 0 |  | 0.00884 | 0.0604 | 0.0675 | -0.0332 | 1 | [1216 empty] | 0
P | 2018-02-01T10:03:05+08:00 | 2018-02-01T12:48:50+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI389G | CPSM Final Test | PASS | 129 | 170.6 | -0.02 | -1.07 | 56.43 | 59.61 | 1.74 | 0.13 | 1 | 3 | -4 | -0.882 | 0.499 | 6.4 | -4 |  | 0 | 0.1369 | 0.0953 | -0.0296 | 1 | [996 empty] | 0
P | 2018-02-01T09:14:01+08:00 | 2018-02-01T12:53:47+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI380G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -1.22 | 56.35 | 54.27 | 1.21 | 0.13 | 0 | 2 | 0 | -0.9377 | 0.48 | 0.7 | 8 |  | 0.00792 | 0.1304 | 0.1936 | 0.0151 | 0.99998 | [1217 empty] | 0
P | 2018-02-01T14:29:04+08:00 | 2018-02-01T17:25:52+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI394G | CPSM Final Test | PASS | 129 | 166.73 | 0.37 | -1.45 | 56.54 | 53.76 | 1.51 | 0.13 | 0 | -5 | 0 | -0.874 | 0.4832 | 4 | -7 |  | 0 | 0.0911 | 0.4149 | -0.0666 | 0.99998 | [549 empty] | 0
P | 2018-02-01T14:24:07+08:00 | 2018-02-01T17:26:40+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | A18SDI396G | CPSM Final Test | PASS | 129 | 167 | -0.02 | -0.99 | 56.76 | 55.2 | 1.21 | 0.13 | 0 | -11 | 3 | -0.9417 | 0.5144 | 0.3 | 22 |  | 0 | 0.0648 | 0.457 | 0.0056 | 1 | [456 empty] | 0
P | 2018-02-01T14:22:49+08:00 | 2018-02-01T17:31:30+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI393G | CPSM Final Test | PASS | 129 | 164.57 | -0.3 | -1.15 | 56.23 | 51.82 | 1.9 | 0.13 | 0 | -11 | 0 | -0.9021 | 0.5232 | 3.8 | 26 |  | 0 | 0.1743 | 0.2988 | 0.0129 | 0.99999 | [618 empty] | 0
P | 2018-02-01T14:15:44+08:00 | 2018-02-01T17:35:02+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI392G | CPSM Final Test | PASS | 129 | 163.94 | -0.49 | -0.99 | 56.52 | 53.26 | 1.44 | 0.13 | 1 | -12 | -1 | -0.9099 | 0.5014 | 4.3 | 3 |  | 0 | 0.1066 | 0.0648 | 0.0011 | 1 | [1218 empty] | 0
P | 2018-02-01T10:48:22+08:00 | 2018-02-01T17:35:43+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI391G | CPSM Final Test | FAIL
P | 2018-02-01T14:12:02+08:00 | 2018-02-01T17:36:25+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI391G | CPSM Final Test | PASS | 129 | 168.71 | -0.02 | -0.54 | 57.11 | 51.82 | 2.2 | 0.13 | 1 | -11 | 0 | -0.9219 | 0.4972 | 2.4 | 28 |  | 0 | 0.0876 | 0.0492 | 0.0085 | 1 | [951 empty] | 0
P | 2018-02-01T12:35:13+08:00 | 2018-02-01T17:40:03+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI390G | CPSM Final Test | PASS | 129 | 168.44 | -0.02 | -0.84 | 56.59 | 54.61 | 1.74 | 0.13 | 0 | -4 | -1 | -0.9456 | 0.5492 | 5.6 | 18 |  | 0.01677 | 0.1595 | 0.2291 | 0.0117 | 0.99998 | [1219 empty] | 0
P | 2018-02-01T14:09:07+08:00 | 2018-02-01T17:48:11+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI395G | CPSM Final Test | PASS | 129 | 167.45 | 0.27 | -1.22 | 56.51 | 53.51 | 1.82 | 0.13 | 1 | -2 | 1 | -0.8639 | 0.498 | 5.5 | 8 |  | 0 | 0.2897 | 0.1073 | 0.0244 | 0.99998 | [1193 empty] | 0
P | 2018-02-01T18:27:01+08:00 | 2018-02-01T19:16:44+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI397G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-02-01T18:45:17+08:00 | 2018-02-01T22:19:54+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI400G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -0.69 | 57.19 | 51.82 | 1.51 | 0.13 | -1 | -2 | 1 | -0.9218 | 0.497 | 4.3 | -6 |  | 0 | 0.0873 | 0.2445 | 0.0373 | 0.99997 | [361 empty] | 0
P | 2018-02-01T05:09:09+08:00 | 2018-02-02T08:56:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | A18SDI401G | CPSM Final Test | FAIL
P | 2018-02-01T19:02:40+08:00 | 2018-02-02T08:56:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | A18SDI401G | CPSM Final Test | PASS | 129 | 165.74 | 0.97 | 0 | 56.77 | 52.24 | 1.21 | 0.13 | -1 | 12 | 0 | -0.9079 | 0.537 | 4.3 | 11 |  | 0 | 0.2648 | 0.6066 | -0.0824 | 1 | [338 empty] | 0
P | 2018-01-29T22:41:39+08:00 | 2018-02-02T11:14:16+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | A18SDI384G | CPSM Final Test | FAIL
P | 2018-02-02T08:22:42+08:00 | 2018-02-02T12:12:58+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI403G | CPSM Final Test | FAIL | [13 empty] | 24.2
P | 2018-01-31T16:57:57+08:00 | 2018-02-02T12:17:46+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | A18SDI384G | CPSM Final Test | PASS | 129 | 165.38 | 0.27 | -1.45 | 56.42 | 59.1 | 0.83 | 0.13 | -1 | 2 | 1 | -0.9917 | 0.491 | 1.2 | 28 |  | 0.04729 | 0.0489 | 0.2117 | -0.0005 | 0.99999 | [445 empty] | 0
P | 2018-02-01T01:08:19+08:00 | 2018-02-02T12:35:26+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI397G | CPSM Final Test | FAIL
P | 2018-02-02T10:04:54+08:00 | 2018-02-02T13:12:27+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI402G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -0.46 | 56.71 | 56.98 | 1.97 | 0.13 | -1 | 1 | 0 | -0.9139 | 0.5024 | 1.8 | 6 |  | 0 | 0.114 | 0.1279 | -0.0203 | 1 | [951 empty] | 0
P | 2018-02-02T10:10:12+08:00 | 2018-02-02T13:13:03+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI404G | CPSM Final Test | PASS | 129 | 163.76 | 0.27 | -0.99 | 56.45 | 51.39 | 1.51 | 0.13 | 0 | 3 | -1 | -0.9019 | 0.487 | 6.3 | -9 |  | 0 | 0.0987 | 0.1137 | -0.006 | 0.99999 | [566 empty] | 0
P | 2018-02-02T09:56:03+08:00 | 2018-02-02T13:16:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI398G | CPSM Final Test | PASS | 129 | 166.01 | -0.02 | -0.92 | 56.52 | 52.92 | 1.82 | 0.13 | 0 | -2 | 0 | -0.8919 | 0.509 | 4.2 | 5 |  | 0.0253 | 0.1657 | 0.1787 | -0.0005 | 1 | [1212 empty] | 0
P | 2018-02-02T06:39:35+08:00 | 2018-02-02T13:22:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI399G | CPSM Final Test | FAIL
P | 2018-02-02T10:17:50+08:00 | 2018-02-02T14:12:46+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI403G | CPSM Final Test | PASS | 129 | 166.64 | -0.02 | -0.99 | 56.5 | 55.03 | 1.67 | 0.13 | 0 | -11 | 3 | -0.9138 | 0.499 | 1.9 | 44 |  | 0.00502 | 0.0725 | 0.3092 | 0.0383 | 0.99996 | [926 empty] | 0
P | 2018-02-02T10:01:06+08:00 | 2018-02-02T15:20:14+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI399G | CPSM Final Test | PASS | 129 | 166.01 | -0.02 | -0.84 | 56.6 | 53.59 | 1.67 | 0.13 | 0 | -2 | -1 | -0.8519 | 0.4924 | 2.8 | 8 |  | 0.00741 | 0.233 | 0.3355 | 0.0526 | 0.99983 | [1220 empty] | 0
P | 2018-02-02T14:15:51+08:00 | 2018-02-02T17:33:22+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | A18SDI409G | CPSM Final Test | PASS | 129 | 163.04 | -0.02 | -0.92 | 56.82 | 54.1 | 1.29 | 0.13 | 0 | 2 | 1 | -0.868 | 0.4678 | 1 | 13 |  | 0 | 0.1757 | 0.2578 | 0.0148 | 0.99999 | [704 empty] | 0
P | 2018-02-02T14:06:36+08:00 | 2018-02-02T17:47:39+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI407G | CPSM Final Test | PASS | 129 | 169.25 | -0.02 | -0.92 | 56.41 | 59.35 | 1.51 | 0.13 | -1 | -12 | 3 | -0.87 | 0.5038 | 1.5 | 47 |  | 0 | 0.0468 | 0.1174 | -0.0029 | 0.99999 | [1184 empty] | 0
P | 2018-02-02T09:34:50+08:00 | 2018-02-02T18:22:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI405G | CPSM Final Test | FAIL
P | 2018-02-02T13:58:29+08:00 | 2018-02-02T18:22:47+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI405G | CPSM Final Test | PASS | 129 | 165.65 | 0.27 | -1.07 | 56.22 | 52.07 | 1.21 | 0.13 | -1 | -3 | -1 | -0.8798 | 0.5076 | 1.8 | -7 |  | 0.02557 | 0.1846 | 0.2285 | 0.0349 | 0.99997 | [1221 empty] | 0
P | 2018-01-31T19:18:45+08:00 | 2018-02-02T18:25:00+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI397G | CPSM Final Test | FAIL
P | 2018-02-02T09:35:13+08:00 | 2018-02-02T18:30:07+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI406G | CPSM Final Test | FAIL
P | 2018-02-02T14:02:40+08:00 | 2018-02-02T18:36:43+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI406G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -1.15 | 56.4 | 56.05 | 1.51 | 0.13 | 1 | -2 | -1 | -0.9337 | 0.512 | 4 | 2 |  | 0 | 0.1386 | 0.1764 | -0.0311 | 1 | [384 empty] | 0
P | 2018-02-02T09:17:01+08:00 | 2018-02-02T19:04:40+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | A18SDI408G | CPSM Final Test | FAIL | [466 empty] | 1671
P | 2018-02-02T17:24:07+08:00 | 2018-02-02T19:56:02+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI411G | CPSM Final Test | FAIL
P | 2018-02-01T18:51:52+08:00 | 2018-02-02T20:09:10+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI397G | CPSM Final Test | PASS | 129 | 173.66 | -0.02 | -0.92 | 56.74 | 54.69 | 1.21 | 0.13 | -1 | -4 | -1 | -0.88 | 0.508 | 2.8 | 10 |  | 0 | 0.1048 | 0.1948 | 0.0103 | 0.99999 | [1222 empty] | 0
P | 2018-02-02T14:10:51+08:00 | 2018-02-02T20:15:28+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | A18SDI408G | CPSM Final Test | PASS | 129 | 161.33 | 0.27 | -1.53 | 56.2 | 54.78 | 0.98 | 0.13 | 0 | -1 | 3 | -0.9198 | 0.5424 | 2.8 | 8 |  | 0.00852 | 0.0346 | 0.3484 | -0.0551 | 0.99998 | [445 empty] | 0
P | 2018-02-02T18:41:44+08:00 | 2018-02-02T21:43:54+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | A18SDI412G | CPSM Final Test | PASS | 129 | 158.36 | -0.02 | -0.84 | 56.92 | 56.05 | 1.51 | 0.13 | -2 | -13 | -2 | -0.9318 | 0.4982 | 1.1 | 10 |  | 0 | 0.0546 | 0.176 | -0.038 | 0.99999 | [1223 empty] | 0
P | 2018-02-02T18:40:05+08:00 | 2018-02-02T21:55:36+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI411G | CPSM Final Test | PASS | 129 | 161.6 | -0.02 | -0.99 | 56.4 | 56.73 | 1.36 | 0.13 | 1 | -2 | -1 | -0.9119 | 0.543 | 1.7 | 7 |  | 0.01216 | 0.0522 | 0.245 | -0.0337 | 0.99999 | [511 empty] | 0
P | 2018-02-02T19:53:13+08:00 | 2018-02-02T22:30:41+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI415G | CPSM Final Test | PASS | 129 | 163.58 | 0.46 | -1.22 | 56.54 | 55.63 | 1.13 | 0.13 | 0 | 3 | -1 | -0.8838 | 0.5234 | 0.9 | 2 |  | 0 | 0.0909 | 0.0623 | 0.0027 | 0.99998 | [1199 empty] | 0
P | 2018-02-03T11:06:21+08:00 | 2018-02-03T13:53:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | A18SDI414G | CPSM Final Test | PASS | 129 | 165.74 | -0.02 | -1.22 | 56.49 | 52.92 | 1.13 | 0.13 | -1 | -1 | -1 | -0.9298 | 0.5086 | 8.3 | 8 |  | 0.02317 | 0.1343 | 0.2115 | -0.0263 | 0.99998 | [1224 empty] | 0
P | 2018-02-03T08:06:10+08:00 | 2018-02-03T15:22:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI419G | CPSM Final Test | FAIL
P | 2018-02-03T11:40:52+08:00 | 2018-02-03T15:44:59+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI418G | CPSM Final Test | PASS | 129 | 166.28 | 0.37 | -0.99 | 56.49 | 53.09 | 1.59 | 0.13 | 1 | -2 | 1 | -0.896 | 0.5018 | 6.1 | 16 |  | 0.00755 | 0.0995 | 0.071 | -0.0289 | 1 | [1196 empty] | 0
P | 2018-02-03T11:44:14+08:00 | 2018-02-03T16:28:15+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI419G | CPSM Final Test | PASS | 129 | 161.24 | -0.02 | -1.15 | 56.75 | 53.26 | 1.21 | 0.13 | 1 | 2 | 3 | -0.8879 | 0.4976 | 2.5 | 2 |  | 0 | 0.0638 | 0.1647 | 0.0144 | 1 | [1225 empty] | 0
P | 2018-02-03T14:59:54+08:00 | 2018-02-03T17:53:18+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | A18SDI420G | CPSM Final Test | PASS | 129 | 163.49 | -0.02 | -0.92 | 56.81 | 52.83 | 2.2 | 0.13 | 0 | -2 | 1 | -0.8461 | 0.4826 | 3.8 | 0 |  | 0 | 0.058 | 0.083 | 0.0333 | 0.99999 | [951 empty] | 0
P | 2018-02-03T14:52:30+08:00 | 2018-02-03T17:57:28+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | A18SDI422G | CPSM Final Test | PASS | 129 | 169.16 | -0.02 | -0.99 | 56.83 | 55.54 | 1.59 | 0.13 | 1 | 4 | 3 | -0.9836 | 0.501 | 2.3 | 0 |  | 0 | 0.2489 | 0.2209 | -0.0065 | 0.99997 | [1226 empty] | 0
P | 2018-02-03T14:51:46+08:00 | 2018-02-03T18:04:30+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | A18SDI421G | CPSM Final Test | PASS | 129 | 166.55 | -0.02 | -1.15 | 56.33 | 53.26 | 1.21 | 0.13 | 0 | -3 | -1 | -0.9019 | 0.5126 | 5.6 | 3 |  | 0.00631 | 0.0628 | 0.2891 | -0.0055 | 0.99998 | [1187 empty] | 0
P | 2018-02-04T10:20:34+08:00 | 2018-02-04T12:16:16+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI417G | CPSM Final Test | FAIL
P | 2018-02-04T11:17:09+08:00 | 2018-02-04T14:09:09+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | A18SDI410G | CPSM Final Test | PASS | 129 | 165.29 | -0.02 | -0.84 | 56.66 | 55.37 | 1.51 | 0.13 | 0 | -11 | 0 | -0.8959 | 0.4938 | 3.2 | 1 |  | 0.01351 | 0.0521 | 0.0741 | -0.028 | 1 | [1227 empty] | 0
P | 2018-02-04T11:14:54+08:00 | 2018-02-04T14:12:02+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | A18SDI416G | CPSM Final Test | PASS | 129 | 161.69 | -0.02 | -0.99 | 56.71 | 59.01 | 1.51 | 0.13 | 0 | -11 | 1 | -0.8481 | 0.502 | 4.4 | 54 |  | 0 | 0.0448 | 0.2081 | 0.0003 | 1 | [1225 empty] | 0
P | 2018-02-04T11:55:56+08:00 | 2018-02-04T15:28:00+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | A18SDI413G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -0.77 | 56.76 | 56.64 | 1.59 | 0.13 | 0 | 2 | 1 | -0.892 | 0.5004 | 8.5 | 20 |  | 0.00085 | 0.0716 | 0.2297 | 0.0004 | 1 | [445 empty] | 0
P | 2018-02-04T01:00:50+08:00 | 2018-02-05T08:57:32+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI417G | CPSM Final Test | FAIL
P | 2018-02-03T23:56:46+08:00 | 2018-02-05T10:01:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI417G | CPSM Final Test | FAIL
P | 2018-02-04T16:59:11+08:00 | 2018-02-05T11:45:34+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | A18SDI417G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -1.3 | 56.51 | 55.03 | 1.29 | 0.13 | 0 | -11 | 0 | -1.0296 | 0.5122 | 8.8 | 5 |  | 0.02545 | 0.1491 | 0.2142 | -0.0523 | 0.99925 | [1148 empty] | 0
P | 2018-02-05T14:14:03+08:00 | 2018-02-05T17:21:27+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | B18USM811G | CPSM Final Test | PASS |  | 169.43 | -0.02 | -0.61 | 56.51 | 53.34 | 1.67 | 0.13 | -1 | 4 | 0 | -0.8999 | 0.5188 | 3 | -7 |  |  | 0.085 | 0.3758 | 0.0127 | 1 |  |  |  |  | 133 |  | 0 | [1221 empty] | 0
P | 2018-02-05T14:31:15+08:00 | 2018-02-05T18:02:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM810G | CPSM Final Test | PASS |  | 164.21 | -0.02 | -1.07 | 56.46 | 53.76 | 1.44 | 0.13 | -1 | -2 | 1 | -0.8818 | 0.4846 | 0.8 | 4 |  |  | 0.1128 | 0.1188 | -0.0592 | 1 |  |  |  |  | 133 |  | 0 | [1222 empty] | 0
P | 2018-02-05T15:23:11+08:00 | 2018-02-05T18:29:44+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | B18USM813G | CPSM Final Test | PASS |  | 164.03 | 0.17 | -0.54 | 56.82 | 51.9 | 1.21 | 0.13 | 1 | -12 | -1 | -0.834 | 0.5182 | 4.3 | 10 |  |  | 0.0484 | 0.042 | -0.0092 | 1 |  |  |  |  | 133 |  | 0 | [365 empty] | 0
P | 2018-02-05T14:04:30+08:00 | 2018-02-05T18:35:19+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | B18USM814G | CPSM Final Test | FAIL | [13 empty] | 21.4
P | 2018-02-05T12:38:07+08:00 | 2018-02-05T19:40:24+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | B18USM812G | CPSM Final Test | FAIL | [1251 empty] | 0
P | 2018-02-05T16:19:55+08:00 | 2018-02-05T20:39:01+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | B18USM814G | CPSM Final Test | PASS |  | 177.35 | 0.08 | -0.08 | 57.01 | 51.39 | 1.97 | 0.13 | 1 | -12 | 0 | -0.8999 | 0.5492 | 5.6 | 26 |  |  | 0.0609 | 0.2082 | -0.0068 | 0.99999 | 0 |  |  |  | 133 |  | 0
P | 2018-02-05T16:09:16+08:00 | 2018-02-05T20:45:54+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | B18USM812G | CPSM Final Test | PASS |  | 166.55 | 0.08 | -0.61 | 56.92 | 55.29 | 1.51 | 0.13 | 0 | 2 | 2 | -0.8918 | 0.531 | 3.4 | 62 |  |  | 0.1345 | 0.0905 | 0.0028 | 0.99993 |  |  |  |  | 133 |  | 0.00056 | [1223 empty] | 0
P | 2018-02-05T17:14:40+08:00 | 2018-02-05T21:06:32+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | B18USM815G | CPSM Final Test | PASS |  | 160.7 | 0.27 | -0.99 | 56.7 | 57.57 | 0.98 | 0.13 | 1 | -2 | 3 | -0.8919 | 0.5156 | 8.3 | 16 |  |  | 0.1102 | 0.143 | 0.0025 | 0.99998 |  |  |  |  | 133 |  | 0 | [808 empty] | 0
P | 2018-02-05T18:36:23+08:00 | 2018-02-05T21:26:33+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM817G | CPSM Final Test | PASS |  | 163.58 | 0.37 | -0.69 | 56.5 | 52.15 | 1.67 | 0.13 | 1 | 2 | -1 | -0.9038 | 0.541 | 1.5 | -14 |  |  | 0.0517 | 0.0797 | -0.0267 | 0.99999 |  |  |  |  | 133 |  | 0 | [282 empty] | 0
P | 2018-02-04T19:10:45+08:00 | 2018-02-06T09:36:04+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | B18USM809G | CPSM Final Test | FAIL | [17 empty] | 1.7958
P | 2018-02-05T14:23:25+08:00 | 2018-02-06T10:54:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | B18USM809G | CPSM Final Test | PASS |  | 168.17 | -0.02 | -0.84 | 56.69 | 56.98 | 1.74 | 0.13 | -1 | -11 | -1 | -0.852 | 0.5282 | 2.9 | 27 |  |  | 0.0747 | 0.0773 | 0.0224 | 0.99994 |  |  |  |  | 133 |  | 0 | [244 empty] | 0
P | 2018-02-04T23:20:05+08:00 | 2018-02-06T11:27:55+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | B18USM816G | CPSM Final Test | FAIL
P | 2018-02-06T11:06:28+08:00 | 2018-02-06T14:25:19+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM819G | CPSM Final Test | PASS |  | 163.04 | -0.02 | -0.69 | 56.51 | 53.34 | 1.59 | 0.13 | 1 | 2 | 1 | -0.8101 | 0.5226 | 2.9 | -9 |  |  | 0.0875 | 0.2427 | -0.0136 | 1 |  |  |  |  | 133 |  | 0 | [1224 empty] | 0
P | 2018-02-06T11:15:55+08:00 | 2018-02-06T14:28:18+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | B18USM818G | CPSM Final Test | PASS |  | 158.45 | 0.37 | -0.61 | 56.42 | 57.91 | 1.29 | 0.13 | 1 | -12 | 0 | -0.864 | 0.543 | 4.8 | -23 |  |  | 0.1741 | 0.5285 | -0.0287 | 1 |  |  |  |  | 133 |  | 0 | [360 empty] | 0
P | 2018-02-06T11:56:22+08:00 | 2018-02-06T14:57:54+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | B18USM820G | CPSM Final Test | PASS |  | 167.99 | 0.08 | -0.99 | 56.71 | 52.83 | 1.36 | 0.13 | 0 | -4 | 0 | -0.9158 | 0.5262 | 4.6 | 26 |  |  | 0.0342 | 0.0772 | 0.0044 | 1 |  |  |  |  | 133 |  | 0 | [1225 empty] | 0
P | 2018-02-06T08:09:28+08:00 | 2018-02-06T14:58:44+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | B18USM823G | CPSM Final Test | FAIL
P | 2018-02-06T12:03:00+08:00 | 2018-02-06T15:25:36+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | B18USM821G | CPSM Final Test | PASS |  | 166.28 | 0.08 | -0.77 | 56.76 | 54.1 | 1.74 | 0.13 | 1 | -3 | -1 | -0.8959 | 0.5028 | 0.8 | 30 |  |  | 0.119 | 0.392 | -0.0378 | 0.99999 |  |  |  |  | 133 |  | 0 | [1226 empty] | 0
P | 2018-02-06T11:51:34+08:00 | 2018-02-06T15:49:31+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | B18USM822G | CPSM Final Test | PASS |  | 165.56 | 0.46 | -0.61 | 56.89 | 56.56 | 1.67 | 0.13 | 1 | 4 | 0 | -0.9197 | 0.511 | 7.2 | 33 |  |  | 0.3228 | 0.3398 | -0.0125 | 0.99999 |  |  |  |  | 133 |  | 0 | [24 empty] | 0
P | 2018-02-06T11:34:06+08:00 | 2018-02-06T16:01:46+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | B18USM823G | CPSM Final Test | PASS |  | 165.29 | -0.02 | -0.69 | 56.53 | 55.54 | 1.36 | 0.13 | 1 | -3 | 3 | -0.888 | 0.456 | 3.6 | 101 |  |  | 0.1801 | 0.0616 | -0.0128 | 0.99999 |  |  |  |  | 133 |  | 0.00111 | [1223 empty] | 0
P | 2018-02-06T14:43:39+08:00 | 2018-02-06T17:49:58+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI423G | CPSM Final Test | PASS | 129 | 167.9 | 0.08 | -0.77 | 56.92 | 54.53 | 1.44 | 0.13 | 1 | -10 | 1 | -0.8978 | 0.4716 | 3.6 | 6 |  | 0 | 0.0603 | 0.2346 | 0.0032 | 1 | [7 empty] | 0
P | 2018-02-06T15:13:50+08:00 | 2018-02-06T18:07:47+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI427G | CPSM Final Test | PASS | 129 | 166.1 | -0.02 | -1.3 | 56.9 | 56.39 | 1.67 | 0.13 | -1 | -11 | 2 | -0.9099 | 0.5076 | 3.9 | 0 |  | 0 | 0.0603 | 0.2732 | -0.0397 | 1 | [950 empty] | 0
P | 2018-02-06T15:46:32+08:00 | 2018-02-06T18:42:04+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | B18USM816G | CPSM Final Test | PASS |  | 162.14 | 0.65 | -1.15 | 56.81 | 56.22 | 1.44 | 0.13 | 0 | -2 | 0 | -0.9379 | 0.497 | 3.6 | 132 |  |  | 0.191 | 0.1521 | -0.0154 | 1 |  |  |  |  | 133 |  | 0 | [1227 empty] | 0
P | 2018-02-06T16:45:05+08:00 | 2018-02-06T20:20:58+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI425G | CPSM Final Test | PASS | 129 | 165.74 | 0.27 | -0.92 | 56.5 | 53.17 | 1.67 | -0.14 | -1 | -1 | 3 | -0.8979 | 0.5498 | 9.3 | 12 |  | 0 | 0.0952 | 0.2495 | -0.0228 | 1 | [1235 empty] | 0
P | 2018-02-06T17:14:23+08:00 | 2018-02-06T20:58:11+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI428G | CPSM Final Test | PASS | 129 | 161.87 | -0.02 | -0.84 | 56.7 | 58.93 | 1.67 | 0.13 | 0 | -2 | -1 | -0.9398 | 0.4658 | 1.6 | 28 |  | 0 | 0.0621 | 0.2947 | -0.001 | 0.99999 | [599 empty] | 0
P | 2018-02-06T17:07:49+08:00 | 2018-02-06T20:58:52+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI426G | CPSM Final Test | PASS | 129 | 166.19 | -0.02 | -1.3 | 56.16 | 51.9 | 1.67 | 0.13 | 0 | -3 | 1 | -0.9218 | 0.4538 | 4.8 | 49 |  | 0.00097 | 0.1108 | 0.1941 | -0.0243 | 1 | [1230 empty] | 0
P | 2018-02-07T08:25:38+08:00 | 2018-02-07T11:23:36+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI433G | CPSM Final Test | FAIL | [13 empty] | 20.5
P | 2018-02-07T09:59:01+08:00 | 2018-02-07T13:04:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI432G | CPSM Final Test | PASS | 129 | 165.38 | -0.02 | -0.99 | 56.48 | 51.9 | 1.51 | 0.13 | 0 | 3 | -1 | -0.8879 | 0.498 | 0.9 | -17 |  | 0 | 0.1192 | 0.1772 | -0.0187 | 1 | [1236 empty] | 0
P | 2018-02-07T10:16:04+08:00 | 2018-02-07T13:17:15+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI430G | CPSM Final Test | PASS | 129 | 171.59 | -0.02 | -0.84 | 56.52 | 54.1 | 1.82 | 0.13 | 0 | -2 | 0 | -0.8639 | 0.495 | 2.5 | 15 |  | 0 | 0.1026 | 0.5086 | 0.0357 | 0.99999 | [1237 empty] | 0
P | 2018-02-07T06:45:48+08:00 | 2018-02-07T13:20:18+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI431G | CPSM Final Test | FAIL
P | 2018-02-07T10:03:03+08:00 | 2018-02-07T13:20:50+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI431G | CPSM Final Test | PASS | 129 | 169.07 | -0.02 | -0.84 | 56.52 | 52.49 | 1.59 | 0.13 | -1 | 3 | 1 | -0.872 | 0.4898 | 3.2 | 68 |  | 0 | 0.0879 | 0.0573 | -0.0106 | 1 | [422 empty] | 0
P | 2018-02-07T09:54:37+08:00 | 2018-02-07T13:23:26+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI433G | CPSM Final Test | PASS | 129 | 155.66 | -0.02 | -0.99 | 56.44 | 54.19 | 1.74 | 0.13 | 1 | -2 | -1 | -0.834 | 0.4504 | 0 | 12 |  | 0 | 0.0346 | 0.0468 | 0.0166 | 0.99999 | [1238 empty] | 0
P | 2018-02-07T05:16:55+08:00 | 2018-02-07T14:58:00+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI434G | CPSM Final Test | FAIL
P | 2018-02-07T12:19:58+08:00 | 2018-02-07T15:50:37+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI435G | CPSM Final Test | FAIL
P | 2018-02-07T03:32:07+08:00 | 2018-02-07T16:42:48+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI434G | CPSM Final Test | FAIL | [1260 empty] | 0
P | 2018-02-07T03:28:55+08:00 | 2018-02-07T16:46:00+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI434G | CPSM Final Test | FAIL | [1260 empty] | 0
P | 2018-02-05T14:42:30+08:00 | 2018-02-07T17:03:00+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI424G | CPSM Final Test | FAIL | [20 empty] | 0.99731
P | 2018-02-07T14:20:54+08:00 | 2018-02-07T17:17:42+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI438G | CPSM Final Test | PASS | 129 | 169.34 | 0.17 | -0.92 | 56.49 | 51.82 | 1.51 | 0.13 | -1 | -1 | -1 | -0.9338 | 0.4934 | 1.8 | 10 |  | 0.01107 | 0.1299 | 0.3036 | 0.0058 | 0.99999 | [1240 empty] | 0
P | 2018-02-07T14:19:28+08:00 | 2018-02-07T17:26:37+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI439G | CPSM Final Test | PASS | 129 | 163.04 | 0.27 | -0.84 | 56.61 | 54.61 | 1.67 | 0.13 | 0 | -2 | 0 | -0.9378 | 0.505 | 4.5 | -5 |  | 0 | 0.1462 | 0.2577 | 0.0293 | 0.99999 | [1241 empty] | 0
P | 2018-02-07T14:09:41+08:00 | 2018-02-07T17:27:14+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI437G | CPSM Final Test | PASS | 129 | 164.39 | -0.02 | -0.92 | 56.7 | 58.67 | 1.67 | 0.13 | -1 | -2 | -1 | -0.91 | 0.495 | 0.6 | -38 |  | 0 | 0.2573 | 0.3363 | -0.0268 | 0.99999 | [123 empty] | 0
P | 2018-02-07T13:59:48+08:00 | 2018-02-07T17:35:40+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI436G | CPSM Final Test | PASS | 129 | 163.58 | -0.02 | -1.15 | 56.49 | 51.65 | 1.59 | 0.13 | -1 | -11 | 1 | -0.9298 | 0.533 | 4.7 | 32 |  | 0 | 0.0977 | 0.0905 | 0.0003 | 0.99998 | [803 empty] | 0
P | 2018-02-07T10:07:28+08:00 | 2018-02-07T17:51:34+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI434G | CPSM Final Test | PASS | 129 | 164.48 | 0.27 | -1.15 | 56.42 | 57.07 | 1.44 | 0.13 | 1 | -2 | -1 | -0.886 | 0.4908 | 7.2 | 15 |  | 0 | 0.1315 | 0.2789 | -0.0052 | 1 | [1239 empty] | 0
P | 2018-02-07T14:05:18+08:00 | 2018-02-07T17:52:27+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI435G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -0.92 | 56.48 | 52.83 | 1.44 | 0.13 | 0 | -1 | 1 | -0.8919 | 0.497 | 6.2 | 9 |  | 0 | 0.0572 | 0.0638 | -0.0011 | 1 | [364 empty] | 0
P | 2018-02-04T20:57:38+08:00 | 2018-02-08T10:47:52+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI424G | CPSM Final Test | FAIL | [17 empty] | 2.0972
P | 2018-02-08T07:06:09+08:00 | 2018-02-08T13:46:06+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI429G | CPSM Final Test | FAIL
P | 2018-02-08T09:19:35+08:00 | 2018-02-08T14:10:26+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI440G | CPSM Final Test | FAIL
P | 2018-02-08T10:50:10+08:00 | 2018-02-08T14:18:19+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI447G | CPSM Final Test | PASS | 129 | 164.66 | 0.08 | -0.84 | 56.54 | 51.82 | 1.9 | 0.21 | 1 | -11 | -1 | -0.854 | 0.4766 | 4.9 | 18 |  | 0 | 0.0599 | 0.0959 | 0.0067 | 0.99999 | [107 empty] | 0
P | 2018-02-08T11:00:59+08:00 | 2018-02-08T14:20:52+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI449G | CPSM Final Test | PASS | 129 | 159.71 | 0.17 | -0.84 | 56.77 | 57.91 | 1.36 | 0.13 | 1 | -11 | 3 | -0.7922 | 0.536 | 0.8 | 11 |  | 0 | 0.0776 | 0.2852 | -0.0142 | 1 | [128 empty] | 0
P | 2018-02-08T11:30:18+08:00 | 2018-02-08T14:29:49+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI445G | CPSM Final Test | PASS | 129 | 164.12 | 0.17 | -1.15 | 56.18 | 53.51 | 1.29 | 0.13 | -1 | -2 | -1 | -0.87 | 0.5432 | 9.3 | -8 |  | 0 | 0.3097 | 0.1146 | -0.0006 | 0.99999 | [627 empty] | 0
P | 2018-02-08T11:13:56+08:00 | 2018-02-08T14:35:04+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI446G | CPSM Final Test | PASS | 129 | 166.55 | -0.02 | -1.15 | 56.72 | 57.4 | 1.51 | 0.13 | 1 | -11 | -1 | -0.86 | 0.4768 | 4.9 | 36 |  | 0.00072 | 0.0578 | 0.1357 | -0.0186 | 1 | [1239 empty] | 0
P | 2018-02-08T11:57:16+08:00 | 2018-02-08T14:37:51+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI441G | CPSM Final Test | PASS | 129 | 159.08 | 0.49 | -1.15 | 56.49 | 54.86 | 1.36 | 0.13 | 0 | -11 | 0 | -0.8259 | 0.486 | 1.1 | 14 |  | 0 | 0.2673 | 0.2307 | 0.0264 | 0.99998 | [1242 empty] | 0
P | 2018-02-08T11:45:01+08:00 | 2018-02-08T15:16:06+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI440G | CPSM Final Test | PASS | 129 | 166.1 | 0.97 | -0.92 | 56.6 | 54.61 | 1.21 | 0.13 | 0 | 2 | -1 | -0.8261 | 0.539 | 0 | 6 |  | 0 | 0.0457 | 0.0896 | -0.0322 | 0.99998 | [1176 empty] | 0
P | 2018-02-08T10:26:08+08:00 | 2018-02-08T15:38:46+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI429G | CPSM Final Test | PASS | 129 | 157.28 | 0.46 | -1.3 | 56.45 | 56.13 | 1.13 | 0.13 | 0 | -11 | 0 | -0.9438 | 0.488 | 8.3 | 2 |  | 0 | 0.026 | 0.2401 | -0.0086 | 1 | [876 empty] | 0
P | 2018-02-04T15:53:32+08:00 | 2018-02-08T15:51:58+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI424G | CPSM Final Test | FAIL | [17 empty] | 1.4997
P | 2018-02-08T13:17:03+08:00 | 2018-02-08T16:06:34+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI443G | CPSM Final Test | FAIL
P | 2018-02-06T15:52:45+08:00 | 2018-02-08T16:55:16+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI424G | CPSM Final Test | PASS | 129 | 161.51 | 0.46 | -1.38 | 56.2 | 50.55 | 1.21 | 0.13 | -1 | 2 | 2 | -0.9159 | 0.5216 | 3.8 | 0 |  | 0 | 0.1942 | 0.1678 | 0.0168 | 1 | [168 empty] | 0
P | 2018-02-08T14:33:43+08:00 | 2018-02-08T17:53:41+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI442G | CPSM Final Test | PASS | 129 | 167 | 0.17 | -1.15 | 56.2 | 52.49 | 1.21 | 0.13 | 1 | 2 | -1 | -0.9199 | 0.477 | 0.2 | -113 |  | 0 | 0.1267 | 0.1417 | 0.0071 | 1 | [1153 empty] | 0
P | 2018-02-08T14:54:31+08:00 | 2018-02-08T18:00:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI453G | CPSM Final Test | PASS | 129 | 161.78 | 0.08 | -0.77 | 56.84 | 56.56 | 1.82 | 0.13 | 1 | -8 | 2 | -0.9239 | 0.528 | 1.8 | 0 |  | 0 | 0.1186 | 0.1615 | 0.0058 | 0.99999 | [175 empty] | 0
P | 2018-02-08T14:41:49+08:00 | 2018-02-08T18:13:10+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI443G | CPSM Final Test | PASS | 129 | 158.27 | 0.27 | -0.84 | 56.72 | 57.74 | 1.59 | 0.13 | -1 | 2 | 1 | -0.884 | 0.4532 | 1.4 | 7 |  | 0 | 0.0844 | 0.0972 | -0.0017 | 0.99999 | [995 empty] | 0
P | 2018-02-08T14:15:14+08:00 | 2018-02-08T19:39:51+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI454G | CPSM Final Test | FAIL
P | 2018-02-08T16:25:08+08:00 | 2018-02-08T20:23:50+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI452G | CPSM Final Test | PASS | 129 | 170.15 | -0.02 | -0.69 | 56.11 | 52.41 | 1.82 | 0.13 | -1 | 3 | 1 | -0.9419 | 0.497 | 3.5 | 16 |  | 0.00091 | 0.0895 | 0.1517 | -0.0072 | 1 | [1243 empty] | 0
P | 2018-02-08T17:20:49+08:00 | 2018-02-08T20:45:30+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI456G | CPSM Final Test | PASS | 129 | 163.22 | 0.37 | -1.22 | 56.5 | 53.09 | 1.21 | 0.13 | 0 | -12 | -2 | -0.9218 | 0.5276 | 3.9 | -10 |  | 0 | 0.0331 | 0.2171 | -0.0056 | 0.99997 | [213 empty] | 0
P | 2018-02-08T17:33:27+08:00 | 2018-02-08T20:58:24+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI455G | CPSM Final Test | PASS | 129 | 163.58 | 0.37 | -1.22 | 56.5 | 52.92 | 1.59 | 0.13 | 0 | -12 | 1 | -0.9398 | 0.496 | 5.2 | 14 |  | 0 | 0.1221 | 0.1923 | -0.0126 | 1 | [1176 empty] | 0
P | 2018-02-08T16:57:32+08:00 | 2018-02-08T21:41:57+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI454G | CPSM Final Test | PASS | 129 | 166.19 | 0.68 | -1.22 | 56.47 | 53.59 | 1.13 | 0.13 | 0 | -2 | -1 | -0.8599 | 0.479 | 4.9 | 15 |  | 0.00864 | 0.0632 | 0.1856 | -0.019 | 0.99999 | [1244 empty] | 0
P | 2018-02-08T18:51:38+08:00 | 2018-02-08T21:47:51+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI444G | CPSM Final Test | PASS | 129 | 159.44 | 0.27 | -0.69 | 56.78 | 55.03 | 1.82 | 0.13 | 0 | -3 | 0 | -0.9418 | 0.5226 | 1.4 | 34 |  | 0 | 0.0151 | 0.1988 | 0.0024 | 1 | [1245 empty] | 0
P | 2018-02-08T19:16:33+08:00 | 2018-02-08T21:59:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI448G | CPSM Final Test | PASS | 129 | 164.84 | -0.02 | -0.99 | 56.75 | 54.69 | 1.51 | 0.13 | -1 | -2 | 0 | -0.918 | 0.5078 | 0.3 | 1 |  | 0 | 0.0604 | 0.1007 | -0.0421 | 1 | [1181 empty] | 0
P | 2018-02-09T13:15:36+08:00 | 2018-02-09T14:09:41+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI464G | CPSM Final Test | FAIL | [5 empty] | 39.54
P | 2018-02-09T09:16:23+08:00 | 2018-02-09T14:31:34+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI463G | CPSM Final Test | FAIL
P | 2018-02-09T12:17:40+08:00 | 2018-02-09T16:24:28+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI458G | CPSM Final Test | PASS | 129 | 167.99 | -0.02 | -1.22 | 56.46 | 56.05 | 1.13 | 0.16 | 1 | -12 | -1 | -0.8759 | 0.5278 | 3.1 | 17 |  | 0 | 0.0994 | 0.1306 | 0.0015 | 1 | [1246 empty] | 0
P | 2018-02-09T12:30:58+08:00 | 2018-02-09T16:28:16+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI457G | CPSM Final Test | PASS | 129 | 159.17 | 0.37 | -1.15 | 56.72 | 59.52 | 1.51 | 0.16 | 1 | -11 | -1 | -0.9059 | 0.47 | 2.3 | -44 |  | 0 | 0.2186 | 0.3833 | -0.0115 | 1 | [1226 empty] | 0
P | 2018-02-09T12:08:42+08:00 | 2018-02-09T16:29:02+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI462G | CPSM Final Test | PASS | 129 | 165.47 | -0.02 | -0.77 | 56.91 | 51.39 | 1.44 | 0.13 | 1 | 2 | 3 | -0.84 | 0.4842 | 2.1 | 5 |  | 0 | 0.1424 | 0.0883 | 0.0122 | 0.99999 | [163 empty] | 0
P | 2018-02-09T11:53:58+08:00 | 2018-02-09T16:32:14+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI463G | CPSM Final Test | PASS | 129 | 154.76 | -0.02 | -0.77 | 56.73 | 58.59 | 1.9 | 0.13 | 1 | -2 | -1 | -0.8619 | 0.4988 | 1.9 | 11 |  | 0 | 0.0726 | 0.201 | -0.0124 | 0.99999 | [829 empty] | 0
P | 2018-02-09T13:42:38+08:00 | 2018-02-09T17:04:15+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI464G | CPSM Final Test | PASS | 129 | 158.63 | 0.37 | -0.84 | 56.54 | 52.75 | 1.44 | 0.13 | 1 | -3 | 1 | -0.826 | 0.51 | 1.8 | 26 |  | 0 | 0.1252 | 0.194 | 0.0361 | 0.99998 | [151 empty] | 0
P | 2018-02-09T14:27:11+08:00 | 2018-02-09T17:30:29+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI459G | CPSM Final Test | PASS | 129 | 164.75 | -0.11 | -0.84 | 56.84 | 53.59 | 1.59 | 0.13 | 0 | 6 | 1 | -0.89 | 0.5312 | 8.1 | -4 |  | 0 | 0.0759 | 0.2408 | 0.0102 | 0.99999 | [1196 empty] | 0
P | 2018-02-08T19:02:36+08:00 | 2018-02-09T19:39:50+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI451G | CPSM Final Test | PASS | 129 | 163.94 | -0.11 | -0.46 | 57.08 | 55.63 | 1.9 | 0.13 | 1 | -2 | 0 | -0.8999 | 0.538 | 1.5 | 2 |  | 0 | 0.105 | 0.1023 | 0.0022 | 0.99999 | [995 empty] | 0
P | 2018-02-09T04:24:00+08:00 | 2018-02-09T19:41:59+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI460G | CPSM Final Test | FAIL | [410 empty] | 6142
P | 2018-02-09T12:03:00+08:00 | 2018-02-09T20:47:35+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI460G | CPSM Final Test | PASS | 129 | 171.59 | 0.49 | -1.07 | 56.73 | 56.56 | 1.13 | 0.13 | 0 | -1 | -1 | -0.9199 | 0.4976 | 2.6 | 5 |  | 0 | 0.2599 | 0.1984 | -0.0393 | 1 | [389 empty] | 0
P | 2018-02-09T17:30:42+08:00 | 2018-02-09T20:57:44+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI461G | CPSM Final Test | PASS | 129 | 168.89 | 0.37 | -0.99 | 56.77 | 54.78 | 1.74 | 0.13 | 0 | -2 | 0 | -0.9638 | 0.4978 | 7.2 | 19 |  | 0.11506 | 0.1242 | 0.2871 | -0.017 | 0.99999 | [141 empty] | 0
P | 2018-02-09T17:11:06+08:00 | 2018-02-09T21:07:14+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI466G | CPSM Final Test | PASS | 129 | 168.8 | -0.11 | -1.15 | 56.71 | 59.1 | 1.51 | 0.13 | 1 | -2 | 4 | -0.852 | 0.539 | 0.8 | 43 |  | 0 | 0.1114 | 0.1683 | 0.0408 | 0.99998 | [1247 empty] | 0
P | 2018-02-09T17:01:55+08:00 | 2018-02-09T21:09:06+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI467G | CPSM Final Test | PASS | 129 | 169.16 | -0.02 | -0.38 | 56.8 | 52.92 | 1.82 | 0.13 | 0 | -2 | 2 | -0.8879 | 0.4714 | 8.8 | 22 |  | 0 | 0.0708 | 0.0677 | -0.0092 | 1 | [1248 empty] | 0
P | 2018-02-09T18:39:38+08:00 | 2018-02-09T21:29:37+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI450G | CPSM Final Test | PASS | 129 | 161.78 | 0.37 | -0.84 | 56.54 | 52.75 | 1.74 | 0.13 | 0 | -3 | 1 | -0.9339 | 0.549 | 0.6 | -23 |  | 0 | 0.1466 | 0.1165 | -0.0068 | 0.99999 | [604 empty] | 0
P | 2018-02-09T01:20:21+08:00 | 2018-02-10T09:19:34+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI465G | CPSM Final Test | FAIL
P | 2018-02-09T01:18:21+08:00 | 2018-02-10T09:21:34+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI465G | CPSM Final Test | FAIL
P | 2018-02-09T17:19:58+08:00 | 2018-02-10T10:24:36+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI465G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.84 | 56.61 | 53.26 | 1.51 | 0.13 | 0 | -2 | 1 | -0.9398 | 0.51 | 0.8 | 0 |  | 0.00085 | 0.1006 | 0.2044 | -0.001 | 1 | [1249 empty] | 0
P | 2018-02-11T13:23:04+08:00 | 2018-02-11T17:34:33+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM827G | CPSM Final Test | FAIL
P | 2018-02-11T14:53:15+08:00 | 2018-02-11T18:10:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | B18USM824G | CPSM Final Test | PASS |  | 161.78 | -0.02 | -0.54 | 56.75 | 53.34 | 1.51 | 0.13 | 0 | -11 | 1 | -0.9039 | 0.506 | 4.1 | 8 |  |  | 0.1403 | 0.112 | -0.0167 | 1 |  |  |  |  | 133 |  | 0 | [351 empty] | 0
P | 2018-02-11T15:25:08+08:00 | 2018-02-11T18:28:53+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM826G | CPSM Final Test | PASS |  | 165.29 | 0.46 | -0.54 | 57.02 | 54.19 | 1.51 | 0.13 | 1 | -12 | 1 | -0.8939 | 0.4714 | 0.8 | 20 |  |  | 0.0392 | 0.07 | -0.0228 | 1 |  |  |  |  | 133 |  | 0 | [382 empty] | 0
P | 2018-02-11T14:58:10+08:00 | 2018-02-11T18:31:46+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | B18USM825G | CPSM Final Test | PASS |  | 164.21 | -0.02 | -0.16 | 56.9 | 55.12 | 1.59 | 0.13 | -1 | -12 | 1 | -0.8699 | 0.5036 | 3.1 | 28 |  |  | 0.2646 | 0.086 | -0.0056 | 0.99998 |  |  |  |  | 133 |  | 0.0144 | [1243 empty] | 0
P | 2018-02-11T15:28:48+08:00 | 2018-02-11T18:39:59+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM827G | CPSM Final Test | PASS |  | 165.65 | 0.17 | -0.61 | 56.76 | 55.8 | 1.59 | 0.13 | 0 | -3 | 0 | -0.872 | 0.5328 | 1.3 | -6 |  |  | 0.0957 | 0.092 | -0.0247 | 1 |  |  |  |  | 133 |  | 0 | [842 empty] | 0
P | 2018-02-12T08:56:14+08:00 | 2018-02-12T12:00:45+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | B18USM832G | CPSM Final Test | FAIL
P | 2018-02-12T07:35:34+08:00 | 2018-02-12T13:21:25+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | B18USM832G | CPSM Final Test | FAIL
P | 2018-02-12T10:24:25+08:00 | 2018-02-12T13:26:49+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM835G | CPSM Final Test | PASS |  | 162.5 | 0.27 | -0.61 | 56.71 | 57.91 | 1.44 | 0.13 | -1 | 4 | 3 | -0.8102 | 0.5468 | 7.5 | 19 |  |  | 0.0798 | 0.1193 | -0.0041 | 0.99999 |  |  |  |  | 133 |  | 0 | [842 empty] | 0
P | 2018-02-12T10:20:37+08:00 | 2018-02-12T13:35:19+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM833G | CPSM Final Test | PASS |  | 163.85 | 0.37 | -0.61 | 56.53 | 53.42 | 1.82 | 0.13 | 1 | 2 | -1 | -0.878 | 0.51 | 4.1 | -6 |  |  | 0.0536 | 0.067 | -0.022 | 1 |  |  |  |  | 133 |  | 0 | [917 empty] | 0
P | 2018-02-12T10:12:38+08:00 | 2018-02-12T13:36:14+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | B18USM829G | CPSM Final Test | PASS |  | 162.32 | -0.02 | -0.54 | 56.63 | 54.02 | 1.74 | 0.13 | 0 | 4 | 1 | -0.836 | 0.5038 | 9.1 | 16 |  |  | 0.1521 | 0.1012 | -0.0399 | 0.99999 |  |  |  |  | 133 |  | 0 | [215 empty] | 0
P | 2018-02-12T10:37:27+08:00 | 2018-02-12T13:44:43+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | B18USM834G | CPSM Final Test | PASS |  | 162.59 | 0.17 | -0.31 | 57.04 | 53.17 | 1.67 | 0.13 | 1 | -2 | 1 | -0.8799 | 0.4998 | 7.2 | 2 |  |  | 0.0409 | 0.2362 | 0.0098 | 0.99999 |  |  |  |  | 133 |  | 0 | [1244 empty] | 0
P | 2018-02-12T10:16:36+08:00 | 2018-02-12T13:47:52+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | B18USM828G | CPSM Final Test | PASS |  | 164.93 | 0.46 | -0.54 | 56.94 | 55.37 | 1.36 | 0.13 | 1 | -3 | -3 | -0.8001 | 0.52 | 4.6 | 54 |  |  | 0.1154 | 0.0569 | -0.0314 | 0.99998 |  |  |  |  | 133 |  | 0 | [1245 empty] | 0
P | 2018-02-12T10:28:29+08:00 | 2018-02-12T15:30:23+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | B18USM832G | CPSM Final Test | PASS |  | 163.85 | -0.02 | -0.61 | 56.63 | 54.02 | 1.29 | 0.13 | 0 | -12 | 1 | -0.874 | 0.5446 | 1.3 | 26 |  |  | 0.0486 | 0.1683 | -0.0252 | 0.99999 |  |  |  |  | 133 |  | 0.00051 | [1246 empty] | 0
P | 2018-02-12T11:02:38+08:00 | 2018-02-12T16:40:05+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM831G | CPSM Final Test | FAIL
P | 2018-02-12T13:29:53+08:00 | 2018-02-12T16:53:20+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | B18USM830G | CPSM Final Test | PASS |  | 156.83 | 0.65 | -0.77 | 56.79 | 55.71 | 1.21 | 0.13 | -1 | -3 | 1 | -0.9678 | 0.544 | 1.6 | 53 |  |  | 0.1096 | 0.1673 | -0.0071 | 0.99999 |  |  |  |  | 133 |  | 0 | [822 empty] | 0
P | 2018-02-12T13:43:57+08:00 | 2018-02-12T16:58:25+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM836G | CPSM Final Test | PASS |  | 161.33 | 0.46 | -0.61 | 56.5 | 50.38 | 1.36 | 0.13 | 1 | 2 | -2 | -0.9139 | 0.4966 | 4.4 | 13 |  |  | 0.1304 | 0.1306 | -0.025 | 0.99999 |  |  |  |  | 133 |  | 0 | [729 empty] | 0
P | 2018-02-12T13:51:22+08:00 | 2018-02-12T18:31:59+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM831G | CPSM Final Test | PASS |  | 165.2 | 0.37 | -0.61 | 56.73 | 57.4 | 1.36 | 0.13 | 0 | -12 | 0 | -0.9059 | 0.511 | 0.9 | -94 |  |  | 0.193 | 0.1493 | -0.0039 | 0.99999 |  |  |  |  | 133 |  | 0 | [1247 empty] | 0
P | 2018-02-12T16:35:40+08:00 | 2018-02-12T20:29:31+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI469G | CPSM Final Test | PASS | 129 | 160.25 | 0.49 | -0.99 | 56.47 | 54.95 | 1.9 | 0.13 | 1 | -3 | -1 | -0.8639 | 0.5014 | 0.7 | -5 |  | 0 | 0.1374 | 0.2349 | 0.0353 | 0.99997 | [793 empty] | 0
P | 2018-02-12T16:30:33+08:00 | 2018-02-12T20:35:15+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI468G | CPSM Final Test | PASS | 129 | 166.91 | 0.17 | -0.38 | 56.93 | 54.78 | 2.12 | 0.13 | -1 | -3 | -1 | -0.9079 | 0.5112 | 2.8 | -5 |  | 0 | 0.0443 | 0.261 | -0.0196 | 0.99999 | [210 empty] | 0
P | 2018-02-12T16:26:38+08:00 | 2018-02-12T20:34:57+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI471G | CPSM Final Test | PASS | 129 | 167.72 | 0.37 | -1.15 | 56.6 | 54.95 | 1.06 | 0.13 | 1 | 2 | 1 | -0.9737 | 0.4836 | 1 | 24 |  | 0.0006 | 0.0809 | 0.1142 | 0.0087 | 1 | [1255 empty] | 0
P | 2018-02-12T16:22:23+08:00 | 2018-02-12T20:37:10+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI473G | CPSM Final Test | PASS | 129 | 164.39 | -0.02 | -1.15 | 56.41 | 60.37 | 1.59 | 0.13 | 1 | -2 | 1 | -0.88 | 0.5266 | 7.8 | -6 |  | 0 | 0.108 | 0.1846 | 0.0402 | 0.99996 | [1256 empty] | 0
P | 2018-02-12T16:18:15+08:00 | 2018-02-12T20:39:04+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI470G | CPSM Final Test | PASS | 129 | 166.64 | 0.27 | -0.99 | 56.8 | 56.73 | 1.74 | 0.13 | 0 | 11 | 3 | -0.9338 | 0.4986 | 4.5 | 9 |  | 0 | 0.1894 | 0.7775 | -0.0391 | 0.99999 | [1257 empty] | 0
P | 2018-02-12T17:33:56+08:00 | 2018-02-12T20:56:20+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI475G | CPSM Final Test | PASS | 129 | 166.19 | -0.02 | -0.77 | 56.63 | 54.61 | 2.2 | 0.13 | 0 | -2 | 3 | -0.9139 | 0.473 | 1.3 | -21 |  | 0 | 0.0439 | 0.3391 | 0.0159 | 0.99999 | [1258 empty] | 0
P | 2018-02-11T23:47:02+08:00 | 2018-02-13T11:11:01+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI476G | CPSM Final Test | FAIL
P | 2018-02-13T07:12:48+08:00 | 2018-02-13T12:16:00+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI477G | CPSM Final Test | FAIL | [312 empty] | 0
P | 2018-02-13T09:40:47+08:00 | 2018-02-13T12:57:13+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI474G | CPSM Final Test | PASS | 129 | 160.79 | 0.27 | -1.07 | 56.62 | 53.17 | 1.74 | 0.13 | 0 | -2 | -1 | -0.7901 | 0.4938 | 0.1 | -3 |  | 0 | 0.0357 | 0.1132 | 0.0205 | 0.99996 | [1259 empty] | 0
P | 2018-02-13T09:51:45+08:00 | 2018-02-13T13:01:01+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI478G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -0.99 | 56.83 | 52.32 | 1.44 | 0.13 | 0 | -11 | 0 | -0.816 | 0.5078 | 4.1 | 1 |  | 0 | 0.0568 | 0.1868 | -0.0067 | 0.99999 | [760 empty] | 0
P | 2018-02-13T06:21:01+08:00 | 2018-02-13T13:24:45+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI472G | CPSM Final Test | FAIL | [1281 empty] | 0
P | 2018-02-13T09:52:53+08:00 | 2018-02-13T14:30:49+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI472G | CPSM Final Test | PASS | 129 | 165.56 | 0.27 | -1.3 | 56.77 | 54.78 | 1.67 | 0.13 | 0 | -11 | 1 | -0.8461 | 0.5044 | 0.8 | 16 |  | 0 | 0.0842 | 0.2731 | -0.0022 | 1 | [1260 empty] | 0
P | 2018-02-13T04:36:34+08:00 | 2018-02-13T14:52:14+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI477G | CPSM Final Test | FAIL | [312 empty] | 0
P | 2018-02-11T19:16:24+08:00 | 2018-02-13T15:41:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI476G | CPSM Final Test | FAIL | [17 empty] | 2.2116
P | 2018-02-13T11:36:08+08:00 | 2018-02-13T15:53:28+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI479G | CPSM Final Test | PASS | 129 | 163.94 | 0.46 | -1.07 | 56.52 | 55.71 | 1.51 | 0.13 | 1 | -2 | 1 | -0.8161 | 0.537 | 1.8 | 9 |  | 0 | 0.1805 | 0.325 | -0.0208 | 0.99999 | [1261 empty] | 0
P | 2018-02-13T11:40:17+08:00 | 2018-02-13T15:55:28+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI480G | CPSM Final Test | PASS | 129 | 161.24 | -0.21 | -1.07 | 56.21 | 53.93 | 1.36 | 0.13 | 0 | -3 | -1 | -0.8101 | 0.4872 | 0.9 | 8 |  | 0.01824 | 0.1335 | 0.0993 | 0.0348 | 0.99999 | [1262 empty] | 0
P | 2018-02-13T11:43:59+08:00 | 2018-02-13T15:57:11+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI481G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.92 | 56.72 | 56.3 | 2.05 | 0.13 | 1 | -3 | -2 | -0.7603 | 0.4902 | 0.3 | 8 |  | 0 | 0.0672 | 0.0969 | -0.0144 | 0.99997 | [950 empty] | 0
P | 2018-02-13T09:44:24+08:00 | 2018-02-13T16:05:56+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI477G | CPSM Final Test | PASS | 129 | 165.38 | 0.37 | -1.3 | 56.44 | 54.95 | 1.21 | 0.13 | 0 | -11 | -1 | -0.864 | 0.4952 | 6.7 | -4 |  | 0.01358 | 0.0997 | 0.729 | -0.014 | 1 | [291 empty] | 0
P | 2018-02-12T17:29:02+08:00 | 2018-02-13T16:44:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI476G | CPSM Final Test | PASS | 129 | 161.33 | 0.46 | -0.99 | 56.6 | 54.27 | 1.21 | 0.13 | 0 | 2 | -1 | -0.9159 | 0.4992 | 6.7 | 1 |  | 0.00795 | 0.0445 | 0.037 | -0.0185 | 0.99999 | [217 empty] | 0
P | 2018-02-13T13:42:46+08:00 | 2018-02-13T16:45:17+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI483G | CPSM Final Test | PASS | 129 | 166.91 | -0.02 | -1.07 | 56.52 | 53.34 | 1.67 | 0.13 | 1 | 2 | -1 | -0.7702 | 0.4992 | 10.1 | -52 |  | 0 | 0.1032 | 0.1554 | 0.0513 | 0.99998 | [854 empty] | 0
P | 2018-02-13T13:48:58+08:00 | 2018-02-13T16:52:23+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI484G | CPSM Final Test | PASS | 129 | 159.26 | 0.17 | -0.92 | 56.76 | 53.85 | 1.59 | 0.13 | -1 | 4 | 3 | -0.9238 | 0.4932 | 0.3 | 11 |  | 0 | 0.074 | 0.2189 | 0.0279 | 0.99999 | [618 empty] | 0
P | 2018-02-13T13:56:52+08:00 | 2018-02-13T19:21:10+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI488G | CPSM Final Test | FAIL
P | 2018-02-13T16:46:56+08:00 | 2018-02-13T20:37:57+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI486G | CPSM Final Test | PASS | 129 | 163.4 | 0.37 | -0.92 | 56.77 | 54.53 | 1.44 | 0.13 | -1 | -10 | -1 | -0.9339 | 0.5 | 11.8 | 9 |  | 0 | 0.2029 | 0.1081 | -0.0072 | 0.99999 | [1263 empty] | 0
P | 2018-02-13T16:42:34+08:00 | 2018-02-13T20:40:20+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI487G | CPSM Final Test | PASS | 129 | 168.98 | 0.46 | -0.99 | 56.61 | 54.19 | 1.67 | 0.13 | -1 | -3 | -2 | -0.9239 | 0.519 | 1.4 | 16 |  | 0.0093 | 0.1286 | 0.2503 | 0.0053 | 0.99998 | [586 empty] | 0
P | 2018-02-13T16:50:26+08:00 | 2018-02-13T20:51:07+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI482G | CPSM Final Test | PASS | 129 | 164.66 | -0.02 | -1.53 | 56.23 | 51.99 | 1.29 | 0.13 | -1 | -1 | -1 | -0.9379 | 0.4792 | 4.2 | 29 |  | 0 | 0.1427 | 0.0948 | -0.0358 | 1 | [1260 empty] | 0
P | 2018-02-13T16:56:37+08:00 | 2018-02-13T20:58:06+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI485G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -1.15 | 56.23 | 53.59 | 1.44 | 0.13 | 1 | 2 | -1 | -0.9917 | 0.491 | 0.8 | -14 |  | 0 | 0.0469 | 0.4644 | -0.0018 | 1 | [291 empty] | 0
P | 2018-02-13T16:39:01+08:00 | 2018-02-13T21:05:56+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI488G | CPSM Final Test | PASS | 129 | 166.91 | -0.02 | -0.61 | 56.61 | 53.42 | 1.67 | 0.13 | 0 | 3 | 0 | -0.9158 | 0.5314 | 13 | -11 |  | 0 | 0.0574 | 0.0576 | 0.0102 | 0.99999 | [1264 empty] | 0
P | 2018-02-13T18:28:43+08:00 | 2018-02-13T21:18:03+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI489G | CPSM Final Test | PASS | 129 | 164.75 | 0.17 | -0.54 | 57.1 | 55.37 | 1.74 | 0.13 | 0 | -13 | 1 | -0.9718 | 0.55 | 5 | 44 |  | 0 | 0.0963 | 0.0664 | -0.0176 | 1 | [217 empty] | 0
P | 2018-02-14T07:40:48+08:00 | 2018-02-14T12:13:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI491G | CPSM Final Test | FAIL | [13 empty] | 19.3
P | 2018-02-14T10:00:31+08:00 | 2018-02-14T13:05:48+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI492G | CPSM Final Test | PASS | 129 | 164.21 | 0.37 | -1.3 | 56.33 | 54.02 | 1.13 | 0.13 | 1 | -2 | 1 | -0.894 | 0.513 | 1.3 | -9 |  | 0 | 0.1983 | 0.6403 | -0.0489 | 0.99999 | [1265 empty] | 0
P | 2018-02-14T10:07:24+08:00 | 2018-02-14T13:13:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI493G | CPSM Final Test | PASS | 129 | 164.93 | 0.17 | -1.07 | 56.47 | 54.1 | 1.59 | 0.13 | -1 | 1 | 0 | -0.834 | 0.4978 | 3.7 | 4 |  | 0 | 0.0608 | 0.4542 | -0.0172 | 0.99999 | [539 empty] | 0
P | 2018-02-14T10:26:43+08:00 | 2018-02-14T13:17:03+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI495G | CPSM Final Test | PASS | 129 | 163.22 | 0.56 | -1.07 | 56.46 | 53.85 | 1.36 | 0.13 | 0 | -2 | 2 | -0.8739 | 0.519 | 0.4 | 6 |  | 0 | 0.1142 | 0.1159 | -0.0118 | 0.99999 | [738 empty] | 0
P | 2018-02-14T10:30:26+08:00 | 2018-02-14T13:18:57+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI494G | CPSM Final Test | PASS | 129 | 159.89 | 0.37 | -0.84 | 56.8 | 54.02 | 1.36 | 0.13 | -1 | 2 | 3 | -0.9059 | 0.4958 | 8.1 | 16 |  | 0 | 0.1347 | 0.118 | -0.0274 | 1 | [13 empty] | 0
P | 2018-02-14T10:07:54+08:00 | 2018-02-14T13:38:15+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI490G | CPSM Final Test | PASS | 129 | 163.13 | -0.21 | -0.54 | 56.84 | 55.2 | 2.05 | 0.13 | 0 | -2 | 0 | -0.8641 | 0.498 | 0.8 | 35 |  | 0 | 0.1523 | 0.2785 | 0.0248 | 0.99997 | [1266 empty] | 0
P | 2018-02-14T09:57:21+08:00 | 2018-02-14T14:14:23+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI491G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -1.07 | 56.45 | 53.09 | 1.59 | 0.13 | 1 | 3 | 1 | -0.8161 | 0.5002 | 9.7 | -12 |  | 0 | 0.0786 | 0.3429 | 0.0308 | 0.99999 | [824 empty] | 0
P | 2018-02-14T11:02:28+08:00 | 2018-02-14T17:04:52+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI496G | CPSM Final Test | FAIL | [1288 empty] | 12535
P | 2018-02-14T14:10:32+08:00 | 2018-02-14T17:17:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI499G | CPSM Final Test | PASS | 129 | 159.62 | 0.37 | -0.99 | 56.52 | 54.61 | 1.06 | 0.13 | 0 | 3 | -1 | -0.888 | 0.501 | 10 | 21 |  | 0 | 0.1093 | 0.1625 | -0.0353 | 1 | [1268 empty] | 0
P | 2018-02-14T14:07:02+08:00 | 2018-02-14T17:21:07+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI497G | CPSM Final Test | PASS | 129 | 163.67 | -0.11 | -0.61 | 56.91 | 55.96 | 1.59 | 0.13 | -1 | -3 | 1 | -0.9099 | 0.4894 | 2.2 | 79 |  | 0 | 0.074 | 0.0851 | -0.008 | 1 | [189 empty] | 0
P | 2018-02-14T13:59:24+08:00 | 2018-02-14T17:28:34+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI498G | CPSM Final Test | PASS | 129 | 170.15 | 0.08 | -1.15 | 56.19 | 52.58 | 1.82 | 0.13 | 0 | -2 | -1 | -0.8979 | 0.513 | 4.8 | 13 |  | 0 | 0.0919 | 0.0987 | 0.044 | 0.99997 | [186 empty] | 0
P | 2018-02-14T16:06:55+08:00 | 2018-02-14T19:02:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI500G | CPSM Final Test | PASS | 129 | 157.46 | 0.37 | -0.99 | 56.63 | 53.68 | 1.59 | 0.13 | -1 | -2 | -1 | -0.9318 | 0.502 | 4.4 | -3 |  | 0 | 0.0554 | 0.2981 | 0.0031 | 0.99998 | [1269 empty] | 0
P | 2018-02-14T16:10:50+08:00 | 2018-02-14T19:02:56+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI502G | CPSM Final Test | PASS | 129 | 165.47 | 0.37 | -0.84 | 56.82 | 55.12 | 1.51 | 0.13 | -1 | -2 | -1 | -0.9498 | 0.4974 | 2.1 | 29 |  | 0 | 0.1294 | 0.3332 | 0.0117 | 0.99999 | [1270 empty] | 0
P | 2018-02-14T16:00:44+08:00 | 2018-02-14T19:08:13+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI501G | CPSM Final Test | PASS | 129 | 163.31 | 0.17 | -0.84 | 57.07 | 53.85 | 1.82 | 0.22 | 0 | -2 | -1 | -0.9597 | 0.4964 | 1.4 | 74 |  | 0 | 0.2523 | 0.1459 | -0.0119 | 1 | [1271 empty] | 0
P | 2018-02-11T13:18:22+08:00 | 2018-02-17T14:48:58+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI496G | CPSM Final Test | FAIL | [19 empty] | -0.6592
P | 2018-02-14T14:03:40+08:00 | 2018-02-17T15:52:06+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI496G | CPSM Final Test | PASS | 129 | 165.11 | 0.37 | -0.84 | 56.5 | 53 | 1.44 | 0.13 | 0 | -2 | 4 | -0.8541 | 0.5068 | 1.3 | 0 |  | 0 | 0.1782 | 0.2364 | 0.0079 | 0.99999 | [1267 empty] | 0
P | 2018-02-21T16:17:06+08:00 | 2018-02-21T19:26:53+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | B18USM841G | CPSM Final Test | PASS |  | 164.12 | 0.37 | -0.54 | 56.79 | 54.69 | 1.82 | 0.16 | -1 | -11 | -1 | -0.9418 | 0.5396 | 5.1 | 13 |  |  | 0.1457 | 0.2746 | 0.0132 | 1 |  |  |  |  | 133 |  | 0 | [265 empty] | 0
P | 2018-02-21T16:02:43+08:00 | 2018-02-21T19:24:05+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM839G | CPSM Final Test | PASS |  | 162.95 | 0.46 | -1.15 | 56.64 | 54.1 | 0.91 | 0.13 | 1 | 3 | 1 | -0.9638 | 0.502 | 0.9 | 97 |  |  | 0.2066 | 0.2067 | 0.0061 | 0.99999 |  |  |  |  | 133 |  | 0.03336 | [801 empty] | 0
P | 2018-02-21T15:58:54+08:00 | 2018-02-21T19:24:31+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM840G | CPSM Final Test | PASS |  | 161.06 | 0.37 | -0.54 | 56.7 | 58.76 | 0.98 | 0.13 | 1 | -2 | -1 | -0.9818 | 0.5438 | 0.3 | -1 | 0 |  | 0.0879 | 0.0861 | -0.0012 | 0.99999 |  |  |  |  | 133 |  | 0
P | 2018-02-21T16:06:18+08:00 | 2018-02-21T19:24:27+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | B18USM842G | CPSM Final Test | PASS |  | 160.79 | -0.02 | -0.54 | 56.91 | 55.63 | 1.82 | 0.13 | 0 | 2 | 3 | -0.9558 | 0.4982 | 4.2 | 54 |  |  | 0.0662 | 0.1435 | -0.0609 | 0.99999 |  |  |  |  | 133 |  | 0 | [1265 empty] | 0
P | 2018-02-21T15:54:53+08:00 | 2018-02-21T19:24:46+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | B18USM838G | CPSM Final Test | PASS |  | 159.98 | 0.37 | -0.31 | 57.22 | 54.19 | 1.67 | 0.13 | 0 | -12 | -1 | -0.9398 | 0.508 | 3 | 44 |  |  | 0.1948 | 0.0514 | 0.0286 | 0.99999 |  |  |  |  | 133 |  | 0 | [1260 empty] | 0
P | 2018-02-21T12:09:22+08:00 | 2018-02-21T19:33:33+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | B18USM837G | CPSM Final Test | FAIL
P | 2018-02-21T15:51:27+08:00 | 2018-02-21T19:34:11+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | B18USM837G | CPSM Final Test | PASS |  | 162.77 | 0.56 | -0.84 | 56.77 | 54.61 | 1.67 | 0.13 | 0 | 1 | 1 | -0.9477 | 0.509 | 1.4 | -9 |  |  | 0.054 | 0.1548 | -0.0107 | 0.99999 |  |  |  |  | 133 |  | 0 | [1266 empty] | 0
P | 2018-02-22T09:22:51+08:00 | 2018-02-22T10:45:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM848G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-02-22T07:50:46+08:00 | 2018-02-22T12:31:27+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | B18USM846G | CPSM Final Test | FAIL
P | 2018-02-22T10:04:01+08:00 | 2018-02-22T13:13:25+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM848G | CPSM Final Test | PASS |  | 158.81 | 0.59 | -0.16 | 57.12 | 53.26 | 1.67 | 0.13 | 0 | 3 | 1 | -0.858 | 0.4582 | 0.5 | 12 |  |  | 0.0705 | 0.3659 | -0.028 | 0.99999 |  |  |  |  | 133 |  | 0 | [1267 empty] | 0
P | 2018-02-22T06:49:29+08:00 | 2018-02-22T13:26:01+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM847G | CPSM Final Test | FAIL
P | 2018-02-22T06:32:31+08:00 | 2018-02-22T13:28:13+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | B18USM844G | CPSM Final Test | FAIL
P | 2018-02-22T10:00:22+08:00 | 2018-02-22T13:31:57+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | B18USM844G | CPSM Final Test | PASS |  | 170.42 | -0.02 | -0.46 | 56.77 | 53.85 | 1.9 | 0.13 | 1 | -10 | 0 | -0.8701 | 0.5012 | 9.7 | 13 |  |  | 0.0767 | 0.1013 | -0.0361 | 1 |  |  |  |  | 133 |  | 0 | [1268 empty] | 0
P | 2018-02-22T10:18:07+08:00 | 2018-02-22T13:35:35+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | B18USM845G | CPSM Final Test | PASS |  | 165.92 | 0.46 | -0.46 | 56.75 | 56.13 | 1.21 | 0.13 | -1 | -10 | 2 | -0.8659 | 0.5048 | 0.8 | -11 |  |  | 0.107 | 0.2329 | 0.0065 | 0.99999 |  |  |  |  | 133 |  | 0 | [1269 empty] | 0
P | 2018-02-22T06:50:02+08:00 | 2018-02-22T13:32:11+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | B18USM846G | CPSM Final Test | FAIL
P | 2018-02-22T10:07:45+08:00 | 2018-02-22T13:38:49+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM847G | CPSM Final Test | PASS |  | 168.71 | 0.27 | -0.54 | 57.03 | 55.12 | 1.97 | 0.13 | -1 | -2 | 0 | -0.856 | 0.501 | 8.3 | 108 |  |  | 0.0357 | 0.1574 | 0.0018 | 1 |  |  |  |  | 133 |  | 0 | [1270 empty] | 0
P | 2018-02-22T06:35:33+08:00 | 2018-02-22T14:01:52+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | B18USM849G | CPSM Final Test | FAIL | [27 empty] | 0.05879
P | 2018-02-22T05:02:29+08:00 | 2018-02-22T14:51:24+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | B18USM843G | CPSM Final Test | FAIL | [584 empty] | 0
P | 2018-02-22T10:18:42+08:00 | 2018-02-22T15:04:50+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | B18USM849G | CPSM Final Test | PASS |  | 165.02 | 0.56 | -0.46 | 56.77 | 53.51 | 1.44 | 0.13 | 1 | -2 | 1 | -0.91 | 0.493 | 1 | -5 |  |  | 0.1727 | 0.464 | -0.0451 | 0.99998 |  |  |  |  | 133 |  | 0 | [1271 empty] | 0
P | 2018-02-22T10:11:06+08:00 | 2018-02-22T15:46:47+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | B18USM846G | CPSM Final Test | PASS |  | 168.17 | 0.37 | -0.46 | 56.94 | 54.36 | 1.74 | 0.13 | 0 | -3 | 0 | -0.844 | 0.4956 | 1.9 | 54 |  |  | 0.087 | 0.1187 | -0.0345 | 0.99999 |  |  |  |  | 133 |  | 0 | [1272 empty] | 0
P | 2018-02-22T09:56:56+08:00 | 2018-02-22T15:57:06+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | B18USM843G | CPSM Final Test | PASS |  | 162.14 | 0.37 | -0.54 | 57 | 54.78 | 1.44 | 0.13 | -1 | -11 | -1 | -0.8939 | 0.498 | 1.5 | 9 |  |  | 0.0794 | 0.2702 | -0.0077 | 0.99998 |  |  |  |  | 133 |  | 0.00613 | [556 empty] | 0
P | 2018-02-22T11:38:27+08:00 | 2018-02-22T16:57:55+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM851G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-02-22T14:21:45+08:00 | 2018-02-22T17:46:19+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | B18USM852G | CPSM Final Test | PASS |  | 164.03 | 0.46 | -0.54 | 56.92 | 57.83 | 1.59 | 0.13 | -1 | 2 | 1 | -0.8382 | 0.5076 | 0 | 46 |  |  | 0.0897 | 0.0815 | -0.0075 | 0.99999 |  |  |  |  | 133 |  | 0 | [229 empty] | 0
P | 2018-02-22T14:33:09+08:00 | 2018-02-22T17:49:31+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI505G | CPSM Final Test | PASS | 129 | 164.3 | 0.17 | -0.61 | 57.05 | 50.97 | 1.59 | 0.13 | -1 | 4 | 1 | -0.9078 | 0.5 | 6.3 | 7 |  | 0 | 0.0899 | 0.1333 | 0.0195 | 0.99997 | [367 empty] | 0
P | 2018-02-22T14:26:56+08:00 | 2018-02-22T17:53:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI503G | CPSM Final Test | PASS | 129 | 163.04 | 0.37 | 0 | 56.91 | 54.44 | 1.36 | 0.13 | 1 | -2 | 1 | -0.7842 | 0.5104 | 4.9 | 28 |  | 0 | 0.065 | 0.3122 | 0.0065 | 0.99994 | [749 empty] | 0
P | 2018-02-22T14:14:09+08:00 | 2018-02-22T18:00:11+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | B18USM850G | CPSM Final Test | PASS |  | 167.18 | 0.17 | -0.31 | 56.76 | 54.95 | 1.82 | 0.13 | 0 | 2 | 1 | -0.9019 | 0.5328 | 9.1 | 17 |  |  | 0.1661 | 0.1176 | -0.0436 | 0.99997 |  |  |  |  | 133 |  | 0 | [806 empty] | 0
P | 2018-02-22T14:18:11+08:00 | 2018-02-22T20:24:03+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | B18USM851G | CPSM Final Test | PASS |  | 160.16 | -0.02 | -0.54 | 56.76 | 54.36 | 1.59 | 0.13 | -1 | 3 | 2 | -0.9498 | 0.5412 | 1.7 | 22 |  |  | 0.042 | 0.0531 | -0.0315 | 0.99995 |  |  |  |  | 133 |  | 0 | [1273 empty] | 0
P | 2018-02-22T17:33:27+08:00 | 2018-02-22T21:01:18+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI508G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.54 | 57.34 | 56.13 | 1.59 | 0.13 | 0 | 2 | -1 | -0.9578 | 0.5046 | 2.2 | 12 |  | 0.00607 | 0.0778 | 0.4715 | -0.0265 | 1 | [1281 empty] | 0
P | 2018-02-22T17:29:46+08:00 | 2018-02-22T21:12:50+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI509G | CPSM Final Test | PASS | 129 | 161.06 | -0.02 | -0.61 | 57.01 | 54.27 | 1.9 | 0.13 | -1 | 2 | 2 | -0.9617 | 0.5092 | 10 | 8 |  | 0.01197 | 0.2402 | 0.2766 | 0.0371 | 0.99997 | [563 empty] | 0
P | 2018-02-22T17:37:38+08:00 | 2018-02-22T21:15:17+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI507G | CPSM Final Test | PASS | 129 | 167.72 | -0.02 | -0.54 | 56.81 | 51.39 | 1.36 | 0.13 | 0 | -2 | 1 | -0.9318 | 0.505 | 7 | 26 |  | 0.00204 | 0.3051 | 0.2378 | -0.0036 | 0.99999 | [1279 empty] | 0
P | 2018-02-23T09:39:43+08:00 | 2018-02-23T13:40:14+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI511G | CPSM Final Test | PASS | 129 | 165.47 | 0.17 | -0.08 | 57.43 | 54.95 | 1.97 | 0.13 | 0 | -2 | 1 | -0.8899 | 0.5262 | 2.7 | -1 |  | 0.01429 | 0.1919 | 0.216 | 0.0419 | 0.99996 | [1282 empty] | 0
P | 2018-02-23T09:47:02+08:00 | 2018-02-23T13:48:23+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI513G | CPSM Final Test | PASS | 129 | 160.88 | 0.37 | -0.61 | 57.05 | 55.12 | 1.44 | 0.13 | 1 | -13 | 2 | -0.872 | 0.534 | 4.6 | 17 |  | 0 | 0.0903 | 0.1826 | -0.0107 | 0.99999 | [700 empty] | 0
P | 2018-02-23T09:43:35+08:00 | 2018-02-23T13:52:05+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI512G | CPSM Final Test | PASS | 129 | 167.18 | -0.02 | -0.23 | 57.12 | 51.82 | 1.67 | 0.13 | 1 | 4 | 1 | -0.886 | 0.4918 | 0.8 | 16 |  | 0 | 0.1484 | 0.2213 | -0.0166 | 0.99999 | [1283 empty] | 0
P | 2018-02-23T09:58:18+08:00 | 2018-02-23T13:57:41+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI504G | CPSM Final Test | PASS | 129 | 161.33 | -0.02 | -0.54 | 56.7 | 59.27 | 1.97 | 0.13 | 0 | -11 | 1 | -0.8639 | 0.511 | 1.2 | 18 |  | 0.01138 | 0.1098 | 0.2765 | 0.0273 | 1 | [551 empty] | 0
P | 2018-02-23T04:26:30+08:00 | 2018-02-23T14:44:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI510G | CPSM Final Test | FAIL | [1305 empty] | 0
P | 2018-02-23T09:35:29+08:00 | 2018-02-23T15:50:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI510G | CPSM Final Test | PASS | 129 | 157.37 | 0.27 | -1.07 | 56.52 | 57.66 | 1.13 | 0.13 | 0 | -2 | 1 | -0.8858 | 0.5034 | 3.6 | 5 |  | 0 | 0.1338 | 0.7074 | -0.0221 | 1 | [1284 empty] | 0
P | 2018-02-23T12:34:09+08:00 | 2018-02-23T16:06:06+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI515G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | -0.23 | 57.11 | 53.59 | 1.82 | 0.13 | 0 | -11 | 2 | -0.9299 | 0.4898 | 1.7 | 1 |  | 0 | 0.0584 | 0.7342 | 0.0034 | 0.99999 | [1285 empty] | 0
P | 2018-02-23T12:30:44+08:00 | 2018-02-23T16:10:23+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI514G | CPSM Final Test | PASS | 129 | 166.46 | -0.02 | -0.46 | 57.49 | 55.71 | 2.2 | 0.13 | 1 | 2 | 1 | -0.844 | 0.501 | 6.3 | -28 |  | 0 | 0.2006 | 0.0785 | 0.0213 | 0.99996 | [148 empty] | 0
P | 2018-02-23T14:41:10+08:00 | 2018-02-23T17:31:25+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI516G | CPSM Final Test | PASS | 129 | 167.72 | -0.02 | -0.31 | 57.07 | 55.46 | 1.82 | 0.13 | -1 | 1 | 1 | -0.9537 | 0.509 | 1.6 | -10 |  | 0 | 0.054 | 0.0606 | -0.021 | 1 | [1286 empty] | 0
P | 2018-02-23T11:10:48+08:00 | 2018-02-23T18:24:46+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI506G | CPSM Final Test | FAIL
P | 2018-02-23T10:49:27+08:00 | 2018-02-23T18:25:24+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI517G | CPSM Final Test | FAIL
P | 2018-02-23T14:47:47+08:00 | 2018-02-23T20:10:50+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI506G | CPSM Final Test | PASS | 129 | 162.41 | -0.4 | -0.77 | 57.2 | 53.93 | 1.67 | 0.13 | 1 | -10 | 0 | -0.9736 | 0.516 | 4.8 | 39 |  | 0 | 0.0472 | 0.1505 | 0.0037 | 1 | [1287 empty] | 0
P | 2018-02-23T17:25:42+08:00 | 2018-02-23T21:00:27+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI522G | CPSM Final Test | PASS | 129 | 164.93 | -0.02 | -0.46 | 57.12 | 56.13 | 1.59 | 0.13 | 1 | -3 | 1 | -0.8938 | 0.548 | 8.3 | 19 |  | 0 | 0.0456 | 0.1711 | 0.0297 | 0.99999 | [1284 empty] | 0
P | 2018-02-23T17:29:23+08:00 | 2018-02-23T21:04:27+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI524G | CPSM Final Test | PASS | 129 | 164.75 | 0.17 | -0.54 | 56.82 | 55.12 | 1.59 | 0.13 | 0 | 2 | 0 | -0.894 | 0.491 | 10.4 | 10 |  | 0 | 0.0878 | 0.1655 | -0.0481 | 1 | [1288 empty] | 0
P | 2018-02-23T17:38:27+08:00 | 2018-02-23T21:09:10+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI520G | CPSM Final Test | PASS | 129 | 161.06 | -0.02 | -0.77 | 56.9 | 54.1 | 1.67 | 0.13 | 0 | -8 | 1 | -0.9198 | 0.4938 | 0.3 | 9 |  | 0 | 0.1213 | 0.1774 | -0.0245 | 0.99999 | [1289 empty] | 0
P | 2018-02-23T17:37:33+08:00 | 2018-02-23T21:13:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI521G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.92 | 56.84 | 54.19 | 1.97 | 0.13 | -1 | -12 | 1 | -0.9558 | 0.501 | 5.9 | 6 |  | 0 | 0.0585 | 0.432 | 0.0107 | 0.99999 | [1290 empty] | 0
P | 2018-02-23T17:42:55+08:00 | 2018-02-23T21:13:57+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI523G | CPSM Final Test | PASS | 129 | 160.25 | 0.27 | -0.77 | 57.11 | 55.71 | 1.44 | 0.13 | 0 | 4 | 1 | -0.8221 | 0.4852 | 4.3 | 12 |  | 0 | 0.0694 | 0.204 | 0.0314 | 0.99998 | [1291 empty] | 0
P | 2018-02-22T17:45:31+08:00 | 2018-02-24T11:29:20+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI517G | CPSM Final Test | FAIL | [20 empty] | 0.9975
P | 2018-02-24T08:05:14+08:00 | 2018-02-24T11:50:15+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI518G | CPSM Final Test | FAIL | [13 empty] | 16.8
P | 2018-02-23T14:37:25+08:00 | 2018-02-24T12:33:22+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI517G | CPSM Final Test | PASS | 129 | 167.54 | -0.02 | -0.31 | 56.82 | 52.92 | 1.82 | 0.13 | 0 | -2 | 1 | -0.9059 | 0.498 | 3.8 | -6 |  | 0 | 0.1551 | 0.1838 | 0.1157 | 0.99993 | [1292 empty] | 0
P | 2018-02-24T06:29:41+08:00 | 2018-02-24T12:55:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI526G | CPSM Final Test | FAIL
P | 2018-02-24T09:50:43+08:00 | 2018-02-24T13:02:14+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI519G | CPSM Final Test | PASS | 129 | 165.56 | -0.02 | -0.31 | 57.21 | 52.24 | 1.82 | 0.13 | 0 | -2 | 1 | -0.8779 | 0.496 | 0.1 | 71 |  | 0 | 0.189 | 0.1444 | -0.0001 | 0.99999 | [1293 empty] | 0
P | 2018-02-24T09:46:28+08:00 | 2018-02-24T13:07:15+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI525G | CPSM Final Test | PASS | 129 | 165.83 | -0.02 | -0.16 | 57.52 | 53.09 | 2.2 | 0.13 | 0 | -3 | 1 | -0.86 | 0.492 | 2.8 | 1 |  | 0 | 0.3218 | 0.3365 | -0.0107 | 0.99996 | [1294 empty] | 0
P | 2018-02-24T09:36:54+08:00 | 2018-02-24T13:26:49+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI527G | CPSM Final Test | PASS | 129 | 162.86 | -0.21 | -0.16 | 57.23 | 54.1 | 1.67 | 0.13 | 0 | 2 | 1 | -0.8919 | 0.509 | 10.4 | 3 |  | 0 | 0.0992 | 0.1888 | 0.0008 | 0.99998 | [1295 empty] | 0
P | 2018-02-24T09:42:26+08:00 | 2018-02-24T14:39:19+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI526G | CPSM Final Test | PASS | 129 | 160.7 | -0.21 | -0.69 | 56.76 | 54.36 | 1.82 | 0.13 | -1 | 2 | 1 | -0.9119 | 0.5046 | 3.7 | 17 |  | 0 | 0.0676 | 0.1145 | -0.0051 | 0.99999 | [1296 empty] | 0
P | 2018-02-23T09:53:06+08:00 | 2018-02-25T10:02:23+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI518G | CPSM Final Test | FAIL | [13 empty] | 15.8
P | 2018-02-24T16:33:25+08:00 | 2018-02-25T11:55:13+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI528G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.61 | 57.07 | 53.09 | 1.82 | 0.13 | 1 | 2 | 0 | -0.9278 | 0.5004 | 0.9 | -8 |  | 0 | 0.0626 | 0.3083 | 0.0462 | 0.99998 | [375 empty] | 0
P | 2018-02-24T09:57:45+08:00 | 2018-02-25T12:02:21+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI518G | CPSM Final Test | PASS | 129 | 160.34 | -0.21 | -0.84 | 57 | 55.03 | 2.05 | 0.13 | 1 | -11 | 1 | -0.9118 | 0.496 | 3.5 | 57 |  | 0 | 0.1093 | 0.4252 | 0.0342 | 0.99999 | [669 empty] | 0
P | 2018-02-24T16:36:16+08:00 | 2018-02-25T12:09:14+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI529G | CPSM Final Test | PASS | 129 | 168.89 | -0.02 | -0.92 | 56.61 | 54.61 | 1.21 | 0.13 | 1 | -3 | 1 | -0.864 | 0.4828 | 8.7 | 8 |  | 0 | 0.1367 | 0.5076 | 0.0074 | 1 | [1297 empty] | 0
P | 2018-02-25T06:25:28+08:00 | 2018-02-25T12:17:26+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI535G | CPSM Final Test | FAIL
P | 2018-02-25T09:35:50+08:00 | 2018-02-25T12:33:59+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI530G | CPSM Final Test | PASS | 129 | 161.42 | -0.02 | -0.69 | 56.84 | 53.59 | 1.59 | 0.13 | 0 | -2 | 1 | -0.7963 | 0.49 | 3 | 17 |  | 0 | 0.2091 | 0.202 | -0.0066 | 1 | [721 empty] | 0
P | 2018-02-25T09:31:28+08:00 | 2018-02-25T12:37:23+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI534G | CPSM Final Test | PASS | 129 | 161.96 | -0.21 | -0.61 | 56.92 | 51.99 | 1.9 | 0.13 | 0 | -12 | 0 | -0.8022 | 0.501 | 3.1 | 0 |  | 0 | 0.087 | 0.0902 | 0.0017 | 1 | [685 empty] | 0
P | 2018-02-25T09:26:08+08:00 | 2018-02-25T12:46:25+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI533G | CPSM Final Test | PASS | 129 | 165.47 | -0.02 | -0.69 | 56.83 | 53.09 | 1.82 | 0.13 | 0 | 1 | -3 | -0.86 | 0.4972 | 3.7 | 9 |  | 0 | 0.0579 | 0.0598 | 0.0337 | 0.99997 | [1298 empty] | 0
P | 2018-02-25T09:39:07+08:00 | 2018-02-25T13:09:02+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI531G | CPSM Final Test | PASS | 129 | 164.75 | -0.49 | -0.77 | 56.81 | 53.76 | 1.67 | 0.13 | 0 | -3 | 1 | -0.8081 | 0.498 | 4.3 | 97 |  | 0 | 0.0709 | 0.1322 | 0.0002 | 0.99998 | [1299 empty] | 0
P | 2018-02-25T14:56:13+08:00 | 2018-02-25T17:14:59+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI536G | CPSM Final Test | FAIL
P | 2018-02-25T15:03:55+08:00 | 2018-02-25T17:19:06+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI537G | CPSM Final Test | FAIL
P | 2018-02-25T15:47:54+08:00 | 2018-02-25T18:50:57+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI539G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.84 | 56.75 | 53.59 | 1.44 | 0.13 | 1 | -11 | 1 | -0.872 | 0.5356 | 8.5 | 56 |  | 0 | 0.1701 | 0.147 | -0.0019 | 1 | [700 empty] | 0
P | 2018-02-25T15:43:24+08:00 | 2018-02-25T18:55:27+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI542G | CPSM Final Test | PASS | 129 | 166.55 | -0.02 | -0.84 | 56.92 | 55.29 | 1.97 | 0.13 | -1 | 3 | 1 | -0.8978 | 0.5124 | 3.6 | 12 |  | 0 | 0.2332 | 0.2376 | -0.0575 | 0.99999 | [1300 empty] | 0
P | 2018-02-25T15:51:23+08:00 | 2018-02-25T19:00:00+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI541G | CPSM Final Test | PASS | 129 | 167.9 | -0.02 | -0.92 | 56.54 | 52.07 | 1.67 | 0.13 | 1 | 3 | 1 | -0.9259 | 0.4888 | 3 | 48 |  | 0 | 0.2858 | 0.1733 | -0.0332 | 1 | [1301 empty] | 0
P | 2018-02-25T15:39:50+08:00 | 2018-02-25T19:00:19+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI538G | CPSM Final Test | PASS | 129 | 163.13 | -0.02 | -0.69 | 57.06 | 54.44 | 1.51 | 0.13 | 0 | -2 | 1 | -0.9139 | 0.4904 | 1.3 | 9 |  | 0 | 0.0574 | 0.1072 | -0.0226 | 0.99997 | [10 empty] | 0
P | 2018-02-25T16:05:36+08:00 | 2018-02-25T19:15:57+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI536G | CPSM Final Test | PASS | 129 | 163.76 | -0.02 | -0.54 | 56.78 | 57.32 | 2.05 | 0.13 | 1 | -11 | 0 | -0.9159 | 0.507 | 1.5 | 35 |  | 0 | 0.0418 | 0.4058 | -0.0038 | 0.99999 | [1302 empty] | 0
P | 2018-02-25T16:11:30+08:00 | 2018-02-25T19:18:54+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI537G | CPSM Final Test | PASS | 129 | 165.29 | -0.02 | -0.77 | 56.78 | 51.22 | 1.59 | 0.13 | -1 | 4 | -1 | -0.9298 | 0.5014 | 0.3 | 10 |  | 0 | 0.0618 | 0.2785 | -0.0216 | 1 | [373 empty] | 0
P | 2018-02-26T09:47:32+08:00 | 2018-02-26T10:52:07+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI545G | CPSM Final Test | FAIL | [204 empty] | FALSE
P | 2018-02-24T19:53:42+08:00 | 2018-02-26T12:04:35+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI532G | CPSM Final Test | FAIL
P | 2018-02-26T10:33:52+08:00 | 2018-02-26T13:36:46+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | B18SDI540G | CPSM Final Test | PASS | 129 | 160.16 | -0.02 | -0.99 | 56.77 | 54.19 | 1.9 | 0.13 | 1 | -3 | 1 | -0.9 | 0.55 | 3.9 | 27 |  | 0 | 0.0971 | 0.164 | 0.0078 | 0.99999 | [1190 empty] | 0
P | 2018-02-26T10:26:48+08:00 | 2018-02-26T13:38:35+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI546G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.92 | 56.73 | 59.35 | 1.36 | 0.13 | 0 | -1 | -2 | -0.8919 | 0.5108 | 1 | 24 |  | 0 | 0.0994 | 0.1842 | -0.002 | 1 | [1303 empty] | 0
P | 2018-02-26T10:23:17+08:00 | 2018-02-26T13:44:21+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | B18SDI544G | CPSM Final Test | PASS | 129 | 165.65 | 0.27 | -0.84 | 56.62 | 53.85 | 1.67 | 0.13 | 1 | -3 | 3 | -0.8899 | 0.5064 | 4.3 | 35 |  | 0 | 0.0507 | 0.0886 | -0.0398 | 1 | [1194 empty] | 0
P | 2018-02-26T10:19:49+08:00 | 2018-02-26T13:46:51+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | B18SDI545G | CPSM Final Test | PASS | 129 | 166.55 | -0.02 | -0.92 | 56.6 | 54.02 | 1.51 | 0.13 | 1 | 2 | 1 | -0.8978 | 0.5462 | 4.9 | 5 |  | 0 | 0.1258 | 0.038 | 0.0059 | 1 | [1304 empty] | 0
P | 2018-02-25T15:59:08+08:00 | 2018-02-26T13:48:13+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | B18SDI532G | CPSM Final Test | PASS | 129 | 161.51 | -0.02 | -1.3 | 56.72 | 58.42 | 1.51 | 0.13 | -1 | 3 | 1 | -0.844 | 0.4748 | 2.3 | 10 |  | 0 | 0.2107 | 0.4654 | 0.0223 | 0.99999 | [1305 empty] | 0
P | 2018-02-26T07:04:28+08:00 | 2018-02-26T14:03:16+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI543G | CPSM Final Test | FAIL
P | 2018-02-24T03:58:30+08:00 | 2018-02-26T14:44:24+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI535G | CPSM Final Test | FAIL | [16 empty] | 1.0195
P | 2018-02-26T10:33:52+08:00 | 2018-02-26T15:51:10+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015017 | B18SDI543G | CPSM Final Test | PASS | 129 | 157.28 | -0.02 | -0.92 | 56.83 | 52.15 | 1.67 | 0.13 | 0 | -1 | 0 | -0.864 | 0.4902 | 1.4 | 21 |  | 0 | 0.0791 | 0.2742 | 0.0256 | 0.99999 | [1306 empty] | 0
P | 2018-02-26T12:22:46+08:00 | 2018-02-26T16:00:01+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | B18SDI547G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.92 | 56.78 | 54.1 | 1.9 | 0.13 | -1 | -12 | 1 | -0.88 | 0.53 | 0.5 | 144 |  | 0 | 0.105 | 0.2621 | -0.0125 | 0.99999 | [1307 empty] | 0
P | 2018-02-24T01:16:36+08:00 | 2018-02-26T17:26:18+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | B18SDI535G | CPSM Final Test | FAIL | [16 empty] | 0.43377
P | 2018-02-27T08:50:16+08:00 | 2018-02-27T11:33:57+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | B18SDI535G | CPSM Final Test | PASS | 129 | 165.65 | -0.02 | 0.07 | 56.88 | 55.2 | 2.12 | 0.13 | 1 | -12 | -1 | -1.0057 | 0.4904 | 2.7 | -1 |  | 0 | 0.0334 | 0.1127 | -0.022 | 1 | [943 empty] | 0
P | 2018-02-28T16:53:36+08:00 | 2018-02-28T20:38:23+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | B18USM854G | CPSM Final Test | PASS |  | 170.78 | 0.97 | -0.61 | 57.1 | 51.39 | 1.51 | 0.13 | 1 | -3 | 1 | -0.9018 | 0.4558 | 2.4 | -24 |  |  | 0.0684 | 0.2348 | -0.0125 | 0.99999 |  |  |  |  | 133 |  | 0 | [1301 empty] | 0
P | 2018-02-28T17:13:14+08:00 | 2018-02-28T20:52:03+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | B18USM860G | CPSM Final Test | PASS |  | 164.84 | 0.65 | -0.31 | 56.82 | 55.12 | 1.44 | 0.13 | 1 | 3 | 2 | -0.8361 | 0.5462 | 0.2 | 35 |  |  | 0.0846 | 0.1255 | -0.0242 | 0.99999 |  |  |  |  | 133 |  | 0 | [509 empty] | 0
P | 2018-02-28T02:51:36+08:00 | 2018-03-01T08:29:07+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | B18USM855G | CPSM Final Test | FAIL | [14 empty] | 614
P | 2018-02-28T01:37:15+08:00 | 2018-03-01T08:30:07+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM853G | CPSM Final Test | FAIL
P | 2018-02-28T17:03:41+08:00 | 2018-03-01T08:31:09+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM853G | CPSM Final Test | PASS |  | 166.55 | 0.65 | -0.61 | 56.78 | 52.15 | 1.82 | 0.13 | 0 | -1 | 0 | -0.924 | 0.529 | 3.3 | 4 |  |  | 0.0576 | 0.0593 | -0.0126 | 1 |  |  |  |  | 133 |  | 0 | [1203 empty] | 0
P | 2018-02-28T17:40:21+08:00 | 2018-03-01T10:19:51+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | B18USM855G | CPSM Final Test | PASS |  | 162.05 | -0.02 | -0.31 | 57.1 | 55.54 | 1.97 | 0.13 | 1 | 2 | 2 | -0.9697 | 0.4982 | 10.7 | 64 |  |  | 0.1285 | 0.1084 | 0.0025 | 1 |  |  |  |  | 133 |  | 0 | [1302 empty] | 0
P | 2018-03-01T08:36:39+08:00 | 2018-03-01T13:18:29+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | B18USM871G | CPSM Final Test | FAIL
P | 2018-03-01T10:01:51+08:00 | 2018-03-01T13:59:43+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM864G | CPSM Final Test | PASS |  | 166.91 | 0.27 | -0.54 | 56.81 | 55.12 | 1.51 | 0.13 | -1 | 4 | 1 | -0.88 | 0.525 | 4 | 15 |  |  | 0.0223 | 0.0316 | -0.0346 | 0.99999 |  |  |  |  | 133 |  | 0 | [194 empty] | 0
P | 2018-03-01T10:19:57+08:00 | 2018-03-01T14:04:11+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | B18USM857G | CPSM Final Test | PASS |  | 160.88 | 0.08 | -0.31 | 57.01 | 54.53 | 1.74 | 0.13 | -1 | -1 | 1 | -0.9138 | 0.523 | 0.9 | 10 |  |  | 0.1031 | 0.0667 | 0.0023 | 1 |  |  |  |  | 133 |  | 0 | [121 empty] | 0
P | 2018-03-01T10:10:18+08:00 | 2018-03-01T14:08:57+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | B18USM867G | CPSM Final Test | PASS |  | 168.26 | 0.46 | -0.69 | 56.83 | 55.03 | 1.44 | 0.13 | -1 | 3 | 1 | -0.8639 | 0.4974 | 3.2 | 26 |  |  | 0.0657 | 0.1056 | -0.0155 | 0.99999 |  |  |  |  | 133 |  | 0 | [1153 empty] | 0
P | 2018-03-01T11:09:44+08:00 | 2018-03-01T14:14:33+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | B18USM870G | CPSM Final Test | PASS |  | 165.29 | 0.27 | -0.54 | 56.61 | 52.58 | 1.74 | 0.13 | -1 | 4 | 1 | -0.7941 | 0.5018 | 6.3 | 47 |  |  | 0.2513 | 0.1476 | 0.0297 | 0.99997 |  |  |  |  | 133 |  | 0 | [121 empty] | 0
P | 2018-03-01T10:48:41+08:00 | 2018-03-01T14:28:58+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | B18USM863G | CPSM Final Test | PASS |  | 165.56 | -0.21 | -0.46 | 56.72 | 59.94 | 1.67 | 0.13 | 0 | -2 | 2 | -0.9018 | 0.533 | 3.3 | -7 |  |  | 0.0489 | 0.2182 | -0.0026 | 0.99998 |  |  |  |  | 133 |  | 0 | [54 empty] | 0
P | 2018-03-01T07:31:17+08:00 | 2018-03-01T14:23:51+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | B18USM871G | CPSM Final Test | FAIL
P | 2018-03-01T05:06:32+08:00 | 2018-03-01T14:40:37+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | B18USM865G | CPSM Final Test | FAIL
P | 2018-03-01T10:57:34+08:00 | 2018-03-01T14:41:43+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | B18USM871G | CPSM Final Test | PASS |  | 151.79 | 0.37 | -0.69 | 56.76 | 55.37 | 1.36 | 0.13 | -1 | 2 | 0 | -0.9337 | 0.5006 | 1.2 | 4 |  |  | 0.1165 | 0.3962 | -0.0046 | 1 |  |  |  |  | 133 |  | 0 | [1303 empty] | 0
P | 2018-03-01T09:53:34+08:00 | 2018-03-01T15:46:23+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | B18USM865G | CPSM Final Test | PASS |  | 167 | 0.87 | -0.38 | 56.82 | 53.93 | 1.74 | 0.13 | 0 | -2 | -1 | -0.8341 | 0.471 | 0.7 | 11 |  |  | 0.0743 | 0.2023 | -0.0434 | 1 |  |  |  |  | 133 |  | 0 | [1030 empty] | 0
P | 2018-03-01T06:42:19+08:00 | 2018-03-01T16:17:37+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015046 | B18USM858G | CPSM Final Test | FAIL
P | 2018-03-01T14:07:41+08:00 | 2018-03-01T17:33:21+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM861G | CPSM Final Test | PASS |  | 166.73 | 0.17 | -0.61 | 56.48 | 54.69 | 1.59 | 0.13 | -1 | -2 | 1 | -0.9358 | 0.493 | 1.8 | 36 |  |  | 0.0192 | 0.0588 | -0.0103 | 1 |  |  |  |  | 133 |  | 0 | [121 empty] | 0
P | 2018-03-01T11:01:19+08:00 | 2018-03-01T18:17:15+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | B18USM866G | CPSM Final Test | FAIL | [963 empty] | 0
P | 2018-03-01T15:18:00+08:00 | 2018-03-01T18:47:07+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015046 | B18USM862G | CPSM Final Test | PASS |  | 162.95 | -0.02 | -0.54 | 56.63 | 54.69 | 1.82 | 0.13 | 0 | 3 | 1 | -0.7941 | 0.5238 | 2.8 | 33 |  |  | 0.0941 | 0.2203 | 0.0058 | 0.99998 |  |  |  |  | 133 |  | 0 | [229 empty] | 0
P | 2018-03-01T14:39:17+08:00 | 2018-03-01T19:23:09+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015046 | B18USM866G | CPSM Final Test | PASS |  | 166.91 | 0.17 | -0.54 | 56.93 | 55.03 | 1.74 | 0.13 | 1 | -12 | -1 | -0.9158 | 0.504 | 2.1 | 10 |  |  | 0.1028 | 0.6668 | -0.01 | 0.99998 |  |  |  |  | 133 |  | 0 | [935 empty] | 0
P | 2018-03-01T16:02:21+08:00 | 2018-03-01T19:24:09+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015046 | B18USM856G | CPSM Final Test | PASS |  | 160.97 | -0.02 | -0.38 | 57.03 | 54.02 | 1.59 | 0.13 | 0 | 3 | 0 | -0.8759 | 0.545 | 1.3 | 2 |  |  | 0.0734 | 0.2647 | -0.0107 | 0.99999 |  |  |  |  | 133 |  | 0 | [810 empty] | 0
P | 2018-03-01T17:33:20+08:00 | 2018-03-01T19:46:34+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI548G | CPSM Final Test | FAIL
P | 2018-03-01T09:01:09+08:00 | 2018-03-01T19:53:41+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | B18USM869G | CPSM Final Test | FAIL
P | 2018-03-01T18:28:01+08:00 | 2018-03-01T21:19:53+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015046 | B18USM858G | CPSM Final Test | PASS |  | 166.1 | 0.27 | -0.16 | 57.04 | 53.76 | 1.59 | 0.13 | 0 | 2 | 0 | -0.906 | 0.513 | 7.9 | 30 |  |  | 0.1291 | 0.0842 | -0.0155 | 1 |  |  |  |  | 133 |  | 0 | [1030 empty] | 0
P | 2018-03-01T17:43:32+08:00 | 2018-03-01T21:22:51+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015046 | B18USM859G | CPSM Final Test | PASS |  | 171.5 | 0.17 | -0.08 | 56.84 | 52.92 | 2.05 | 0.13 | 0 | 3 | 1 | -0.9558 | 0.462 | 9.8 | 9 |  |  | 0.1381 | 0.082 | -0.0293 | 1 |  |  |  |  | 133 |  | 0 | [1304 empty] | 0
P | 2018-03-01T18:39:57+08:00 | 2018-03-01T21:47:14+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI548G | CPSM Final Test | PASS | 129 | 166.01 | 0.27 | -0.99 | 56.6 | 53.59 | 1.67 | 0.13 | -1 | -11 | 1 | -0.9458 | 0.4684 | 0.8 | 30 |  | 0 | 0.0565 | 0.1982 | 0.0177 | 1 | [817 empty] | 0
P | 2018-03-01T18:51:57+08:00 | 2018-03-01T21:48:37+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015046 | B18USM868G | CPSM Final Test | PASS |  | 167.18 | -0.02 | -0.46 | 56.63 | 52.75 | 1.67 | 0.13 | -1 | 1 | -1 | -0.9278 | 0.5446 | 0.9 | 13 |  |  | 0.0275 | 0.0427 | -0.0298 | 1 |  |  |  |  | 133 |  | 0.00989 | [1305 empty] | 0
P | 2018-03-01T14:27:25+08:00 | 2018-03-01T22:01:01+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015046 | B18USM869G | CPSM Final Test | PASS |  | 160.7 | 0.65 | -0.38 | 56.92 | 54.95 | 1.82 | 0.13 | 0 | -11 | -1 | -0.814 | 0.506 | 2.5 | 69 |  |  | 0.1208 | 0.1181 | -0.0294 | 0.99999 |  |  |  |  | 133 |  | 0 | [652 empty] | 0
P | 2018-03-02T10:32:24+08:00 | 2018-03-02T13:32:01+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI555G | CPSM Final Test | PASS | 129 | 157.46 | -0.02 | 0.91 | 57.2 | 57.66 | 1.82 | 0.13 | 0 | 4 | -1 | -0.8061 | 0.4932 | 1.4 | 2 |  | 0 | 0.1187 | 0.098 | -0.0019 | 1 | [367 empty] | 0
P | 2018-03-02T10:25:27+08:00 | 2018-03-02T13:35:23+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI554G | CPSM Final Test | PASS | 129 | 164.21 | 0.27 | -0.16 | 57 | 55.12 | 1.74 | 0.13 | -1 | 4 | -1 | -0.8161 | 0.5024 | 0.5 | 11 |  | 0.01263 | 0.0257 | 0.0563 | -0.002 | 0.99999 | [72 empty] | 0
P | 2018-03-02T10:17:26+08:00 | 2018-03-02T13:54:13+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI558G | CPSM Final Test | PASS | 129 | 160.88 | -0.02 | -0.08 | 57.53 | 54.95 | 1.9 | 0.13 | 0 | -13 | 1 | -0.816 | 0.4638 | 5.6 | -4 |  | 0 | 0.0672 | 0.1177 | -0.006 | 1 | [719 empty] | 0
P | 2018-03-02T11:04:15+08:00 | 2018-03-02T14:12:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI560G | CPSM Final Test | PASS | 129 | 164.3 | 0.27 | -0.23 | 56.92 | 52.83 | 1.44 | 0.13 | -1 | -11 | -1 | -0.846 | 0.4784 | 7 | -8 |  | 0 | 0.0916 | 0.1778 | 0.0257 | 0.99999 | [30 empty] | 0
P | 2018-03-02T10:49:12+08:00 | 2018-03-02T14:17:34+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI553G | CPSM Final Test | PASS | 129 | 162.32 | 0.08 | -0.38 | 56.93 | 52.24 | 1.82 | 0.13 | 1 | 3 | 3 | -0.9278 | 0.5116 | 4.8 | 48 |  | 0.01472 | 0.1358 | 0.1604 | -0.0152 | 1 | [775 empty] | 0
P | 2018-03-02T11:50:31+08:00 | 2018-03-02T14:32:01+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI556G | CPSM Final Test | PASS | 129 | 161.51 | 0.49 | -0.38 | 56.91 | 53 | 1.74 | 0.13 | 1 | 2 | 0 | -0.81 | 0.478 | 5.8 | 4 |  | 0.02357 | 0.1483 | 0.2745 | 0.0062 | 0.99994 | [33 empty] | 0
P | 2018-03-02T11:47:45+08:00 | 2018-03-02T17:19:35+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI565G | CPSM Final Test | FAIL
P | 2018-03-02T14:49:19+08:00 | 2018-03-02T18:02:29+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI566G | CPSM Final Test | PASS | 129 | 165.56 | 0.17 | -0.38 | 57.07 | 57.66 | 2.05 | 0.13 | 0 | 4 | 1 | -0.874 | 0.5342 | 0.8 | -6 |  | 0.01308 | 0.0283 | 0.0974 | 0.0115 | 0.99999 | [847 empty] | 0
P | 2018-03-02T14:47:43+08:00 | 2018-03-02T18:04:14+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI564G | CPSM Final Test | PASS | 129 | 165.92 | 0.08 | -0.23 | 57.02 | 50.12 | 1.59 | 0.13 | 0 | -2 | 1 | -0.828 | 0.4904 | 3.6 | 18 |  | 0 | 0.2165 | 0.21 | 0.0005 | 1 | [155 empty] | 0
P | 2018-03-02T14:46:40+08:00 | 2018-03-02T18:06:23+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI561G | CPSM Final Test | PASS | 129 | 157.37 | -0.02 | -0.31 | 56.79 | 55.54 | 1.67 | 0.13 | 0 | 2 | 1 | -0.9079 | 0.502 | 1.3 | 40 |  | 0 | 0.1269 | 0.2092 | -0.004 | 1 | [1313 empty] | 0
P | 2018-03-02T14:54:34+08:00 | 2018-03-02T18:11:01+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI550G | CPSM Final Test | PASS | 129 | 164.12 | -0.11 | -0.61 | 56.8 | 51.65 | 1.9 | 0.13 | 0 | 4 | -1 | -0.856 | 0.465 | 3 | 17 |  | 0 | 0.099 | 0.1762 | 0.0279 | 0.99999 | [1314 empty] | 0
P | 2018-03-02T14:57:16+08:00 | 2018-03-02T18:12:01+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI551G | CPSM Final Test | PASS | 129 | 163.67 | 0.37 | -0.69 | 56.79 | 53.17 | 1.29 | 0.13 | 0 | -11 | 1 | -0.826 | 0.5084 | 2.3 | 8 |  | 0 | 0.1219 | 0.1698 | 0.0497 | 0.99998 | [1288 empty] | 0
P | 2018-03-02T14:33:40+08:00 | 2018-03-02T19:16:07+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI565G | CPSM Final Test | PASS | 129 | 170.24 | 0.27 | -0.84 | 56.81 | 55.12 | 1.82 | 0.13 | 0 | -2 | -1 | -0.8281 | 0.5484 | 3 | 35 |  | 0.00852 | 0.0772 | 0.1328 | -0.0241 | 1 | [625 empty] | 0
P | 2018-03-02T03:09:42+08:00 | 2018-03-02T19:26:17+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI559G | CPSM Final Test | FAIL
P | 2018-03-02T11:18:00+08:00 | 2018-03-02T21:19:43+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI559G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -0.16 | 57.09 | 53.17 | 2.43 | 0.13 | 0 | -1 | 1 | -0.9198 | 0.5168 | 1.9 | -12 |  | 0.0167 | 0.0795 | 0.5735 | 0.0302 | 0.99999 | [381 empty] | 0
P | 2018-03-02T04:12:28+08:00 | 2018-03-03T10:56:31+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI568G | CPSM Final Test | FAIL | [14 empty] | 639
P | 2018-03-02T04:23:23+08:00 | 2018-03-03T11:26:23+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI562G | CPSM Final Test | FAIL
P | 2018-03-02T19:48:08+08:00 | 2018-03-03T11:57:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI563G | CPSM Final Test | PASS | 129 | 164.75 | -0.02 | -0.23 | 56.94 | 51.22 | 1.97 | 0.13 | -1 | -12 | -1 | -0.9158 | 0.532 | 5.4 | 16 |  | 0 | 0.079 | 0.1719 | -0.036 | 0.99999 | [1117 empty] | 0
P | 2018-03-02T19:41:29+08:00 | 2018-03-03T11:59:17+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI569G | CPSM Final Test | PASS | 129 | 156.47 | -0.02 | -0.08 | 57.03 | 54.95 | 2.05 | 0.13 | -1 | 1 | 1 | -0.9338 | 0.4714 | 0.8 | 87 |  | 0.01853 | 0.0635 | 0.0846 | -0.0054 | 0.99999 | [804 empty] | 0
P | 2018-03-03T08:38:31+08:00 | 2018-03-03T12:04:35+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI552G | CPSM Final Test | PASS | 129 | 161.78 | 0.37 | -0.38 | 57.08 | 53.51 | 1.82 | 0.13 | 1 | 2 | 1 | -0.9539 | 0.5012 | 0.9 | 16 |  | 0 | 0.2519 | 0.0659 | 0.013 | 0.99997 | [1190 empty] | 0
P | 2018-03-03T05:19:09+08:00 | 2018-03-03T12:37:51+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI567G | CPSM Final Test | FAIL | [1336 empty] | 0
P | 2018-03-02T19:34:29+08:00 | 2018-03-03T12:56:57+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI568G | CPSM Final Test | PASS | 129 | 157.1 | 0.08 | -0.46 | 57.08 | 54.95 | 1.9 | 0.13 | 1 | 4 | 1 | -0.9218 | 0.544 | 4.3 | 14 |  | 0 | 0.1249 | 0.3676 | 0.0228 | 0.99997 | [956 empty] | 0
P | 2018-03-02T19:54:53+08:00 | 2018-03-03T13:20:51+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI562G | CPSM Final Test | PASS | 129 | 161.96 | -0.02 | 0.38 | 56.62 | 51.9 | 1.59 | 0.13 | 0 | 3 | 0 | -0.9557 | 0.5118 | 5.5 | 5 |  | 0 | 0.0851 | 0.1683 | -0.0177 | 1 | [883 empty] | 0
P | 2018-03-03T03:55:03+08:00 | 2018-03-03T13:34:07+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI549G | CPSM Final Test | FAIL | [17 empty] | 63.5327
P | 2018-03-03T08:58:30+08:00 | 2018-03-03T13:43:25+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI567G | CPSM Final Test | PASS | 129 | 167.09 | -0.02 | -0.61 | 56.71 | 58.93 | 1.44 | 0.13 | 1 | 4 | 1 | -0.9217 | 0.507 | 6.6 | 55 |  | 0 | 0.0415 | 0.4526 | -0.0258 | 1 | [1315 empty] | 0
P | 2018-03-03T08:44:35+08:00 | 2018-03-03T14:37:05+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI549G | CPSM Final Test | PASS | 129 | 166.64 | 0.46 | -0.38 | 56.9 | 54.27 | 1.59 | 0.13 | -1 | -11 | -1 | -0.9298 | 0.519 | 0.8 | 37 |  | 0 | 0.0697 | 0.1771 | 0.0224 | 0.99997 | [1153 empty] | 0
P | 2018-03-03T13:23:43+08:00 | 2018-03-03T14:56:41+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI573G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-03T11:15:46+08:00 | 2018-03-03T15:24:51+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI575G | CPSM Final Test | FAIL
P | 2018-03-03T13:40:18+08:00 | 2018-03-03T17:14:56+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI571G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.08 | 57.29 | 55.12 | 1.9 | 0.13 | -1 | 2 | 1 | -0.9497 | 0.516 | 8.4 | 5 |  | 0.00615 | 0.107 | 0.157 | -0.0033 | 1 | [551 empty] | 0
P | 2018-03-03T13:27:16+08:00 | 2018-03-03T17:15:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI570G | CPSM Final Test | PASS | 129 | 165.92 | -0.02 | 0.84 | 57.13 | 53.76 | 2.58 | 0.13 | -1 | -2 | 1 | -0.892 | 0.459 | 4.1 | 196 |  | 0.01281 | 0.0208 | 0.2796 | 0.0117 | 0.99999 | [227 empty] | 0
P | 2018-03-03T14:26:42+08:00 | 2018-03-03T17:21:21+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI574G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.54 | 56.81 | 52.83 | 1.67 | 0.13 | 0 | -2 | 0 | -0.9697 | 0.525 | 2 | 26 |  | 0 | 0.0826 | 0.4086 | -0.0285 | 1 | [1315 empty] | 0
P | 2018-03-03T13:20:18+08:00 | 2018-03-03T17:25:15+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI575G | CPSM Final Test | PASS | 129 | 161.42 | 0.46 | -0.31 | 57.06 | 54.1 | 1.59 | 0.13 | 0 | 2 | 2 | -0.8581 | 0.4586 | 2.6 | 50 |  | 0 | 0.1242 | 0.0696 | -0.0062 | 1 | [1316 empty] | 0
P | 2018-03-03T13:50:08+08:00 | 2018-03-03T17:50:21+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI572G | CPSM Final Test | PASS | 129 | 162.14 | -0.02 | -0.46 | 57.08 | 55.37 | 1.36 | 0.13 | 0 | -2 | 1 | -0.9417 | 0.4664 | 9.6 | 35 |  | 0 | 0.0593 | 0.1615 | 0.0022 | 1 | [1313 empty] | 0
P | 2018-03-03T14:10:12+08:00 | 2018-03-03T17:54:33+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI573G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | 0.45 | 56.75 | 54.19 | 1.74 | 0.13 | 1 | -12 | 1 | -0.9498 | 0.521 | 1.7 | 39 |  | 0 | 0.2232 | 0.1305 | 0.007 | 0.99999 | [1311 empty] | 0
P | 2018-03-03T15:32:56+08:00 | 2018-03-03T18:19:51+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI578G | CPSM Final Test | PASS | 129 | 165.83 | -0.21 | -0.77 | 56.93 | 55.2 | 1.21 | 0.13 | 0 | 4 | 3 | -0.9457 | 0.471 | 1.8 | 27 |  | 0 | 0.0657 | 0.3572 | 0.0008 | 0.99997 | [1317 empty] | 0
P | 2018-02-27T09:38:07+08:00 | 2018-03-05T10:42:25+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI557G | CPSM Final Test | FAIL
P | 2018-03-02T10:10:16+08:00 | 2018-03-05T12:27:13+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI557G | CPSM Final Test | PASS | 129 | 164.66 | 0.08 | -0.23 | 57.3 | 56.3 | 1.9 | 0.13 | 0 | -11 | 1 | -0.9238 | 0.526 | 2.2 | 23 |  | 0 | 0.1085 | 0.1178 | 0.0131 | 0.99999 | [942 empty] | 0
P | 2018-03-05T12:11:05+08:00 | 2018-03-05T16:08:51+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI582G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -0.61 | 57.04 | 55.88 | 1.74 | 0.13 | 0 | -2 | 1 | -0.8679 | 0.5044 | 1.3 | 18 |  | 0.01746 | 0.1529 | 0.3123 | -0.0397 | 0.99999 | [826 empty] | 0
P | 2018-03-05T12:31:14+08:00 | 2018-03-05T16:11:17+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI581G | CPSM Final Test | PASS | 129 | 166.73 | -0.02 | -0.84 | 56.51 | 52.75 | 1.36 | 0.13 | -1 | -2 | 1 | -0.9118 | 0.5398 | 1.8 | 30 |  | 0 | 0.1138 | 0.3151 | -0.0266 | 0.99999 | [883 empty] | 0
P | 2018-03-05T07:54:54+08:00 | 2018-03-05T16:14:25+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI577G | CPSM Final Test | FAIL | [125 empty] | 0
P | 2018-03-05T12:20:45+08:00 | 2018-03-05T16:16:14+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI580G | CPSM Final Test | PASS | 129 | 167.36 | -0.02 | -0.08 | 57.03 | 52.92 | 1.51 | 0.13 | 0 | -11 | 0 | -0.82 | 0.5312 | 1.1 | -2 |  | 0 | 0.1172 | 0.1054 | 0.0123 | 1 | [258 empty] | 0
P | 2018-03-05T13:30:19+08:00 | 2018-03-05T16:20:23+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI579G | CPSM Final Test | PASS | 129 | 168.35 | 0.27 | -0.54 | 56.79 | 53 | 1.67 | 0.13 | 1 | -11 | 1 | -0.854 | 0.5116 | 6 | -21 |  | 0 | 0.0488 | 0.4025 | 0.0134 | 0.99997 | [780 empty] | 0
P | 2018-03-05T13:50:29+08:00 | 2018-03-05T16:43:29+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI585G | CPSM Final Test | PASS | 129 | 163.04 | 0.17 | -0.31 | 57.08 | 54.95 | 1.67 | 0.13 | -1 | -10 | 1 | -0.87 | 0.5374 | 0.9 | 7 |  | 0 | 0.0641 | 0.2797 | 0.0131 | 1 | [782 empty] | 0
P | 2018-03-05T12:42:41+08:00 | 2018-03-05T16:56:07+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI576G | CPSM Final Test | FAIL
P | 2018-03-05T14:24:45+08:00 | 2018-03-05T17:50:21+08:00 | CPSM Final Test TE83064-15 | 25705657 | Autosampler | 186015017 | B18SDI587G | CPSM Final Test | PASS | 129 | 165.65 | 0.08 | -0.54 | 56.93 | 52.83 | 1.82 | 0.13 | -1 | -10 | 1 | -0.834 | 0.482 | 2.4 | 14 |  | 0 | 0.0766 | 0.4485 | 0.0216 | 0.99995 | [1244 empty] | 0
P | 2018-03-05T14:49:24+08:00 | 2018-03-05T18:41:05+08:00 | CPSM Final Test TE83064-11 | 25705657 | Autosampler | 186015017 | B18SDI576G | CPSM Final Test | PASS | 129 | 162.95 | 0.75 | 0.38 | 56.78 | 54.69 | 1.06 | 0.13 | 0 | 2 | -1 | -0.9159 | 0.5272 | 0.6 | 12 |  | 0 | 0.1398 | 0.1173 | 0.0131 | 0.99998 | [1161 empty] | 0
P | 2018-03-05T05:10:48+08:00 | 2018-03-05T18:58:31+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI577G | CPSM Final Test | FAIL | [18 empty] | 1.6066
P | 2018-03-05T12:04:39+08:00 | 2018-03-05T20:30:31+08:00 | CPSM Final Test TE83064-10 | 25705657 | Autosampler | 186015017 | B18SDI577G | CPSM Final Test | PASS | 129 | 160.61 | 0.08 | -0.23 | 57 | 53.26 | 1.9 | 0.13 | 1 | 2 | 1 | -0.8381 | 0.534 | 10.6 | 39 |  | 0 | 0.2976 | 0.6815 | -0.0154 | 0.99998 | [104 empty] | 0
P | 2018-03-05T17:14:03+08:00 | 2018-03-05T21:04:39+08:00 | CPSM Final Test TE83064-13 | 25705657 | Autosampler | 186015017 | B18SDI589G | CPSM Final Test | PASS | 129 | 166.55 | 0.08 | -0.69 | 56.81 | 52.24 | 1.67 | 0.13 | 1 | -11 | -1 | -0.9299 | 0.545 | 6.3 | -48 |  | 0.01689 | 0.041 | 0.081 | -0.0299 | 1 | [1178 empty] | 0
P | 2018-03-05T17:00:48+08:00 | 2018-03-05T21:07:05+08:00 | CPSM Final Test TE83064-16 | 25705657 | Autosampler | 186015017 | B18SDI588G | CPSM Final Test | PASS | 129 | 157.1 | -0.02 | -0.84 | 56.82 | 54.69 | 1.21 | 0.13 | 0 | -2 | 1 | -0.9098 | 0.5414 | 1.3 | 35 |  | 0 | 0.239 | 0.3519 | -0.0041 | 1 | [134 empty] | 0
P | 2018-03-05T17:27:10+08:00 | 2018-03-05T21:11:33+08:00 | CPSM Final Test TE83064-14 | 25705657 | Autosampler | 186015017 | B18SDI584G | CPSM Final Test | PASS | 129 | 164.39 | 0.17 | 0.15 | 57.12 | 55.12 | 1.97 | 0.13 | -1 | -1 | 1 | -0.9179 | 0.489 | 2.4 | 6 |  | 0 | 0.2105 | 0.3109 | 0.0206 | 0.99997 | [822 empty] | 0
P | 2018-03-05T04:11:09+08:00 | 2018-03-06T08:28:16+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI586G | CPSM Final Test | FAIL
P | 2018-03-05T18:19:42+08:00 | 2018-03-06T10:13:26+08:00 | CPSM Final Test TE83064-12 | 25705657 | Autosampler | 186015017 | B18SDI586G | CPSM Final Test | PASS | 129 | 160.25 | 0.17 | -0.69 | 56.82 | 55.37 | 1.59 | 0.13 | 1 | -11 | 0 | -0.9118 | 0.4964 | 4.4 | -1 |  | 0 | 0.2759 | 0.1153 | 0.001 | 0.99999 | [13 empty] | 0
P | 2018-03-04T23:42:24+08:00 | 2018-03-06T11:31:27+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI583G | CPSM Final Test | FAIL
P | 2018-03-05T17:36:55+08:00 | 2018-03-06T12:35:07+08:00 | CPSM Final Test TE83064-17 | 25705657 | Autosampler | 186015017 | B18SDI583G | CPSM Final Test | PASS | 129 | 159.89 | -0.02 | -0.61 | 56.74 | 56.22 | 1.82 | 0.13 | 1 | 4 | 1 | -0.9019 | 0.518 | 1.3 | -3 |  | 0 | 0.0833 | 0.2476 | 0.0322 | 0.99999 | [1318 empty] | 0
P | 2018-03-07T13:07:42+08:00 | 2018-03-07T16:02:07+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | C18USM877G | CPSM Final Test | FAIL | [13 empty] | 0
P | 2018-03-07T15:59:30+08:00 | 2018-03-07T16:20:27+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM879G | CPSM Final Test | FAIL | [48 empty] | FALSE
P | 2018-03-07T15:54:02+08:00 | 2018-03-07T16:25:55+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM879G | CPSM Final Test | FAIL | [87 empty] | FALSE
P | 2018-03-07T11:19:37+08:00 | 2018-03-07T17:24:42+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM874G | CPSM Final Test | FAIL
P | 2018-03-07T14:18:14+08:00 | 2018-03-07T17:37:41+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM875G | CPSM Final Test | PASS |  | 161.24 | 0.17 | -0.84 | 56.61 | 55.29 | 1.06 | 0.13 | -1 | -11 | 1 | -0.876 | 0.5088 | 4.1 | 18 |  |  | 0.135 | 0.8723 | 0.043 | 0.9999 |  |  |  |  | 133 |  | 0 | [355 empty] | 0
P | 2018-03-07T14:29:55+08:00 | 2018-03-07T17:44:42+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM878G | CPSM Final Test | PASS |  | 163.76 | -0.02 | -0.61 | 56.93 | 53.93 | 1.67 | 0.13 | -1 | 3 | 1 | -0.8061 | 0.484 | 2 | 3 |  |  | 0.0652 | 0.223 | -0.0274 | 1 |  |  |  |  | 133 |  | 0 | [205 empty] | 0
P | 2018-03-07T14:26:01+08:00 | 2018-03-07T17:49:53+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM873G | CPSM Final Test | PASS |  | 165.74 | 0.46 | 0.23 | 56.61 | 53.68 | 0.91 | 0.16 | 0 | -11 | 1 | -0.864 | 0.489 | 5.7 | 2 |  |  | 0.0717 | 0.0727 | -0.0266 | 1 |  |  |  |  | 133 |  | 0.02275 | [380 empty] | 0
P | 2018-03-07T14:42:09+08:00 | 2018-03-07T18:16:38+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM876G | CPSM Final Test | PASS |  | 168.08 | -0.02 | 0.23 | 56.51 | 53.68 | 1.59 | 0.13 | 0 | -11 | 1 | -0.9377 | 0.4892 | 2.5 | 10 |  |  | 0.1089 | 0.2323 | -0.0159 | 1 |  |  |  |  | 133 |  | 0 | [45 empty] | 0
P | 2018-03-07T09:42:03+08:00 | 2018-03-07T19:27:45+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | C18USM877G | CPSM Final Test | FAIL | [14 empty] | 657
P | 2018-03-07T12:50:24+08:00 | 2018-03-07T19:29:33+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM879G | CPSM Final Test | FAIL
P | 2018-03-07T14:22:09+08:00 | 2018-03-07T19:30:42+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM874G | CPSM Final Test | PASS |  | 162.59 | 0.17 | -0.69 | 56.53 | 53.51 | 1.67 | 0.13 | 0 | -4 | 1 | -0.8878 | 0.5038 | 10.6 | 7 |  |  | 0.1061 | 0.1276 | 0.0137 | 0.99997 |  |  |  |  | 133 |  | 0.00453 | [97 empty] | 0
P | 2018-03-07T16:11:08+08:00 | 2018-03-07T19:34:15+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | C18USM880G | CPSM Final Test | PASS |  | 161.96 | -0.02 | -0.77 | 56.81 | 55.12 | 1.06 | 0.13 | 0 | -2 | 1 | -0.814 | 0.542 | 6.8 | 5 |  |  | 0.1166 | 0.1638 | 0.0237 | 0.99999 |  |  |  |  | 133 |  | 0 | [1312 empty] | 0
P | 2018-03-07T14:34:54+08:00 | 2018-03-07T21:19:55+08:00 | CPSM Final Test TE83064-16 | 25705598 | Autosampler | 186015046 | C18USM877G | CPSM Final Test | PASS |  | 167.54 | -0.02 | -0.46 | 56.7 | 57.83 | 1.51 | 0.13 | 0 | -3 | 1 | -0.9238 | 0.5116 | 6.9 | 39 |  |  | 0.047 | 0.2013 | 0.0299 | 0.99996 |  |  |  |  | 133 |  | 0 | [1313 empty] | 0
P | 2018-03-07T16:09:59+08:00 | 2018-03-07T21:25:53+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM879G | CPSM Final Test | PASS |  | 159.17 | -0.02 | -0.77 | 56.42 | 59.27 | 1.06 | 0.13 | 0 | -2 | -1 | -0.8818 | 0.4944 | 7.2 | 14 |  |  | 0.0607 | 0.315 | -0.0247 | 0.99999 |  |  |  |  | 133 |  | 0 | [1314 empty] | 0
P | 2018-03-07T18:54:49+08:00 | 2018-03-07T22:07:21+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM872G | CPSM Final Test | PASS |  | 159.8 | -0.02 | -0.69 | 56.71 | 58 | 1.29 | 0.13 | 0 | -2 | -1 | -0.8181 | 0.516 | 2.3 | 57 |  |  | 0.0516 | 0.1422 | -0.006 | 0.99999 |  |  |  |  | 133 |  | 0.01827 | [221 empty] | 0
P | 2018-03-07T19:25:41+08:00 | 2018-03-07T22:09:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM882G | CPSM Final Test | PASS |  | 160.7 | -0.02 | -0.54 | 56.33 | 51.99 | 0.98 | 0.13 | 0 | -2 | 0 | -0.9258 | 0.5424 | 9.6 | 2 |  |  | 0.1436 | 0.1371 | -0.0276 | 1 |  |  |  |  | 133 |  | 0 | [886 empty] | 0
P | 2018-03-08T10:24:03+08:00 | 2018-03-08T13:33:45+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM885G | CPSM Final Test | PASS |  | 162.05 | -0.02 | -0.31 | 56.93 | 54.95 | 1.67 | 0.13 | 1 | -3 | -2 | -0.7701 | 0.4834 | 4.7 | 9 |  |  | 0.1177 | 0.123 | -0.0111 | 0.99999 |  |  |  |  | 133 |  | 0.00632 | [1030 empty] | 0
P | 2018-03-08T10:11:52+08:00 | 2018-03-08T13:45:41+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015046 | C18USM883G | CPSM Final Test | PASS |  | 164.39 | -0.02 | -0.54 | 57.01 | 54.78 | 1.51 | 0.13 | 1 | -2 | 1 | -0.828 | 0.4972 | 2.7 | 2 |  |  | 0.152 | 0.1223 | -0.0218 | 0.99997 |  |  |  |  | 133 |  | 0 | [1315 empty] | 0
P | 2018-03-08T10:19:37+08:00 | 2018-03-08T13:46:11+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM881G | CPSM Final Test | PASS |  | 161.15 | 0.27 | -0.61 | 56.77 | 56.22 | 1.36 | 0.13 | 1 | 3 | -1 | -0.8559 | 0.5014 | 3.1 | 36 |  |  | 0.0465 | 0.0797 | -0.0252 | 1 |  |  |  |  | 133 |  | 0.017 | [934 empty] | 0
P | 2018-03-08T10:15:43+08:00 | 2018-03-08T13:50:32+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM884G | CPSM Final Test | PASS |  | 167.99 | -0.02 | -0.77 | 56.52 | 54.61 | 0.98 | 0.13 | 1 | 2 | -1 | -0.7662 | 0.501 | 1.1 | 44 |  |  | 0.1238 | 0.2693 | -0.0092 | 1 |  |  |  |  | 133 |  | 0.00614 | [753 empty] | 0
P | 2018-03-08T12:35:36+08:00 | 2018-03-08T16:05:01+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015046 | C18USM888G | CPSM Final Test | PASS |  | 164.57 | -0.3 | -0.54 | 56.63 | 52.75 | 1.59 | 0.13 | -1 | 3 | 1 | -0.878 | 0.5216 | 2.9 | 12 |  |  | 0.1271 | 0.4248 | 0.0061 | 1 |  |  |  |  | 133 |  | 0.00731 | [31 empty] | 0
P | 2018-03-08T12:28:39+08:00 | 2018-03-08T16:07:55+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015046 | C18USM887G | CPSM Final Test | PASS |  | 161.6 | -0.02 | -0.54 | 56.93 | 54.95 | 1.82 | 0.21 | -1 | -2 | 1 | -0.844 | 0.519 | 3 | 23 |  |  | 0.0973 | 0.258 | -0.0046 | 0.99997 |  |  |  |  | 133 |  | 0 | [741 empty] | 0
P | 2018-03-08T12:35:24+08:00 | 2018-03-08T16:15:14+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015046 | C18USM886G | CPSM Final Test | PASS |  | 164.93 | -0.02 | -0.61 | 56.77 | 54.69 | 1.67 | 0.13 | 1 | 2 | -1 | -0.8879 | 0.4828 | 3.4 | 8 |  |  | 0.1272 | 0.2418 | 0.026 | 0.99998 |  |  |  |  | 133 |  | 0 | [161 empty] | 0
P | 2018-03-08T15:25:20+08:00 | 2018-03-08T18:20:47+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015046 | C18USM889G | CPSM Final Test | PASS |  | 166.28 | -0.02 | -0.31 | 56.79 | 54.02 | 1.9 | 0.13 | 1 | 4 | 0 | -0.872 | 0.5418 | 2.5 | 16 |  |  | 0.104 | 0.0844 | -0.0272 | 1 |  |  |  |  | 133 |  | 0.0183 | [205 empty] | 0
P | 2018-03-08T15:29:24+08:00 | 2018-03-08T18:24:07+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015046 | C18USM891G | CPSM Final Test | PASS |  | 164.21 | -0.21 | -0.54 | 56.64 | 55.2 | 1.44 | 0.13 | 0 | 3 | 1 | -0.82 | 0.491 | 2.8 | -3 |  |  | 0.0703 | 0.1661 | -0.0179 | 1 |  |  |  |  | 133 |  | 0 | [395 empty] | 0
P | 2018-03-08T15:22:01+08:00 | 2018-03-08T18:28:44+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015046 | C18USM890G | CPSM Final Test | PASS |  | 160.43 | -0.21 | -0.31 | 56.75 | 53.59 | 1.51 | 0.13 | 0 | 4 | 1 | -0.8738 | 0.498 | 5.5 | 12 |  |  | 0.1971 | 0.1332 | 0.0033 | 0.99999 |  |  |  |  | 133 |  | 0.01691 | [157 empty] | 0
P | 2018-03-08T15:16:20+08:00 | 2018-03-08T19:03:31+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI591G | CPSM Final Test | FAIL
P | 2018-03-08T16:06:57+08:00 | 2018-03-08T19:26:29+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI590G | CPSM Final Test | PASS | 129 | 169.07 | -0.02 | -0.38 | 57.07 | 53.68 | 1.67 | 0.13 | -1 | -2 | 1 | -0.8261 | 0.505 | 4.3 | 14 |  | 0 | 0.2464 | 0.0934 | -0.0468 | 1 | [1323 empty] | 0
P | 2018-03-08T17:05:52+08:00 | 2018-03-08T20:40:07+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI592G | CPSM Final Test | PASS | 129 | 162.86 | -0.02 | -0.69 | 56.8 | 55.71 | 1.67 | 0.13 | 0 | -3 | 1 | -0.832 | 0.5328 | 8.6 | 7 |  | 0 | 0.1694 | 0.2613 | -0.0259 | 0.99996 | [302 empty] | 0
P | 2018-03-08T17:09:56+08:00 | 2018-03-08T21:03:11+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI591G | CPSM Final Test | PASS | 129 | 162.32 | -0.02 | -0.61 | 56.52 | 54.53 | 1.67 | 0.13 | -1 | 4 | 1 | -0.8281 | 0.5388 | 4.8 | 21 |  | 0 | 0.1028 | 0.2183 | -0.0459 | 1 | [846 empty] | 0
P | 2018-03-09T10:35:07+08:00 | 2018-03-09T13:23:17+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI595G | CPSM Final Test | PASS | 129 | 162.59 | -0.02 | -0.54 | 57.03 | 52.32 | 1.82 | 0.13 | 0 | 3 | 1 | -0.8839 | 0.529 | 12.2 | -7 |  | 0 | 0.0765 | 0.2662 | -0.0366 | 1 | [736 empty] | 0
P | 2018-03-09T10:42:00+08:00 | 2018-03-09T13:34:12+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI593G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -0.54 | 56.5 | 53.68 | 1.51 | 0.13 | 0 | 3 | -2 | -0.86 | 0.496 | 9.8 | 49 |  | 0 | 0.0822 | 0.3733 | 0.0456 | 0.99999 | [61 empty] | 0
P | 2018-03-09T10:30:19+08:00 | 2018-03-09T13:38:17+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI594G | CPSM Final Test | PASS | 129 | 160.97 | -0.02 | -0.31 | 57.05 | 56.81 | 1.82 | 0.13 | 0 | -2 | 1 | -0.796 | 0.4984 | 10.1 | 10 |  | 0 | 0.3747 | 0.1411 | 0.0158 | 0.99995 | [372 empty] | 0
P | 2018-03-09T12:04:11+08:00 | 2018-03-09T15:41:05+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI596G | CPSM Final Test | PASS | 129 | 165.2 | 0.37 | -0.69 | 56.79 | 54.53 | 1.59 | 0.13 | 1 | -3 | -1 | -0.866 | 0.4988 | 12.1 | 3 |  | 0 | 0.1047 | 0.2854 | -0.0318 | 0.99998 | [1324 empty] | 0
P | 2018-03-09T12:08:31+08:00 | 2018-03-09T15:48:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI597G | CPSM Final Test | PASS | 129 | 161.06 | 0.37 | -0.31 | 56.83 | 54.19 | 1.67 | 0.13 | 0 | -2 | 1 | -0.804 | 0.492 | 2.7 | 91 |  | 0.00512 | 0.1852 | 0.4161 | 0.0009 | 0.99998 | [200 empty] | 0
P | 2018-03-09T13:28:48+08:00 | 2018-03-09T16:31:53+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI598G | CPSM Final Test | PASS | 129 | 160.25 | -0.21 | -0.31 | 56.82 | 54.69 | 1.67 | 0.13 | 1 | 4 | -1 | -0.826 | 0.4988 | 1.4 | 15 |  | 0.01247 | 0.103 | 0.1615 | -0.0029 | 1 | [591 empty] | 0
P | 2018-03-09T13:24:08+08:00 | 2018-03-09T16:33:51+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI599G | CPSM Final Test | PASS | 129 | 163.4 | -0.02 | -0.46 | 56.93 | 53.09 | 1.67 | 0.13 | 0 | 2 | 1 | -0.844 | 0.4962 | 5.5 | -7 |  | 0.0183 | 0.0327 | 0.0542 | -0.0128 | 1 | [367 empty] | 0
P | 2018-03-09T15:35:12+08:00 | 2018-03-09T18:35:18+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI601G | CPSM Final Test | PASS | 129 | 156.92 | 0.17 | -0.46 | 57.42 | 54.86 | 1.59 | 0.13 | 1 | 2 | 1 | -0.9019 | 0.499 | 5.7 | 3 |  | 0 | 0.1081 | 0.1168 | 0.0026 | 0.99999 | [812 empty] | 0
P | 2018-03-09T15:28:26+08:00 | 2018-03-09T18:37:27+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI600G | CPSM Final Test | PASS | 129 | 163.85 | -0.21 | 0.07 | 57.22 | 55.03 | 1.59 | 0.13 | 0 | -11 | 1 | -0.9397 | 0.4958 | 7.7 | -5 |  | 0 | 0.0625 | 0.467 | 0.0097 | 1 | [806 empty] | 0
P | 2018-03-08T22:25:49+08:00 | 2018-03-10T08:48:19+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI602G | CPSM Final Test | FAIL
P | 2018-03-08T23:59:11+08:00 | 2018-03-10T10:12:28+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI603G | CPSM Final Test | FAIL | [14 empty] | 796
P | 2018-03-09T17:10:18+08:00 | 2018-03-10T10:59:45+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI604G | CPSM Final Test | PASS | 129 | 161.78 | -0.02 | -0.54 | 57.03 | 52.75 | 1.82 | -0.17 | 0 | -2 | 1 | -0.8838 | 0.4998 | 4.4 | 9 |  | 0 | 0.1252 | 0.2292 | -0.015 | 0.99999 | [1222 empty] | 0
P | 2018-03-08T19:34:05+08:00 | 2018-03-10T11:40:03+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI602G | CPSM Final Test | FAIL
P | 2018-03-08T19:29:31+08:00 | 2018-03-10T11:44:37+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI602G | CPSM Final Test | FAIL | [16 empty] | 0
P | 2018-03-09T17:05:50+08:00 | 2018-03-10T11:56:54+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI603G | CPSM Final Test | PASS | 129 | 176.18 | -0.02 | -0.54 | 57.39 | 54.53 | 1.51 | 0.18 | -1 | 2 | 1 | -0.9318 | 0.536 | 2.2 | -4 |  | 0.00228 | 0.4443 | 0.4464 | -0.0543 | 0.99997 | [447 empty] | 0
P | 2018-03-09T15:37:04+08:00 | 2018-03-10T12:47:29+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI602G | CPSM Final Test | PASS | 129 | 164.48 | -0.02 | -0.61 | 56.77 | 55.03 | 1.82 | 0.13 | 0 | -12 | -2 | -0.798 | 0.5214 | 6.9 | 8 |  | 0 | 0.0596 | 0.1936 | 0.0316 | 0.99999 | [139 empty] | 0
P | 2018-03-10T11:08:04+08:00 | 2018-03-10T15:11:28+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI607G | CPSM Final Test | PASS | 129 | 169.34 | -0.02 | -0.61 | 56.77 | 55.96 | 1.51 | 0.13 | 0 | 2 | 2 | -0.8619 | 0.5056 | 0.7 | 9 |  | 0 | 0.1082 | 0.2052 | 0.0077 | 0.99999 | [1191 empty] | 0
P | 2018-03-10T06:38:35+08:00 | 2018-03-10T15:23:47+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI606G | CPSM Final Test | FAIL | [242 empty] | 0
P | 2018-03-10T11:01:11+08:00 | 2018-03-10T16:29:39+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI606G | CPSM Final Test | PASS | 129 | 164.93 | -0.02 | -0.84 | 56.77 | 55.46 | 1.21 | 0.13 | -1 | -2 | 0 | -0.87 | 0.4988 | 2.3 | 53 |  | 0.01722 | 0.0559 | 0.0647 | -0.0174 | 1 | [221 empty] | 0
P | 2018-03-10T05:11:19+08:00 | 2018-03-10T16:42:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI605G | CPSM Final Test | FAIL | [1257 empty] | 0
P | 2018-03-10T13:54:31+08:00 | 2018-03-10T17:16:53+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI610G | CPSM Final Test | PASS | 129 | 160.61 | -0.21 | -1.22 | 56.53 | 55.29 | 1.67 | 0.13 | -1 | -2 | 1 | -0.7563 | 0.533 | 1.6 | 7 |  | 0 | 0.0861 | 0.2685 | 0.0362 | 0.99997 | [286 empty] | 0
P | 2018-03-10T10:57:09+08:00 | 2018-03-10T17:48:59+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI605G | CPSM Final Test | PASS | 129 | 160.61 | -0.02 | -0.61 | 57.05 | 53.93 | 1.36 | 0.13 | 0 | -3 | 2 | -0.84 | 0.499 | 1.4 | 2 |  | 0 | 0.1643 | 0.4589 | -0.0154 | 1 | [1236 empty] | 0
P | 2018-03-10T15:27:43+08:00 | 2018-03-10T18:16:36+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI608G | CPSM Final Test | PASS | 129 | 168.89 | -0.02 | -0.77 | 57.07 | 55.37 | 1.29 | 0.13 | 1 | -2 | 1 | -0.8679 | 0.508 | 5.3 | -43 |  | 0.03556 | 0.0344 | 0.1927 | -0.0011 | 0.99998 | [1325 empty] | 0
P | 2018-03-10T15:31:36+08:00 | 2018-03-10T18:17:07+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI611G | CPSM Final Test | PASS | 129 | 162.23 | -0.02 | -1.15 | 56.81 | 55.88 | 1.13 | 0.13 | 0 | 4 | 2 | -0.848 | 0.504 | 4.1 | 10 |  | 0.00658 | 0.0513 | 0.3867 | -0.029 | 1 | [856 empty] | 0
P | 2018-03-08T16:49:57+08:00 | 2018-03-12T11:06:25+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI609G | CPSM Final Test | FAIL | [16 empty] | 0.16924
P | 2018-03-08T16:20:05+08:00 | 2018-03-12T11:36:17+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI609G | CPSM Final Test | FAIL
P | 2018-03-12T09:45:15+08:00 | 2018-03-12T12:36:13+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI613G | CPSM Final Test | PASS | 129 | 163.22 | -0.02 | -0.84 | 56.74 | 58.25 | 1.74 | 0.13 | 0 | -2 | -1 | -0.8381 | 0.5368 | 0.3 | 52 |  | 0.02966 | 0.1145 | 0.088 | -0.0282 | 1 | [221 empty] | 0
P | 2018-03-12T09:40:56+08:00 | 2018-03-12T12:40:41+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI612G | CPSM Final Test | PASS | 129 | 162.95 | -0.02 | -1.53 | 56.46 | 58.42 | 0.83 | 0.13 | 1 | -2 | 1 | -0.8401 | 0.49 | 1.3 | 20 |  | 0 | 0.1676 | 0.3012 | -0.0114 | 0.99999 | [1236 empty] | 0
P | 2018-03-10T13:58:11+08:00 | 2018-03-12T12:42:39+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI609G | CPSM Final Test | PASS | 129 | 160.97 | -0.02 | -0.92 | 56.53 | 53.93 | 1.97 | 0.13 | 1 | 2 | 0 | -0.9099 | 0.4948 | 0.3 | 28 |  | 0.01192 | 0.1736 | 0.1801 | -0.0292 | 0.99999 | [1326 empty] | 0
P | 2018-03-12T11:29:45+08:00 | 2018-03-12T15:28:25+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI615G | CPSM Final Test | PASS | 129 | 159.98 | 0.37 | -0.77 | 56.82 | 54.61 | 1.44 | 0.13 | 0 | 2 | 1 | -0.8219 | 0.509 | 1.3 | 22 |  | 0.00825 | 0.0925 | 0.2153 | -0.0179 | 0.99998 | [744 empty] | 0
P | 2018-03-12T11:43:34+08:00 | 2018-03-12T15:31:03+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI614G | CPSM Final Test | PASS | 129 | 161.87 | -0.02 | -1.22 | 56.49 | 57.32 | 1.59 | 0.13 | 0 | 2 | 0 | -0.8461 | 0.5188 | 1.8 | 25 |  | 0 | 0.1273 | 0.2145 | 0.0666 | 0.99996 | [1327 empty] | 0
P | 2018-03-12T11:37:05+08:00 | 2018-03-12T15:36:34+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI616G | CPSM Final Test | PASS | 129 | 164.3 | -0.02 | -0.61 | 56.76 | 53.85 | 1.36 | 0.13 | 1 | -11 | 1 | -0.806 | 0.513 | 8.3 | -2 |  | 0 | 0.1018 | 0.1972 | 0.0093 | 0.99998 | [1328 empty] | 0
P | 2018-03-12T11:41:18+08:00 | 2018-03-12T15:46:05+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI617G | CPSM Final Test | PASS | 129 | 176.72 | 0.17 | -0.69 | 56.78 | 54.78 | 1.74 | -0.21 | 0 | -3 | 1 | -0.7741 | 0.527 | 6.7 | 29 |  | 0.01365 | 0.0715 | 0.2687 | 0.0736 | 0.99994 | [801 empty] | 0
P | 2018-03-12T14:38:55+08:00 | 2018-03-12T17:30:29+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI619G | CPSM Final Test | PASS | 129 | 155.48 | -0.02 | -0.54 | 57.04 | 53.93 | 1.44 | 0.13 | 1 | -2 | 1 | -0.8759 | 0.5334 | 2.1 | 43 |  | 0.02226 | 0.0973 | 0.0783 | -0.0317 | 1 | [210 empty] | 0
P | 2018-03-12T14:35:22+08:00 | 2018-03-12T17:33:53+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI618G | CPSM Final Test | PASS | 129 | 162.14 | -0.21 | -0.99 | 56.44 | 54.61 | 1.51 | 0.13 | 1 | -2 | 1 | -0.902 | 0.5194 | 3.7 | 11 |  | 0 | 0.3023 | 0.2795 | -0.038 | 1 | [860 empty] | 0
P | 2018-03-12T13:37:06+08:00 | 2018-03-12T18:35:47+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI621G | CPSM Final Test | FAIL | [13 empty] | 23.2
P | 2018-03-12T16:19:52+08:00 | 2018-03-12T19:15:28+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI620G | CPSM Final Test | PASS | 129 | 171.77 | -0.02 | -0.92 | 56.71 | 59.01 | 1.67 | 0.13 | 1 | -2 | 1 | -0.8899 | 0.486 | 4.6 | -3 |  | 0 | 0.0841 | 0.3683 | 0.0248 | 0.99998 | [1191 empty] | 0
P | 2018-03-12T12:49:56+08:00 | 2018-03-12T19:36:19+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI622G | CPSM Final Test | FAIL
P | 2018-03-12T16:13:08+08:00 | 2018-03-12T19:36:47+08:00 | CPSM Final Test TE83064-17 | 25705598 | Autosampler | 186015017 | C18SDI622G | CPSM Final Test | PASS | 129 | 167.63 | -0.02 | -0.54 | 57.12 | 54.36 | 1.59 | 0.13 | 0 | 1 | 0 | -0.8979 | 0.5192 | 1.9 | 5 |  | 0.00776 | 0.0958 | 0.2227 | 0.0068 | 1 | [146 empty] | 0
P | 2018-03-12T16:06:26+08:00 | 2018-03-12T20:38:35+08:00 | CPSM Final Test TE83064-15 | 25705598 | Autosampler | 186015017 | C18SDI621G | CPSM Final Test | PASS | 129 | 173.57 | -0.02 | -0.92 | 56.5 | 54.19 | 0.98 | 0.13 | 1 | 11 | 1 | -0.838 | 0.4938 | 1.2 | -14 |  | 0 | 0.1322 | 0.1604 | -0.0065 | 0.99996 | [364 empty] | 0
P | 2018-03-12T18:44:27+08:00 | 2018-03-12T21:28:23+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI624G | CPSM Final Test | PASS | 129 | 163.67 | -0.02 | -0.54 | 56.84 | 54.53 | 1.44 | 0.13 | 1 | 3 | 0 | -0.8739 | 0.5386 | 8.2 | -5 |  | 0 | 0.0977 | 0.1256 | -0.0084 | 0.99999 | [827 empty] | 0
P | 2018-03-12T18:41:05+08:00 | 2018-03-12T21:38:23+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI623G | CPSM Final Test | PASS | 129 | 163.49 | 0.17 | -1.07 | 56.49 | 54.61 | 1.06 | 0.13 | 1 | 2 | 1 | -0.84 | 0.4974 | 1.1 | 13 |  | 0 | 0.1367 | 0.2231 | 0.0194 | 1 | [826 empty] | 0
P | 2018-03-13T07:40:41+08:00 | 2018-03-13T12:29:43+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI625G | CPSM Final Test | FAIL
P | 2018-03-13T10:13:35+08:00 | 2018-03-13T13:08:34+08:00 | CPSM Final Test TE83064-14 | 25705598 | Autosampler | 186015017 | C18SDI628G | CPSM Final Test | PASS | 129 | 168.53 | -0.02 | -0.92 | 56.8 | 54.19 | 1.36 | 0.13 | 1 | 4 | -3 | -0.82 | 0.5448 | 0.9 | 20 |  | 0 | 0.1577 | 0.0894 | 0.0111 | 0.99999 | [345 empty] | 0
P | 2018-03-13T10:01:21+08:00 | 2018-03-13T13:30:45+08:00 | CPSM Final Test TE83064-13 | 25705598 | Autosampler | 186015017 | C18SDI630G | CPSM Final Test | PASS | 129 | 161.69 | -0.21 | -0.84 | 56.81 | 54.86 | 0.91 | 0.13 | -1 | 2 | 0 | -0.8082 | 0.4962 | 7.6 | 1 |  | 0.0204 | 0.0816 | 0.0476 | -0.0021 | 0.99999 | [955 empty] | 0
P | 2018-03-13T09:57:36+08:00 | 2018-03-13T13:38:39+08:00 | CPSM Final Test TE83064-11 | 25705598 | Autosampler | 186015017 | C18SDI629G | CPSM Final Test | PASS | 129 | 162.32 | -0.4 | -0.61 | 57.08 | 52.07 | 1.67 | 0.13 | 0 | -2 | 2 | -0.838 | 0.4938 | 1.6 | 16 |  | 0 | 0.0939 | 0.1566 | -0.0292 | 0.99998 | [1329 empty] | 0
P | 2018-03-13T09:54:30+08:00 | 2018-03-13T13:54:31+08:00 | CPSM Final Test TE83064-10 | 25705598 | Autosampler | 186015017 | C18SDI627G | CPSM Final Test | PASS | 129 | 162.77 | -0.02 | -0.84 | 56.78 | 54.02 | 1.21 | 0.13 | 1 | -2 | -1 | -0.8021 | 0.4994 | 2.4 | 6 |  | 0.0084 | 0.4347 | 0.218 | -0.0146 | 1 | [1330 empty] | 0
P | 2018-03-13T10:05:12+08:00 | 2018-03-13T14:14:13+08:00 | CPSM Final Test TE83064-12 | 25705598 | Autosampler | 186015017 | C18SDI625G | CPSM Final Test | PASS | 129 | 157.91 | -0.02 | -0.69 | 56.54 | 55.12 | 1.67 | 0.13 | -1 | -2 | 1 | -0.7881 | 0.481 | 3.2 | -15 |  | 0 | 0.1249 | 0.3983 | -0.0038 | 0.99999 | [368 empty] | 0