## Minutes Of Meeting-CLOSED (hidden)
ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | ASP/TSP Matrix Development - REFER HISTORY from 20th Jun 2018 email 30 May 2018 - To update and maintain the tracker to be on current state (6 Jun 2018 onwards) 11 Jul 2018 - target ASP BSM\QSM update by 31 July 2018. 15 Aug 2018: Pending ECR/ECO . 17 OCt 2018 : 1 ECO submitted and rest are pending/ to trigger Gary once approve 16 Jan 2019: Details are updated at "ASP TSP Summary" | Gary/Yee Hwa | 2017-07-06 | TBD | Closed | Low
7 | Implentation for Alliance Template 1.9 - Hua Zhen to verify the chemical part number as per WAT271129 Rev ? - after resolving the Carryover Test (High Blank) - 18 Jul 2018 24 Oct 2018 ONHOLD ACTIVITY 16 Jan 2019: No Template 1.9 will be impletemented. New testers are expected to be send to flex - TBA | Hua Zhen | 2018-06-20 | 6/27/2018 7/11/2018 12/30/2018 | Closed | Low
8 | Consigned Equipment Health Review Identify the list of consigned equipment.Total 56 equipment. Identify the downtime data on the consigned equipment Based on Jan-Jun data. Establish a checklist to review the health of the equipment.Avaialable Start the review of the health of the equipment. Review the status of the health with Waters team. Take actions and close on the problematic testers. Continue monitor for effectiveness - Dashbord update in progress for Alliance - 18 Jul 2018 - Alliance Dashboard : completed for submissiont 01 Aug 2018 (time frame CEHR) - Pump Dashboard completed - submitted on 31 Aug 2018 (resend the original file) - 17 Oct 2018 Identify and try to source locally available parts (PUMPS Testers) - 21 Nov 2018 Prioritize PUMP Tester (WPPL to source for PCBA and Pump Control Module). Target 07 Dec 2018 28 Nov 2018: Joel to follow up with a OESR (Pump) so that WPPL can work on the sourcing. 05 Dec 2018: Joel will send out the OESR by today 19 Dec 2019: To send Detector Health review 16 Jan 2019: Will be review through Project meeting | TAN CK/JOEL | 2018-07-07 | 2018-08-31 | Closed | Medium
9 | eDHR into Engg discussion Kick off meeting and align with CFT team on proposal Work with IT on configuration setup Matrix preparation Proposal and change request to customer Verify in test environment Release to production continue monitor and update 16 Jan 2019: Will be review through Project meeting | TAN CK | 2018-07-07 | 2018-11-30 | Closed | Low
10 | Monitor chemical for Carryover Test as per Rev 1.8 - 19 Sep 18 Hua Zhen to follow-up with Milford the implementation of Alliance Template 1.9 - Flex to continue the tracking and record potential issue with new Seal Pack, no details from Eng'g to share but higher suspect. Once design review finish with new seal pack and may have FW upgrade later. If FT performance did well in Flex, Engg will accelerate the upgrade process. - 17 Oct 2018 - To trigger 2 months ahead / During weekly Eng'g meeting and through email - 24 Oct 2018 refer to "CONSUMABLE TAB" - 21 Nov 2018 - To follow up with MCotter for chemical - Yee Hwa - 28 Nov 2018 - Yee Hwa will follow up to get additional 16 boxes that can last till early Jan19 - 05 Dec 2018: Joel to check and update on the qty. - 12 Dec 2018: return 1 box YH 16 Jan 2019: To review based on sheet "Consumables" | JOEL MD | 2018-07-25 | TBD | Closed | Medium
11 | Consigned Health review on Tooling Identify the list of tooling (approx. 275) from Su Fern Priority by family (alliance, pumps, BSM and Detectors) 2 weeks / family - 19 Sep 2018 Alliance looks OK, Test Equipment will be to prioritized - 26 Sep 2018 Onhold until tester - 09 Oct 2018 to discuss further the bandwidth? - 17 Oct 2018 - may need to push thrugh because of Oct'18 Audit - 31 Oct 2018 - resume for tooling with plan and GANTT Chart - 28 Nov 2018 - Discussed on the GANTT Chart - 12 Nov 2018 Item 1 completed (refer to CJFT Tab) - 19 Dec 2018 To provide details CK Tan as per CJFT Tab " Analyze downtime data " 16 Jan 2019: Will be review through Project meeting | Gary /CK | 8/15/2018 01 / 11/ 18 | 2018-12-31 | Closed | Medium
16 | Pump eEvent Log 05 Dec 2018: To generate a configuration file for Pump family 10 Jan 2019 : Awaiting review from YH / file submitted 27 Dec 2018 16 Jan 2019: Will be review through Project meeting | CK/ Joel | 2018-12-05 | 2018-12-31 | Closed | Medium
17 | CPU 2000 – 210000520 - REFER HISTORY from 20th Jun 2018 email - 210000520 to try on other model and pass is shippable. (YH) - to prepare a plan for 7 pcba return to Flex 05 Dec 2018: Yee Hwa to review with Melody to close this action item 19 Dec 2018 To prepare slide comparission for Year to Year CPU 2000 Performance 16 Jan 2019: Joel to consolidate the performance and close the monitoring YYYY - 2016 43/12604 (0.3411%) YYYY - 2017 32/12591 (0.2541%) YYYY - 2018 20/12315 (0.1624%) as of Nov 2018 (Dec 2018 Not Available) | Team | 2017-07-27 | 7/27/2018 8/31/2018 12/31/2018 02/28/2019 | Close | Medium

## ASP TSP Summary (hidden)
Pumps | ASP | 15 | 0 | 15 | 4 | 11 | 26.67% | 7x ASP belong to sub assy & Main Line 4x ASP belong to Packing
| TSP | 2 | 0 | 2 | 1 | 1 | 50% | 1x TSP belong to 515 pump
Last update: 26 Sep 2018
Product Family | ASP/ TSP | ASP/TSP Compliance Check by Flex/Waters | [6 empty] | ASP/TSP Compliance Implementation by Flex/Waters |  |  |  | PKI | Arc Bio
|  | Total | Not in scope -Low runners (Velox/5034/5053/Arc Bio) | In scope (High Runners) | Completed Review | Pending Review | % Status | Pending Document Remarks | Completed Implementation | Pending Implementation | % Status | Remarks
eAlliance | ASP | 30 | 4 | 26 | 26 | 0 | 100% | NA | 26 | 0 | 100% | completed
| TSP | 10 | 0 | 10 | 10 | 0 | 100% | NA | 10 | 0 | 100% | completed
BSM/QSM | ASP | 15 | 3 | 12 | 12 | 0 | 100% | NA | 11 | 1 | 91.67% | 7x Red line documents (word format)submitted to WPPL for ECO. 4x competed review. No ECO required. 2x pending ECO. - 19 Sep 18 - 1X ECO release today - 31 Oct 18 - 1X ECO release today 16Jan19 - In process for ECR Target closure: 31Jan19
| TSP | 3 | 0 | 3 | 3 | 0 | 100% | NA | 3 | 0 | 100% | completed
TUV Detectors | ASP | 26 | 7 | 19 | 19 | 0 | 100% | NA | 11 | 8 | 57.89% | 8x Red line documents (pdf format)submitted to WPPL for ECO. Target 30 Nov 18 | 8 | 2
| TSP | 6 | 0 | 6 | 6 | 0 | 100% | NA | 6 | 0 | 100% | completed | 2
PDA Detectors | ASP | 39 | 13 | 26 | 26 | 0 | 100% | NA | 7 | 19 | 26.92% | 19x Red line documents (pdf format)submitted to WPPL for ECO. Target 31 Oct 18, 31 Dec 2018 | 8 | 2
| TSP | 2 | 1 | 1 | 1 | 0 | 100% | NA | 1 | 0 | 100% | completed
Pump | ASP | 16 | 0 | 16 | 14 | 2 | 87.5% | NA | 6 | 8 | 42.86% | 5x Redline done (PDF format), 3x Redline done (Word format) 7 - In process of ECR 1 - In process of verification with Waters Target closure:31 Jan 2019
| TSP | 3 | 0 | 3 | 3 | 0 | 100% | NA | 3 | 0 | 100% | completed

## OESR_v00 (hidden)
OESR ACTION ITEMS DISCUSSED
S/N | DESCRIPTION | ACTION ITEM TO : | DATE STARTED | TARGET CLOSURE | STATUS | PRIORITY
1 | OESR # 193460 Additional Capacity for Flow cell Tester (Pressure Decay Tester - TE82475) 14Aug19 to put th TE number (JMD) 28Aug19 : Q3- Q4: 74.1% Prime / 89.3 with Prime and Spare Q1-Q2 2020:Prime 79.3% / Prime + Spare 91.56% (HZ to update) 04Sep19 - HZL send reminder to MJ Milford 16Sep19 - MJ update the Milford is working concept to modernize this tester as duplicating it is impractical due to parts obsolecence 10Oct19: HZL to check for update from MJ for the concept of modernization of this tester 13Nov19:HZL to checked for update from Milford 27Nov19: HZL updated flowcell PN to MJ and prepare CEA for review. 04Dec19: Pending for HZL comfirmation | JMD HZL | 2019-08-14 | 2020-01-31 | Open | Medium
2 | OESR # 193461 Additional Capacity for Alliance CCD Tester (TE82520 / Tool 119828) 14Aug19 to put th TE number (JMD) 28Aug19 : Q3- Q4 66.1% Prime / 74.7 with Prime and Spare Q1-Q2 2020: Prime 81.1% / Prime + Spare 85.5% (HZ to update) 04Sep19 : TE82520-3 CEA Approval loop 09Sep19: TE82520-3 CEA Approved/ HZL sent the CEA to Milford 10Oct19: HZL to check and advise the progress from Milford and respective completion date 16Oct19: ETD on Q2'20 advice from Chau Nguyen 04Dec19: Pending for HZL comfirmation | JMD HZL | 2019-08-14 | 31-Jan-20 31-Jul-20 | Open | Medium
3 | OESR # 193124 Additional Capacity for Alliance Pump Tray Tester (TE82526) 14Aug19: to put the TE number (JMD) (TE82526) 28Aug19: Milford in progress in building and updating the CPU to Windows 10 (HZ to update) 03Sep19: HZL check with Matt Brule for the tester BOM and pricing / CEA 11Sep19: Finance AT mentioned that able to use existing CEA 10Oct19: HZL to check and advise the progress from Milford 13Nov19: HZL to check for update 04Dec19: Pending for HZL comfirmation 11Dec19: On-track as per HZL feedback from Milford | JMD HZL | 2019-08-14 | 2020-01-31 | Open | Medium
4 | OESR # 190059 Alliance Final Test 3 Frequently Freeze During Final Test OESR # 193244 Milford OESR - 03 Apr 2019 04 Apr 2019 - Matt Brule request Flex to provide detail on Hang issue 05 Apr 2019 - Joel provide detail to Matt Brule for detail of Hang issue 04 Sep 2019 - Joel to monitor and update by 06 Sep 2019 06 Sep 2019 - JMD sent monitoring data to HZL (refer rto attach Alliance FT3 System Freeze Tracking Record) 09 Sep 2019 : HZL updated Matt Brule with attached Alliance FT3 System Freeze Tracking Record 18 Sep 2019: HZL last occurrence is 3 Sep 19, slow build for Alliance follows thus need to extend monitoring / with respective volume (min 100 units), HZL to inform Matt that Flex to further monitor this issue. 25 Sep 2019: JMD issue the tracker for 100 units to Prod'n and monitoring in progress | JMD HZL | 2019-09-04 | 13 Sep 19 27 Sep 19 30 Oct 19 | Monitoring | High
5 | OESR # 190079 TE82943-11 AND HR200 Electronics Balance Has LCD Missing Segment. 15May2019 JMD raised OESR for LCD Missing Segment. Its be replace by Spare Balance (after calibration by ITS ) 04Sep2019 JMD arrange recalibration of spare balance to manufacture specs by 13 Sep 2019 11Sep2019 Ming Deng Calibrate the balance as full range and waiting for final report 13 Sep2019 18Sep2019 : JMD will inform YHC once EQID 0497 is available 19Sep2019: JMD updated OESR that EQID 0497 has been calibrated and release, awaiting for VAF from YHC 25Sep2019 VAF not required as long as same model and same parameter is in used as per Process Validation Procedure 730001835 (section 8.2.2.1) the replacement of identical components is a low risk change and does not have any associated validation requirements. | JMD HZL | 2019-09-04 | 2019-09-27 | Close | High
6 | OESR # 190102 - TE82988-2 Power Supply 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to HZL that Senai need below as per investigation with Matt Brule 40-612-002 Hi Density MUX TE83004 Self Test Fixture | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
7 | OESR 190104 - TE82988-2 Spare Parts holding at Senai - to request return back to WPPL 04 Sep 2019 FLX Kallang JMD to follow-up in Senai 05 Sep 2019 Erzad to inventories the item to return to WPPL by 06 Sep 2019 06 Sep 2019 Erzad update to return below and ask HZL for shipping address, 2x unit Lambda Power Supply 1x unit Load Module (Chrome) HZL replied to return to WPPL office. 10 Sep 2019 - Erzad update ETA back to WPPL by 23 Sep 2019 | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
8 | OESR # 190108 - TE82882-2 Request for Spare Part CPU for tester - Edwin to review and advise for below spare parts Tester/Fixture# CPU Card # PCBA # 1) TE 83044 210000273 Rev G 210000345 2) TE 83002 N/A 210000414 3) TE82718 487 CPU 081060 279002660 4) TE82961-2 PCB 210000182 210000565 5) TE82940 PCB 210000333 210000355 & 210000404 04 Sep 2019 FLX Kallang JMD to follow-up to Edwin 05 Sep 2019 WPPL Edwin gave the email approval to allocate spares for CPU Cards and asked HZL to secure these spares | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | Medium
9 | OESR # S190109 - TSP having 735002531 with 735001613 does not have consistent content 04 Sep19: reviewed with HZL/EK/DA and JMD and agreed to update the OESR content and its Model PN. 05Sep19: JMD updated the OESR 190109 and send to HZL 18Sep19: HZL to checked for update 24Sep19: HZL to follow-up with E Turner 27Sep19: HZL to follow-up with E Turner reminder due as Waters Annual Audit for Flex 08Oct19 : HZL send reminder to E.T 08Oct19 - HZL ask ET if need to revise procedure/monitoring or uprev the next revision. 13Nov19: HZL to check for update | JMD HZL | 2019-09-04 | 06-Sep-19 10-Oct-19 30-Dec-19 | Monitoring | Medium
10 | OESR # 190117 - TE82988-2 GTS to check if Alliance Logo to test or read a checksum using GTS Tyler recommend to reverse back to old tester (HZL login to WPPL Folder) Flex Senai to verify if need to verify Alliance logo as per TSP 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai on when they implement the reverse back to old tester as per email from Tyler 05 Sep 2019 Start 20 Aug 2019 used in 210000411 SN 010919110M 11 Sep 2019 - U53 is program at ICT - every chip is program at ICT and verified at ICT - Paul Deschenes | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
11 | OESR #190118/ #190119 Reticule Plate Broken Target Mirror 19Aug19 Prod'n feedback Reticule Plate/Target Mirror Broken. Raised OESR (JMD) 30Aug19 Flex source quotation at US$255 per pc. SFC will source from AVL (ETA 30Dec19) 04Sep19: Need to update PM to include cosmetic check, example nick, small crack, scratches etc. 23Sep19: SFC relelase PO to vendor 10Oct19 PM documentation approval in-progress 16Oct19: Part lead time is 4week from now (16Nov19) 20Oct19: GLY assembled and release to Production. | JMD SFC | 2019-09-04 | 2019-12-30 | Close | Medium
12 | OESR #190124 - 193544 Request for TE# for ICT Fixture 210001344 04 Sep 2019 FLX Kallang JMD to follow-up from FLX Senai the request documentation from Petracarbon such as block, wiring, schematic, BOM 05 Sep 2019 Erzad to send the documentation to JMD 06 Sep 2019 Senai send available information such as Probe Slide Location/Wiring Side/Test Head Wiring List/Test Head Power List to share to WPPL 11 Sep 2019 , use this TE # for below: 210000562 TE0082540-1 210001164 / 1203 TE0084038-1 210001344 TE0084055-1 Respective documentation requested by WPPL to Petracarbon | JMD HZL Senai | 2019-09-04 | 2019-09-06 | Close | High
13 | OESR S190130 -193550 TE82882 PDA Final Test System Missing Spare OB PN 700011652 02Sep2019 - PDA Final Test System TE82882-8,9,10,11,13, and 14 have missing Spare OB PN 700011652. (Raise OESR - JMD) 04Sep2019 - JMD shown that PDAlin.txt to HZL / EK / DA have typing error and agreed to give before and aftet updated file of PDAlin.txt to HZL 05Sep2019 - JMD send to HZL the corrected PDAini.txt for follow-up to E.Turner 11Sep2019 - HZL updated that Flex need to build 1X of 700011652 for validation after E.Turner return on 18 Sep 2019 18Sep2019 - HZL and Milford advise is to run using the higher level PA and strike off the PN and follow GDP / Flex advise to wait until validation complete. JMD to follow-up internal requirement (QMS) Refer to attached email from KS Chua, its OK to Matt Franklin suggesstion, Flex have no concern that Matt to do correction and sign off on the DHR, before we release the Product for shipment. 25Sep2019 - TE82882-8, 9, 10, 11, 13, and 14 Final Test report with PN 700011652 has been submitted to HZL (24Sep19) 26Sep2019 - TE82882-10SEP19-VAF_Approved and released tester for usage. | JMD HZL | 2019-09-04 | 06 Sep 19 30 Sep 19 | Close | High
14 | OESR # S190132 - TE82943-13 USB Device Error 03Sep19 JMD raised OESR for TE82943-13 USB Device Error 04Sep19 Team agreed to replace fist the USB port and monitor until 06 Sep 2019 06Sep19 No re-occurrence of USB Device Error at the moment but will continue to monitor until 14:19 PM. Replaced the USB cable and monitor for a week. 10Sep19 08:40 AM USB Device Error pop-out on the screen again. We restart the CPU and its OK. 11Sep19 08:40 AM Monitor USB Device Error if pop-out on the screen again 18Sep19: Close all program and click Instacal icon on the desktop, If device could not see, plug and replug the USB cable, re-launch the Instacal icon, it then found the USB Device, click OK and exit the Instacal program. Monitor and highlight if surface again. 25Sep19: To verify the USB Device part number and get a qoutation with leadtime/Shutdown the tester and open the box to source and replace 09Oct19: Quotation submitted to HZL (SGD561.75 - 04 Oct 2019), need approval before release of PO, need to wait from performace check from other CM before approval 10Oct19: Bliss Pte Ltd do not accept Credit Card payments at this time, other than payment via cash or cheque, vendor also accept payment via PayNow (vendor provide the QR code once confirm an order), Total cost adjusted by vendor to SGD585.07. 10Oct19: Informed purchaser to release pay using Paypal. 16Oct19: JMD to follow-up E.T.A from buyer 23Oct19: Bliss target ETA to Flex is on 29Oct19 30Oct19: To install today (Shamsul) and update status 06Nov19: Installed and update that no issue of USB Device Error surface after installation (from 30 Oct 19 to 04 Nov 19) | JMD HZL | 2019-09-04 | 06 Sep 19 18 Sep 19 25 Sep 19 30 Oct 19 | Close | High
15 | OESR # 190133 Label Printing 04 Sep 2019 GLY to follow-up with DA | GLY / DA | 2019-09-04 | 2019-09-06 | Close | High
16 | OESR # S190156 TE82503-5 Injector Seal Pack Ramp and Decay Test Fail 06Oct19: HZL ask if able to fabricate locally but need to provide drawing and specification 07Oct19: JMD to advise the part number of this transducer 2 and check locally see if can find a replacement part. 10Oct19: JMD submitted to HZL the transducer the pressure transducer PN: 9224401 CN: ABH10KPSC1B, using Honeywell reference 31DTMAIN, no adapter / manifold available from Honeywell. 16Oct19: HZL advise to try using teflon tape if able to run test without leaking. 17Oct19: Applied teflon tape to settle the Transducer Valve adapter leak. Observed that ASFVO to waste is leaking during Ramp and Decay Test. 18Oct19: Installed the ASFVO Flow valve and able to pass 4 units. Release to Production under monitoring and close 21Oct19 21Oct19: Tester release to Production. | JMD HZL | 2019-10-03 | 2019-10-21 | Close | High
17 | OESR # S190194 TE82943 Truncated (test time) on the Test Report | JMD HZL | 2019-12-17 | TBD | Open | Medium
18 | OESR # S190195 TE82882 Truncated (test time) on the Test Report | JMD HZL | 2019-12-17 | TBD | Open | Medium

## Pump Tray Setup Schedule
UPDATE : | 2020-02-28
Pump Tray Test Capacity Enhancement (OESR 193124-TRF MJUE-BJYJL4) |  |  | WW 0320 13-17 Jan'20 | WW 0420 20-24 Jan'20 | WW 0520 27-31Jan'20 | WW 0620 03-07 Feb'20 | WW 0720 10-14 Feb'20 | WW 0820 17-21 Feb'20 | WW 0920 24-28 Feb'20 | WW 10'20 02-06 Mar'20 | WW 11'20 09-13 Mar '20 | WW 12'20 16-20 Mar'20 | WW 13'20 23-27 Mar'20 | WW 14'20 30 Mar-03 Apr '20
TE 82526 (new) | Plan | Tester E.T.A. to Flex and verified complete | [5 empty] | FLEX PHYSICAL INVENTORY 17-20 Feb 2020
| Actual |  | 16 Jan 2020 Do Not Open
| Plan | Build Creform |  | Creform Table Setup
Actual
| Plan | Mount and Install the Pump Tray Tester
Actual
| Plan | Validation
Actual
TE 82526 Upgrade | Plan | Tester E.T.A. to Flex and verified complete
| Actual |  | 16 Jan 2020 Do Not Open
| Plan | Mount and Install the Pump Tray Tester upgrade
Actual
| Plan | Validation
| Actual | [8 empty] | 05 Mar 2020 IQR Report submitted to WPPL | Validation done
TE 82526-1 Upgrade | Plan | Tester E.T.A. to Flex and verified complete
| Actual |  | 16 Jan 2020 Do Not Open
| Plan | Mount and Install the Pump Tray Tester upgrade
| Actual | [9 empty] | Upgrading in progress
| Plan | Validation
Actual
TE 82526-4 Upgrade | Plan | Tester E.T.A. to Flex and verified complete
| Actual |  | 16 Jan 2020 Do Not Open
| Plan | Mount and Install the Pump Tray Tester upgrade
Actual
| Plan | Validation
Actual

## Pump Head (HotPart) (hidden)
Update as if 04 Dec 2019
Below data per latest sound measurement conducted today.
Present : Aldwin. Ivy, Desmond, Guan Yee, Eugene
As per discussion today the data gathered today shall be used as reference final reference for the Sound comparison of Hot Part vs Known Good Part..
Observations:
The sound level of the Hot Part Motor (68.2 dB) is still high when installed to production unit (Hot Part Motor + New#6)
There is an average of 5.0 dB difference between the Hot Part compared to 6 samples of production units.
After re-sitting ang tightening the pump coupling shaft (WAT271467) of the Hot Part Motor, the sound level is still higher.
SN | Date | Comparison | General Sound Level ( A ) dB (min) | General Sound Level ( A ) dB (max)
1 | 2019-12-04 | Hot Part (1st try) | 66.4 | 68.4
2 | 2019-12-04 | New#1-1 | 62 | 64.6
3 | 2019-12-04 | New#2-1 | 62.6 | 65.2
4 | 2019-12-04 | New#3-1 | 61.4 | 63.1
5 | 2019-12-04 | Hot Part (2nd try) | 68.4 | 70
6 | 2019-12-04 | Hot Part (3rd try) | 68.2 | 69.5
7 | 2019-12-04 | New#4-1 | 65.1 | 67.6
8 | 2019-12-04 | New#5-1 | 61.8 | 62.5
9 | 2019-12-04 | New#5-2 | 61.7 | 62.3
10 | 2019-12-04 | New#6-1 | 64 | 64.7
11 | 2019-12-04 | Hot Part (4th try) | 67.4 | 70.8
12 | 2019-12-04 | New 6 Motor + Hot Part Body (1st try) | 62.8 | 64.5
13 | 2019-12-04 | New#6-2 | 66.8 | 67.7
14 | 2019-12-04 | Hot Part Motor + Hot Part Body (after Tightened) | 69.3 | 70
15 | 2019-12-04 | Hot Part Motor + Hot Part Body (test 1) | 68.1 | 69.4
16 | 2019-12-04 | Hot Part Motor + Hot Part Body (test 2) | 67.9 | 69.5
17 | 2019-12-04 | Hot Part Motor + New#6 | 66.8 | 68.2
18 | 2019-12-04 | New#6 Motor + Hot Part Body | 63.7 | 64.2
2019-12-09 | Final report submitted to WPPL QA Team. Pending final dispotion. Eng'g Team activity Close

## Alliance FT3 Tracker (hidden)
Alliance Final Tester 3 Freeze OESR S190059/OESR 193244 Alliance Final Test 3 Frequently Freeze During Final Test
Item | Serial Number | Freeze (Y/N) | Time | Prepared By | Remarks
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50

## IQR TE82943 UPDATE (hidden)
Used Known Good Unit QSM+ (F19QSP275A) and BSM (F19UPB765A)
Week 0f (estimate) | System Number | 2019-06-03 | 2019-06-03 | 2019-06-10 | 2019-06-17 | 2019-06-24 | 2019-07-01 | 2019-07-08 | 2019-07-15 | 2019-07-22 | 2019-07-29 | 2019-08-05 | 2019-08-12 | 2019-08-19 | 2019-08-26 | 2019-09-02 | 2019-09-09 | 2019-09-16 | 2019-09-23
WPPL Flextronics | TE82943-11 | 2019-06-03: TMD 010161 v01; 2019-06-03: IQ
| ACTUAL | 2019-06-03: TEST SOFTWARE UPDATE 07 Jun 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-06-10: 10 Jun 2019 186018080 F19QSP275A 14 Jun 2019 186015001 F19UPB765A; 2019-06-17: PREPARED BY: CYH 19 Jun 2019; 2019-06-24: REVIEWED BY: EM 19 Jun 2019 EJ 20 Jun 2019 APROVED BY: TKC 28 Jun 2019
| TE82943-12 | 2019-06-03: PLAN; 2019-06-10: IQ
| ACTUAL | 2019-06-17: TEST SOFTWARE UPDATE 25 Jun 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-06-24: 26 Jun 2019 186018080 F19QSP275A 26 Jun 2019 186015001 F19UPB765A; 2019-07-08: PREPARED BY: CYH 09 Jul 2019 REVIEWED BY: EJ 12 Jul 2019; 2019-07-15: APROVED BY: TKC 15 Jul 2019
| TE82943-13 | 2019-06-03: PLAN; 2019-06-17: IQ
| ACTUAL | 2019-07-01: TEST SOFTWARE UPDATE 09 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-07-08: 10 Jul 2019 186018080 F19QSP275A 11 Jul 2019 186015001 F19UPB765A PREPARED BY: CYH 12 Jul 2019 REVIEWED BY: EJ 12 Jul 2019; 2019-07-15: APROVED BY: TKC 15 Jul 2019
| TE82943-14 | 2019-06-03: PLAN; 2019-06-24: IQ
| ACTUAL | 2019-07-15: TEST SOFTWARE UPDATE 17 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 18 Jul 2019 186018080 F19QSP275A 19 Jul 2019 186015001 F19UPB765A; 2019-07-22: PREPARED BY: CYH 25 Jul 2019 REVIEWED BY: LHZ 26 Jul 2019; 2019-07-29: APPROVED BY: TKC 29 Jul 2019
| TE82943-15 | 2019-06-03: PLAN; 2019-07-01: IQ
| ACTUAL | 2019-07-15: TEST SOFTWARE UPDATE 17 Jul 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software 18 Jul 2019 186015001 F19UPB765A; 2019-07-22: 22 Jul 2019 186018080 F19QSP275A PREPARED BY: CYH 25 Jul 2019 REVIEWED BY: LHZ 26 Jul 2019; 2019-07-29: APPROVED BY: TKC 29 Jul 2019
| TE82943-16 | 2019-06-03: PLAN; 2019-07-01: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 07 Aug 2019 186015080 F19QSP275A 08 Aug 2019 186015001 F19UPB765A; 2019-08-12: APPROVED BY: TKC 14 Aug 2019
| TE82943-17 | 2019-06-03: PLAN; 2019-07-08: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8 08 Aug 2019 186015080 F19QSP275A 07 Aug 2019 186015001 F19UPB765A; 2019-08-12: APPROVED BY: TKC 14 Aug 2019
| TE82943-18 | 2019-06-03: PLAN; 2019-07-15: IQ
| ACTUAL | 2019-08-05: TEST SOFTWARE UPDATE 06 Aug 2019 1.BSM Test Software from 2.2 to 2.3 2.QSM Plus Test Software from 1.7 to 1.8; 2019-08-12: 14 Aug 2019 - testing completed; 2019-08-19: 21 Aug 2019 Resubmit : EQID 0497 Test report with updated calibration report and Evauation Result - Not allowed to used due to somer error on 90g reading during calibration; 2019-08-26: TMD010661 v 01 To be extended until end Sept 19 29 Aug 2019 Submitted TE89243-18 test report to YHC using AND HR 200 SN 12337042 (EQID 2092) 30 Aug 2019 WPPL ISOISO13485 Audit; 2019-09-02: WPPL ISO13485 Audit 30 Aug 2019-02 Sep 2019 TMD010661 v 02 Start at Flex 02 Sep 2019 extended until end Sept 19 AND HR 200 ITS verification completed - within specs. 02 Sep 2019 AND HR 200 Miing Deng re-calibration within specs 05 Sep 2019; 2019-09-09: KC Tan on business trip until 18 Sep 19; 2019-09-16: KC Tan sign-off 18 Sep 19

## Consined Alliance Training Plan (hidden)
Known Good Alliance Transfer Plan To Flex Senai
Mounting and dismantling of PCBA to UUT - 2695, Alarm Message (Related to Allianc 2695 only) - Instrument Specification (Basic related to PCBA) | WW 40 | WW 41 | WW 42 | WW 43 | WW 44 | WW 45 | REMARKS
Plan
Actual
Generic Preventive Maintenance Procedure For Test Equipment (735004259)
Plan
Actual
Transfer of Alliance Known Good Unit to Senai
Plan
Actual
Flex Senai to use Known Good Unit for Troubleshooting
Plan
Actual
Flex Senai to source Spare Part such as cables to enhance toubleshooting
Plan
Actual

## TUV Shunt Energy Level (hidden)
TUV Flow Cell | 55% to 150 %
Energy Setting as per ASP
Shunt ID

## Pump Spares Part (hidden)
Waters ID | Description | Part Number | Spare Part Description | OESR | SOURCE BY: |  | Remarks
[5 empty] | FLEX (Yes/No) | WPPL (Yes/No)
TE82705-2 | 1500 Series Final Test Station | AND HR 200 | Electronics Balance |  |  | Yes
|  | WAT024024 | 0.020 x 0.062 SST |  | Yes
|  | WAT024025 | 0.040 X 0.062 SST |  | Yes
|  | ASFVO | Air switch Flow Valve |  |  | Yes
|  | M401-37 | Solenoid Valve |  |  | Yes
|  | BSDA025T-NC-PV | Flow Valve |  |  | Yes
|  |  | Waters Pump Control Module |  |  | Yes
|  | SC2072 ASSY181310-02 | National Instrument Controller Card |  |  | Yes
TE82534 | 515 Final Tester |  |  |  |  | Yes
|  | AND HR 200 | Electronics Balance |  |  | Yes
|  | WAT024024 | 0.020 x 0.062 SST |  | Yes | No
|  | WAT024025 | 0.040 X 0.062 SST |  | Yes | No
|  | ASFVO | Air switch Flow Valve |  |  | Yes
|  | M401-37 | Solenoid Valve |  |  | Yes
|  | BSDA025T-NC-PV | Flow Valve |  |  | Yes
TE82613-2 | Pump Mechanic Run In Module |  | TE82613 PUMP MECHANIC RUN IN MODULE |  |  | Yes
TE82851 | 515 Bezel Assemby Tester |  | TE82851 , 515 BEZEL ASSEMBLY TESTER Both Waters customized PCBA PN is not available |  |  | Yes

## Alliance Pump HDD Imaging (hidden)
Alliance / pump |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.50 | TE 82534 - 2 | 515 Pump Final Test | 2019-02-12 | Completed
10.249.26.51 | TE 82705 - 2 | 1500 pump Final Test | 2019-03-12 | Completed
10.249.26.20 | Final Test - 1 | Waters Alliance Test Station | 2019-02-12 | Completed
10.249.26.21 | Final Test - 2 | Waters Alliance Test Station | 2019-03-13 | Completed
10.249.26.22 | Final Test - 3 | Waters Alliance Test Station | 2019-02-11 | Completed
10.249.26.23 | Final Test - 4 | Waters Alliance Test Station | 2019-02-12 | Completed
10.249.26.24 | Final Test - 5 | Waters Alliance Test Station | 2019-03-13 | Completed
10.249.26.25 | Final Test - 6 | Waters Alliance Test Station | 2019-02-11 | Completed
10.249.26.26 | TE 82526 - 1 | Waters Alliance Test Station ( Pumptray Test - 1 ) | 2019-03-13 | Completed
10.249.26.27 | TE 82526 - 2 | Waters Alliance Test Station ( Pumptray Test - 2 ) | 2019-03-13 | Completed
10.249.26.28 | TE 82526 - 3 | Waters Alliance Test Station ( Pumptray Test - 3 ) | 2019-03-13 | Completed
10.249.26.29 | TE 82519 | Waters Alliance Test Station ( Pump Head Mechanics ) | Not able to backup, CPU auto shutdown once insert the CD Rom
10.249.26.32 | TE 82573 - 1 | E-Alliance Rear Panel Test | 2019-03-12 | Completed
10.249.26.34 | AutoLoader - 2 | Alliance AutoLoader Station 2 | 2019-03-12 | Completed
10.249.26.33 | AutoLoader - 3 | Alliance AutoLoader Station 3 | 2019-03-12 | Completed
10.249.26.35 | TE 82520 - 1 | Alliance Carousel/Carrier Alignment Station -1 | Not able to backup, CPU auto shutdown once insert the CD Rom
10.249.26.36 | TE 82520 - 2 | Alliance Carousel/Carrier Alignment Station -2 | 2019-03-13 | Backup thru Windows OS, not able to use Acronis Imaging Software
10.249.26.37 | TE 82546 | Front Panel Station | 2019-03-13 | Completed
10.249.26.38 | TE 0084018-1 | Heater Cooler Tester | 2019-03-28
10.249.26.90 | QA Workstation | QA Workstation/SQT.ORT | 2019-02-11 | Completed
NA | PMS | Particle Measurement |  | Plan : 30 Apr 19
PLAN : Completion Date

## BSM.QSM HDD Imaging (hidden)
BSM /QSM |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.2 | Waters SISNL1 | Solectron Instrument Serial Number 1 | 2019-03-12 | Completed
10.249.26.3 | Waters SPNL1 | Solectron Part Number Labeling 1 | 2019-03-12 | Completed
10.249.26.4 | WATERSCOFP1 | WATERS Certificate Of Funtional Performance 1 | 2019-03-12 | Completed
10.249.26.9 | TE 82800 - 2 | Acquity Actuator Tester #2 | 2019-02-13 | Completed
10.249.26.10 | TE 82800 - 3 | Acquity Actuator Tester #3 | 2019-03-12 | Completed
10.249.26.11 | TE 82943 - 11 | Acquity SDS Tester #11 | 2019-02-13 | Completed
10.249.26.12 | TE 82943 - 12 | Acquity SDS Tester #12 | 2019-03-12 | Completed
10.249.26.13 | TE 82943 - 13 | Acquity SDS Tester #13 | 2019-02-13 | Completed
10.249.26.14 | TE 82943 - 14 | Acquity SDS Tester #14 | 2019-02-13 | Completed
10.249.26.15 | TE 82943 - 15 | Acquity SDS Tester #15 | 2019-02-13 | Completed
10.249.26.16 | TE 82943 - 16 | Acquity SDS Tester #16 | 2019-02-13 | Completed
10.249.26.17 | TE 82943 - 17 | Acquity SDS Tester #17 | 2019-03-12 | Completed
10.249.26.18 | TE 82943 - 18 | Acquity SDS Tester #18 | 2019-03-13 | Completed
10.249.26.55 | Genealogy SDS | Acquity SDS Genealogy Stn (BSM/QSM) | 2019-03-12 | Completed
PLAN : Completed

## Detector HDD Imaging (hidden)
Detectors / Cleanroom |  |  | Image Created
IP Address | Name | Description | Date | Remarks
10.249.26.40 | TE 82903 - 6 | TUV Final Test Station | N/A | White Lace \ CD ROM Not able to activate Acronis Imaging Software
10.249.26.42 | TE 82903 - 8 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.43 | TE 82903 - 9 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.44 | TE 82903 - 10 | TUV Final Test Station | 2019-03-14 | Completed
10.249.26.45 | TE 82964 - 1 | TUV Pretest System | 2019-03-13 | Completed
10.249.26.75 | TE 82969 - 1 | PDA Linearity Tester | 2019-03-13 | Completed
10.249.26.76 | TE 83057-4 | PDA Download Station | 2019-03-13 | Completed
10.249.26.78 | TE 82882 - 2 | 2998/PDA Final Test System | 2019-03-13 | Completed
10.249.26.79 | TE 82882 - 8 | 2998/PDA Final Test 8 | 2019-03-13 | Completed
10.249.26.80 | TE 82882 - 9 | 2998/PDA Final Test 9 | 2019-03-13 | Completed
10.249.26.81 | TE 82882 - 10 | 2998/PDA Final Test 10 | 2019-03-13 | Completed
10.249.26.82 | TE 82882 - 11 | 2998/PDA Final Test 11 | 2019-03-14 | Completed: Backup thru Windows OS
10.249.26.84 | TE 82882 -13 | PDA Final Test System 13 | 2019-03-14 | Completed: Backup thru Windows OS
10.249.26.85 | TE 82882 -14 | PDA Final Test System 14 | 2019-03-14 | Completed
10.249.26.63 | TE 82918 - 2 | 2998 Fine Alignment Sys | 2019-03-14 | Completed
10.249.26.64 | TE 82868 - 2 | 2998 Aperture Alignment System | 2019-03-14 | Completed
10.249.26.65 | TE 82951 - 2 | 2998 Mask Alignment Sys | 2019-03-14 | Completed
10.249.26.66 | TE 82916 - 2 | Flowcell Alignment Fixture | 2019-03-14 | Completed
10.249.26.68 | TE 82973 - 1 | 2998 Flowcell Noise & Drift System | NA | CD-ROM unable to read
10.249.26.72 | TE 82917 - 1 | 2998 Rough Alignment | 2019-03-14 | Completed
10.249.26.73 | TE 82880 - 2 | PDA Rough Alignment Sys | 2019-03-14 | Completed
10.249.26.74 | TE 82881 - 2 | PDA Fine Alignment Stn | 2019-03-14 | Completed
10.249.26.56 | Genealogy SDS | Acquity SDS Genealogy Stn (Detector) | 2019-03-15 | Completed: Backup thru Windows OS
10.249.26.59 | TE83096-2 | Universal PDA Alignment | 2019-03-16 | Completed: Backup thru Windows OS
PLAN : Completion Date