Ripple test

swon eventinjectstart (switch power on)

reset  pumpMotorCtlr (resets pumptray)

swon eventsw2 (direct to waste valve)

swon eventsw4 (solvent supply)

pumphome (initializes pump mechanics)

getstate (should be idle)

swon eventsw4 (solvent supply)

swon eventsw2 (direct to waste valve)

swon eventsw1

swon eventsw3

setflow 1

swoff eventsw2

xread accum (check for 1700 psi 1600 – 1800)

