To Check Restrictors from Hyperterminal
Commands are in bold letters.

Hit enter after typing each command.
swon eventinjectstart (switch power on)

reset  pumpMotorCtlr (resets pumptray)

swon eventsw2 (FV2 direct to waste valve)

swon eventsw4 (solvent supply)

pumphome (initializes pump mechanics)

getstate (should be idle)

swon eventsw4 (solvent supply)

swon eventsw2 (FV2 direct to waste valve)

swon eventsw1( FV1 allow s flow to restrictors through FV3)
swon eventsw3 ( FV3 select  1700 psi restrictor)
setflow 1

swoff eventsw2 (close FV2  flow to waste)
xread accum (check for 1700 psi 1600 – 1800) 
swoff eventsw3 ( FV3 select  600 psi restrictor)

xread accum (check for 600 psi 500 - 700) 
