## TE82943-1
SERIALNUMBER | Tester ID | FAILURES MODE | Month | FAILING SUB-ASSY | LSL | Measurement | USL
A18VQS563G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-05 | Retest | 0.297 | 0.257 | 0.303
A18VQS601G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-19 | Retest | 0.297 | 0.312 | 0.303
A18VQS639G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-31 | Re-Tighten Tubing Connected to Vent Valve | 0.297 | 0.309 | 0.303
A18VQS650G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-02-02 | Retest | 0.297 | 0.305 | 0.303
A18VQS655G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-02-03 | Retest | 0.297 | 0.289 | 0.303
B18VQS667G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-02-11 | Retest | 0.297 | 0.305 | 0.303
B18VQS723G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-02-25 | Retest | 0.297 | 0.322 | 0.303
B18VQS743G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-02 | Retest | 0.297 | 0.296 | 0.303
C18VQS757G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-07 | Priming and Retest | 0.297 | 0.328 | 0.303
C18VQS763G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-08 | Priming and Retest | 0.297 | 0.306 | 0.303
C18VQS769G | TE82943-1 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-08 | Priming and Retest | 0.297 | 0.335 | 0.303
SERIALNUMBER | Tester ID | FAILURES MODE | Month | FAILING SUB-ASSY | LSL | Measurement | USL
B18VQS707G | TE82943-1 | Test Flow Rate Accuracy 2.5ml/min | 2018-02-23 | Retest | 2.475 | 2.59 | 2.525
B18VQS721G | TE82943-1 | Test Flow Rate Accuracy 2.5ml/min | 2018-02-24 | Re-Tighten Tubing Connected to Vent Valve | 2.475 | 2.535 | 2.525
B18VQS749G | TE82943-1 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-02 | Priming and Retest | 2.475 | 2.529 | 2.525
C18VQS760G | TE82943-1 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-07 | Priming and Retest | 2.475 | 2.616 | 2.525
C18VQS753G | TE82943-1 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-07 | Priming and Retest | 2.475 | 2.743 | 2.525
C18VQS807G | TE82943-1 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-20 | Priming and Retest | 2.475 | 2.59 | 2.525

## TE82943-24
SERIALNUMBER | Tester ID | FAILURES MODE | Month | FAILING SUB-ASSY | LSL | Measurement | USL
A18VQS560G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-04 | Retest | 0.297 | 0.276 | 0.303
A18VQS571G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-07 | Retest | 0.297 | 0.295 | 0.303
A18VQS573G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-08 | Retest | 0.297 | 0.281 | 0.303
A18VQS581G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-11 | Retest | 0.297 | 0.28 | 0.303
A18VQS594G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-13 | Retest | 0.297 | 0.288 | 0.303
A18VQS608G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-21 | Retest | 0.297 | 0.288 | 0.303
A18VQS625G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-01-24 | Retest | 0.297 | 0.292 | 0.303
B18VQS699G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-02-19 | Priming and Retest | 0.297 | 0.312 | 0.303
B18VQS712G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-02-25 | Changed Check valve (W1205001670-R) | 0.297 | 0.309 | 0.303
C18VQS768G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-09 | Retighten Check Valve W1205001670-R | 0.297 | 0.294 | 0.303
C18VQS785G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-13 | Change Check Valve W1205001670-R | 0.297 | 0.296 | 0.303
C18VQS793G | TE82943-24 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-14 | Priming and Retest | 0.297 | 0.304 | 0.303
C18VQS834G | TE82943-21 | Test Flow Rate Accuracy 0.3ml/min | 2018-03-28 | Priming and Retest | 0.297 | 0.293 | 0.303
SERIALNUMBER | Tester ID | FAILURES MODE | Month | FAILING SUB-ASSY | LSL | Measurement | USL
A18VQS614G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min | 2018-01-22 | Retest | 2.475 | 0 | 2.525
A18VQS619G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min | 2018-01-24 | Retest | 2.475 | 2.454 | 2.525
B18VQS691G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min | 2018-02-18 | Retest | 2.475 | 2.549 | 2.525
C18VQS767G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-08 | Priming and Retest | 2.475 | -0.002 | 2.525
C18VQS801G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-19 | Retighten Check Valve W1205001670-R | 2.475 | 0.261 | 2.525
C18VQS821G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-26 | Retighten Check Valve W1205001670-R | 2.475 | 2.596 | 2.525
C18VQS828G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min; | 2018-03-27 | Priming and Retest | 2.475 | 0.088 | 2.525
C18VQS841G | TE82943-24 | Test Flow Rate Accuracy 2.5ml/min | 2018-03-28 | Retighten Check Valve W1205001670-R | 2.475 | 2.526 | 2.525