## Slide 1
CM – 2019 Failure

## Slide 2
No | Failure | % Fail 2019 | % Improve 2020 | Actions & Results |  |  | 
 |  |  |  | Venture | PIC / Date | Waters | PIC / Date
1 2 | Heater Cooler Failure (5039, 5049) Immediately cut off test and fail at heater cooler 1 & 2. Ecord Failure (5038, 5039, 5043, 5049) Ecord fail although i-button is connected properly | 50.5% CM 10.8% CM-Aux 11.0% | To be monitored and advised in Monthly Basic | Heat up column temperature above 22deg C. It helps for heater cooler 1 but may still fail heater cooler 2 Propose to change the software setting for start up temperature at 18degC instead of 20 degC Restart UUT just before Ecord 3 testing. It is suspected software enumeration issue Need Waters Design Team to look into software coding | Chan Since Feb 2019 Chan Since Feb 2019 |  | 
CM – 2020 FPY Improvement Plans

## Slide 3
No | Failure | % Fail 2019 | % Improve 2020 | Actions & Results |  |  | 
 |  |  |  | Venture | PIC / Date | Waters | PIC / Date
3 | Leak Detector Interface Test Failure (5038, 5039, 5043, 5049) Stop at “dump leakdetector” and no response from software | CM 15.5% CM-Aux 22.0% | To be monitored and advised in Monthly Basic | Check PC driver in “InstaCal” , the USB device is detected. Restart PC, retest and passed Swapped leak detector switch box with another tester , the results still the same It is suspected software communication issue Need Waters Design Team to look into software coding. | Chan Since Feb 2019 |  | 
CM – 2020 FPY Improvement Plans