## Slide 1: First Pass Yield Discussion on Column Manager – Others failures
11
 | TSP: 735002236 CM-Aux (186015049, 186015039) | Frequency Occurs | Observations
Ecord 1 | Could not detected Ecord although i-button is connected properly | ~1/10 | Communication is ok if 2 E-cords are connected to UUT since LED test in console
Leak Detector Interface Test | Only two clicking sound from sensor box instead of 3 times. | ~1/10 | Communication lost or slow during test. Retest and Pass

## Slide 2: First Pass Yield Discussion on Column Manager – Others failures
11
 | TSP: 735002229 CM (186015043, 186015038 | Frequency Occurs | Observations
Firmware Download | Hang at firmware download. Cannot reset UUT by test application | ~1/10 | Retest only
Ecord 3 & 6 | Could not detected Ecord although i-button is connected properly | ~2/10 | Communication is ok if all 4 E-cords are connected to UUT at the beginning of test. and Manually reset UUT once at the start of Ecord 3.
Leak Detector Interface Test | No return value from test software. Fail test at “dump leakdetector” | ~2/10 | Check driver under InstaCal and it is detected. Retest and Pass
Left valve/Right valve | No issue in debug mode but fail in production mode | ~1/10 | Retest and Pass