## Waters Equipment Loc
Person In Charge : Joel
Locations | STATIONS | Customer ID | Equipment | Maker | Model | Serial No. | Calibration Procedure | Calibration Date | Calibration Cycle (Months/ Days) | Calibration Due Date | Calibration By Internal | Calibration By External | Calibration Report No | Reason of Failure and Length of downtime | Replacement Item / Spare Parts | Replacement Cost (US$) | STATUS | TEST/PROCESS | Cycle Time (mins) | % YEILD @ FINAL TEST STATION ONLY | Number of common testers | Criticalilty | Risk | No of occurences of Down time | Action Taken | Schematics avialable? (Yes/No) | Visual Condition? (Pass/Fail) | ObsolateParts? (Yes/No) | Abnormality/ Intermittent? (Pass/Fail) | Critical to Line Down? (Yes/No) | Checked By | Reviewed By | Remarks | Additional Remarks
eAlliance Sub Assembly Stations | 2690 eAlliance Front Panel Cell | TE82546 | 2690 eAlliance Front Panel Test Station | Waters | eAlliance | TE82546 | NA | NA | NA | NA | NA | NA | NA | NA |  |  | In-Use | TEST | 15 | 98% | 1 | 3 | 3 | 0 | NA | No | Fail | Yes | Pass | No | David | Gary | Able to test using Final Test in case tester down. | 1.) No ON/OFF Light Indication. 2.) No Proper Cover,When Not In Use. 3.) Blue Connector Broken. 4.) Wire Soldered 4 Places For Extension. 5.) Front Panel PCBA (056370) Rev AD is Obsolete.
| Pump Head Tester | TE 82519 | 2690 Pump Head Tester | WATERS | eAlliance | TE 82519 | 069004 | 2018-06-01 | 183 | 2018-12-01 | Shamsul | NA | CR/530 | Down On 13/07 Continue Beeping |  |  | In-Use | TEST | 10 | 90% | 1 | 3 | 3 | 4 | 1.) Replace Motor Of Degasser 2.) Replace Tubings and Change New Filter. 3.) Replaced tubing from transducer to filter and retightened all tubings 4.) Replaced Dc power cable,reconnect cables to PCBand replace air switch for valve 1 | No | Pass | Yes | Pass | No | David | Gary | Use Pump Tester To Testing | 1.) Frequent Change Flow Valve. 2.) Change Inlet/Outlet Tubing Every 2-4Days. 3.) Pump Head Tester PCBA (056451) Rev.0 is Obsolete.
| eAlliance Pump Tray 1 | TE82526 | 2690 Pump Tray Tester | Waters | eAlliance | TE82526-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 85% | 3 | 2 | 3 | 1 | 1.) Restart Tester And CPU | No | Fail | Yes | Fail | No | David | Gary | NA | 1.) Pump Tray PCBA (056370) Rev AD is Obsolete. 2.) Light Indicator Not Work. 3.) ACCUM Connector Intermittent. 4.) ASFV0 - Air Switch Flow Valve (0) is Obsolete. 5.) Interface Fitting. 6.) Frequent Change Flow Value Unable To Ramp.
|  | 90891 | Electronic Weighing Balance | WATERS | HR200 | 12305757 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | MDF120074-3 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2 | NA | [9 empty] | Obsolate
| eAlliance Pump Tray 2 | TE82526 | 2690 Pump Tray Tester | Waters | eAlliance | TE82526-2 | NA | NA | NA | NA | NA | NA | NA | 14 Jun 18 - Present Failed on ripple test TE82526 Failed on ripple test(pump tray tester) PUMP TRAY TUBING LEAKING UNBLE TO RAMP UP 5000 PSI TE82526-1 (PUMP TRAY TESTER) UNABLE TO START UP | May need a new vent valve p/n:7716-053 due to leakage found in between head and body of vent valve | Consigned | In-Use | TEST | 35 | 85% | 3 | 2 | 2 | 4 | 1.) Interface Wear And Tear 2.) Replace Interface Tubing 3.) Restart Tester And CPU 4.) Repair In Progress | No | Fail | Yes | Fail | No | David | Gary | NA | 1.) Frequent Charge Flow Value Unable To Ramp. 2.) Light Indicator Not Work. 3.) Pump Tray PCBA (056370) Rev AD is Obsolete. 4.) ASFV0 - Air Switch Flow Valve (0) is Obsolete. 5.) Interface Fitting.
|  | 95206 | Electronic Weighing Balance | AND | HR200 | 12316922 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | CR/MDF120074-1 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2 | NA | [9 empty] | Obsolate
| eAlliance Pump Tray 3 | TE82526 | 2690 Pump Tray Tester |  |  | TE82526-4 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 85% | 3 | 2 | 2 | 1 | 1.) Replace New Tester Fan | No | Fail | Yes | Pass | No | David | Gary | NA | 1.) Pump Tray PCBA (056370) Rev AD is Obsolete. 2.) ACCUM Connector Broken. 3.) ASFV0 - Air Switch Flow Valve (0) is Obsolete. 4.) Interface Fitting. 5.) UUT Power Light Indicator Not Work. 6.) Frequent Change Flow Valve.
|  | 90038 | Electronic Weighing Balance |  |  | 12200404 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | CR/MDF140265-1 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2 | NA | [9 empty] | Obsolate
| eAlliance HPMV Cell | TE 82625-1 | HPMV Tester | WATERS | NA | TE 82625-1 | 735000288 | 2018-01-19 | 181 | 2018-07-19 | Shamsul | NA | CR/561 | Tester failed |  |  | In-Use | TEST | 10 | 90% | 1 | 3 | 3 | 3 | 1.) Repair Interface Tubing 2.) Reset Power 3.) Clean The Restrictor Line | No | Fail | Yes | Pass | No | David | Gary | Can Use FT To Do Testing | 1.) FittingSpoiled. Replaced With Compatible One. 2.) Frequent Changing Of Tubing Every 2 - 3 Days. 3.) Replace In-line Filter 3 - 6 Month. 4.) Fixture Modified With Tape. 5.) Keypad Not In Good Condition. 6.) Replace Connector PIN When Required. 7.) All of HPMV Tester PCBA is Obsolete. 8.) Zilog Microcontroller Is Obsolete.
| 2690/2790 Syringe Drive Cell | TE82656 | 2690/2790 Syringe Drive Tester | Waters | eAlliance | TE82656 | NA | NA | NA | NA | NA | NA | NA | NA |  |  | In-Use | TEST | 10 | 95% | 1 | 3 | 2 | 0 | NA | No | Pass | Yes | Pass | NO | David | Gary | NA | 1.) Connectors Broken & Replaced. 2.) No Longer Building 2790 So 2790 Function Not Required. 3.) Controller IC (UCN5804B) Is Obsolete. 4.) All Syringe Tester PCBA is Obsolete.
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-1 | Carousel Carrier Drive Alignment Station | Waters | NA | TE82520-1 | NA | NA | NA | NA | NA | NA | NA | NA |  |  | In-Use | TEST | 35 | 90% | 1 | 3 | 3 | 1 | 1.) Reassemble and retighten the shaft and the union | No | Pass | No | Pass | No | Joel | CK | NA | NA
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-2 | Carousel Carrier Drive Alignment Station | Waters | NA | TE82520-2 | NA | NA | NA | NA | NA | NA | NA | NA |  |  | In-Use | TEST | 35 | 90% | 1 | 3 | 3 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| Fluidics Tray Tester | TE82523 | Fluidics Tray Tester |  |  | TE82523 | NA | NA | NA | NA | NA | NA | NA | NA |  |  | In-Use | TEST | 5 | 90% | 1 | 3 | 3 | 2 | 1.) Rewiring Of Cables Due To Wear And Tear | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Keypad Not In Good Condition. 2.) All Fluidics TrayTester PCBA Is Obsolete. 3.) Zilog Microcontorller Is Obsolete.
| Injector Seal Pack Cell/Engineering Lab | TE 82503-04 | Injector Seal Pack Tester | Waters | NA | TE 82503-04 | WAT043816 | 2018-01-18 | 184 | 2018-07-21 | Shamsul | NA | CR/562 | Fail on ramp and decay testTE82503-5 unable to startup tester |  |  | In-Use | TEST | 15 | 85% | 2 | 3 | 3 | 4 | 1.) Replace Interface tubings 2.) re-calibrate and re-tighten tubings and transducers 3.) Replaced 2 filters,tubings and resistor. re-tightened 4.) replaced interface tubing | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Keypad Not In Good Condition. 2.) Poor Joints For Inlet/Outlet. 3.) All injector Seal Wash Tester PCBA is Obsolate. 4.) Interface Fitting. 5.) Zilog Microcontroller is Obsolete.
|  | TE 82503-05 | Injector Seal Pack Tester |  | NA | TE 82503-05 | WAT043816 | 2018-03-02 | 184 | 2018-09-02 | Shamsul | NA | CR/562 | 14 Jun 18 - 14 Jun 18 Failed on decay test | Replace flow valve ASFVO at valve 3 due to leakage | Consigned | In-Use | TEST | 15 | 85% | 2 | 3 | 3 | 3 | 1.) Repair done on broken connector 2.) replace valve and tubings 3.) Retighten transducer fitings | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Keypad Not In Good Condition. 2.) Poor Joints For Inlet/Outlet. 3.) All injector Seal Wash Tester PCBA is Obsolete. 4.) Interface Fitting.
eAlliance Final Test | Alliance Programming Station | NA | Autoloader 2 |  |  | Lace 32 B08L32122N | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 25 | 95% |  | 1 |  | 0 | NA | No | Pass | Yes | Pass | Yes | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP)
| Rear Panel Station | TE82573-1 | NA |  |  | Lace 32 C08L32093W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 7 | 95% |  | 1 |  | 0 | NA | No | Pass | Yes | Pass | Yes | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP)
eAlliance Final Test Station 1 | eAlliance Final Test 1 CPU | IBM | CPU | KKPRA0M | CPU | C12L32255W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% | 5 | 1 | 1 | 1 | 1.) Missing (corrupted) \Windows\system32\ config\system | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP) 2.) Shortage Of TCM 1
| eAlliance Final Test 1 | 81029 | Temperature Control Module | WATERS | NA | H01TCM643M | 735001271 | 2018-01-23 | 365 | 2019-01-23 | Shamsul | NA | CR/572 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 1 port 1 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E726A | WAT-RW-SKI01717 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/514 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1 | 1 | 1.) Repair Done | No | Pass | No | Pass | No | Joel | CK | NA | NA
|  | NA | Column Heater | WATERS | WAT038040 | D15CHM641R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/510 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 1 port 2 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | D1387E759A | WAT-RW-SKI01717 | 2018-04-10 | 365 | 2019-04-10 | Shamsul | NA | CR/515 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81014 | Column Heater | WATERS | WAT091727 | F00CHM531R | NA | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/558 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Test 1 Port 3 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | D1387E765A | WAT-RW-SKI01717 | 2018-04-12 | 365 | 2019-04-12 | Shamsul | NA | CR/518 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81021 | Column Heater | WATERS | WAT038040 | M00CHM910M | WAT38172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/563 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 1 Port 4 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | H1387E729A | WAT-RW-SKI01717 | 2017-07-22 | 365 | 2018-07-22 | Shamsul | NA | CR/515 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81020 | Column Heater | WATERS | WAT038040 | H01CHM352M | WAT38172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/564 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
eAlliance Final Test Station 2 | eAlliance Final Test 2 CPU | M50 | CPU | WATERS | 2489 | KKPRA3T | NA | NA | 365 | NA |  | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1 |  | 0 | NA | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP) 2.) Shortage Of TCM 1
| eAlliance Final Test 2 | 81024 | Temperature Control Module | WATERS | 2489 | J99TCM867R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  |  | 1 |  | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Tester 2 port 1 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | C1387E135M | WAT-RW-SKI01717 | 2018-04-12 | 365 | 2019-04-12 | Shamsul | NA | CR/516 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81008 | Column Heater | WATERS |  | M03CHM015R | WAT38172 | 2018-04-04 | 365 | 2019-04-04 | Shamsul | NA | CR/509 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 2 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E710A | WAT-RW-SKI01717 | 2017-07-11 | 365 | 2018-07-11 | Shamsul | NA | CR/545 | 11 Jul 18 - 11 Jul 18 Lamp expiry hours exceed 2000hrs | Replaced Lamp p/n: 402000351 | 154.4 | In-Use | TEST |  |  |  | 1
|  | 81018 | Column Heater | WATERS |  | MX6JM7863M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/563 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E731A | WAT-RW-SKI01717 | 2018-06-12 | 365 | 2019-06-12 | Shamsul | NA | CR/545 | 12 Jun 18 - 12 Jun 18 Lamp expiry hours exceed 2000hrs | Replaced Lamp p/n: 402000351 | 154.4 | In-Use | TEST |  |  |  | 1
|  | 81009 | Column Heater | WATERS |  | E01CHM203R | WAT38172 | 2018-04-03 | 365 | 2019-04-03 | Shamsul | NA | CR/508 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E727A | WAT-RW-SKI01717 | 2017-07-11 | 365 | 2018-07-11 | Shamsul | NA | CR/537 | 11 Jul 18 - 11 Jul 18 Lamp expiry hours exceed 2000hrs | Replaced Lamp p/n: 402000351 | 154.4 | In-Use | TEST |  |  |  | 1
|  | 81016 | Column Heater |  |  | MX6JM7867M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/569 |  |  |  | In-Use | TEST |  |  |  | 1
eAlliance Final Test Station 3 | eAlliance Final Tester 3 CPU | Lace 32 | CPU | KKTCN1W | CPU | M11L32097W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1 |  | 0 | NA | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP) 2.) Shortage Of TCM 1
| eAlliance Final Tester 3 | 81030 | Temperature Control Module | WATERS | WAT091728 | MX5MM4884M | 735001271 | 2018-01-26 | 365 | 2019-01-26 | Shamsul | NA | CR/573 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Tester 3 port 1 | 81013 | Column Heater | WATERS | WAT091727 | F00CHM528R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/568 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E735A | WAT-RW-SKI01717 | 2018-07-11 | 365 | 2019-07-11 | Shamsul | NA | CR/544 | 11 Jul 18 - 11 Jul 18 Lamp expiry hours exceed 2000hrs | Replaced Lamp p/n: 402000351 | 154.4 | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Tester 3 port 2 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E728A | WAT-RW-SKI01717 | 2018-06-28 | 365 | 2019-06-28 | Shamsul | NA | CR/541 | 28 Jun 18 - 28 Jun 18 Lamp expiry hours exceed 2000hrs | Replaced lamp p/n:402000351 | 154.4 | In-Use | TEST |  |  | 18 | 1 | 1 | 1 | 1.) Repair Done | No | Pass | No | Pass | No | Joel | CK | NA | NA
|  | 81012 | Column Heater | WATERS | WAT091727 | B03CHM551R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/570 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Tester 3 port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E737A | WAT-RW-SKI01717 | 2017-07-12 | 365 | 2018-07-12 | Shamsul | NA | CR/542 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81010 | Column Heater | WATERS | WAT091727 | L97CHM711M | WAT038172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/559 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Tester 3 port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E734A | WAT-RW-SKI01717 | 2017-07-21 | 365 | 2018-07-21 | Shamsul | NA | CR/536 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81019 | Column Heater | Waters | WAT091727 | D01CHM128R | WAT038172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/565 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 3 | 81026 | Temperature Control Module | WATERS | 2489 | G01TCM633M | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/574 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
eAlliance Final Test Station 4 | eAlliance Final Test 4 CPU | NA | CPU | Waters | NA | H10L32039W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1 |  | 0 | NA | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP) 2.) Shortage Of TCM 1
| eAlliance Final Test 4 | 81033 | Temperature Control Module | WATERS | WAT091728 | E97TCM153R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 4 Port 1 | 81022 | Column Heater | WATERS | WAT091727 | H01CHM335M | WAT038172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/567 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | D1387E766A | WAT-RW-SKI01717 | 2018-04-06 | 365 | 2019-04-06 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1 | 1 | 1.) Repair done on vacuum pump | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 4 Port 2 | 81017 | Column Heater | WATERS | WAT038040 | MX6KM9444M | WAT38172 | 2018-04-02 | 365 | 2019-04-02 | Shamsul | NA | CR/507 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E733A | WAT-RW-SKI01717 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/535 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1 | 1 | 1.) Repair done on vacuum pump | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 4 | 81024 | Temperature Control Module | WATERS | WAT091728 | J99TCM867R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 4 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E732A | WAT-RW-SKI01717 | 2018-06-12 | 365 | 2019-06-12 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81011 | Column Heater | WATERS | WAT038040 | K03CHM938R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/560 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 4 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E736A | WAT-RW-SKI01717 | 2017-06-21 | 365 | 2018-06-21 | Shamsul | NA | CR/537 | 21 Jun 18 - 21 Jun 18 Lamp expiry hours exceed 2000hrs | Replaced lamp p/n:402000351 | 154.4 | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81023 | Column Heater | WATERS | WAT038040 | MX5MM5993M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/504 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
eAlliance Final Test Station 5 | eAlliance Final Test 5 CPU | NA | CPU | Waters | NA | C12L32205W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  |  |  | 1 | 1.) Repaired Done On CPU | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP) 2.) Shortage Of TCM 1
| eAlliance Final Test 5 | NA | Temperature Control Module | WATERS | WAT091728 | J99TCM867R | 735001271 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 5 Port 1 | NA | Column Heater | WATERS | WAT091727 | C13CHM795R | WAT038172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/567 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | B1587E610A | WAT-RW-SKI01717 | 2018-02-20 | 365 | 2019-02-20 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 5 Port 2 | NA | Column Heater | WATERS | WAT038040 | D12CHM299R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/507 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E607R | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/535 | 07 Jun 18 - 07 Jun 18 Lamp expiry hours exceed 2000hrs High fallout of UUT Station 5 computer unable to power up | Replaced lamp p/n:402000351 | 154.4 | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 5 | NA | Temperature Control Module | WATERS | WAT091728 | E97TCM153R | 735001271 | 2018-01-26 | 365 | 2019-01-26 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 5 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E608A | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | K14CHM398R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/560 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 5 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E609A | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/537 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1 | 1 | 1.) Repair Done | No | Pass | No | Pass | No | Joel | CK | NA | NA
|  | NA | Column Heater | WATERS | WAT038040 | L14CHM438R | WAT38172 | 2017-07-20 | 365 | 2018-07-20 | Shamsul | NA | CR/504 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 6 CPU | NA | CPU | Waters | NA | A14L32340W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  |  |  | 0 | NA | No | Pass | Yes | Pass | No | Joel | CK | NA | 1.) Operating System is Obsolete (Window XP) 2.) Shortage Of TCM 1
| eAlliance Final Test 6 | NA | Temperature Control Module 2 | WATERS | WAT091728 | A18TC2481G | NA | 2018-01-29 | 365 | 2019-01-29 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1 | 0 | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
| eAlliance Final Test 6 Port 1 | NA | Column Heater | WATERS | WAT091727 | M17CHM840M | WAT038172 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/567 | 13 Jun 18 - 13 Jun 18 First filter installation | Install inline filter p/n: WAT035190 | 13.26 | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | C1887E101A | WAT-RW-SKI01717 | 2018-03-21 | 365 | 2019-03-21 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 6 Port 2 | NA | Column Heater | WATERS | WAT038040 | M17CHM834M | WAT38172 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/507 | 13 Jun 18 - 13 Jun 18 First filter installation | Install inline filter p/n: WAT035190 | 13.26 | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | C1887E103A | WAT-RW-SKI01717 | 2018-03-21 | 365 | 2019-03-21 | Shamsul | NA | CR/535 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 6 | NA | Temperature Control Module | WATERS | WAT091728 | B18TC2495G | NA | 2018-02-06 | 365 | 2019-02-06 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 6 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | C1887E106A | WAT-RW-SKI01717 | 2018-03-20 | 365 | 2019-03-20 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | M17CHM839M | WAT038040 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/560 | 13 Jun 18 - 13 Jun 18 First filter installation | Install inline filter p/n: WAT035190 | 13.26 | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 6 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | D1887E275A | WAT-RW-SKI01717 | 2018-04-09 | 365 | 2019-04-09 | Shamsul | NA | CR/537 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | M17CHM835M | WAT038040 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/504 | 13 Jun 18 - 13 Jun 18 First filter installation | Install inline filter p/n: WAT035190 | 13.26 | In-Use | TEST |  |  | 16 | 1 | 1
| Hipot | 95037 | Electrical Safety Analyzer |  |  | 5037655 | NA | 2017-11-09 | 365 | 2018-11-09 | [6 empty] | In-Use | TEST | 7 |  | 4 | 2 | 2 | NA | NA | No | Pass | No | Pass | No | Joel | CK | NA | NA
Not-In-Use
QA WORKSTATION | QA WORKSTATION | NA | CPU |  |  | C12L32205W | NA | NA | NA | NA | Shamsul | NA | NA |  |  |  |  | TEST | 270 | 80% | 7 | 1 | 1
| QA WORKSTATION | NA | Column Heater |  |  | H05SMH075E | WAT38172 | NA | NA | NA | Shamsul | NA | NA |  |  |  |  | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector |  |  | D1387E763A | WAT-RW-SKI01717 | 2018-04-11 | 365 | 2019-04-11 | Shamsul | NA | CR/541 |  |  |  |  | TEST |  |  | 18 | 1 | 1
PUMP CELLS | PUMP Cell | NA | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 6069766 | NA | 2017-12-06 | 365 | 2018-12-06 | NA | SGS | CR/10336649-2.1 |  |  |  | In-Use | TEST | 7 |  | 4 | 2 | 2
| 1500 SERIES CELL STATION | 90273 | Electronic Weighing Balance | AND | HR200 | 12308385 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 80 | 90% | 6 | 2 | 2
|  | TE 82705-2 | 1500 Series Final Test Station | WATERS | NA | TE 82705-2 | 735000477 | 2018-07-22 | 181 | 2019-01-19 | Shamsul | NA | CR/546 |  |  |  | In-Use | TEST |  |  | 1 | 3 | 3
| 515 PUMP CELL STATION | 90018 | Electronic Weighing Balance | AND | HR200 | 12303279 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 35 | 90% | 6 | 2 | 2
|  | TE 82534 | 515 Final Tester | WATERS | NA | TE 82534 | 070801 | 2018-05-22 | 181 | 2018-11-19 | Shamsul | NA | CR/547 | 22 Jun 18 - 22 Jun 18 Failed on priming(Pressure too high) | Replaced Air switch (Humphrey) p/n: M40137 | Consign | In-Use | TEST |  |  | 1 | 3 | 3
BSM/QSM FINAL TESTER | BSM/QSM Cell | 98206 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 6069766 | NA | 2017-12-06 | 365 | 2018-12-06 | NA | SGS | NA |  |  |  | In-Use | TEST | 7 |  | 4 | 2 | 2
| BSM/QSM TE82943-11 | 90606 | Electronic Weighing Balance | AND | HR-200 | 90606 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-11 | BSM Final Function Tester | WATERS | TE82943-11 | TE82943-11 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | Waters | CR/549 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84010 | QSM_TUV DETECTOR | WATERS | TE 82943-11 | M09UPT 364A | 735001828 rev B | 2018-03-02 | 365 | 2019-03-02 | Shamsul | NA | CR/506 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-12 | 90605 | Electronic Weighing Balance | AND | HR-200 | 90605 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-12 | BSM Final Function Tester | WATERS | TE82943-12 | TE82943-12 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | na | CR/550 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84007 | QSM_TUV DETECTOR | WATERS | TE 82943-12- | M09UPT 361A | 735001828 rev B | 2018-05-04 | 365 | 2019-05-04 | Shamsul | NA | CR/523 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-13 | 90608 | Electronic Weighing Balance | AND | HR-200 | 90608 | NA | 2018-09-18 | 181 | 2019-03-18 | NA | Ming Deng | NA | 05 Jul 18 - Present Error 1 shown on display (Unable to get a stable reading) | To be send for repair | NA | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-13 | BSM Final Function Tester | WATERS | TE82943-13 | TE82943-13 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/551 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84011 | QSM_TUV DETECTOR | WATERS | TE 82943-13- | K09UPT 215A | 735001828 rev B | 2018-05-03 | 365 | 2019-05-03 | Shamsul | NA | CR/525 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-14 | 90604 | Electronic Weighing Balance | AND | HR-200 | 90604 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-14 | BSM Final Function Tester | WATERS | TE82943-14 | TE82943-14 | NA | 2018-01-15 | 365 | 2019-01-15 | Shamsul | NA | CR/552 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84008 | QSM_TUV DETECTOR | WATERS | TE 82943-14- | M09UPT 362A | 735001828 rev B | 2018-04-20 | 365 | 2019-04-20 | Shamsul | NA | CR/522 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-15 | 90603 | Electronic Weighing Balance | AND | HR-200 | 90603 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-15 | BSM Final Function Tester | WATERS | TE82943-15 | TE82943-15 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/553 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84009 | QSM_TUV DETECTOR | WATERS | TE 82943-15- | M09UPT 363A | 735001828 rev B | 2018-04-21 | 365 | 2019-04-21 | Shamsul | NA | CR/521 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-16 | 90600 | Electronic Weighing Balance | AND | HR-200 | 90600 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-16 | BSM Final Function Tester | WATERS | TE82943-16 | TE82943-16 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/554 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84006 | QSM_TUV DETECTOR | WATERS | TE 82943-16- | M09UPT 360A | 735001828 rev B | 2018-05-04 | 365 | 2019-05-04 | Shamsul | NA | CR/524 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-17 | 90607 | Electronic Weighing Balance | AND | HR-200 | 90607 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA | 05 Jul 18 - Present No display on weighing balance | To be send for repair | NA | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-17 | BSM Final Function Tester | WATERS | TE82943-17 | TE82943-17 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/555 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84005 | QSM_TUV DETECTOR | WATERS | TE 82943-17 | M09UPT 359A | 735001828 rev B | 2018-04-21 | 365 | 2019-04-21 | Shamsul | NA | CR/520 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-18 | 90610 | Electronic Weighing Balance | AND | HR-200 | 90610 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-18 | BSM Final Function Tester | WATERS | TE82943-18 | TE82943-18 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/556 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | NA | QSM_TUV DETECTOR | WATERS | TE 82943-18 | M09UPT 366A | 735001828 rev B | 2018-04-21 | 365 | 2019-04-21 | Shamsul | NA | CR/519 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| Actuator Tester TE 82800-2 | TE 82800-2 | Aquity Actuator Test | WATERS | TE 82800-2 | TE 82800-2 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 90% | 2 | 2 | 2
| Actuator Tester TE 82800-3 | TE 82800-3 | Aquity Actuator Test | WATERS | TE 82800-3 | TE 82800-3 | NA | NA | NA | NA | NA | NA | NA | 07 Jun 18 - 07 Jun 18 Failed on decay test | Replaced vent valve p/n: 279000916 | $362.58 | In-Use | TEST | 15 | 90% | 2 | 2 | 2
Cleanroom | Clean Room | NA | Picoammeter | Keithley | 485 | 435110 | NA | 2018-03-28 | 365 | 2019-03-28 | NA | Keithley | NA |  |  |  | In-Use | TEST | NA | 90% | 2 | 2 | 1
| Clean Room | NA | Light Meter |  |  | NA | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | 90% | 2 | 2 | 1
| Clean Room | TE 82475-2 | Flowcell Tester | WATER | NA | TE 82475-2 | TE82475-2 | 2018-04-11 | 365 | 2019-04-10 | Shamsul | NA | CR/513 | 11 Jun 18 - 11 Jun 18 Failed on decay test | Retightened all fittings and wash the Check Valve (butler Valve) with methanol | NA | In-Use | TEST | NA | 90% | 2 | 1 | 1
| Clean Room | TE82881-2 | Optic Bench fine alignment Tester (Box Build) | Waters | 2998 | K15UPD782R | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE82881-2 | PDA Fine Alignment CPU |  |  | D08L32121W / C08BLM058C | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE82880-2 | PDA ROUGH ALIGNMENT CPU | Waters |  | L07L32070W | NA | NA | NA | NA |  | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE# 82874 | TUV lamp supply | Waters | 2998 | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE # 82965 | Filter Wheel Alignment | Waters | TE 82965 | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE9990293134 | Mercury lamp power Supply | Waters |  | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | Test | NA | 90% |  | 1
| Clean Room | TE829002998 | Flowcell Tester CPU | Waters |  | NA | NA | NA | NA | NA |  | NA | NA |  |  |  | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82918-2 TE# 999029981 | 2998 fine alignment(box build) |  |  | E08998580N | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82918-2 | 2998 fine alignment CPU |  |  | B08L32065W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82917-1 | 2998 rough alignment CPU | Waters |  | B08L32068W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE186002998 | Flowcell tester(2998) | Waters |  | G08998943N | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82951-1 | PDA Assembly CPU |  |  | J06L32014W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82868-2 | PDA Aperture Alignment System | Waters |  | K07L32122W PN: 668000351 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | PROCESS | NA | 90% | 1 | 3 | 3
| Clean Room | 95515 | Starlite 150 |  |  | RSL1002609 | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE83096-2 | PDA,2998 Coarse and fine Align |  |  | C17L32063W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
Eng's Area | Engineering Lab | Pressure calibrator | Pressure Calibrator | AMETEK | APC10KGINDG | 9093205 | NA | 2017-09-13 | 365 | 2018-09-12 | NA | SGS | CR/10311178-1.1 |  |  |  | In-Use | TEST | As per usage on evaluation only | NA | 1 | 1 | 1
| Engineering Lab | FLUKE179 | Multimeter | FLUKE | 179 | 90240092 | NA | 2018-01-31 | 365 | 2019-01-30 | NA | External | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | High Volrage Probe | High Voltage Probe | FLUKE | 80K-6 | 89200011 | NA | 2017-09-13 | 365 | 2018-09-12 | NA | SGS | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | Picometer | Picometer |  |  | 310788 | NA | Calibrate only 1 month before the other one due date ( Cost Management) |  |  |  | Keithly | NA |  |  |  | In-Use |  |  | NA
| Engineering Lab | Clamp Meter | Clamp Meter | KYORITSU | 2033 | 66386 | NA | 2017-08-05 | 365 | 2018-08-04 | NA | SGS | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | Fluke 51 II | Digital Thermometer with probe | Fluke | 51 II | 89280141 | NA | 2017-08-10 | 365 | 2018-08-10 | NA | SGS | FMS/323 |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | NA | PDA detector |  |  | B08UPD563M | NA | NA | NA | NA | [6 empty] | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | FLUKE110 | Multimeter |  |  | 85730619 | NA | 2018-01-05 | 365 | 2019-01-05 | [11 empty] | 1
| Engineering Lab | 90406 | Liquilaz | Particle Measuring Systems | E20P | 51053 | NA | 2017-08-17 | 365 | 2018-08-16 | NA | Particle Measuring Systems | 110130306 |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | 95037 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 5037655 | NA | 2018-01-09 | 365 | 2019-01-09 | [11 empty] | 1
| Engineering Lab | 92207 | Electronic Weighing Balance |  |  | 92207 | NA | 2018-03-08 | 182 | 2018-09-06
| Engineering Lab | 90282 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 1140660 | NA | 2013-06-21 | 365 | 2014-06-21 | NA | SGS | NA | Inactive-Remove from line (RE/094/13) Date 03/12/13. |  |  | In-Use | TEST | 7 | NA | 4 | 2 | 2
Detector | DOWNLOADING AND PROGRAMMING | TE 82964-1 | Download and programme(2489 and TUV) | Waters | NA | TE 82964-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 90% | 3 | 2 | 3
In-Use
|  | TE82972-1 | Download and programme(2998, PDA and elamda) | Waters | NA | TE82972-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 |  |  | 2 | 3
In-Use
|  | TE 82969-1 | Calibration and Linearity test for PDA family |  |  | TE 82969-1 | NA | NA | NA | NA | [11 empty] | 1
| 2489/TUV | TE 82903-10 | Final Tester | Waters | NA | TE 82903-10 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 150 | 90% | 4 | 2 | 2
In-Use
|  | TE 82903-8 | Final Tester | Waters | NA | TE 82903-8 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  |  | 4 | 2 | 2
In-Use
| 2489/TUV | TE 82903-9 | Final Tester | Waters | NA | TE 82903-9 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  | 90% | 4 | 2 | 2
In-Use
|  | TE 82903-6 | Final Tester | Waters | NA | TE 82903-6 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  |  | 4 | 2 | 2
In-Use
| 2998, PDA eLambda | TE 82882-8 | Final Tester | Waters | NA | TE 82882-8 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 150 | 90% | 16 | 2 | 1
|  | TE 82882-9 | Final Tester | Waters | NA | TE 82882-9 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 1 | 1
|  | TE 82882-10 | Final Tester | Waters | NA | TE 82882-10 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 2 | 1
|  | TE 82882-11 | Final Tester | Waters | NA | TE 82882-11 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 2 | 1
|  | TE 82943 - 13 | 2998 Final Tester |  |  | TE 82943 - 13 | NA | NA | NA | NA | [11 empty] | 1
|  | TE 82943 - 14 | 2998 Final Tester |  |  | TE 82943 - 14 |  | NA | NA | NA | [11 empty] | 1
| Printer | NA | HP Laserjet P3005n |  |  | SN: JPKGJ01792 | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | NA | NA | NA | NA
|  | NA | Brother |  |  | E69530C4N142942 |  | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | NA | NA | NA | NA
| Hipot | NA | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 08280865 | NA | 2018-10-13 | 365 | 2019-10-13 | NA | SGS | CR/10339085/1.1 |  |  |  |  | TEST | 7 | NA | 4 | 2 | 2
Production Floor | Label Printing Area | WATERS | WATERSISNL1 |  |  | 54177579 | NA | NA | NA | NA | [6 empty] | In-Use | PROCESS | NA | NA | 1 | 3 | 3
Production Floor | Label Printing Area | WATERS | WATERSPNL1 |  |  | 54177587 | NA | NA | NA | NA | [6 empty] | In-Use | PROCESS | NA | NA | 1 | 3 | 3
| TE82943-14 | NA | Electronic Weighing Balance | AND | HR-200 | 12336741 | NA | 2017-12-08 | 182 | 2018-06-08 | NA | Ming Deng | NA |  |  |  | In-Use | TEST
Formulas:
Calibration Due Date = [Calibration Date]+[Calibration Cycle (Months/ Days)]
Calibration Due Date = [Calibration Date]+365
Note on P4: skwjdedu: Period 1st Jan 2011 to 31st Dec 2011
Note on R4: skwjdedu: Green - Fixture working Red - Fixture not working
Note on W4: user: 3 = unique tester 2 = multiple tester but will affect output when down 1 = standard tester
Note on X4: user: 3 = High Risk, production volume high, will be affected once tester down 2 = Medium Risk, tester down will not affect output for max 2 weeks 1 = Low Risk, tester down will not affect output when down for more than 2 weeks.
Note on V9: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V11: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on G13: SGSJDEDU: EQID 0216
Note on V13: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V19: skwjdedu: 1 down
Note on V25: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V26: skwjdedu: 12 working, 4 spare
Note on V27: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V30: skwjdedu: 12 working, 4 spare
Note on V32: skwjdedu: 12 working, 4 spare
Note on G43: SGSMSBAf: original CPU KKTCN1W
Note on V45: skwjdedu: 12 working, 4 spare
Note on V46: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V47: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V48: skwjdedu: 12 working, 4 spare
Note on V49: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V50: skwjdedu: 12 working, 4 spare
Note on V51: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V52: skwjdedu: 12 working, 4 spare
Note on V56: skwjdedu: 12 working, 4 spare
Note on V57: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V58: skwjdedu: 12 working, 4 spare
Note on V59: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V60: skwjdedu: 4 working, 3 spare
Note on V61: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V62: skwjdedu: 12 working, 4 spare
Note on V63: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V64: skwjdedu: 12 working, 4 spare
Note on V67: skwjdedu: 12 working, 4 spare
Note on V68: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V69: skwjdedu: 12 working, 4 spare
Note on V70: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V71: skwjdedu: 4 working, 3 spare
Note on V72: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V73: skwjdedu: 12 working, 4 spare
Note on V74: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V75: skwjdedu: 12 working, 4 spare
Note on V78: skwjdedu: 12 working, 4 spare
Note on V79: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V80: skwjdedu: 12 working, 4 spare
Note on V81: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V82: skwjdedu: 4 working, 3 spare
Note on V83: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V84: skwjdedu: 12 working, 4 spare
Note on V85: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V86: skwjdedu: 12 working, 4 spare
Note on V87: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V90: skwjdedu: 12 working, 4 spare
Note on V91: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V93: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V94: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V96: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V99: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V102: skwjdedu: BSM no need detectors
Note on V105: skwjdedu: BSM no need detectors
Note on V108: skwjdedu: BSM no need detectors
Note on V111: skwjdedu: BSM no need detectors
Note on V114: skwjdedu: BSM no need detectors
Note on V117: skwjdedu: BSM no need detectors
Note on V120: skwjdedu: BSM no need detectors
Note on V123: skwjdedu: BSM no need detectors
Note on V127: skwjdedu: 1 - clean room, 1 - eng area but not calibrated. Express calibration - 2 weeks, normal - 1 mth.
Note on V129: skwjdedu: 1 - clean room, 1 spare
Note on C148: SGSJDEDU: EQID 0279
Note on C151: SGSJDEDU: EQID 0196
Note on V183: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors

## Column Inventory
P/N 1860002699 ( Column )
Date | Qty In | Qty Out | Balance | Remarks
2016-06-15 | 20 |  |  | Received on 15 Jun 16
2016-06-28 |  | 1 | 19 | Replaced column from S/N:010932104111 to 01093507912602 on FT 4 Port 2 (CHM 81017)
2016-09-23 |  | 1 | 18 | Replaced column from s/n:010932104128 to S/n:01103524612501 on FT 1 Port 3 (CHM 81020)
2016-11-23 |  | 1 | 17 | Replaced column from s/n:010932104115 to s/n: 01093524015005 on FT 2 Port 2 (CHM 81018)
2016-12-15 |  | 1 | 16 | Replaced column from s/n:010933275130 to s/n: 01093524014901 on FT 1 Port 1 (CHM S/n:D15CHM641R)
2016-12-16 |  | 1 | 15 | Replaced column from s/n:010933275126 to s/n: 01093524015004 on FT 5 Port 4 (CHM S/n:L14CHM438R)
2017-04-10 |  | 1 | 14 | Replaced Column from s/n: 01073013010321 to s/n:01093524014910 FT 1 port 3 (CHM 81021)
2017-07-07 |  | 1 | 13 | Replaced from s/n:010135074109 to s/n: 01093524015001 FT 4 port 3 (CHM 81011)
2017-10-19 |  | 1 | 12 | Replaced column from s/n:010933275126 to s/n:01093524014909 FT 2 port 3 (CHM 81009)
2017-10-25 |  | 1 | 11 | Replaced column from s/n:01093507912601 to s/n:01093524015003 FT 2 port 1 (CHM 81008)
2017-11-21 |  | 1 | 10 | Replaced column p/n:186002699 from s/n:010932104120 to s/n:01103524612507 FT 2 port 4 (CHM 81016)
2018-01-10 |  | 1 | 9 | Replaced column p/n:186002699 from s/n:010730188123 to s/n:01093524015006 FT 1 port 2 (CHM 81014)