## Waters Equipment Loc (hidden)
Person In Charge : Joel
Locations | STATIONS | Customer ID | Equipment | Maker | Model | Serial No. | Calibration Procedure | Calibration Date | Calibration Cycle (Months/ Days) | Calibration Due Date | Calibration By Internal | Calibration By External | Calibration Report No | Reason of Failure and Length of downtime | Replacement Item / Spare Parts | Replacement Cost (US$) | STATUS | TEST/PROCESS | Cycle Time (mins) | % YEILD @ FINAL TEST STATION ONLY | Number of common testers | Criticalilty | Risk
eAlliance Sub Assembly Stations | 2690 eAlliance Front Panel Cell | TE82546 | 2690 eAlliance Front Panel Test Station | Waters | eAlliance | TE82546 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 98% | 1 | 3 | 3
| Pump Head Tester | TE 82519 | 2690 Pump Head Tester | WATERS | eAlliance | TE 82519 | 069004 | 2018-06-01 | 183 | 2018-12-01 | Shamsul | NA | CR/530 | 16 Jul 18 -16 Jul 18 Fail to initialize | Replaced power cable p/n: WAT270617-01 due to wear and tear and solenoid valve p/n: | Consigned | In-Use | TEST | 10 | 90% | 1 | 3 | 3
| eAlliance Pump Tray 1 | TE82526-1 | 2690 Pump Tray Tester | Waters | eAlliance | TE82526-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 85% | 3 | 2 | 3
|  | 90891 | Electronic Weighing Balance | WATERS | HR200 | 12305757 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | MDF120074-3 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2
| eAlliance Pump Tray 2 | TE82526 | 2690 Pump Tray Tester | Waters | eAlliance | TE82526 | NA | NA | NA | NA | NA | NA | NA | 14 Jun 18 - Present Failed on ripple test | N/A | Consigned | In-Use | TEST | 35 | 85% | 3 | 2 | 2
|  | 95206 | Electronic Weighing Balance | AND | HR200 | 12316922 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | CR/MDF120074-1 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2
| eAlliance Pump Tray 3 | TE82526-4 | 2690 Pump Tray Tester |  |  | TE82526-4 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 85% | 3 | 2 | 2
|  | 90038 | Electronic Weighing Balance |  |  | 12200404 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | CR/MDF140265-1 |  |  |  | In-Use | TEST |  |  | 6 | 2 | 2
| eAlliance HPMV Cell | TE 82625-1 | HPMV Tester | WATERS | NA | TE 82625-1 | 735000288 | 2018-01-19 | 181 | 2018-07-19 | Shamsul | NA | CR/561 | 01 Aug 18 - 01 Aug 18 Failed on ramp and decay test | Replaced fitting plug p/n:WAT025566 | 5.79 | In-Use | TEST | 10 | 90% | 1 | 3 | 3
| 2690/2790 Syringe Drive Cell | TE82656 | 2690/2790 Syringe Drive Tester | Waters | eAlliance | TE82656 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 10 | 95% | 1 | 3 | 2
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-1 | Carousel Carrier Drive Alignment Station | Waters | NA | TE82520-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 90% | 1 | 3 | 3
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-2 | Carousel Carrier Drive Alignment Station | Waters | NA | TE82520-2 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 35 | 90% | 1 | 3 | 3
| Fluidics Tray Tester | TE82523 | Fluidics Tray Tester |  |  | TE82523 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 5 | 90% | 1 | 3 | 3
| Injector Seal Pack Cell/Engineering Lab | TE 82503-04 | Injector Seal Pack Tester | Waters | NA | TE 82503-04 | WAT043816 | 2018-07-17 | 184 | 2019-01-17 | Shamsul | NA | CR/562 |  |  |  | In-Use | TEST | 15 | 85% | 2 | 3 | 3
|  | TE 82503-05 | Injector Seal Pack Tester |  | NA | TE 82503-05 | WAT043816 | 2018-03-02 | 184 | 2018-09-02 | Shamsul | NA | CR/562 |  |  |  | In-Use | TEST | 15 | 85% | 2 | 3 | 3
| eAlliance Heater Chiller Cell | TE82620 | Heater Chiller Tester | Waters | NA | TE 82620 | 735000206 | 2013-03-15 | 365 | 2014-03-15 | Shamsul | NA | CR/454 | Inactive since 14 Mar 14 (ref no: RE/040/14) |  |  | Not-In-Use | TEST | 45 | 95% | 1 | 3 | 3
|  | NA | Multimeter | FLUKE | 110 | 85730619 | NA | 2018-01-05 | 365 | 2019-01-05 | NA | SGS | NA | [7 empty] | 1 | 3 | 3
|  | NA | RS 232 C Precision DC Module | Keithley | KNM-RTD32 | 000693213 | NA | 2010-11-19 | 60 Mths | 2015-11-19 | NA | ESSCO | CR/422086 | Inactive since 14 Mar 14 (ref no: RE/130/15) | [6 empty] | 1 | 3 | 3
eAlliance Final Test | Alliance Programming Station | NA | Autoloader 2 |  |  | Lace 32 B08L32122N | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 25 | 95% |  | 1
| Rear Panel Station | TE82573-1 | NA |  |  | Lace 32 C08L32093W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 7 | 95% |  | 1
| eAlliance Sample Station | 90913 | Electronic Weighing Balance | WATERS | HR200 | 12315717 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 5 | 95% | 6 | 2 | 2
eAlliance Final Test Station 1 | eAlliance Final Test 1 CPU | IBM | CPU | KKPRA0M | CPU | C12L32255W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% | 5 | 1 | 1
| eAlliance Final Test 1 | 81029 | Temperature Control Module | WATERS | NA | H01TCM643M | 735001271 | 2018-01-23 | 365 | 2019-01-23 | Shamsul | NA | CR/572 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 1 port 1 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E726A | WAT-RW-SKI01717 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/514 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | D15CHM641R | WAT38172 | 2018-07-20 | 365 | 2019-07-20 | Shamsul | NA | CR/510 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 1 port 2 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | D1387E759A | WAT-RW-SKI01717 | 2018-04-10 | 365 | 2019-04-10 | Shamsul | NA | CR/515 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81014 | Column Heater | WATERS | WAT091727 | F00CHM531R | NA | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/558 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Test 1 Port 3 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | D1387E765A | WAT-RW-SKI01717 | 2018-04-12 | 365 | 2019-04-12 | Shamsul | NA | CR/518 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81021 | Column Heater | WATERS | WAT038040 | M00CHM910M | WAT38172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/563 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 1 Port 4 | NA | Dual Wavelength Absorbance Detector | Waters | 2489 | H1387E729A | WAT-RW-SKI01717 | 2018-07-19 | 365 | 2019-07-19 | Shamsul | NA | CR/515 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81020 | Column Heater | WATERS | WAT038040 | H01CHM352M | WAT38172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/564 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
eAlliance Final Test Station 2 | eAlliance Final Test 2 CPU | M50 | CPU | WATERS | 2489 | KKPRA3T | NA | NA | 365 | NA |  | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1
| eAlliance Final Test 2 | 81024 | Temperature Control Module | WATERS | 2489 | J99TCM867R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 1 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | C1387E135M | WAT-RW-SKI01717 | 2018-04-12 | 365 | 2019-04-12 | Shamsul | NA | CR/516 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81008 | Column Heater | WATERS |  | M03CHM015R | WAT38172 | 2018-04-04 | 365 | 2019-04-04 | Shamsul | NA | CR/509 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 2 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E710A | WAT-RW-SKI01717 | 2018-07-11 | 365 | 2019-07-11 | Shamsul | NA | CR/545 | 11 Jul 18 - 11 Jul 18 Expiry of lamp Hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  |  | 1
|  | 81018 | Column Heater | WATERS |  | MX6JM7863M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/563 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E731A | WAT-RW-SKI01717 | 2018-06-12 | 365 | 2019-06-12 | Shamsul | NA | CR/545 |  |  |  | In-Use | TEST |  |  |  | 1
|  | 81009 | Column Heater | WATERS |  | E01CHM203R | WAT38172 | 2018-04-03 | 365 | 2019-04-03 | Shamsul | NA | CR/508 |  |  |  | In-Use | TEST |  |  |  | 1
| eAlliance Final Tester 2 port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E727A | WAT-RW-SKI01717 | 2018-07-11 | 365 | 2019-07-11 | Shamsul | NA | CR/537 | 11 Jul 18 - 11 Jul 18 Expiry of lamp Hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  |  | 1
|  | 81016 | Column Heater |  |  | MX6JM7867M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/569 |  |  |  | In-Use | TEST |  |  |  | 1
eAlliance Final Test Station 3 | eAlliance Final Tester 3 | Lace 32 | CPU | KKTCN1W | CPU | M11L32097W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1
| eAlliance Final Tester 3 | 81030 | Temperature Control Module | WATERS | WAT091728 | MX5MM4884M | 735001271 | 2018-01-26 | 365 | 2019-01-26 | Shamsul | NA | CR/573 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Tester 3 port 1 | 81013 | Column Heater | WATERS | WAT091727 | F00CHM528R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/568 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E735A | WAT-RW-SKI01717 | 2018-07-11 | 365 | 2019-07-11 | Shamsul | NA | CR/544 | 11 Jul 18 - 11 Jul 18 Expiry of lamp Hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Tester 3 port 2 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E728A | WAT-RW-SKI01717 | 2018-06-28 | 365 | 2019-06-28 | Shamsul | NA | CR/541 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81012 | Column Heater | WATERS | WAT091727 | B03CHM551R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/570 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Tester 3 port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E737A | WAT-RW-SKI01717 | 2018-07-12 | 365 | 2019-07-12 | Shamsul | NA | CR/542 | 12 Jul 18 - 12 Jul 18 Expiry of lamp Hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81010 | Column Heater | WATERS | WAT091727 | L97CHM711M | WAT038172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/559 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Tester 3 port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | H1387E734A | WAT-RW-SKI01717 | 2018-07-19 | 365 | 2019-07-19 | Shamsul | NA | CR/536 | 19 Jul 18 - 19 Jul 18 Expiry of lamp Hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81019 | Column Heater | Waters | WAT091727 | D01CHM128R | WAT038172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/565 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 3 | 81026 | Temperature Control Module | WATERS | 2489 | G01TCM633M | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/574 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
eAlliance Final Test Station 4 | eAlliance Final Test 4 | NA | CPU | Waters | NA | H10L32039W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80% |  | 1
| eAlliance Final Test 4 | 81033 | Temperature Control Module | WATERS | WAT091728 | E97TCM153R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 4 Port 1 | 81022 | Column Heater | WATERS | WAT091727 | H01CHM335M | WAT038172 | 2018-02-13 | 365 | 2019-02-13 | Shamsul | NA | CR/567 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | D1387E766A | WAT-RW-SKI01717 | 2018-04-06 | 365 | 2019-04-06 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 4 Port 2 | 81017 | Column Heater | WATERS | WAT038040 | MX6KM9444M | WAT38172 | 2018-04-02 | 365 | 2019-04-02 | Shamsul | NA | CR/507 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E733A | WAT-RW-SKI01717 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/535 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 4 | 81024 | Temperature Control Module | WATERS | WAT091728 | J99TCM867R | 735001271 | 2018-01-24 | 365 | 2019-01-24 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 4 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E732A | WAT-RW-SKI01717 | 2018-06-12 | 365 | 2019-06-12 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81011 | Column Heater | WATERS | WAT038040 | K03CHM938R | WAT38172 | 2018-02-14 | 365 | 2019-02-14 | Shamsul | NA | CR/560 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 4 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | H1387E736A | WAT-RW-SKI01717 | 2018-06-21 | 365 | 2019-06-21 | Shamsul | NA | CR/537 | 21 Jun 18 - 21 Jun 18 Expiry of lamp Hours | Replaced lamp p/n:402000351 | 155.94 | In-Use | TEST |  |  | 18 | 1 | 1
|  | 81023 | Column Heater | WATERS | WAT038040 | MX5MM5993M | WAT38172 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/504 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
eAlliance Final Test Station 5 | eAlliance Final Test 5 | NA | CPU | Waters | NA | C12L32205W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80%
| eAlliance Final Test 5 | NA | Temperature Control Module | WATERS | WAT091728 | J99TCM867R | 735001271 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 5 Port 1 | NA | Column Heater | WATERS | WAT091727 | C13CHM795R | WAT038172 | 2018-07-20 | 365 | 2019-07-20 | Shamsul | NA | CR/567 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | B1587E610A | WAT-RW-SKI01717 | 2018-02-20 | 365 | 2019-02-20 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 5 Port 2 | NA | Column Heater | WATERS | WAT038040 | D12CHM299R | WAT38172 | 2018-07-20 | 365 | 2019-07-20 | Shamsul | NA | CR/507 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E607R | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/535 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 5 | NA | Temperature Control Module | WATERS | WAT091728 | E97TCM153R | 735001271 | 2018-01-26 | 365 | 2019-01-26 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 5 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E608A | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | K14CHM398R | WAT38172 | 2018-07-20 | 365 | 2019-07-20 | Shamsul | NA | CR/560 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 5 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | B1587E609A | WAT-RW-SKI01717 | 2018-02-15 | 365 | 2019-02-15 | Shamsul | NA | CR/537 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | L14CHM438R | WAT38172 | 2018-07-20 | 365 | 2019-07-20 | Shamsul | NA | CR/504 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 6 | NA | CPU | Waters | NA | A14L32340W | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 299 | 80%
| eAlliance Final Test 6 | NA | Temperature Control Module 2 | WATERS | WAT091728 | A18TC2481G | NA | 2018-01-29 | 365 | 2019-01-29 | Shamsul | NA | CR/571 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 6 Port 1 | NA | Column Heater | WATERS | WAT091727 | M17CHM840M | WAT038172 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/567 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2489 | C1887E101A | WAT-RW-SKI01717 | 2018-03-21 | 365 | 2019-03-21 | Shamsul | NA | CR/534 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 6 Port 2 | NA | Column Heater | WATERS | WAT038040 | M17CHM834M | WAT38172 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/507 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | C1887E103A | WAT-RW-SKI01717 | 2018-03-21 | 365 | 2019-03-21 | Shamsul | NA | CR/535 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
| eAlliance Final Test 6 | NA | Temperature Control Module | WATERS | WAT091728 | B18TC2495G | NA | 2018-02-06 | 365 | 2019-02-06 | Shamsul | NA | CR/575 |  |  |  | In-Use | TEST |  |  | 7 | 1 | 1
| eAlliance Final Test 6 Port 3 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | C1887E106A | WAT-RW-SKI01717 | 2018-03-20 | 365 | 2019-03-20 | Shamsul | NA | CR/543 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | M17CHM839M | WAT038040 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/560 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| eAlliance Final Test 6 Port 4 | NA | Dual Wavelength Absorbance Detector | WATERS | 2487 | D1887E275A | WAT-RW-SKI01717 | 2018-04-09 | 365 | 2019-04-09 | Shamsul | NA | CR/537 |  |  |  | In-Use | TEST |  |  | 18 | 1 | 1
|  | NA | Column Heater | WATERS | WAT038040 | M17CHM835M | WAT038040 | 2018-06-13 | 365 | 2019-06-13 | Shamsul | NA | CR/504 |  |  |  | In-Use | TEST |  |  | 16 | 1 | 1
| Hipot | 95037 | Electrical Safety Analyzer |  |  | 5037655 | NA | 2017-11-09 | 365 | 2018-11-09 | [6 empty] | In-Use | TEST | 7 |  | 4 | 2 | 2
Not-In-Use
QA WORKSTATION | QA WORKSTATION | NA | CPU |  |  | C12L32205W | NA | NA | NA | NA | Shamsul | NA | NA |  |  |  |  | TEST | 270 | 80% | 7 | 1 | 1
| QA WORKSTATION | NA | Column Heater |  |  | H05SMH075E | WAT38172 | NA | NA | NA | Shamsul | NA | NA |  |  |  |  | TEST |  |  | 16 | 1 | 1
|  | NA | Dual Wavelength Absorbance Detector |  |  | D1387E763A | WAT-RW-SKI01717 | 2018-04-11 | 365 | 2019-04-11 | Shamsul | NA | CR/541 |  |  |  |  | TEST |  |  | 18 | 1 | 1
PUMP CELLS | PUMP Cell | NA | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 6069766 | NA | 2017-12-06 | 365 | 2018-12-06 | NA | SGS | CR/10336649-2.1 |  |  |  | In-Use | TEST | 7 |  | 4 | 2 | 2
| 1500 SERIES CELL STATION | 90273 | Electronic Weighing Balance | AND | HR200 | 12308385 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 80 | 90% | 6 | 2 | 2
|  | TE 82705-2 | 1500 Series Final Test Station | WATERS | NA | TE 82705-2 | 735000477 | 2018-05-22 | 181 | 2018-11-19 | Shamsul | NA | CR/546 |  |  |  | In-Use | TEST |  |  | 1 | 3 | 3
| 515 PUMP CELL STATION | 90018 | Electronic Weighing Balance | AND | HR200 | 12303279 | NA | 2018-03-08 | 181 | 2018-09-05 | NA | SGS | NA |  |  |  | In-Use | TEST | 35 | 90% | 6 | 2 | 2
|  | TE 82534 | 515 Final Tester | WATERS | NA | TE 82534 | 070801 | 2018-05-22 | 181 | 2018-11-19 | Shamsul | NA | CR/547 | 22 Jun 18 - 22 Jun 18 Failed on HPLC priming | Replaced solenoid valve (Humphrey) p/n:M401 37 due to unable to activate Flow valve 3 | Consigned | In-Use | TEST |  |  | 1 | 3 | 3
BSM/QSM FINAL TESTER | BSM/QSM Cell | 98206 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 6069766 | NA | 2017-12-06 | 365 | 2018-12-06 | NA | SGS | NA |  |  |  | In-Use | TEST | 7 |  | 4 | 2 | 2
| BSM/QSM TE82943-11 | 90606 | Electronic Weighing Balance | AND | HR-200 | 90606 | NA | 2018-01-15 | 182 | 2018-07-16 | NA | Ming Deng | NA | To send for calibration | To send for calibration |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-11 | BSM Final Function Tester | WATERS | TE82943-11 | TE82943-11 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | Waters | CR/549 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84010 | QSM_TUV DETECTOR | WATERS | TE 82943-11 | M09UPT 364A | 735001828 rev B | 2018-03-02 | 365 | 2019-03-02 | Shamsul | NA | CR/506 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-12 | 90605 | Electronic Weighing Balance | AND | HR-200 | 90605 | NA | 2018-07-10 | 182 | 2019-01-08 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-12 | BSM Final Function Tester | WATERS | TE82943-12 | TE82943-12 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | na | CR/550 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84007 | QSM_TUV DETECTOR | WATERS | TE 82943-12- | M09UPT 361A | 735001828 rev B | 2018-05-04 | 365 | 2019-05-04 | Shamsul | NA | CR/523 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-13 | 90608 | Electronic Weighing Balance | AND | HR-200 | 12326184 | NA | 2018-07-10 | 181 | 2019-01-07 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-13 | BSM Final Function Tester | WATERS | TE82943-13 | TE82943-13 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/551 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84011 | QSM_TUV DETECTOR | WATERS | TE 82943-13- | K09UPT 215A | 735001828 rev B | 2018-05-03 | 365 | 2019-05-03 | Shamsul | NA | CR/525 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-14 | NA | Electronic Weighing Balance | AND | HR-200 | 12336741 | NA | 2018-05-15 | 182 | 2018-11-13 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-14 | BSM Final Function Tester | WATERS | TE82943-14 | TE82943-14 | NA | 2018-01-15 | 365 | 2019-01-15 | Shamsul | NA | CR/552 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84008 | QSM_TUV DETECTOR | WATERS | TE 82943-14- | M09UPT 362A | 735001828 rev B | 2018-04-20 | 365 | 2019-04-20 | Shamsul | NA | CR/522 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-15 | 90603 | Electronic Weighing Balance | AND | HR-200 | 90603 | NA | 2018-07-10 | 182 | 2019-01-08 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-15 | BSM Final Function Tester | WATERS | TE82943-15 | TE82943-15 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/553 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84009 | QSM_TUV DETECTOR | WATERS | TE 82943-15- | M09UPT 363A | 735001828 rev B | 2018-04-21 | 365 | 2019-04-21 | Shamsul | NA | CR/521 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-16 | 90600 | Electronic Weighing Balance | AND | HR-200 | 90600 | NA | 2018-07-10 | 182 | 2019-01-08 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-16 | BSM Final Function Tester | WATERS | TE82943-16 | TE82943-16 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/554 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84006 | QSM_TUV DETECTOR | WATERS | TE 82943-16- | M09UPT 360A | 735001828 rev B | 2018-05-04 | 365 | 2019-05-04 | Shamsul | NA | CR/524 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-17 | 90607 | Electronic Weighing Balance | AND | HR-200 | 90607 | NA | 2018-07-10 | 182 | 2019-01-08 | NA | Ming Deng | NA | To send for calibration | To send for calibration |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-17 | BSM Final Function Tester | WATERS | TE82943-17 | TE82943-17 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/555 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | 84005 | QSM_TUV DETECTOR | WATERS | TE 82943-17 | M09UPT 359A | 735001828 rev B | 2018-04-21 | 365 | 2019-04-21 | Shamsul | NA | CR/520 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| BSM/QSM TE82943-18 | 90610 | Electronic Weighing Balance | AND | HR-200 | 90610 | NA | 2018-07-10 | 182 | 2019-01-08 | NA | Ming Deng | NA |  |  |  | In-Use | TEST | 180 | 90% | 8 | 2 | 2
|  | TE82943-18 | BSM Final Function Tester | WATERS | TE82943-18 | TE82943-18 | NA | 2018-01-05 | 365 | 2019-01-05 | Shamsul | NA | CR/556 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
|  | NA | QSM_TUV DETECTOR | WATERS | TE 82943-18 | M09UPT 366A | 735001828 rev B | 2018-04-21 | 365 | 2019-04-21 | Shamsul | NA | CR/519 |  |  |  | In-Use | TEST |  |  | 8 | 2 | 3
| Actuator Tester TE 82800-2 | TE 82800-2 | Aquity Actuator Test | WATERS | TE 82800-2 | TE 82800-2 | NA | NA | NA | NA | NA | NA | NA | 12 Jul 18 - 12 Jul 18 Failed to initialize | Rework on power cable | Consigned | In-Use | TEST | 15 | 90% | 2 | 2 | 2
| Actuator Tester TE 82800-3 | TE 82800-3 | Aquity Actuator Test | WATERS | TE 82800-3 | TE 82800-3 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 90% | 2 | 2 | 2
Cleanroom | Clean Room | NA | Picoammeter | Keithley | 485 | 435110 | NA | 2018-03-28 | 365 | 2019-03-28 | NA | Keithley | NA | 12 Jul 18 - 12 Jul 18 Fail to read energy | Replaced Oscilloscope cable | Consigned | In-Use | TEST | NA | 90% | 2 | 2 | 1
| Clean Room | NA | Light Meter |  |  | NA | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | 90% | 2 | 2 | 1
| Clean Room | TE 82475-2 | Flowcell Tester | WATER | NA | TE 82475-2 | TE82475-2 | 2018-04-11 | 365 | 2019-04-10 | Shamsul | NA | CR/513 |  |  |  | In-Use | TEST | NA | 90% | 2 | 1 | 1
| Clean Room | TE82881-2 | Optic Bench fine alignment Tester (Box Build) | Waters | 2998 | K15UPD782R | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE82881-2 | PDA Fine Alignment CPU |  |  | D08L32121W / C08BLM058C | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE82880-2 | PDA ROUGH ALIGNMENT CPU | Waters |  | L07L32070W | NA | NA | NA | NA |  | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE# 82874 | TUV lamp supply | Waters | 2998 | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE # 82965 | Filter Wheel Alignment | Waters | TE 82965 | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | 90% | 1 | 3 | 3
| Clean Room | TE9990293134 | Mercury lamp power Supply | Waters |  | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | Test | NA | 90% |  | 1
| Clean Room | TE829002998 | Flowcell Tester CPU | Waters |  | NA | NA | NA | NA | NA |  | NA | NA |  |  |  | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82918-2 TE# 999029981 | 2998 fine alignment(box build) |  |  | E08998580N | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82918-2 | 2998 fine alignment CPU |  |  | B08L32065W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82917-1 | 2998 rough alignment CPU | Waters |  | B08L32068W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE186002998 | Flowcell tester(2998) | Waters |  | G08998943N | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82951-1 | PDA Assembly CPU |  |  | J06L32014W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE82868-2 | PDA Aperture Alignment System | Waters |  | K07L32122W PN: 668000351 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | PROCESS | NA | 90% | 1 | 3 | 3
| Clean Room | 95515 | Starlite 150 |  |  | RSL1002609 | NA | 2018-04-25 | 365 | 2018-04-24 | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
| Clean Room | TE83096-2 | PDA,2998 Coarse and fine Align |  |  | C17L32063W | NA | NA | NA | NA | [6 empty] | In-Use | Test | NA | 90% | 1 | 3 | 3
Eng's Area | Engineering Lab | Pressure calibrator | Pressure Calibrator | AMETEK | APC10KGINDG | 9093205 | NA | 2017-09-13 | 365 | 2018-09-12 | NA | SGS | CR/10311178-1.1 |  |  |  | In-Use | TEST | As per usage on evaluation only | NA | 1 | 1 | 1
| Engineering Lab | FLUKE179 | Multimeter | FLUKE | 179 | 90240092 | NA | 2018-01-31 | 365 | 2019-01-30 | NA | External | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | High Volrage Probe | High Voltage Probe | FLUKE | 80K-6 | 89200011 | NA | 2017-09-13 | 365 | 2018-09-12 | NA | SGS | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | Picometer | Picometer |  |  | 310788 | NA | Calibrate only 1 month before the other one due date ( Cost Management) |  |  |  | Keithly | NA |  |  |  | In-Use |  |  | NA
| Engineering Lab | Clamp Meter | Clamp Meter | KYORITSU | 2033 | 66386 | NA | 2017-08-05 | 365 | 2018-08-04 | NA | SGS | NA |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | Fluke 51 II | Digital Thermometer with probe | Fluke | 51 II | 89280141 | NA | 2017-08-10 | 365 | 2018-08-10 | NA | SGS | FMS/323 |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | NA | PDA detector |  |  | B08UPD563M | NA | NA | NA | NA | [6 empty] | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | FLUKE110 | Multimeter |  |  | 85730619 | NA | 2018-01-05 | 365 | 2019-01-05 | [11 empty] | 1
| Engineering Lab | 90406 | Liquilaz | Particle Measuring Systems | E20P | 51053 | NA | 2017-08-17 | 365 | 2018-08-16 | NA | Particle Measuring Systems | 110130306 |  |  |  | In-Use | TEST |  | NA | 1 | 1 | 1
| Engineering Lab | 95037 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 5037655 | NA | 2018-01-09 | 365 | 2019-01-09 | [11 empty] | 1
| Engineering Lab | 92207 | Electronic Weighing Balance |  |  | 92207 | NA | 2018-03-08 | 182 | 2018-09-06
| Engineering Lab | 90282 | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 1140660 | NA | 2013-06-21 | 365 | 2014-06-21 | NA | SGS | NA | Inactive-Remove from line (RE/094/13) Date 03/12/13. |  |  | In-Use | TEST | 7 | NA | 4 | 2 | 2
Detector | DOWNLOADING AND PROGRAMMING | TE 82964-1 | Download and programme(2489 and TUV) | Waters | NA | TE 82964-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 | 90% | 3 | 2 | 3
In-Use
|  | TE82972-1 | Download and programme(2998, PDA and elamda) | Waters | NA | TE82972-1 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 15 |  |  | 2 | 3
In-Use
|  | TE 82969-1 | Calibration and Linearity test for PDA family |  |  | TE 82969-1 | NA | NA | NA | NA | [11 empty] | 1
| 2489/TUV | TE 82903-10 | Final Tester | Waters | NA | TE 82903-10 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 150 | 90% | 4 | 2 | 2
In-Use
|  | TE 82903-8 | Final Tester | Waters | NA | TE 82903-8 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  |  | 4 | 2 | 2
In-Use
| 2489/TUV | TE 82903-9 | Final Tester | Waters | NA | TE 82903-9 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  | 90% | 4 | 2 | 2
In-Use
|  | TE 82903-6 | Final Tester | Waters | NA | TE 82903-6 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST |  |  | 4 | 2 | 2
In-Use
| 2998, PDA eLambda | TE 82882-8 | Final Tester | Waters | NA | TE 82882-8 | NA | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | 150 | 90% | 16 | 2 | 1
|  | TE 82882-9 | Final Tester | Waters | NA | TE 82882-9 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 1 | 1
|  | TE 82882-10 | Final Tester | Waters | NA | TE 82882-10 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 2 | 1
|  | TE 82882-11 | Final Tester | Waters | NA | TE 82882-11 | NA | NA | NA | NA | NA | NA | NA | [7 empty] | 16 | 2 | 1
|  | TE 82943 - 13 | 2998 Final Tester |  |  | TE 82882 - 13 | NA | NA | NA | NA | [11 empty] | 1
|  | TE 82943 - 14 | 2998 Final Tester |  |  | TE 82882 - 14 |  | NA | NA | NA | [11 empty] | 1
| Printer | NA | HP Laserjet P3005n |  |  | SN: JPKGJ01792 | NA | NA | NA | NA | [6 empty] | In-Use | TEST | NA | NA | NA | NA | NA
|  | NA | Brother |  |  | E69530C4N142942 |  | NA | NA | NA | NA | NA | NA |  |  |  | In-Use | TEST | NA | NA | NA | NA | NA
| Hipot | NA | Electrical Safety Analyzer | QUADTECH | Guardian 6000 | 08280865 | NA | 2018-10-13 | 365 | 2019-10-13 | NA | SGS | CR/10339085/1.1 |  |  |  |  | TEST | 7 | NA | 4 | 2 | 2
Production Floor | Label Printing Area | WATERS | WATERSISNL1 |  |  | 54177579 | NA | NA | NA | NA | [6 empty] | In-Use | PROCESS | NA | NA | 1 | 3 | 3
Production Floor | Label Printing Area | WATERS | WATERSPNL1 |  |  | 54177587 | NA | NA | NA | NA | [6 empty] | In-Use | PROCESS | NA | NA | 1 | 3 | 3
Formulas:
Calibration Due Date = [Calibration Date]+[Calibration Cycle (Months/ Days)]
Calibration Due Date = [Calibration Date]+365
Note on P4: skwjdedu: Period 1st Jan 2011 to 31st Dec 2011
Note on R4: skwjdedu: Green - Fixture working Red - Fixture not working
Note on W4: user: 3 = unique tester 2 = multiple tester but will affect output when down 1 = standard tester
Note on X4: user: 3 = High Risk, production volume high, will be affected once tester down 2 = Medium Risk, tester down will not affect output for max 2 weeks 1 = Low Risk, tester down will not affect output when down for more than 2 weeks.
Note on V9: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V11: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on G13: SGSJDEDU: EQID 0216
Note on V13: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V19: skwjdedu: 1 down
Note on V26: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V29: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V30: skwjdedu: 12 working, 4 spare
Note on V31: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V34: skwjdedu: 12 working, 4 spare
Note on V36: skwjdedu: 12 working, 4 spare
Note on G47: SGSMSBAf: original CPU KKTCN1W
Note on V49: skwjdedu: 12 working, 4 spare
Note on V50: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V51: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V52: skwjdedu: 12 working, 4 spare
Note on V53: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V54: skwjdedu: 12 working, 4 spare
Note on V55: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V56: skwjdedu: 12 working, 4 spare
Note on V60: skwjdedu: 12 working, 4 spare
Note on V61: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V62: skwjdedu: 12 working, 4 spare
Note on V63: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V64: skwjdedu: 4 working, 3 spare
Note on V65: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V66: skwjdedu: 12 working, 4 spare
Note on V67: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V68: skwjdedu: 12 working, 4 spare
Note on V71: skwjdedu: 12 working, 4 spare
Note on V72: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V73: skwjdedu: 12 working, 4 spare
Note on V74: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V75: skwjdedu: 4 working, 3 spare
Note on V76: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V77: skwjdedu: 12 working, 4 spare
Note on V78: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V79: skwjdedu: 12 working, 4 spare
Note on V82: skwjdedu: 12 working, 4 spare
Note on V83: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V84: skwjdedu: 12 working, 4 spare
Note on V85: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V86: skwjdedu: 4 working, 3 spare
Note on V87: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V88: skwjdedu: 12 working, 4 spare
Note on V89: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V90: skwjdedu: 12 working, 4 spare
Note on V91: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V94: skwjdedu: 12 working, 4 spare
Note on V95: skwjdedu: 12 working testers, 3 spares, 3 inactive, need to replace optic bench
Note on V97: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V98: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V100: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on V103: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors
Note on V106: skwjdedu: BSM no need detectors
Note on V109: skwjdedu: BSM no need detectors
Note on V112: skwjdedu: BSM no need detectors
Note on V115: skwjdedu: BSM no need detectors
Note on V118: skwjdedu: BSM no need detectors
Note on V121: skwjdedu: BSM no need detectors
Note on V124: skwjdedu: BSM no need detectors
Note on V127: skwjdedu: BSM no need detectors
Note on V131: skwjdedu: 1 - clean room, 1 - eng area but not calibrated. Express calibration - 2 weeks, normal - 1 mth.
Note on V133: skwjdedu: 1 - clean room, 1 spare
Note on C152: SGSJDEDU: EQID 0279
Note on C155: SGSJDEDU: EQID 0196
Note on V187: skwjdedu: 1 - pumps, 1 - BSM, 1 - alliance, 1 - detectors

## Summary
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 82 | 1 | 82 | 0 | 82 | 3 | 82 | 1 | 82 | 7 | 82 | 6
2 | 51 | 5 | 51 | 1 | 51 | 3 | 51 | 3 | 51 | 5 | 51 | 1
3 | 28 | 2 | 28 | 1 | 28 | 3 | 28 | 3 | 28 | 1 | 28 | 1
Total: | 161 | 8 | 161 | 2 | 161 | 9 | 161 | 7 | 161 | 13 | 161 | 8
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 82 | 1 | 82 | 0 | 82 |  | 82 |  | 82 |  | 82
2 | 51 | 5 | 51 | 3 | 51 |  | 51 |  | 51 |  | 51
3 | 28 | 3 | 28 | 3 | 28 |  | 28 |  | 28 |  | 28
Total: | 161 | 9 | 161 | 6 | 161 | 0 | 161 | 0 | 161 | 0 | 161 | 0

## Detail
Locations | STATIONS | Customer ID | Serial Number | Equipment | Number of common testers | Criticalilty | Cycle Time (hrs) | % YIELD @ FINAL TEST STATION ONLY | Unit Per Day | Utilization (%) | Calibration/PM Due Date | Reason of Failure and Length of downtime | Action | Replacement Cost (US$) | STATUS | Risk
Alliance Sub Assembly Stations | 2690 Alliance Front Panel Cell | TE82546 | NA | 2690 Alliance Front Panel Test Station | 1 | 3 | 0.085 | 95% | 89.4118 | 22.37% | 2019-02-08 | Tester display memory check Fail, Address 4606CF Val:ff 13 Aug 2018 - 13 Aug 2018 | 1. Replaced interface cable 270084 Bought from Sim Lim Tower Interface Cable with Molex Connector (male \ female) | S$48 | In-Use | 3
| Pump Head Tester | TE 82519 | NA | 2690 Pump Head Tester | 1 | 3 | 0.16 | 90% | 45 | 26.67% | 2018-12-01 |  |  |  | In-Use | 3
| Alliance Pump Tray 1 | TE82526-1 | NA | 2690 Pump Tray Tester | 3 | 2 | 0.65 | 80% | 9.8462 | 74.07% | 2019-02-08 |  |  |  | In-Use | 3
|  | 90891 | 12305757 | Electronic Weighing Balance | 6 | 2 |  | 80% |  |  | 2019-02-28 |  |  | NA | Reapair-in-progess | 2
| Alliance Pump Tray 2 | TE82526 | NA | 2690 Pump Tray Tester | 3 | 2 | 0.65 | 80% | 9.8462 |  | 2019-02-08 | Ripple Test Fail 19 Jun 2018 - 28 Sep 2018 | 1. Pin Plug FV2 during Ripple Test - Test Result PASS 2. Pin Plug FV2 during Flow Test - Test Result PASS 3. Removed the Pin Plug during Ripple Test - Test Result FAIL and observed small amount of leak solvent during test. 4. Removed the Pin Plug during Flow Test - Test Result FAIL and observed small amount of leak solvent during test. 5. Therefore I conclude that the root cause is leaky ASFVO Flow Valve. 6. Replaced ASFVO Valve 7. Replaced Vent Valve (01 Aug 2018) still failed Ripple test 8. Repair in Progress as of 28 Aug 2018 due to Ripple Test Fail 8.1 Replaced restrictor (600PSI)(Restrictor leading to the weighing balance) 8.2 Replaced filter at (F2 (2x) and F3 (2x) 8.3 Replaced Flow valve 1 and Flow valve 2 (ASFVO) 8.4 Replaced flow valve 4 (BSDA025T-NC-PV)(To push out the solvent) 8.5 Replaced air switch valve for FV1 and FV4 8.6 Replaced interface tubing(.020 x .062) at the output and input of tester 8.7 Replaced interface tubing(.020 x .062) going to FV 1 and FV 2 8.8 Ensure that there is no leakage on interface tubing that link to vent valve. 8.9 Swap in inline degasser from other pass tester 8.10 Swap in Electronic Weigh balance 8.11 Repair In Progress (awaiting cpu personality board from MJ.) |  | In-Use | 2
|  | 95206 | 12316922 | Electronic Weighing Balance | 6 | 2 |  |  |  |  | 2019-02-28 |  |  |  | In-Use | 2
| Alliance Pump Tray 3 | TE82526-4 | NA | 2690 Pump Tray Tester | 3 | 2 | 0.65 | 80% | 9.8462 |  | 2018-11-10 | Failed to home 24 Sep 18 - 24 Sep 18 | Rework on transducer connector | NA | In-Use | 2
|  | 90038 | 12200404 | Electronic Weighing Balance | 6 | 2 |  |  |  |  | 2019-02-28 |  |  |  | In-Use | 2
| Printer | NA | PHCND52518 | LaserJet Pro M402dn | [6 empty] | 2019-08-15 |  |  |  | In-Use | NA
| Alliance HPMV Cell | TE 82625-1 | NA | HPMV Tester | 1 | 3 | 0.11 | 85% | 61.8182 | 19.41% | 2019-01-17 |  |  |  | In-Use | 3
| 2690/2790 Syringe Drive Cell | TE82656 | NA | 2690/2790 Syringe Drive Tester | 1 | 3 | 0.05 | 95% | 152 | 19.74% | 2019-02-08 |  |  |  | In-Use | 2
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-1 | NA | Carousel Carrier Drive Alignment Station | 1 | 3 | 0.42 | 85% | 16.1905 | 62.5% | 2018-12-08 |  |  |  | In-Use | 3
| 2690 Carousel Carrier Drive Alignment Cell | TE82520-2 | NA | Carousel Carrier Drive Alignment Station | 1 | 3 | 0.42 | 85% | 16.1905 |  | 2018-11-10 |  |  |  | In-Use | 3
| Fluidics Tray Tester | TE82523 | NA | Fluidics Tray Tester | 1 | 3 | 0.25 | 90% | 28.8 | 86.81% | 2019-02-08 |  |  |  | In-Use | 3
| Injector Seal Pack Cell | TE 82503-04 | NA | Injector Seal Pack Tester | 2 | 3 | 0.3 | 85% | 22.6667 | 73.33% | 2019-01-17 |  |  |  | In-Use | 3
|  | TE 82503-05 | NA | Injector Seal Pack Tester | 2 | 3 | 0.3 | 85% | 22.6667 | 73.33% | 2019-03-01 | Ramp and Decay Test Fail 31 Aug 2018 - 04 Aug 2018 | TE82503-5 Injector Seal Pack Not able to hold pressure has been recovered. Below were observation and action taken to recover the Injector station. Observation: Ramp and Decay Failed - able to ramp pressure to 5000 PSI but not able to hold-off pressure for the last 3 mins. Action Taken : Verify Calibration : OK, able to calibrate Verify Fittings / tubing : No visible leak from fittings / tubing. Verify Filters : No visible dirt on filters / replaced filters Verify all Flow valve p/n: ASFVO(Position FV1,FV2 and FV3). Verify the output of the air switch for (position FV1,FV2 and FV3). Found out that there is air leaking from FV2 TE82503-5 Injector Seal Pack has been recovered. It is recommended to have standby ASFVO Flow valve and air switch. | ASFV Consign Part | In-Use | 3
| Alliance Heater Chiller Cell | TE82620 | NA | Heater Chiller Tester | N/A | N/A | Return to Milford | Return to Milford | NA | Return to Milford | Return to Milford | Inactive since 14 Mar 14 (ref no: RE/040/14) |  |  | Not-In-Use | N/A
|  | NA | NA | RS 232 C Precision DC Module | N/A | N/A | Return to Milford | Return to Milford | NA | Return to Milford | Return to Milford | Inactive since 14 Mar 14 (ref no: RE/130/15) |  |  |  | N/A
Alliance Final Test | Alliance Programming Station | NA | B08L32122N | Autoloader 2 | 1 | 1 | 0.41 | 95% | 18.5366 | 1.0789 | 2018-09-22 |  |  |  | In-Use | 3
| Rear Panel Station | TE82573-1 | C08L32093W | Rear panel test | 1 | 2 | 0.12 | 95% | 63.3333 | 0.31579 | 2018-09-22 |  |  |  | In-Use
| Alliance Sample Station | 90913 | 12315717 | Electronic Weighing Balance | 6 | 2 | 0.08 | 95% | 95 | 0.21053 | 2019-02-28 |  |  |  | In-Use | 2
Alliance Final Test Station 1 | Alliance Final Test 1 CPU | N/A | C12L32255W | CPU | 5 | 1 | 5.72 | 80% | NA | NA | 2018-12-06 | CPU Run Out of Virtual Memory 12 Sep 2018 to 12 Sep 2018 | Restart CPU | NA | In-Use | 1
| Alliance Final Test 1 | 81029 | H01TCM643M | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2018-01-23 |  |  |  | In-Use | 1
| Alliance Final Test 1 port 1 | 2489 | H1387E726A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 | 74.4% | 2019-06-13 |  |  |  | In-Use | 1
|  | WAT038040 | D15CHM641R | Column Heater | 16 | 1 |  |  |  |  | 2019-07-20 |  |  |  | In-Use | 1
| Alliance Final Test 1 port 2 | NA | D1387E759A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-04-10 |  |  |  | In-Use | 1
|  | 81014 | F00CHM531R | Column Heater |  | 1 |  |  |  |  | 2019-02-13 |  |  |  | In-Use
| Alliance Final Test 1 Port 3 | NA | D1387E765A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-04-12 |  |  |  | In-Use
|  | 81021 | M00CHM910M | Column Heater | 16 | 1 |  |  |  |  | 2019-02-13 |  |  |  | In-Use | 1
| Alliance Final Test 1 Port 4 | NA | H1387E728A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-06-28 |  |  |  | In-Use
|  | 81020 | H01CHM352M | Column Heater | 16 | 1 |  |  |  |  | 2019-02-13 |  |  |  | In-Use | 1
Alliance Final Test Station 2 | Alliance Final Test 2 CPU | N/A | KKPRA3T | CPU | 6 | 1 | 5.72 | 80% | NA | NA | 2019-05-11 |  |  |  | In-Use
| Alliance Final Test 2 | 81033 | E97TCM153R | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2019-01-24 |  |  |  | In-Use
| Alliance Final Tester 2 port 1 | NA | C1387E135M | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-04-12 |  |  |  | In-Use
|  | 81008 | M03CHM015R | Column Heater | 24 | 1 |  |  |  |  | 2019-04-03 |  |  |  | In-Use
| Alliance Final Tester 2 port 2 | NA | H1387E710A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-11 |  |  |  | In-Use
|  | 81018 | MX6JM7863M | Column Heater | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use
| Alliance Final Tester 2 port 3 | NA | H1387E731A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-06-12 |  |  |  | In-Use
|  | 81009 | E01CHM203R | Column Heater | 24 | 1 |  |  |  |  | 2019-04-03 |  |  |  | In-Use
| Alliance Final Tester 2 port 4 | NA | H1387E727A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-11 |  |  |  | In-Use
|  | 81016 | MX6JM7867M | Column Heater | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use
Alliance Final Test Station 3 | Alliance Final Tester 3 | N/A | M11L32097W | CPU | 6 | 1 | 5.72 | 80% | NA | NA | 2019-05-11 |  |  |  | In-Use
| Alliance Final Tester 3 | 81030 | MX5MM4884M | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2019-01-26 |  |  |  | In-Use | 1
| Alliance Final Tester 3 port 1 | 81013 | F00CHM528R | Column Heater | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-02-14 |  |  |  | In-Use | 1
|  | NA | H1387E735A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-07-11 |  |  |  | In-Use | 1
| Alliance Final Tester 3 port 2 | NA | H1387E730A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-12 |  |  |  | In-Use | 1
|  | 81012 | B03CHM551R | Column Heater | 24 | 1 |  |  |  |  | 2019-02-14 |  |  |  | In-Use | 1
| Alliance Final Tester 3 port 3 | NA | H1387E737A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-12 |  |  |  | In-Use | 1
|  | 81019 | L97CHM711M | Column Heater | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use | 1
| Alliance Final Tester 3 port 4 | NA | H1387E734A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-19 |  |  |  | In-Use | 1
|  | 81019 | D01CHM128R | Column Heater | 24 | 1 |  |  |  |  | 2019-02-14 |  |  |  | In-Use | 1
Alliance Final Test Station 4 | Alliance Final Test 4 | NA | H10L32039W | CPU |  | 1 | 5.72 | 80% | NA | NA | 2019-05-11 |  |  |  | In-Use
| Alliance Final Test 4 | 81033 | MX5DM2250R | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2019-04-03 |  |  |  | In-Use | 1
| Alliance Final Test 4 Port 1 | 81022 | H01CHM335M | Column Heater | 16 | 1 |  |  | 1.1189 | 74.49% | 2019-02-13 |  |  |  | In-Use | 1
|  | NA | D1387E766A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  |  |  | 2019-04-06 |  |  |  | In-Use | 1
| Alliance Final Test 4 Port 2 | 81017 | MX6KM9444M | Column Heater | 16 | 1 |  |  | 1.1189 |  | 2019-02-14 |  |  |  | In-Use | 1
|  | NA | H1387E733A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  |  |  | 2019-06-13 |  |  |  | In-Use | 1
| Alliance Final Test 4 Port 3 | NA | H1387E732A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  | 1.1189 |  | 2019-06-12 |  |  |  | In-Use | 1
|  | 81011 | K03CHM938R | Column Heater | 16 | 1 |  |  |  |  | 2019-02-14 |  |  |  | In-Use | 1
| Alliance Final Test 4 Port 4 | NA | H1387E736A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  | 1.1189 |  | 2019-06-21 |  |  |  | In-Use | 1
|  | 81023 | MX5MM5993M | Column Heater | 16 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use | 1
Alliance Final Test Station 5 | Alliance Final Test 5 | NA | C12L32205W | CPU | 6 | 1 | 5.72 | 80% | NA |  | 2019-05-24 |  |  |  | In-Use
| Alliance Final Test 5 | NA | G01TCM633M | Temperature Control Module | 7 | 1 |  |  | #DIV/0! |  | 2019-01-24 |  |  |  | In-Use | 1
| Alliance Final Test 5 Port 1 | NA | C13CHM795R | Column Heater | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-07-20 |  |  |  | In-Use | 1
|  | NA | B1587E610A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-02-20 |  |  |  | In-Use | 1
| Alliance Final Test 5 Port 2 | NA | D12CHM299R | Column Heater | 24 | 1 |  |  | 1.1189 |  | 2019-07-20 |  |  |  | In-Use | 1
|  | NA | B1587E607A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use | 1
| Alliance Final Test 5 Port 3 | NA | B1587E608A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-02-15 |  |  |  | In-Use | 1
|  | NA | K14CHM398R | Column Heater | 24 | 1 |  |  |  |  | 2019-07-20 |  |  |  | In-Use | 1
| Alliance Final Test 5 Port 4 | NA | B1587E609A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-02-15 |  |  |  | In-Use | 1
|  | NA | L14CHM438R | Column Heater | 24 | 1 |  |  |  |  | 2019-07-20 |  |  |  | In-Use | 1
Alliance Final Test Station 6 | Alliance Final Test 6 | NA | A14L32340W | CPU | 6 | 1 | 5.72 | 80% | NA | NA | 2019-06-13 |  |  |  | In-Use
| Alliance Final Test 6 | NA | A18TC2481G | Temperature Control Module 2 | 7 | 1 |  |  | NA | NA | 2019-01-29 |  |  |  | In-Use | 1
| Alliance Final Test 6 Port 1 | NA | M17CHM840M | Column Heater | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-06-13 |  |  |  | In-Use | 1
|  | NA | C1887E101A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-03-21 |  |  |  | In-Use | 1
| Alliance Final Test 6 Port 2 | NA | M17CHM834M | Column Heater | 24 | 1 |  |  | 1.1189 |  | 2019-06-13 |  |  |  | In-Use | 1
|  | NA | C1887E103A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-03-21 |  |  |  | In-Use | 1
| Alliance Final Test 6 Port 3 | NA | C1887E106A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-03-20 |  |  |  | In-Use | 1
|  | NA | M17CHM835M | Column Heater | 24 | 1 |  |  |  |  | 2019-06-13 |  |  |  | In-Use | 1
| Alliance Final Test 6 Port 4 | NA | D1887E275A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-04-09 |  |  |  | In-Use | 1
|  | NA | M17CHM839M | Column Heater | 24 | 1 |  |  |  |  | 2019-06-13 |  |  |  | In-Use | 1
| Alliance Final Test 6 | NA | B18TC2495G | Temperature Control Module | 7 | 1 |  |  | NA |  | 2019-02-06
Safety Analyzer | Hipot | NA | 19032G000182 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 95% | 69.0909 | 28.95% | 2018-11-09 |  |  |  | In-Use | 2
QA WORKSTATION | QA WORKSTATION | NA | NA | CPU | 7 | 1 | 5.72 | 80% | 1.1189 | NA | 2018-12-22 |  |  |  |  | 1
| QA WORKSTATION | NA | H14SMC938G | Column Heater | 16 | 1 |  |  |  | NA | 2018-12-22 |  |  |  |  | 1
|  | NA | H1387E729A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  |  | NA | 2019-07-19 |  |  |  |  | 1
PUMP CELLS | PUMP Cell | NA | 19032G000182 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 95% | 69.0909 | 28.95% | 2018-11-09 |  |  |  | In-Use | 2
| 1500 SERIES CELL STATION | 90604 | 12326187 | Electronic Weighing Balance | 6 | 2 | 1.58 | 95% | 4.8101 | 103.9% | 2019-01-10 |  |  |  | In-Use | 2
|  | TE 82705-2 | TE 82705-2 | 1500 Series Final Test Station | 1 | 3 |  |  |  |  | 2018-11-22 | 1. Ripple Test Fail 04 Sep 2018 to 06 Sep 2018 | Replaced Restrictor 1K PSI | NA | In-Use | 3
| 515 PUMP CELL STATION | 90018 | 12303279 | Electronic Weighing Balance | 6 | 2 | 0.83 | 95% | 9.1566 | 54.61% |  |  |  |  | In-Use | 2
|  | TE 82534 | TE 82534 | 515 Final Tester | 1 | 3 |  |  |  |  | 2018-11-22 |  |  |  | In-Use | 3
| Printer | NA | JPKGJ01942 | LaserJet P3005n | [6 empty] | 2019-08-15 |  |  |  | In-Use | NA
BSM/QSM FINAL TESTER | BSM/QSM Cell | 98206 | 6069766 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 98% | 71.2727 | 25.26% | 2018-12-04 |  |  |  | In-Use | 2
| BSM/QSM TE82943-11 | 90606 | 12326300 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-02-06 | Missing Segment on Display 20 Jul 2018 to 28 Jul 2018 Reschedule Calibration 28 Jul 18 to 16 Aug 2018 | Calibration – PO raised Jul 2018 PO released 28 Jul 2018 Schedule Calibration : 06 Aug 2018 Calibration Completed Release to Prod’n – 16 Aug 2018 | NA | In-Use | 2
|  | TE82943-11 | TE-82943-11 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 84010 | M09UPT364A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-03-02 |  |  |  | In-Use | 3
| BSM/QSM TE82943-12 | 90605 | 12324733 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
|  | TE82943-12 | TE82943-12 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 84007 | M09UPT361A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-05-04 |  |  |  | In-Use | 3
| BSM/QSM TE82943-13 | 90608 | 12326184 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | Under repair | Unable to get stable reading 09 Jul 2018 - 31 Aug 2018 | 1. Ensure that the bubble spirit level is okay. 2. Re-Zero the weighing balance but to no avail. 3. Ensure that balance stability is not influence by breeze or vibration. 4. Remove weighing pan, Re insert the pan and Re-zero but no avail. 5. Press and hold the Re-Zero key until balance display CAL. Release the RE-Zero Key. Display shows "Error 1". 7. Repair In-progress 8. ETA: WW40'2018 | NA | Not-In-Use | 2
|  | TE82943-13 | TE82943-13 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 84011 | K09UPT215A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-05-03 |  |  |  | In-Use | 3
| BSM/QSM TE82943-14 | NA | 12336741 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2018-11-15 |  |  |  | In-Use | 2
|  | TE82943-14 | TE82943-14 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 84008 | M09UPT362A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-04-20 |  |  |  | In-Use | 3
| BSM/QSM TE82943-15 | 90603 | 12326180 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
|  | TE82943-15 | TE82943-15 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 84009 | M09UPT363A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-04-21 |  |  |  | In-Use | 3
| BSM/QSM TE82943-16 | 90600 | 12326190 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
|  | TE82943-16 | TE82943-16 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 84006 | M09UPT360A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-05-04 |  |  |  | In-Use | 3
| BSM/QSM TE82943-17 | 90607 | 12326279 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-02-06 | 1. Missing Segment on display 20 Jul 2018-28 Jul2018 2. Reschedule calibration 28 Jul 2018 to 16 Aug 2018 | Calibration – PO raised Jul 2018 PO released 28 Jul 2018 Schedule Calibration : 06 Aug 2018 Calibration Completed Release to Prod’n – 16 Aug 2018 |  | In-Use | 2
|  | TE82943-17 | TE82943-17 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 84005 | M09UPT359A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-04-21 |  |  |  | In-Use | 3
| BSM/QSM TE82943-18 | 90610 | 12326186 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
|  | TE82943-18 | TE82943-18 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
|  | 90610 | 12326186 | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-01-10 |  |  |  | In-Use | 3
| Acquity Actuator Homing box TE83076-3 | NA | TE83076-3 | Acquity Actuator Homing box | [6 empty] | 2019-07-27
| Actuator Tester TE 82800-2 | TE 82800-2 | TE 82800-2 | Aquity Actuator Test | 2 | 2 | 0.23 | 85% | 29.5652 | 101.5% | 2019-02-08 | Fail To initialize 10 Sep 18 to 10 Sep 2018 | Replaced power cable p/n: WAT-WAT270313-03 | NA | In-Use | 2
| Actuator Tester TE 82800-3 | TE 82800-3 | TE 82800-3 | Aquity Actuator Test | 2 | 2 | 0.23 | 85% | 29.5652 | 101.5% | 2019-02-08 |  |  |  | In-Use | 2
| Printer | NA | PHCNF49758 | LaserJet Pro M402dn | [6 empty] | 2019-08-15 |  |  |  | In-Use | NA
Cleanroom (Running 2 Shift) | Clean Room | 92258 | 435110 | Picoammeter | 2 | 2 | 0.62 | 95% | 24.5161 | 73.41% | 2019-03-28 |  |  |  | In-Use | 1
| Clean Room | 9120000 | NA | Light Meter | 2 | 2 | 0.62 | 95% | 24.5161 | 73.41% | NA |  |  |  | In-Use | 1
| Clean Room | TE 82475-2 | TE 82475-2 | Flowcell Tester | 2 | 1 | 0.36 | 95% | 42.2222 | 42.63% | 2019-04-11 |  |  |  | In-Use | 1
| Clean Room | TE82881-2 | K15UPD782R | Optic Bench fine alignment Tester (Box Build) | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82881-2 | D08L32121W / C08BLM058C | PDA Fine Alignment CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82880-2 | L07L32070W | PDA ROUGH ALIGNMENT CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82874 | TE82874 | TUV lamp supply | 1 | 3 | 0.28 | 95% | 54.2857 | 33.16% | 2019-09-21 |  |  |  | In-Use | 3
| Clean Room | TE82965 | TE82965 | Filter Wheel Alignment | 1 | 3 | 0.28 | 95% | 54.2857 | 33.16% | 2019-09-21 |  |  |  | In-Use | 3
| Clean Room | TE9990293134 | NA | Mercury lamp power Supply |  | 1 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use
| Clean Room | TE82973-1 (TE829002998) | B08L32248W/ A08BLM491C | Flowcell Tester CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82918-2 | A16VPD196R | 2998 fine alignment(box build) | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82918-2 | B08L32065W | 2998 fine alignment CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82917-1 | B08L32068W | 2998 rough alignment CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE186002998 | G08998943N | Flowcell tester(2998) | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82951-1 | J06L32014W / G06BLM650C | PDA Assembly CPU | 1 | 3 | 0.28 | 95% | 27.1429 | 66.32% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | TE82868-2 | K07L32122W | PDA Aperture Alignment System | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
| Clean Room | 95515 | RSL 1002609 | Starlite 150 | 1 | 3 | 0.36 | 95% | 42.2222 | 42.63% | 2019-04-24 |  |  |  | In-Use | 3
| Clean Room | TE83096-2 | TE83096-2 | PDA Universal Alignment | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | NA |  |  |  | In-Use | 3
Eng's Area | Engineering Lab | NA | 9093205 | Pressure Calibrator | 1 | 1 | As per usage on evaluation only | NA | NA |  | 2018-09-12 |  |  |  | In-Use | 1
| Engineering Lab | FLUKE179 | 90240092 | Multimeter | 1 | 1 |  | NA | NA |  | 2019-01-31 |  |  |  | In-Use | 1
| Engineering Lab | High Voltage Probe | 89200011 | High Voltage Probe | 1 | 1 |  | NA | NA |  | 2019-09-25 |  |  |  | In-Use | 1
| Engineering Lab | 310788 | 85074 | Picometer |  |  |  | NA | NA |  | 2018-04-17 |  |  |  | In-Use
| Engineering Lab | Clamp Meter | 66386 | Clamp Meter | 1 | 1 |  | NA | NA |  | 2019-07-24 |  |  |  | In-Use | 1
| Engineering Lab | Fluke 51 II | 89280141 | Digital Thermometer with probe | 1 | 1 |  | NA | NA |  | 2019-07-23 |  |  |  | In-Use | 1
| Engineering Lab | FLUKE110 | 85730619 | Multimeter |  | 1 |  | NA | NA |  | 2019-01-05
| Engineering Lab | 90406 | 51053 | Liquilaz | 1 | 1 |  | NA | NA |  | 2019-08-17 |  |  |  | In-Use | 1
| Engineering Lab | 95037 | 5037655 | Electrical Safety Analyzer |  | 1 |  | NA | NA |  | 2019-01-10
| Engineering Lab | 90273 | 12308385 | Electronic Weighing Balance |  |  |  | NA | NA |  | 2018-09-08 |  | Return to WPPL
| Engineering Lab | 92207 | 12337042 | Electronic Weighing Balance |  |  |  | NA | NA |  | 2018-09-08 |  | Return to WPPL
| Engineering Lab | 90282 | 1140660 | Electrical Safety Analyzer | 4 | 2 | 7 | SCRAP | #VALUE! |  | (RE/084/14 Removed) 9 jun2014 | Inactive-Remove from line (RE/094/13) Date 03/12/13. |  |  | In-Use | 2
Detector | DOWNLOADING AND PROGRAMMING | TE 82964-1 | H04BLM012C / J04L32108W | Download and programme(2489 and TUV) | 3 | 2 | 0.36 | 90% | 20 | 90% | 2019-08-03 |  |  |  | In-Use | 3
| DOWNLOADING AND PROGRAMMING | TE82972-1 \ TE83057-4 | C08L32078W | Download and programme(2998, PDA and elamda) | 3 | 2 | 0.36 | 90% | 20 | 90% | 2019-08-03 |  |  |  | In-Use | 3
| DOWNLOADING AND PROGRAMMING | TE 82969-1 | C08L32208W | Calibration and Linearity test for PDA family | 3 | 1 | 0.36 | 90% | 20 | 90% | 2019-08-03 | 1. TE82969-1 not configure to run Linearity test for PDA 21 Sep 2018 to 24 Sep 2018 | 1. Raise OESR 2. Tester configure by E.T (ADMIN) to be able to do Linearity test for PDA. |  | In-Use
| 2489/TUV | TE 82903-10 | J04L32102W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
|  | TE 82903-8 | J04L32120W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
|  | TE 82903-9 | J04L32091W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
|  | TE 82903-6 | G05L32009W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
| 2998, PDA eLambda | TE 82882-8 | E13L32018W | Final Tester | 16 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-08-03 |  |  |  | In-Use | 1
|  | TE 82882-9 | NA | Final Tester | 16 | 1 | 3 | 90% | 9.6 | 65.1% | 2019-05-11 |  |  |  | In-Use | 1
|  | TE 82882-10 | M12L32412W | Final Tester | 16 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-05-11 |  |  |  | In-Use | 1
|  | TE 82882-11 | F13L32287W | Final Tester | 16 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-05-11 |  |  |  | In-Use | 1
|  | TE82882-13 | S4ALMVW | 2998 Final Tester | 16 | 1 | 3 | 90% | 9.6 | 65.1% | 2019-03-15 | Test station test software version 4.1 completed test but not able to print report 21 Sep 2018 - 24 Sep 2018 | 1. Raise OESR 2. E.T. (ADMIN) unblock user access for tester account. | NA | In-Use
|  | TE82882-14 | S4ALYMH | 2998 Final Tester | 16 | 1 | 3 | 90% | 9.6 | 65.1% | 2019-03-15 |  |  |  | In-Use
| Printer | NA | PHCNH24483 | LaserJet Pro M402dn | NA | NA | NA | NA | NA | NA | 2019-08-15 |  |  |  | In-Use | NA
| Hipot | 86041 | 08280865 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 98% | 71.2727 | 70.15% | 2018-10-13 |  |  |  |  | 2
Production Floor | Label Printing Area | WATERS |  | WATERSISNL1 | 1 | 3 | NA | NA | NA | NA | NA | NA | NA | NA | In-Use | 3
Production Floor | Label Printing Area | WATERS |  | WATERSPNL1 | 1 | 3 | NA | NA | NA | NA | NA | NA | NA | NA | In-Use | 3
SERVER RACK | SERVER | Power Edge T320 |  | Waters Server | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
| SERVER | NA |  | Flat Panel Monitor | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
| SERVER | NA |  | Mouse | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
| SERVER | NA |  | Keyboard | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
| SERVER | APC UPS |  | UPS | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
Note on G1: user: 3 = unique tester 2 = multiple tester but will affect output when down 1 = standard tester
Note on M1: skwjdedu:: Period 1st Jan 2011 to 31st Dec 2011
Note on N1: skwjdedu: Period 1st Jan 2011 to 31st Dec 2011
Note on P1: skwjdedu: Green - Fixture working Red - Fixture not working
Note on Q1: user: 3 = High Risk, production volume high, will be affected once tester down 2 = Medium Risk, tester down will not affect output for max 2 weeks 1 = Low Risk, tester down will not affect output when down for more than 2 weeks.
Note on F5: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on F7: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on F9: skwjdedu: 3 for punp tray, 2 for piumps and 1 final test station
Note on F16: skwjdedu: 1 down
Note on H18: Joel Mamalateo Deduyo:
Note on F55: skwjdedu: 12 working, 4 spare

## Spare CPU (hidden)
SPARE COMPUTER
ITEM | Description | QTY | REMARKS
1 | SRVR, eLACE/32 WXP#11 | 1 | Eng'n Area (No RS232)
C10L32154W
2 | SRVR, eLACE/32 WXP#13 | 1 | Alliance Final Test 5th Station
C12L32255W
3 | SRVR, eLACE/32 WXP#10 | 1 | Eng'g Area
A04L32012W(not sure working;old white lace)
4 | SRVR, eLACE/32 WXP#14 | 1 | Used at Pump Tray 1 | Previously setup at Clean Rooom for Genealogy
G13L32034W
5 | SRVR, eLACE/32 WXP#14 | 1 | Release to Desmond 27 Oct 2017
A14L32340W