## IT Infrastructure
Servers
IP Address | Computer Name | Computer Description | IP # Assigned | Platform | Serial Number | Mfg Year | RAM (GB) | OS | Location
10.249.26.6 | SOLECTRONMFG | Solectron Server (Flextronics Singapore) | 2006-01-20 | Other | 40NM291 | 2005 | 1 | Server 2003 | Flex SG
10.249.26.1 | SOLECTRONMFGNEW | Solectron Server (Flextronics Singapore) | 2018-07-17 | Other | HXQKN22 | 2014 | 8 | Server 2012 | Flex SG
10.249.49.1 | FLEXTRONICSMFG | FLEXTRONICS Malaysia Server | 2011-05-25 | Other | - | - | 4 | Server 2008 R2 | Flex Senai
10.249.40.1 | VENTUREMFG | VENTURE Server | 2009-02-24 | Other | 5LD6P22 | 2017 | 8 | Server 2012 R2 | Venture SG
10.249.40.233 | UK Server | VENTURE UK Server | - | Other | S4CM0575 | 2017 | 64 | Win7 (64bit) | Venture PCB
10.249.55.1 | CEIMFG | CEI Server, Singapore | 2013-06-20 | Other | 33RTGX1 | NA | 8 | Server 2008 R2 | CEI SG
10.249.50.1 | CEIMFG | CEI Server, Indonesia | 2011-10-04 | Other | D4CVVQ1 | 060611 | 4 | Server 2008 R2 | CEI Batam
10.64.2.246 | WAT-Server | UK Server (006883) | NA | NA | NA | NA | NA | WIN 7 SP1 | CEI Batam
Label Printers
Pinter Model | Manufacturer | Serial Number | Mfg Year | Service Provider | Contract start date | Contract due date |  |  | Location
DMX I-4308 | Datamax | 54177579 | NA | Autopack Pte Ltd | NA | NA |  |  | Flex SG
DMX I-4308 | Datamax | 54177587 | NA | Autopack Pte Ltd | NA | NA |  |  | Flex SG
DMX I-4308 | Datamax | 83463054 | 2000 | Autopack | NA | NA |  |  | Venture SG
DMX I-4309 | Datamax | 83463058 | 2000 | Autopack | NA | NA |  |  | Venture SG
I-4310e | Datamax O'neil | 32353001 | NA | Autopack | NA | NA |  |  | CEI SG
I-4310e | Datamax O'neil | 32353002 | NA | Autopack | NA | NA |  |  | CEI SG
Total items in CMs
| Flex Singapore | Flex Senai | Venture Singapore | Venture PCBA | CEI Singapore | CEI Batam | Richport | Keypoint | Total
PC | 55 | 5 | 44 | 5 | 8 | 6 | 1 | 2 | 126
Tester | 161 | 5 | 47 | 5 | 5 | 6 | 1 | 2 | 232

## Flex SG
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 82 | 1 | 82 | 0 | 82 | 3 | 82 | 1 | 82 | 7 | 82 | 6
2 | 51 | 5 | 51 | 1 | 51 | 3 | 51 | 3 | 51 | 5 | 51 | 1
3 | 28 | 2 | 28 | 1 | 28 | 3 | 28 | 3 | 28 | 1 | 28 | 1
Total: | 161 | 8 | 161 | 2 | 161 | 9 | 161 | 7 | 161 | 13 | 161 | 8
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 82 | 1 | 82 | 0 | 82 |  | 82 |  | 82 |  | 82
2 | 51 | 5 | 51 | 3 | 51 |  | 51 |  | 51 |  | 51
3 | 28 | 3 | 28 | 3 | 28 |  | 28 |  | 28 |  | 28
Total: | 161 | 9 | 161 | 6 | 161 | 0 | 161 | 0 | 161 | 0 | 161 | 0
STATIONS | Customer ID | Serial Number | Equipment | Number of common testers | Criticalilty | Cycle Time (hrs) | % YIELD @ FINAL TEST STATION ONLY | Unit Per Day | Utilization (%) | Calibration/PM Due Date | Reason of Failure and Length of downtime | Action | Replacement Cost (US$) | STATUS | Risk
2690 Alliance Front Panel Cell | TE82546 | NA | 2690 Alliance Front Panel Test Station | 1 | 3 | 0.085 | 95% | 89.4118 | 22.37% | 2019-02-08 | Tester display memory check Fail, Address 4606CF Val:ff 13 Aug 2018 - 13 Aug 2018 | 1. Replaced interface cable 270084 Bought from Sim Lim Tower Interface Cable with Molex Connector (male \ female) | S$48 | In-Use | 3
Pump Head Tester | TE 82519 | NA | 2690 Pump Head Tester | 1 | 3 | 0.16 | 90% | 45 | 26.67% | 2018-12-01 |  |  |  | In-Use | 3
Alliance Pump Tray 1 | TE82526-1 | NA | 2690 Pump Tray Tester | 3 | 2 | 0.65 | 80% | 9.8462 | 74.07% | 2019-02-08 |  |  |  | In-Use | 3
| 90891 | 12305757 | Electronic Weighing Balance | 6 | 2 |  | 80% |  |  | 2019-02-28 |  |  | NA | Reapair-in-progess | 2
Alliance Pump Tray 2 | TE82526 | NA | 2690 Pump Tray Tester | 3 | 2 | 0.65 | 80% | 9.8462 |  | 2019-02-08 | Ripple Test Fail 19 Jun 2018 - 28 Sep 2018 | 1. Pin Plug FV2 during Ripple Test - Test Result PASS 2. Pin Plug FV2 during Flow Test - Test Result PASS 3. Removed the Pin Plug during Ripple Test - Test Result FAIL and observed small amount of leak solvent during test. 4. Removed the Pin Plug during Flow Test - Test Result FAIL and observed small amount of leak solvent during test. 5. Therefore I conclude that the root cause is leaky ASFVO Flow Valve. 6. Replaced ASFVO Valve 7. Replaced Vent Valve (01 Aug 2018) still failed Ripple test 8. Repair in Progress as of 28 Aug 2018 due to Ripple Test Fail 8.1 Replaced restrictor (600PSI)(Restrictor leading to the weighing balance) 8.2 Replaced filter at (F2 (2x) and F3 (2x) 8.3 Replaced Flow valve 1 and Flow valve 2 (ASFVO) 8.4 Replaced flow valve 4 (BSDA025T-NC-PV)(To push out the solvent) 8.5 Replaced air switch valve for FV1 and FV4 8.6 Replaced interface tubing(.020 x .062) at the output and input of tester 8.7 Replaced interface tubing(.020 x .062) going to FV 1 and FV 2 8.8 Ensure that there is no leakage on interface tubing that link to vent valve. 8.9 Swap in inline degasser from other pass tester 8.10 Swap in Electronic Weigh balance 8.11 Repair In Progress (awaiting cpu personality board from MJ.) |  | In-Use | 2
| 95206 | 12316922 | Electronic Weighing Balance | 6 | 2 |  |  |  |  | 2019-02-28 |  |  |  | In-Use | 2
Alliance Pump Tray 3 | TE82526-4 | NA | 2690 Pump Tray Tester | 3 | 2 | 0.65 | 80% | 9.8462 |  | 2018-11-10 | Failed to home 24 Sep 18 - 24 Sep 18 | Rework on transducer connector | NA | In-Use | 2
| 90038 | 12200404 | Electronic Weighing Balance | 6 | 2 |  |  |  |  | 2019-02-28 |  |  |  | In-Use | 2
Printer | NA | PHCND52518 | LaserJet Pro M402dn | [6 empty] | 2019-08-15 |  |  |  | In-Use | NA
Alliance HPMV Cell | TE 82625-1 | NA | HPMV Tester | 1 | 3 | 0.11 | 85% | 61.8182 | 19.41% | 2019-01-17 |  |  |  | In-Use | 3
2690/2790 Syringe Drive Cell | TE82656 | NA | 2690/2790 Syringe Drive Tester | 1 | 3 | 0.05 | 95% | 152 | 19.74% | 2019-02-08 |  |  |  | In-Use | 2
2690 Carousel Carrier Drive Alignment Cell | TE82520-1 | NA | Carousel Carrier Drive Alignment Station | 1 | 3 | 0.42 | 85% | 16.1905 | 62.5% | 2018-12-08 |  |  |  | In-Use | 3
2690 Carousel Carrier Drive Alignment Cell | TE82520-2 | NA | Carousel Carrier Drive Alignment Station | 1 | 3 | 0.42 | 85% | 16.1905 |  | 2018-11-10 |  |  |  | In-Use | 3
Fluidics Tray Tester | TE82523 | NA | Fluidics Tray Tester | 1 | 3 | 0.25 | 90% | 28.8 | 86.81% | 2019-02-08 |  |  |  | In-Use | 3
Injector Seal Pack Cell | TE 82503-04 | NA | Injector Seal Pack Tester | 2 | 3 | 0.3 | 85% | 22.6667 | 73.33% | 2019-01-17 |  |  |  | In-Use | 3
| TE 82503-05 | NA | Injector Seal Pack Tester | 2 | 3 | 0.3 | 85% | 22.6667 | 73.33% | 2019-03-01 | Ramp and Decay Test Fail 31 Aug 2018 - 04 Aug 2018 | TE82503-5 Injector Seal Pack Not able to hold pressure has been recovered. Below were observation and action taken to recover the Injector station. Observation: Ramp and Decay Failed - able to ramp pressure to 5000 PSI but not able to hold-off pressure for the last 3 mins. Action Taken : Verify Calibration : OK, able to calibrate Verify Fittings / tubing : No visible leak from fittings / tubing. Verify Filters : No visible dirt on filters / replaced filters Verify all Flow valve p/n: ASFVO(Position FV1,FV2 and FV3). Verify the output of the air switch for (position FV1,FV2 and FV3). Found out that there is air leaking from FV2 TE82503-5 Injector Seal Pack has been recovered. It is recommended to have standby ASFVO Flow valve and air switch. | ASFV Consign Part | In-Use | 3
Alliance Heater Chiller Cell | TE82620 | NA | Heater Chiller Tester | N/A | N/A | Return to Milford | Return to Milford | NA | Return to Milford | Return to Milford | Inactive since 14 Mar 14 (ref no: RE/040/14) |  |  | Not-In-Use | N/A
| NA | NA | RS 232 C Precision DC Module | N/A | N/A | Return to Milford | Return to Milford | NA | Return to Milford | Return to Milford | Inactive since 14 Mar 14 (ref no: RE/130/15) |  |  |  | N/A
Alliance Programming Station | NA | B08L32122N | Autoloader 2 | 1 | 1 | 0.41 | 95% | 18.5366 | 1.0789 | 2018-09-22 |  |  |  | In-Use | 3
Rear Panel Station | TE82573-1 | C08L32093W | Rear panel test | 1 | 2 | 0.12 | 95% | 63.3333 | 0.31579 | 2018-09-22 |  |  |  | In-Use
Alliance Sample Station | 90913 | 12315717 | Electronic Weighing Balance | 6 | 2 | 0.08 | 95% | 95 | 0.21053 | 2019-02-28 |  |  |  | In-Use | 2
Alliance Final Test 1 CPU | N/A | C12L32255W | CPU | 5 | 1 | 5.72 | 80% | NA | NA | 2018-12-06 | CPU Run Out of Virtual Memory 12 Sep 2018 to 12 Sep 2018 | Restart CPU | NA | In-Use | 1
Alliance Final Test 1 | 81029 | H01TCM643M | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2018-01-23 |  |  |  | In-Use | 1
Alliance Final Test 1 port 1 | 2489 | H1387E726A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 | 74.4% | 2019-06-13 |  |  |  | In-Use | 1
| WAT038040 | D15CHM641R | Column Heater | 16 | 1 |  |  |  |  | 2019-07-20 |  |  |  | In-Use | 1
Alliance Final Test 1 port 2 | NA | D1387E759A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-04-10 |  |  |  | In-Use | 1
| 81014 | F00CHM531R | Column Heater |  | 1 |  |  |  |  | 2019-02-13 |  |  |  | In-Use
Alliance Final Test 1 Port 3 | NA | D1387E765A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-04-12 |  |  |  | In-Use
| 81021 | M00CHM910M | Column Heater | 16 | 1 |  |  |  |  | 2019-02-13 |  |  |  | In-Use | 1
Alliance Final Test 1 Port 4 | NA | H1387E728A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-06-28 |  |  |  | In-Use
| 81020 | H01CHM352M | Column Heater | 16 | 1 |  |  |  |  | 2019-02-13 |  |  |  | In-Use | 1
Alliance Final Test 2 CPU | N/A | KKPRA3T | CPU | 6 | 1 | 5.72 | 80% | NA | NA | 2019-05-11 |  |  |  | In-Use
Alliance Final Test 2 | 81033 | E97TCM153R | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2019-01-24 |  |  |  | In-Use
Alliance Final Tester 2 port 1 | NA | C1387E135M | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-04-12 |  |  |  | In-Use
| 81008 | M03CHM015R | Column Heater | 24 | 1 |  |  |  |  | 2019-04-03 |  |  |  | In-Use
Alliance Final Tester 2 port 2 | NA | H1387E710A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-11 |  |  |  | In-Use
| 81018 | MX6JM7863M | Column Heater | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use
Alliance Final Tester 2 port 3 | NA | H1387E731A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-06-12 |  |  |  | In-Use
| 81009 | E01CHM203R | Column Heater | 24 | 1 |  |  |  |  | 2019-04-03 |  |  |  | In-Use
Alliance Final Tester 2 port 4 | NA | H1387E727A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-11 |  |  |  | In-Use
| 81016 | MX6JM7867M | Column Heater | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use
Alliance Final Tester 3 | N/A | M11L32097W | CPU | 6 | 1 | 5.72 | 80% | NA | NA | 2019-05-11 |  |  |  | In-Use
Alliance Final Tester 3 | 81030 | MX5MM4884M | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2019-01-26 |  |  |  | In-Use | 1
Alliance Final Tester 3 port 1 | 81013 | F00CHM528R | Column Heater | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-02-14 |  |  |  | In-Use | 1
| NA | H1387E735A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-07-11 |  |  |  | In-Use | 1
Alliance Final Tester 3 port 2 | NA | H1387E730A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-12 |  |  |  | In-Use | 1
| 81012 | B03CHM551R | Column Heater | 24 | 1 |  |  |  |  | 2019-02-14 |  |  |  | In-Use | 1
Alliance Final Tester 3 port 3 | NA | H1387E737A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-12 |  |  |  | In-Use | 1
| 81019 | L97CHM711M | Column Heater | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use | 1
Alliance Final Tester 3 port 4 | NA | H1387E734A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-07-19 |  |  |  | In-Use | 1
| 81019 | D01CHM128R | Column Heater | 24 | 1 |  |  |  |  | 2019-02-14 |  |  |  | In-Use | 1
Alliance Final Test 4 | NA | H10L32039W | CPU |  | 1 | 5.72 | 80% | NA | NA | 2019-05-11 |  |  |  | In-Use
Alliance Final Test 4 | 81033 | MX5DM2250R | Temperature Control Module | 7 | 1 |  |  | NA | NA | 2019-04-03 |  |  |  | In-Use | 1
Alliance Final Test 4 Port 1 | 81022 | H01CHM335M | Column Heater | 16 | 1 |  |  | 1.1189 | 74.49% | 2019-02-13 |  |  |  | In-Use | 1
| NA | D1387E766A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  |  |  | 2019-04-06 |  |  |  | In-Use | 1
Alliance Final Test 4 Port 2 | 81017 | MX6KM9444M | Column Heater | 16 | 1 |  |  | 1.1189 |  | 2019-02-14 |  |  |  | In-Use | 1
| NA | H1387E733A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  |  |  | 2019-06-13 |  |  |  | In-Use | 1
Alliance Final Test 4 Port 3 | NA | H1387E732A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  | 1.1189 |  | 2019-06-12 |  |  |  | In-Use | 1
| 81011 | K03CHM938R | Column Heater | 16 | 1 |  |  |  |  | 2019-02-14 |  |  |  | In-Use | 1
Alliance Final Test 4 Port 4 | NA | H1387E736A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  | 1.1189 |  | 2019-06-21 |  |  |  | In-Use | 1
| 81023 | MX5MM5993M | Column Heater | 16 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use | 1
Alliance Final Test 5 | NA | C12L32205W | CPU | 6 | 1 | 5.72 | 80% | NA |  | 2019-05-24 |  |  |  | In-Use
Alliance Final Test 5 | NA | G01TCM633M | Temperature Control Module | 7 | 1 |  |  | #DIV/0! |  | 2019-01-24 |  |  |  | In-Use | 1
Alliance Final Test 5 Port 1 | NA | C13CHM795R | Column Heater | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-07-20 |  |  |  | In-Use | 1
| NA | B1587E610A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-02-20 |  |  |  | In-Use | 1
Alliance Final Test 5 Port 2 | NA | D12CHM299R | Column Heater | 24 | 1 |  |  | 1.1189 |  | 2019-07-20 |  |  |  | In-Use | 1
| NA | B1587E607A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-02-15 |  |  |  | In-Use | 1
Alliance Final Test 5 Port 3 | NA | B1587E608A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-02-15 |  |  |  | In-Use | 1
| NA | K14CHM398R | Column Heater | 24 | 1 |  |  |  |  | 2019-07-20 |  |  |  | In-Use | 1
Alliance Final Test 5 Port 4 | NA | B1587E609A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-02-15 |  |  |  | In-Use | 1
| NA | L14CHM438R | Column Heater | 24 | 1 |  |  |  |  | 2019-07-20 |  |  |  | In-Use | 1
Alliance Final Test 6 | NA | A14L32340W | CPU | 6 | 1 | 5.72 | 80% | NA | NA | 2019-06-13 |  |  |  | In-Use
Alliance Final Test 6 | NA | A18TC2481G | Temperature Control Module 2 | 7 | 1 |  |  | NA | NA | 2019-01-29 |  |  |  | In-Use | 1
Alliance Final Test 6 Port 1 | NA | M17CHM840M | Column Heater | 24 | 1 |  |  | 1.1189 | 74.49% | 2019-06-13 |  |  |  | In-Use | 1
| NA | C1887E101A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-03-21 |  |  |  | In-Use | 1
Alliance Final Test 6 Port 2 | NA | M17CHM834M | Column Heater | 24 | 1 |  |  | 1.1189 |  | 2019-06-13 |  |  |  | In-Use | 1
| NA | C1887E103A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  |  |  | 2019-03-21 |  |  |  | In-Use | 1
Alliance Final Test 6 Port 3 | NA | C1887E106A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-03-20 |  |  |  | In-Use | 1
| NA | M17CHM835M | Column Heater | 24 | 1 |  |  |  |  | 2019-06-13 |  |  |  | In-Use | 1
Alliance Final Test 6 Port 4 | NA | D1887E275A | Dual Wavelength Absorbance Detector | 24 | 1 |  |  | 1.1189 |  | 2019-04-09 |  |  |  | In-Use | 1
| NA | M17CHM839M | Column Heater | 24 | 1 |  |  |  |  | 2019-06-13 |  |  |  | In-Use | 1
Alliance Final Test 6 | NA | B18TC2495G | Temperature Control Module | 7 | 1 |  |  | NA |  | 2019-02-06
Hipot | NA | 19032G000182 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 95% | 69.0909 | 28.95% | 2018-11-09 |  |  |  | In-Use | 2
QA WORKSTATION | NA | NA | CPU | 7 | 1 | 5.72 | 80% | 1.1189 | NA | 2018-12-22 |  |  |  |  | 1
QA WORKSTATION | NA | H14SMC938G | Column Heater | 16 | 1 |  |  |  | NA | 2018-12-22 |  |  |  |  | 1
| NA | H1387E729A | Dual Wavelength Absorbance Detector | 18 | 1 |  |  |  | NA | 2019-07-19 |  |  |  |  | 1
PUMP Cell | NA | 19032G000182 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 95% | 69.0909 | 28.95% | 2018-11-09 |  |  |  | In-Use | 2
1500 SERIES CELL STATION | 90604 | 12326187 | Electronic Weighing Balance | 6 | 2 | 1.58 | 95% | 4.8101 | 103.9% | 2019-01-10 |  |  |  | In-Use | 2
| TE 82705-2 | TE 82705-2 | 1500 Series Final Test Station | 1 | 3 |  |  |  |  | 2018-11-22 | 1. Ripple Test Fail 04 Sep 2018 to 06 Sep 2018 | Replaced Restrictor 1K PSI | NA | In-Use | 3
515 PUMP CELL STATION | 90018 | 12303279 | Electronic Weighing Balance | 6 | 2 | 0.83 | 95% | 9.1566 | 54.61% |  |  |  |  | In-Use | 2
| TE 82534 | TE 82534 | 515 Final Tester | 1 | 3 |  |  |  |  | 2018-11-22 |  |  |  | In-Use | 3
Printer | NA | JPKGJ01942 | LaserJet P3005n | [6 empty] | 2019-08-15 |  |  |  | In-Use | NA
BSM/QSM Cell | 98206 | 6069766 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 98% | 71.2727 | 25.26% | 2018-12-04 |  |  |  | In-Use | 2
BSM/QSM TE82943-11 | 90606 | 12326300 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-02-06 | Missing Segment on Display 20 Jul 2018 to 28 Jul 2018 Reschedule Calibration 28 Jul 18 to 16 Aug 2018 | Calibration – PO raised Jul 2018 PO released 28 Jul 2018 Schedule Calibration : 06 Aug 2018 Calibration Completed Release to Prod’n – 16 Aug 2018 | NA | In-Use | 2
| TE82943-11 | TE-82943-11 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 84010 | M09UPT364A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-03-02 |  |  |  | In-Use | 3
BSM/QSM TE82943-12 | 90605 | 12324733 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
| TE82943-12 | TE82943-12 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 84007 | M09UPT361A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-05-04 |  |  |  | In-Use | 3
BSM/QSM TE82943-13 | 90608 | 12326184 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | Under repair | Unable to get stable reading 09 Jul 2018 - 31 Aug 2018 | 1. Ensure that the bubble spirit level is okay. 2. Re-Zero the weighing balance but to no avail. 3. Ensure that balance stability is not influence by breeze or vibration. 4. Remove weighing pan, Re insert the pan and Re-zero but no avail. 5. Press and hold the Re-Zero key until balance display CAL. Release the RE-Zero Key. Display shows "Error 1". 7. Repair In-progress 8. ETA: WW40'2018 | NA | Not-In-Use | 2
| TE82943-13 | TE82943-13 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 84011 | K09UPT215A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-05-03 |  |  |  | In-Use | 3
BSM/QSM TE82943-14 | NA | 12336741 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2018-11-15 |  |  |  | In-Use | 2
| TE82943-14 | TE82943-14 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 84008 | M09UPT362A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-04-20 |  |  |  | In-Use | 3
BSM/QSM TE82943-15 | 90603 | 12326180 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
| TE82943-15 | TE82943-15 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 84009 | M09UPT363A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-04-21 |  |  |  | In-Use | 3
BSM/QSM TE82943-16 | 90600 | 12326190 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
| TE82943-16 | TE82943-16 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 84006 | M09UPT360A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-05-04 |  |  |  | In-Use | 3
BSM/QSM TE82943-17 | 90607 | 12326279 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-02-06 | 1. Missing Segment on display 20 Jul 2018-28 Jul2018 2. Reschedule calibration 28 Jul 2018 to 16 Aug 2018 | Calibration – PO raised Jul 2018 PO released 28 Jul 2018 Schedule Calibration : 06 Aug 2018 Calibration Completed Release to Prod’n – 16 Aug 2018 |  | In-Use | 2
| TE82943-17 | TE82943-17 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 84005 | M09UPT359A | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-04-21 |  |  |  | In-Use | 3
BSM/QSM TE82943-18 | 90610 | 12326186 | Electronic Weighing Balance | 8 | 2 | 3 | 80% | 2.1333 | 105.6% | 2019-01-10 |  |  |  | In-Use | 2
| TE82943-18 | TE82943-18 | BSM Final Function Tester | 8 | 2 |  |  |  |  | 2019-01-05 |  |  |  | In-Use | 3
| 90610 | 12326186 | QSM_TUV DETECTOR | 8 | 2 |  |  |  |  | 2019-01-10 |  |  |  | In-Use | 3
Acquity Actuator Homing box TE83076-3 | NA | TE83076-3 | Acquity Actuator Homing box | [6 empty] | 2019-07-27
Actuator Tester TE 82800-2 | TE 82800-2 | TE 82800-2 | Aquity Actuator Test | 2 | 2 | 0.23 | 85% | 29.5652 | 101.5% | 2019-02-08 | Fail To initialize 10 Sep 18 to 10 Sep 2018 | Replaced power cable p/n: WAT-WAT270313-03 | NA | In-Use | 2
Actuator Tester TE 82800-3 | TE 82800-3 | TE 82800-3 | Aquity Actuator Test | 2 | 2 | 0.23 | 85% | 29.5652 | 101.5% | 2019-02-08 |  |  |  | In-Use | 2
Printer | NA | PHCNF49758 | LaserJet Pro M402dn | [6 empty] | 2019-08-15 |  |  |  | In-Use | NA
Clean Room | 92258 | 435110 | Picoammeter | 2 | 2 | 0.62 | 95% | 24.5161 | 73.41% | 2019-03-28 |  |  |  | In-Use | 1
Clean Room | 9120000 | NA | Light Meter | 2 | 2 | 0.62 | 95% | 24.5161 | 73.41% | NA |  |  |  | In-Use | 1
Clean Room | TE 82475-2 | TE 82475-2 | Flowcell Tester | 2 | 1 | 0.36 | 95% | 42.2222 | 42.63% | 2019-04-11 |  |  |  | In-Use | 1
Clean Room | TE82881-2 | K15UPD782R | Optic Bench fine alignment Tester (Box Build) | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82881-2 | D08L32121W / C08BLM058C | PDA Fine Alignment CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82880-2 | L07L32070W | PDA ROUGH ALIGNMENT CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82874 | TE82874 | TUV lamp supply | 1 | 3 | 0.28 | 95% | 54.2857 | 33.16% | 2019-09-21 |  |  |  | In-Use | 3
Clean Room | TE82965 | TE82965 | Filter Wheel Alignment | 1 | 3 | 0.28 | 95% | 54.2857 | 33.16% | 2019-09-21 |  |  |  | In-Use | 3
Clean Room | TE9990293134 | NA | Mercury lamp power Supply |  | 1 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use
Clean Room | TE82973-1 (TE829002998) | B08L32248W/ A08BLM491C | Flowcell Tester CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82918-2 | A16VPD196R | 2998 fine alignment(box build) | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82918-2 | B08L32065W | 2998 fine alignment CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82917-1 | B08L32068W | 2998 rough alignment CPU | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE186002998 | G08998943N | Flowcell tester(2998) | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82951-1 | J06L32014W / G06BLM650C | PDA Assembly CPU | 1 | 3 | 0.28 | 95% | 27.1429 | 66.32% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | TE82868-2 | K07L32122W | PDA Aperture Alignment System | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | 2019-08-17 |  |  |  | In-Use | 3
Clean Room | 95515 | RSL 1002609 | Starlite 150 | 1 | 3 | 0.36 | 95% | 42.2222 | 42.63% | 2019-04-24 |  |  |  | In-Use | 3
Clean Room | TE83096-2 | TE83096-2 | PDA Universal Alignment | 1 | 3 | 0.75 | 95% | 20.2667 | 88.82% | NA |  |  |  | In-Use | 3
Engineering Lab | NA | 9093205 | Pressure Calibrator | 1 | 1 | As per usage on evaluation only | NA | NA |  | 2018-09-12 |  |  |  | In-Use | 1
Engineering Lab | FLUKE179 | 90240092 | Multimeter | 1 | 1 |  | NA | NA |  | 2019-01-31 |  |  |  | In-Use | 1
Engineering Lab | High Voltage Probe | 89200011 | High Voltage Probe | 1 | 1 |  | NA | NA |  | 2019-09-25 |  |  |  | In-Use | 1
Engineering Lab | 310788 | 85074 | Picometer |  |  |  | NA | NA |  | 2018-04-17 |  |  |  | In-Use
Engineering Lab | Clamp Meter | 66386 | Clamp Meter | 1 | 1 |  | NA | NA |  | 2019-07-24 |  |  |  | In-Use | 1
Engineering Lab | Fluke 51 II | 89280141 | Digital Thermometer with probe | 1 | 1 |  | NA | NA |  | 2019-07-23 |  |  |  | In-Use | 1
Engineering Lab | FLUKE110 | 85730619 | Multimeter |  | 1 |  | NA | NA |  | 2019-01-05
Engineering Lab | 90406 | 51053 | Liquilaz | 1 | 1 |  | NA | NA |  | 2019-08-17 |  |  |  | In-Use | 1
Engineering Lab | 95037 | 5037655 | Electrical Safety Analyzer |  | 1 |  | NA | NA |  | 2019-01-10
Engineering Lab | 90273 | 12308385 | Electronic Weighing Balance |  |  |  | NA | NA |  | 2018-09-08 |  | Return to WPPL
Engineering Lab | 92207 | 12337042 | Electronic Weighing Balance |  |  |  | NA | NA |  | 2018-09-08 |  | Return to WPPL
Engineering Lab | 90282 | 1140660 | Electrical Safety Analyzer | 4 | 2 | 7 | SCRAP | #VALUE! |  | (RE/084/14 Removed) 9 jun2014 | Inactive-Remove from line (RE/094/13) Date 03/12/13. |  |  | In-Use | 2
DOWNLOADING AND PROGRAMMING | TE 82964-1 | H04BLM012C / J04L32108W | Download and programme(2489 and TUV) | 3 | 2 | 0.36 | 90% | 20 | 90% | 2019-08-03 |  |  |  | In-Use | 3
DOWNLOADING AND PROGRAMMING | TE82972-1 \ TE83057-4 | C08L32078W | Download and programme(2998, PDA and elamda) | 3 | 2 | 0.36 | 90% | 20 | 90% | 2019-08-03 |  |  |  | In-Use | 3
DOWNLOADING AND PROGRAMMING | TE 82969-1 | C08L32208W | Calibration and Linearity test for PDA family | 3 | 1 | 0.36 | 90% | 20 | 90% | 2019-08-03 | 1. TE82969-1 not configure to run Linearity test for PDA 21 Sep 2018 to 24 Sep 2018 | 1. Raise OESR 2. Tester configure by E.T (ADMIN) to be able to do Linearity test for PDA. |  | In-Use
2489/TUV | TE 82903-10 | J04L32102W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
| TE 82903-8 | J04L32120W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
| TE 82903-9 | J04L32091W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
| TE 82903-6 | G05L32009W | Final Tester | 4 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-07-20 |  |  |  | In-Use | 2
2998, PDA eLambda | TE 82882-8 | E13L32018W | Final Tester | 16 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-08-03 |  |  |  | In-Use | 1
| TE 82882-9 | NA | Final Tester | 16 | 1 | 3 | 90% | 9.6 | 65.1% | 2019-05-11 |  |  |  | In-Use | 1
| TE 82882-10 | M12L32412W | Final Tester | 16 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-05-11 |  |  |  | In-Use | 1
| TE 82882-11 | F13L32287W | Final Tester | 16 | 2 | 3 | 90% | 9.6 | 65.1% | 2019-05-11 |  |  |  | In-Use | 1
| TE82882-13 | S4ALMVW | 2998 Final Tester | 16 | 1 | 3 | 90% | 9.6 | 65.1% | 2019-03-15 | Test station test software version 4.1 completed test but not able to print report 21 Sep 2018 - 24 Sep 2018 | 1. Raise OESR 2. E.T. (ADMIN) unblock user access for tester account. | NA | In-Use
| TE82882-14 | S4ALYMH | 2998 Final Tester | 16 | 1 | 3 | 90% | 9.6 | 65.1% | 2019-03-15 |  |  |  | In-Use
Printer | NA | PHCNH24483 | LaserJet Pro M402dn | NA | NA | NA | NA | NA | NA | 2019-08-15 |  |  |  | In-Use | NA
Hipot | 86041 | 08280865 | Electrical Safety Analyzer | 4 | 2 | 0.11 | 98% | 71.2727 | 70.15% | 2018-10-13 |  |  |  |  | 2
Label Printing Area | WATERS |  | WATERSISNL1 | 1 | 3 | NA | NA | NA | NA | NA | NA | NA | NA | In-Use | 3
Label Printing Area | WATERS |  | WATERSPNL1 | 1 | 3 | NA | NA | NA | NA | NA | NA | NA | NA | In-Use | 3
SERVER | Power Edge T320 |  | Waters Server | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
SERVER | NA |  | Flat Panel Monitor | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
SERVER | NA |  | Mouse | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
SERVER | NA |  | Keyboard | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
SERVER | APC UPS |  | UPS | 1 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | Online | 3
Note on F18: Author: 3 = unique tester 2 = multiple tester but will affect output when down 1 = standard tester
Note on L18: skwjdedu:: Period 1st Jan 2011 to 31st Dec 2011
Note on M18: skwjdedu: Period 1st Jan 2011 to 31st Dec 2011
Note on O18: Author: Green - Fixture working Red - Fixture not working
Note on P18: Author: 3 = High Risk, production volume high, will be affected once tester down 2 = Medium Risk, tester down will not affect output for max 2 weeks 1 = Low Risk, tester down will not affect output when down for more than 2 weeks.
Note on E22: Author: 3 for punp tray, 2 for piumps and 1 final test station
Note on E24: Author: 3 for punp tray, 2 for piumps and 1 final test station
Note on E26: Author: 3 for punp tray, 2 for piumps and 1 final test station
Note on E33: Author: 1 down
Note on G35: Author:
Note on E72: Author: 12 working, 4 spare

## Flex Senai
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2 | [6 empty] | 5
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0
IP Address | Computer Name | Computer Description | IP # Assigned | Platform | Serial Number | Mfg Year | RAM (GB) | OS | Symantec AV
10.249.49.10 | TE82720-1 | Detector Front End Board Tester | 2011-10-21 | LAC/E 10 | S018918.1 | - | 1 | Windows XP SP3 | SAV Endpoint 12
10.249.49.11 | TE82940-1 | 2489 Personality Board Debug Tester | 2011-10-21 | LAC/E 12 | 52600013744 | - | 1 | Windows XP SP3 | SAV Endpoint 12
10.249.49.12 | TE82961-2 | 2998 Charge Amp Board Tester | 2011-10-21 | LAC/E 12 | 52600013326 | - | 2 | Windows XP SP2 | SAV Endpoint 12
10.249.49.13 | TE82895-1 | 2998 Personality Board Debug Tester | 2011-10-21 | LAC/E 13 | 52600059037 | - | 2 | Windows XP SP3 | SAV Endpoint 12
10.249.49.19 | TE82988-2 | Waters GTS (Board Functional Tester) | - | GTS | - | - | 2 | Windows XP SP3 | SAV Endpoint 12

## Venture
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 4 | 0 | 4 | 0 | 4 | 0 | 4 | 0 | 4 | 0 | 4 | 0
2 | 25 | 0 | 25 | 0 | 25 | 2 | 25 | 2 | 25 | 0 | 25 | 0
3 | 18 | 0 | 18 | 0 | 18 | 0 | 18 | 1 | 18 | 1 | 18 | 0
Total: | 47 | 0 | 47 | 0 | 47 | 2 | 47 | 3 | 47 | 1 | 47 | 0
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 4 | 0 | 4 | 0 | 4 | 0 | 4 |  | 4 |  | 4
2 | 25 | 2 | 25 | 3 | 25 | 1 | 25 |  | 25 |  | 25
3 | 18 | 1 | 18 | 0 | 18 | 0 | 18 |  | 18 |  | 18
Total: | 47 | 3 | 47 | 3 | 47 | 1 | 47 |  | 47 |  | 47
Locations | STATIONS | Customer ID | Equipment | Number of common testers | Criticalilty | Cycle Time (mins) | % YIELD @ FINAL TEST STATION ONLY | Unit Per Day | Utilization (%) | Calibration Due Date | Reason of Failure and Length of downtime | Actions | Replacement Cost (US$) | STATUS
5042 TEST AREA | FINAL TEST | TE83013-1 | HIGH TEMP COLUMN HEATER TEST | 1 | 3 | 10 | 93 | 48 | 12.2 | 2019-01-02 | Communication test fail, down for 1 hours on 28May18. | Control Board 210000367 | 158.74
5010 TEST AREA | FINAL TEST | TE82956-2 | ACQUITY HTCH TESTER | 1 | 3 | 8 | 94 | 60 | 1.45 | 2019-03-13 | NA | NA | NA
2441 TEST AREA | FINAL TEST | TE82818 | TCM II TESTER | 1 | 3 | 15 | 97 | 32 | 4.2 | 2019-02-28 | NA | NA | NA
2449 TEST AREA | FINAL TEST | TE82817 | PCM II TESTER | 1 | 3 | 15 | 99 | 32 | 1.21 | 2019-02-28 | NA | NA | NA
COLUMN MANAGER TEST AREA | FINAL TEST | TE82984-1 | COLUMN MANAGER TEST SYSTEM | 2 | 3 | 45 | 96 | 10.6667 | 10.38 | 2019-04-19 | NA | NA | NA
COLUMN MANAGER TEST AREA | FINAL TEST | TE82987-1 | EMPOWER WORK STATION | 2 | 3 | 40 | 96 | 12 | 10.38 | 2019-04-19 | NA | NA | NA
1791 TEST AREA | FINAL TEST | 91356 | ELECTRICAL SAFETY ANALYZER | 4 | 1 | 5 | 100 | 96 | NA | 2018-11-15 | NA | NA | NA
1791 TEST AREA | FINAL TEST | TE82774-2 | ALLIANCE COLUMN HEATER/COOLER | 2 | 2 | 30 | 99 | 16 | 46.07 | 2018-10-16 | NA | NA | NA
1791 TEST AREA | FINAL TEST | TE82774-3 | ALLIANCE COLUMN HEATER/COOLER | 2 | 2 | 30 | 99 | 16 | 46.07 | 2018-10-16 | NA | NA | NA
1500 TEST AREA | FINAL TEST | TE82723 | 1500 COLUMN HEATER FINAL TEST | 1 | 3 | 30 | 99 | 16 | 1.67 | 2019-01-02 | NA | NA | NA
RTZ ASSEMBLY AREA | RTZ TEST | TE82995-2 | ACQUITY CPSM RTZ ASSEMBLY TESTER | 2 | 2 | 30 | NA | 16 | NA | 2019-01-16 | UUT cannot complete the testing, down for 4 hours on 13Apr18. | Control Board 210000367 | 158.74
RTZ ASSEMBLY AREA | RTZ TEST | TE82995-3 | ACQUITY CPSM RTZ ASSEMBLY TESTER | 2 | 2 | 30 | NA | 16 | NA | 2019-01-16 | NA | NA | NA
METERING SYRINGE AERA | MSV TEST | TE83006-2 | CPSM METERING SYRINGE TESTER | 1 | 3 | 5 | NA | 96 | NA | 2019-01-02 | NA | NA | NA
FTN ASSEMBLY AREA | FINAL TEST | 91683 | ELECTRICAL SAFETY ANALYZER | 4 | 1 | 5 | 100 | 96 | NA | 2019-01-18 | NA | NA | NA
FTN TEST AREA | FINAL TEST | TE83064-10 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2018-10-04 | NA | NA | NA
FTN TEST AREA | FINAL TEST | TE83064-11 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2018-10-04 | BSM having severe leakage at pump B primary. Problem solve by changing one complete seal set. Station down time is 4 hours on 20Aug18. | WAT076152(CRC O-ring 2.016) 425000930 (Seal Wash .0787 ID) 425001037 (Seal Wash Press Dual Supply) | 133.46
FTN TEST AREA | FINAL TEST | TE83064-12 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2018-11-13 | Due to blockage at Right Valve, Part Number: 289006760 in the Column Manager and leakage at Right Valve Inlet tubing Part Number: 430003154. Station down time is 5 days from 21Aug18 to 29Aug18. | NA | NA
FTN TEST AREA | FINAL TEST | TE83064-13 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2018-10-13 | Severe leakage at BSM pump A accumulator and pump B primary. Down for 4 days from 23Jul18 to 26Jul18. | WAT076152(CRC O-ring 2.016) 425000930 (Seal Wash .0787 ID) 425001037 (Seal Wash Press Dual Supply) 205001672 (Check Valve) 205001476 (ASSY VENT VALVE). | 212.52
FTN TEST AREA | FINAL TEST | TE83064-14 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2019-01-02 | NA | NA | NA
FTN TEST AREA | FINAL TEST | TE83064-15 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2018-12-30 | TE83064-15 was having system pressure unstable on 18Aug18 down for 4 hours due to I2V check valve in the BSM unit at the test stack. | 205001669 (ASSY, CARTRIDGE, I2V, HEXANE/THF) | 59.33
FTN TEST AREA | FINAL TEST | TE83064-16 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2019-01-15 | Pump B of BSM has high leakage replace seal on both Primary and Accumulator pump B. Down for 1 day on 02Jul18. | WAT076152(CRC O-ring 2.016) 425000930 (Seal Wash .0787 ID) 425001037 (Seal Wash Press Dual Supply) | 133.46
FTN TEST AREA | FINAL TEST | TE83064-17 | ACQUITY CPSM FINAL TEST SYSTEM | 8 | 2 | 195 | 55 | 2.4615 | 66.44 | 2019-01-11 | BSM having severe leakage at pump B primary. Problem solve by changing one complete seal set. Station down time is 4 hours on 20Sep18. | WAT076152(CRC O-ring 2.016) 425000930 (Seal Wash .0787 ID) 425001037 (Seal Wash Press Dual Supply) | 133.46
COLUMN MANAGER TEST AREA | FINAL TEST | TE83026-1 | ACQUITY CMA TEST SYSTEM | 2 | 2 | 70 | 82 | 6.8571 | 55.01 | 2019-09-05 | NA | NA | NA
COLUMN MANAGER TEST AREA | FINAL TEST | TE83026-2 | ACQUITY CMA TEST SYSTEM | 2 | 2 | 70 | 82 | 6.8571 | 55.01 | 2019-09-05 | NA | NA | NA
NEW COLUMN MANAGER TEST AREA | FINAL TEST | TE83013-2 | HIGH TEMP COLUMN HEATER TEST | 2 | 2 | 8 | 93 | 60 | 12.2 | 2019-05-31 | NA | NA | NA
W600 TEST AREA | FINAL TEST | TE82824 | W600 TESTER | 1 | 3 | 15 | NA | 32 | NA | 2019-01-02 | NA | NA | NA
DETECTOR TEST AREA | FINAL TEST | 91890 | ELECTRICAL SAFETY ANALYZER | 4 | 1 | 5 | 100 | 96 | NA | 2019-03-14 | NA | NA | NA
CLEAN ROOM | 2475 OPTICS BENCH ALIGNMENT STATION | TE83031 | OPTICS BENCH ALIGNMENT STATION | 1 | 3 | 65 | NA | 7.3846 | NA | 2019-03-30 | NA | NA | NA
CLEAN ROOM | FLR OPTICS BENCH ALIGNMENT STATION | TE83032 | OPTICS BENCH ALIGNMENT STATION | 1 | 3 | 65 | NA | 7.3846 | NA | 2019-03-30 | NA | NA | NA
DETECTOR TEST AREA | FINAL TEST | TE82757 | FIRMWARE DOWNLOAD TESTER | 1 | 2 | 7 | NA | 68.5714 | NA | 2019-01-02 | NA | NA | NA
DETECTOR TEST AREA | FINAL TEST | TE82760-4 | FINAL TEST STATION | 2 | 2 | 75 | 97 | 6.4 | 44.59 | 2019-05-04 | NA | NA | NA
DETECTOR TEST AREA | FINAL TEST | TE83760-5 | FINAL TEST STATION | 2 | 2 | 75 | 97 | 6.4 | 44.59 | 2019-05-11 | NA | NA | NA
BSM/QSM TEST AREA | FINAL TEST | TE82943-1 | ACQUITY SDS FINAL TEST STATION | 6 | 2 | 172 | 56 | 2.7907 | 57.85 | 2018-10-11 | NA | NA | NA
BSM/QSM TEST AREA | FINAL TEST | TE82943-19 | ACQUITY SDS FINAL TEST STATION | 6 | 2 | 172 | 56 | 2.7907 | 57.85 | 2018-10-12 | NA | NA | NA
BSM/QSM TEST AREA | FINAL TEST | TE82943-20 | ACQUITY SDS FINAL TEST STATION | 6 | 2 | 172 | 56 | 2.7907 | 57.85 | 2018-10-10 | NA | NA | NA
BSM/QSM TEST AREA | FINAL TEST | TE82943-21 | ACQUITY SDS FINAL TEST STATION | 6 | 2 | 172 | 56 | 2.7907 | 57.85 | 2018-10-13 | NA | NA | NA
BSM/QSM TEST AREA | FINAL TEST | TE82943-22 | ACQUITY SDS FINAL TEST STATION | 6 | 2 | 172 | 56 | 2.7907 | 57.85 | 2018-10-13 | NA | NA | NA
BSM/QSM TEST AREA | FINAL TEST | TE82943-24 | ACQUITY SDS FINAL TEST STATION | 6 | 2 | 172 | 56 | 2.7907 | 57.85 | 2018-10-12 | NA | NA | NA
ACTUATOR ASSY AREA | SUB TEST | TE83076-1 | ACQUITY ACTUATOR HOMING BOX | 1 | 2 | 2 | NA | 240 | NA | 2018-11-02 | NA | NA | NA
BSM TEST AREA | FINAL TEST | 93467 | ELECTRICAL SAFETY ANALYZER | 4 | 1 | 5 | 100 | 96 | NA | 2019-07-30 | NA | NA | NA
Note on G6: Author: 1: TE 82760-5 down for 2 days from 12 to 13 Mar due to faulty Edge RS232 convertor. 2: Hi-pot tester 95050 during on site calibration discover that AC output is at 1/2 of output setting. No impact to products as AC test is not in use.
Note on I6: Author: 1: Acquity H-Class CH-A Tester, TE83013-1 down for 3 days from 09-Apr to 11-Apr-18 due to faulty PC work station. 2: RTZ tester TE82995-2 down for 4 hours on 13-Apr-18 due to faulty Control Board.
Note on I7: Author: 1: Actuator Homing station, TE83076-1 down for 2 days from 20Ap13 to 23Apr18 due to faulty power supply.
Note on K7: Author: 1: HTCH-A Tester TE83013-1 was having consistent "communication" test failure on 28-May-2018. After verifying, it was confirmed that failure was on 210000367 Rev G in the tester. It was replaced by PCBA 210000367 Rev H. After replacing, "communication" test failure was resolved at the same day.
Note on C13: Author: 1: FTN, TE83064-16, down on 02Jul18, 1 day. Due to BSM high leakage, replace seal on both Primary and Accumulator pump B. 2: FTN, TE83064-13, down on 11Jul18, 1 day. Due to Column 4 in Column Manager Blockage. 3: FTN, TE83064-13, down on 23Jul18 to 26Jul18, 4 days. Due to severe leakage at BSM pump A accumulator and pump B primary.
Note on E13: Author: 1: FTN test station TE83064-15. Station down for 4 hours on 18Aug18. System pressure unstable due to I2V check valve in the BSM unit at the test stack. Change: P/N: 205001669 (289003706) ASSY, CARTRIDGE, I2V, HEXANE/THF. Qty: 1 piece. 2: FTN test station TE83064-11. Station down for 4 hours on 20Aug18. BSM having severe leakage at pump B primary. Problem solve by changing one complete seal set. 3: FTN test station TE83064-12. Station down for 5 days from 21Aug18 to 29Aug18. Due to blockage at Right Valve, Part Number: 289006760 in the Column Manager and leakage at Right Valve Inlet tubing Part Number: 430003154. Part received on 29Aug18.
Note on G13: Author: 1: FTN test station TE83064-17. Station down for 4 hours on 20Sep18. BSM having severe leakage at pump B primary. Problem solve by changing one complete seal set.
Note on C14: Author: CM, TE82987, down on 24Jul18, 1 day. Due to Ethernet Hub faulty
Note on F18: Author: 3 = unique tester 2 = multiple tester but will affect output when down 1 = standard tester
Note on L18: hctan: Period 1st Jan 2018 to 13 Aug 2018.
Note on M18: hctan: Period 01 Jan 2018 to 13 Aug 18.
Note on O18: Author: Green - Fixture working Red - Fixture not working

## Venture PCB
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Criticality | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2 | [6 empty] | 5
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 5 |  | 0 |  | 0
IP # Assigned | Platform | Serial Number | Mfg Year | RAM (GB) | OS | Symantec AV | Remark
2012-06-14 | RACAL | G0300242252-003 | NA | 2 | Windows 2000 | SAV Endpoint 11 | No TeamViewer Installed
2010-05-25 | GTS | NA | NA | 2 | Windows XP SP3 | SAV Endpoint 12 | TeamViewer Account: 551031545 Password: Fct54321
NA | Universal Tester | Old ( faulty): 13726-0202-11-1001 New: 23453-072216-1002 | 5/6/2012 12-Sept-18 | 2 4 | Windows XP SP3 Windows XP SP3 | SAV Endpoint 12 SAV Endpoint 12 | No TeamViewer Installed
NA | Junior GTS | NA | 41079 | 4 | Windows 7 Professional ( SP1 ) | SAV Endpoint 14 | TeamViewer Account:631662116 Password:Jgts54321
NA | GTS | NA | 41079 | 4 | Windows XP SP3 | SAV Endpoint 11 | TeamViewer Account:1036110609 Password:12Gts54321
Note on G6: Author: 1: TE 82760-5 down for 2 days from 12 to 13 Mar due to faulty Edge RS232 convertor. 2: Hi-pot tester 95050 during on site calibration discover that AC output is at 1/2 of output setting. No impact to products as AC test is not in use.
Note on I6: Author: 1: Acquity H-Class CH-A Tester, TE83013-1 down for 3 days from 09-Apr to 11-Apr-18 due to faulty PC work station. 2: RTZ tester TE82995-2 down for 4 hours on 13-Apr-18 due to faulty Control Board.
Note on I7: Author: 1: Actuator Homing station, TE83076-1 down for 2 days from 20Ap13 to 23Apr18 due to faulty power supply.
Note on K7: Author: 1: HTCH-A Tester TE83013-1 was having consistent "communication" test failure on 28-May-2018. After verifying, it was confirmed that failure was on 210000367 Rev G in the tester. It was replaced by PCBA 210000367 Rev H. After replacing, "communication" test failure was resolved at the same day.
Note on C13: Author: 1: FTN, TE83064-16, down on 02Jul18, 1 day. Due to BSM high leakage, replace seal on both Primary and Accumulator pump B. 2: FTN, TE83064-13, down on 11Jul18, 1 day. Due to Column 4 in Column Manager Blockage. 3: FTN, TE83064-13, down on 23Jul18 to 26Jul18, 4 days. Due to severe leakage at BSM pump A accumulator and pump B primary.
Note on E13: Author: 1: FTN test station TE83064-15. Station down for 4 hours on 18Aug18. System pressure unstable due to I2V check valve in the BSM unit at the test stack. Change: P/N: 205001669 (289003706) ASSY, CARTRIDGE, I2V, HEXANE/THF. Qty: 1 piece. 2: FTN test station TE83064-11. Station down for 4 hours on 20Aug18. BSM having severe leakage at pump B primary. Problem solve by changing one complete seal set. 3: FTN test station TE83064-12. Station down for 5 days from 21Aug18 to 29Aug18. Due to blockage at Right Valve, Part Number: 289006760 in the Column Manager and leakage at Right Valve Inlet tubing Part Number: 430003154. Part received on 29Aug18.
Note on G13: Author: 1: FTN test station TE83064-17. Station down for 4 hours on 20Sep18. BSM having severe leakage at pump B primary. Problem solve by changing one complete seal set.
Note on C14: Author: CM, TE82987, down on 24Jul18, 1 day. Due to Ethernet Hub faulty

## CEI SG
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 2 | 0 | 2 | 0 | 2 | 0 | 2 | 0 | 2 | 0 | 2 | 0
2 | 2 | 0 | 2 | 0 | 2 | 0 | 2 | 0 | 2 | 0 | 2 | 0
3 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0
Total: | 5 | 0 | 5 | 0 | 5 | 0 | 5 | 0 | 5 | 0 | 5 | 0
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1 | 2 | 0 | 2 | 0 | 2 | 0
2 | 2 | 0 | 2 | 0 | 2 | 0
3 | 1 | 0 | 1 | 0 | 1 | 0
Total: | 5 | 0 | 5 | 0 | 5 | 0
Locations | STATIONS | Customer ID | Equipment | Number of common testers | Criticalilty | Cycle Time (mins) | % YIELD @ FINAL TEST STATION ONLY | Unit Per Day | Utilization (%) | Calibration Due Date | Reason of Failure and Length of downtime | Replacement Item / Spare Parts | Replacement Cost (US$) | STATUS
SAMPLE ORGANIZER | FINAL TEST | 919032050 | ELECTRICAL SAFETY ANALYZER | 1 | 1 | 5 | 100 | 96 | NA | 2019-03-13
SAMPLE ORGANIZER | FINAL TEST | TE82865-5 | FINAL TESTER | 2 | 2 | 90 | 82 | 5.3333 | 11.15% | N.A.
SAMPLE ORGANIZER | FINAL TEST | TE82865-6 | FINAL TESTER | 2 | 2 | 90 | 82 | 5.3333 | 9.08% | N.A.
SAMPLE ORGANIZER | FINAL TEST | TE82865-2 | FINAL TESTER | 1 | 1 | 120 | 92 | 4 | 5.06% | N.A.
SAMPLE ORGANIZER | Y-AXIS TRANSFER MECHANISM ASSY | TE82869-2 | SHUTTLE TESTER | 1 | 3 | 10 | 100 | 48 | NA | 2019-04-02
SAMPLE ORGANIZER | FINAL TEST | 14354696 | LOAD BOX | 1 | 1 | 5 | 100 | 96 | NA | 2019-04-06
Note on F16: Author: 3 = unique tester 2 = multiple tester but will affect output when down 1 = standard tester
Note on L16: skwjdedu:: Period 1st Jan 2011 to 31st Dec 2011
Note on M16: skwjdedu: Period 1st Jan 2011 to 31st Dec 2011
Note on O16: Author: Green - Fixture working Red - Fixture not working

## CEI Batam
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2 | [6 empty] | 6
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 6
IP Address | Computer Name | Computer Description | IP # Assigned | Platform | Serial Number | Mfg Year | RAM (GB) | OS | Symantec AV
10.249.50.18 | TE82743-2 | RACAL Functional Board Tester | NA | RACAL | G0300242252-004 | 09/26/03 | 2 | Windows 2000 | SAV Endpoint 11
10.249.50.19 | TE82988-4 | Waters GTS (Functional Board Tester) | NA | GTS | 80001101PDSMU99T070 | NA | 2 | Windows XP SP3 | SAV Endpoint 12
192.168.9.15 | Waters GTS | GTS 2 (TE82988-9) | NA | NA | NA | NA | NA | WIN XP SP3 | NA
192.168.9.3 | WAT3 | WAT Tester 3 | NA | NA | NA | NA | NA | WIN 7 SP1 | NA
192.168.9.8 | WAT8 | WAT Tester 8 | NA | NA | NA | NA | NA | WIN 7 SP1 | NA
192.168.9.11 | TE82934-WATPC1 | RF GENERATor Tester (TE82934) | NA | NA | NA | NA | NA | WIN 7 SP1 | NA

## Richport
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2 | [6 empty] | 1
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 1
IP Address | Computer Name | Computer Description | IP # Assigned | Platform | Serial Number | Mfg Year | RAM (GB) | OS | Symantec AV
NA(no internet connection) | RichportAB1 | Richport Auto Bagger 1 | Not connected to net | NA | 76487-OEM-0066135 | 2009 | 1.97 | Ms XP Professional Version 2002, Service pack 3 | NA

## Keypoint
Dashboard Tracker 2018
| JAN |  | FEB |  | MAR |  | APR |  | MAY |  | JUN
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| JUL |  | AUG |  | SEP |  | OCT |  | NOV |  | DEC
Number of common testers | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down | Total Tester | Tester Down
1
2 | [6 empty] | 2
3
Total: | 0 | 0 | 0 | 0 | 0 | 0 | 2
IP Address | Computer Name | Computer Description | IP # Assigned | Platform | Serial Number | Mfg Year | RAM (GB) | OS | Symantec AV
N/A (No Internet connection) | KEYPOINTPNL 1 | KeyPoint Part Number Labelling 1 | NA | Waters Model: D915GUX;Desktop: Intel® Pentium ®4 CPU 3.00GHz x86 | 55274-OEM-0056863-15375 | 2005 | 0.99 | Microsoft Windows XP Professional Version 2002 Service Pack 3 | Ver: 11.0.5002.333
N/A(No Internet connection) | KEYPOINTAB1 | KeyPoint Auto Bagger 1 | NA | Waters Model: LAC/E32 Config#12; Desktop: Intel® Core ™2 CPU 6400@2.13GHz x86 | 76487-OEM-0066135-21407 | 2007 | 1.97 | Microsoft Windows XP Professional Version 2002 Service Pack 3 | Ver: 11.0.5002.333