For the CPSM Final Test System (TE0084050-3)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Rick Davis | Principal Test Engineer |  |  |

The “Reviewed By” signature[s] denotes agreement with the description of the qualification activities.

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Deirdre Donovan | Principal Validation Engineer |  |  |

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Iteration: | 00 | Originator: | Emma Pierce | Change ID: | N/A |
| --- | --- | --- | --- | --- | --- |
| Summary of Change: First document for approval. |  |  |  |  |  |

# PURPOSE

The purpose of this Installation Qualification Report is to detail the activities associated with the execution of the Installation Qualification Protocol 730004135, Ver. 03 for the CPSM Final Test System (TE0084050-3).

## SYSTEM BACKGROUND

The CPSM Final Test System (TE0084050-3) is located in Milford NPI.  The station is used to test FTN, FL, and ARC FTN sample managers listed in the CPSM Final Test Procedure 735002984.  The system performs testing to ensure the product meets the critical to quality specifications in accordance with the appropriate functional specification document (refer to § 9.0 of this report).

The CPSM Final Test System downloads firmware to the UUT, verifies electrical connection and assembly, performs various calibrations, verifies functional mechanical performance and assembly, and verifies chemistry performance.  Upon completion of all testing a Final Test Record is automatically generated and saved to the SigmaSure Data Archive; a copy of the record is printed as part of the Device History Record.

# SCOPE

This Installation Qualification Report details the qualification activities undertaken for the CPSM Final Test System (TE0084050-3), based on the requirements of the executed Installation Qualification Protocol 730004135, Ver. 03 as detailed in the Process Validation Procedure 730001835.

# SUMMARY OF RESULTS – STATEMENT OF FITNESS FOR INTENDED USE

The CPSM Final Test System (TE0084050-3) was tested in accordance with the Installation Qualification Protocol 730004135, Ver. 03.  All test scripts were performed and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the CPSM Final Test System (TE0084050-3) complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Installation Qualification Report, the CPSM Final Test System (TE0084050-3) can be considered fully qualified and fit for intended use.

# ABBREVIATIONS

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| (ARC) FTN | (Acquity Reduced Cost) Flow through Needle |
| BOM | Bill of Materials |
| CAL | Calibration |
| cGMP | Current Good Manufacturing Practice |
| CPSM | Common Platform Sample Manager |
| CTQ | Critical to Quality |
| FL | Fixed Loop |
| IQ(R) | Installation Qualification (Report) |
| N/A | Not Applicable |
| NPI | New Product Introduction |
| OESR | Operations Engineering Service Request |
| P/N | Part Number |
| PM | Preventative Maintenance |
| S/N | Serial Number |
| SAP | Enterprise Resource Planning Software |
| TSP | Test Specification and Procedure |
| UUT | Unit Under Test |

# TRAINING

The executor of this Installation Qualification has been trained to the Installation Qualification Protocol 730004135, Ver. 03.  The operator of the CPSM Final Test System (TE0084050-3) was trained according to the CPSM Final Test Procedure 735002984, Rev. P, Ver. 13.  The executor and operator training records are provided as Attachment 7.1 and Attachment 7.2, respectively.

# QUALIFICATION RESULTS

## Test Script IQ-01: Associated Documentation Verification was performed according to Installation Qualification Procedure 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.3.

#### The CPSM Final Test Procedure 735002984, Rev. P, Ver. 13 and the CPSM Final Test Maintenance Procedure 735003028, Rev. E, Ver. 04 were confirmed to be in SAP with a status of Release General.  Screen captures of each document status are provided as Attachment 7.4.

## Test Script IQ-02: Site Verification was performed according to Installation Qualification Procedure 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.5.

#### The ambient temperature and humidity were measured using a Fluke 971 Temperature/ Humidity Meter (Asset # 99157) and were determined to be within specification as defined by the Acquity Common Platform Sample Manager FTN and FL Functional Specification 721001299, Ver. 12.  The temperature/ humidity meter certificate of calibration is provided as Attachment 7.6.

#### The electrical supply was measured using a Fluke 289 Digital Multimeter (Asset # 99155) and was determined to be within specification as defined by the Acquity Common Platform Sample Manager FTN and FL Functional Specification 721001299, Ver. 12.  The digital multimeter certificate of calibration is provided as Attachment 7.7.

#### The Test Station was confirmed to have the five (5) dedicated electrical outlets and the one (1) network connection required.

#### The space and test bench of the CPSM Final Test System (TE0084050-3) were confirmed to have adequate space for testing and service.  The bench was confirmed to be large and strong enough to support all test equipment.  A picture of the Test Station work area is provided as Attachment 7.8.

## Test Script IQ-03: Instrumentation Calibration Verification was performed according to Installation Qualification Procedure 730004135, Ver. 03.  The completed script is provided as Attachment 7.9.

#### The CPSM Final Test System (TE0084050-3), Waters TUV (Asset # 97038), Waters Acquity Column Manager (Asset # 97037), and Waters Acquity BSM (Asset # 97036) were all confirmed to be in the SAP Calibration Database with a twelve (12) month maintenance cycle. A screen capture from the database entry for each item is provided as Attachment 7.10.

#### The CPSM Final Test Station (TE0084050-3) is newly installed and therefore does not have a recent Calibration/ PM Record.  Note that the date of Cal/ PM Due listed on Test Script IQ-03: Instrumentation Calibration Verification is not a full maintenance cycle from the installation date because the system and its existing stack components are serviced together.

#### The Waters TUV (Asset # 97038), Waters Acquity Column Manager (Asset # 97037), and Waters Acquity BSM (Asset # 97036) in use at this system were used with the previous CPSM Final Test System (TE83064-20) and have had preventative maintenance within the last twelve (12) months.  The most recent Preventative Maintenance Record is provided as Attachment 7.11.

#### Note that the leak test specification is the same as the dynamic leak test specification listed in 735003028. OESR 214531 was submitted to add the limit to the maintenance procedure template upon revision.

#### Note that these components receive preventative maintenance as a stack, not individually, and the date of Last Cal/ PM listed on Test Script IQ-03: Instrumentation Calibration Verification reflects the date the preventative maintenance was completed.

## Test Script IQ-04: Software and Instrument Firmware Version was performed according to Installation Qualification Procedure 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.12.

#### The Test Equipment Software version listed in CPSM Final Test Procedure 735002984, Rev. P, Ver. 13 is Test Software 1.00 and was verified to match the version listed in the header of the Final Test Record.  Refer to Attachment 7.16.

## Test Script IQ-05: General Safety and Alarms was performed according to Installation Qualification Procedure 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.13; refer to the picture provided as Attachment 7.8.

#### It was confirmed that there were no mechanical safety pinch points or sharp edges.

#### It was confirmed that the glass bottles in use at the test station were secured.

#### It was confirmed that the tank in use at the test station was secured.

#### This system is not equipped with an emergency stop or guarding.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 03.  The completed Test Script Data Sheet is provided as Attachment 7.14.

#### One ARC FTN (P/N: 186017001, S/N: A21VSM397N) was tested using the CPSM Final Test System (TE0084050-3) with all tests passing.  The passing test record as printed to the floor and as printed from the SigmaSure Data Archive are provided as Attachment 7.16.

#### The firmware version listed in the BOM in SAP for P/N: 186017001 is version 1.71.164, as shown in the screen capture from SAP provided as Attachment 7.15.  It was verified that the UUT was installed with firmware version 1.71.164; refer to the header information in the Final Test Record provided as Attachment 7.16.

## This table provides a summary of each test script as performed according to 730004135, Ver. 03 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachments 7.3 and 7.4 | None |
| IQ-02 | Pass | Attachments 7.5 through 7.8 | None |
| IQ-03 | Pass | Attachments 7.9, 7.10, and 7.11 | None |
| IQ-04 | Pass | Attachments 7.12 and 7.16 | None |
| IQ-05 | Pass | Attachment 7.13 and 7.8 | None |
| IQ-06 | Pass | Attachments 7.14, 7.15, and 7.16 | None |

# ATTACHMENTS

## Completed Author Training Record

## Completed Operator Training Record

## Completed Test Script IQ-01: Associated Document Verification

## Screen Capture of Each Document Status in SAP

## Completed Test Script IQ-02: Site Verification

## Fluke 971 Temperature/ Humidity Meter (Asset # 99157) Certificate of Calibration

## Fluke 289 Digital Multimeter (Asset # 99155) Certificate of Calibration

## Picture of Test Station Work Area

## Completed Test Script IQ-03: Instrument Calibration Verification

## Screen Captures from the Calibration Database

## Stack Component Preventative Maintenance Record

## Completed Test Script IQ-04: Test Software and Instrument Firmware Version

## Completed Test Script IQ-05: General Safety and Alarms

## Completed Test Script IQ-06: Functional Verification

## BOM in SAP for P/N: 186017001

## Final Test Record as Printed to the Floor and Printed to SigmaSure Data Archive

# DEVIATION FROM THE QUALIFICATION PROTOCOL

There were no deviations generated during the execution of the Installation Qualification Protocol 730004135, Ver. 03.

# REFERENCES

The current approved versions of the following documents are referenced within this document:

| Document Number | Document title |
| --- | --- |
| 735002984 | CPSM Final Test Procedure |
| 730000523 | Good Documentation Practices |
| 730004135 | Installation Qualification Protocol |
| 735003028 | CPSM Final Test Maintenance Procedure |
| 721001299 | Acquity Common Platform Sample Manager FTN and FL Functional Specification |
| 721003992 | Veyron CPSM FTN Functional Specification Document |
| 721007514 | Velox FTN Functional Specification |
| 721013356 | Acquity ARC Bio FTN Functional Specification Document |
| 730000945 | Global Training Policy |
| 730002845 | Validation Report Template |
| 730001835 | Process Validation Procedure |
| OESR 214531 | Service Request for CPSM Maintenance Procedure |