TE0084050 – CPSM Final Test PC Create Procedure

NOTE: Steps 14 and 15 requires the Acquity stack, including an autosampler be connected

- Using Acronis, restore the PC image

- Allow Windows to update

- Update Symantec

- Change passwords

- Add GUID

- Change screen resolution to 1024x768

- Set IE home page to sigmaprod\sigma

- Goto settings and Uncheck “Allow Windows to manage your printers”

- Add the printer you will be using and make it the default printer.

- Open NIMAX, then delete DEV1 and save changes

- Change computer name

- Fix the Oracle Database with file VMOracleRepairTool.exe - must right click on file, then select Run as Administrator

- Configure new Empower Node

  - Login to empower

  - Select <file><new><Chromatographic System>

  - Select Connect to Existing System

  - Select one of the 3 Systems

  - Repeat until you have selected all 3 systems

  - Delete the old Node (only after steps a thru e above)

- Open the test software - – When test software starts an error will be shown. This is OK.

  - Scan for serial ports

  - Change Location

  - Add Operators – this can be performed when deployed at the test location

- Configure Empower Run Samples screens

  - Change stack serial numbers

    - For FTN system A123456789

    - For FL system B123456789

    - For ARC system C123456789

  - Connect an ARC FTN

    - Connect the rear panel cables

    - Connect the pump SS line to the CM SS line with a union

    - Login to Empower as ARC_Autosampler

    - Select Run Samples

    - Make the run samples window almost as large as the screen

    - Line up Instrument Status panels

    - Remove tool bars:

      - Time Remaining

      - Solvent Required

      - Instrument Method

      - Apply Table Preferences

    - Load and run Instrument method Carryover Long Equil

    - When the plots are visible

      - Remove the scrollbars

      - Make the background color black

      - Change the pump Pressure dataline color a bright green

    - Save preferences

  - Repeat for each of the 5 Projects