TE0084050 – CPSM Final Test PC Create Procedure
NOTE: Steps 14 and 15 requires the Acquity stack, including an autosampler be connected

1. Using Acronis, restore the PC image
2. Allow Windows to update
3. Update Symantec 
4. Change passwords
5. Add GUID
6. Change screen resolution to 1024x768
7. Set IE home page to sigmaprod\sigma
8. Goto settings and Uncheck “Allow Windows to manage your printers”
9. Add the printer you will be using and make it the default printer.
10. Open NIMAX, then delete DEV1 and save changes
11. Change computer name
12. Fix the Oracle Database with file VMOracleRepairTool.exe - must right click on file, then select Run as Administrator
13. Configure new Empower Node
a. Login to empower
b. Select <file><new><Chromatographic System>
c. Select Connect to Existing System
d. Select one of the 3 Systems
e. Repeat until you have selected all 3 systems
f. Delete the old Node (only after steps a thru e above)
14. Open the test software - – When test software starts an error will be shown. This is OK.
a. Scan for serial ports
b. Change Location
c. Add Operators – this can be performed when deployed at the test location
15. Configure Empower Run Samples screens
a. Change stack serial numbers
i. For FTN system A123456789
ii. For FL system B123456789
iii. For ARC system C123456789
b. Connect an ARC FTN
i. Connect the rear panel cables
ii. Connect the pump SS line to the CM SS line with a union
iii. Login to Empower as ARC_Autosampler
iv. Select Run Samples
v. Make the run samples window almost as large as the screen
vi. Line up Instrument Status panels
vii. Remove tool bars:
1. Time Remaining
2. Solvent Required
3. Instrument Method
4. Apply Table Preferences
viii. Load and run Instrument method Carryover Long Equil
ix. When the plots are visible 
1. Remove the scrollbars
2. Make the background color black
3. [bookmark: _GoBack]Change the pump Pressure dataline color a bright green
x. Save preferences
c. Repeat for each of the 5 Projects
