TE0084050 – CPSM Final Test PC Create Procedure
NOTE: Steps 14 and 15 requires the Acquity stack, including an autosampler be connected

1. Using Acronis, restore the PC image
2. Configure the Network connection with a defined fixed IP address – use replaced system’s IP address
3. Configure the Instrument LAN – verify, should be part of the image
4. Label the Ethernet ports Instrument LAN and Network
5. Configure Windows for local time zone
6. Allow Windows to update
7. Update Symantec 
8. Change passwords using TE’s conventions and the new computer name
9. Add GUID
10. Change screen resolution to 1024x768 for both users, Tester and TE-1 – verify, should be part of image
11. Set IE home page to sigmaprod\sigma – verify, should be part of image
12. Go to Settings: Devices:Printers & Scanners – verify, should be part of image
a. Uncheck “Allow Windows to manage your printers”
b. Add the printer you will be using and make it the default printer.
c. If listed as a printer, delete OneNote.
13. Open NIMAX, then delete DEV1 and save changes
14. Change computer name
15. Fix the Oracle Database with file VMOracleRepairTool.exe - must right click on file, then select Run as Administrator
16. Configure new Empower Node
a. Login to empower as system
b. Select <file><new><Chromatographic System>
c. Select Connect to Existing System
d. Select one of the 3 Systems
e. Repeat until you have selected all 3 systems
f. Delete the old Node (only after steps a thru e above)
g. Open each system properties
h. Change owner of that system to the System Name
17. Open the test software - – When test software starts an error will be shown. This is OK.
a. Scan for serial ports
b. Change Location in Utilities\Test Specs
c. Add Operators – this should be performed when deployed at the test location
18. Configure Empower Run Samples screens
a. Change stack serial numbers
i. For FTN system A123456789
ii. For FL system B123456789
iii. For ARC system C123456789
b. Connect an ARC FTN
i. Connect the rear panel cables
ii. Connect the pump SS line to the CM SS line with a union
iii. Login to Empower as ARC_Autosampler
iv. Select Run Samples
v. Make the run samples window almost as large as the screen
vi. Line up Instrument Status panels
vii. Remove tool bars:
1. Time Remaining
2. Solvent Required
3. Instrument Method
4. Apply Table Preferences
viii. Load and run Instrument method Carryover Long Equil
ix. When the plots are visible 
1. Remove the scrollbars
2. Make the background color black
3. Change the pump Pressure dataline color a bright green
x. Save preferences
c. Repeat for each of the 5 Projects
