MAIN PANEL

Power up SM via AC relay DaqMx Dig I/O (AC Power control task)

Read in Test Specs from Global File Path

C:\CPSM\Support Files\testspecs.dat

Contents

[Specifications]

HeatingRate = 0.800000

CoolingRate = -0.600000

R Axis LostSteps = 20.000000

Z Axis LostSteps = 20.000000

Zp Axis LostSteps = 20.000000

1stE USL = 0.300000

1stE LSL = -0.300000

Transducer USL = 0.550000

Transducer LSL = 0.450000

DynamicLeakRate USL = 500.000000

DynamicLeakRate LSL = 0.000000

FTN NeedleSeal Force USL = 68.000000

FTN NeedleSeal Force LSL = 48.000000

FTN Sample Syringe USL = 15.000000

FTN Sample Syringe LSL = -15.000000

FTN NeedleSealPosition USL = 58.000000

FTN NeedleSealPosition LSL = 55.000000

FL NeedleVolume USL = 13.000000

FL NeedleVolume LSL = 9.000000

FL LoopVolume USL = 13.000000

FL LoopVolume LSL = 7.500000

FL NeedleSealPosition USL = 61.000000

FL NeedleSealPosition LSL = 57.000000

FL Sample Syringe USL = 40.000000

FL Sample Syringe LSL = -5.000000

Needle Seal Decay USL = 50.000000

Needle Seal Decay LSL = -10.000000

Wash Syringe Decay USL = 30.000000

Wash Syringe Decay LSL = -5.000000

AirSensorVolume USL = 10.000000

AirSensorVolume LSL = 7.000000

BackPressureRegulator USL = 165.000000

BackPressureRegulator LSL = 135.000000

M-Class Needle Volume USL = 13.000000

M-Class Needle Volume LSL = 9.000000

M-Class Loop Volume USL = 7.500000

M-Class Loop Volume LSL = 3.000000

FL System Compliance = 2.000000

H-Class Carryover = 0.128000

H-Class Plus Carryover = 0.060000

I-Class Carryover = 0.040000

M-Class Carryover = 0.128000

FL Carryover = 0.040000

UPC2 Carryover = 0.128000

Rsqrd = 0.999000

RSDArea = 0.500000

RSDHeight = 0.900000

XIntercept USL = 0.200000

XIntercept LSL = -0.200000

ARC Carryover = 0.060000

ARC RSDarea = 0.250000

ARC Dynamic Leak USL = 500.000000

H-Class FTN Jumpers = 129.000000

I-Class FTN Jumpers = 133.000000

Arc FTN Jumpers = 128.000000

M-Class FL Jumpers = 129.000000

I-Class FL Jumpers = 129.000000

Password = "password"

Version = 1.010000

Station Version = "A"

Location = "TE, Milford, US"

Last Changed = "12/10/2020 11:17 AM"

CPSM Serial Port = "COM3"

BSM Serial Port = "COM6"

TUV Serial Port = "COM5"

CM Serial Port = "COM4"

Fill Array with test specs

Fill Global Specs

Fill VISA Resource name Globals (From testspecs.dat)

Fill password station name Globals (From testspecs.dat)

Load system leak rates and fill Globals

C:\CPSM\Support Files\Leak Rates.txt

[ARC]

ARC System Leak Rate = 0.000080

Date = "9/11/2019"

Time = "10:16 AM"

[H-Class]

H-Class System Leak Rate = 0.000180

Date = "9/11/2019"

Time = "10:15 AM"

[I-Class]

I-Class System Leak Rate = 0.000280

Date = "9/11/2019"

Time = "10:14 AM"

Get GUID from probe config file and make Global

Home Solvent Select Valve

Get Station Name

Initialize all serial ports used (from specs file)

Start an event loop that times out every 250 ms

Events:

Timeout - check to see if it's Monday

If it's Monday or no text file(System Status.dat) display status board and build file

C:\CPSM\Support Files\System Status.dat  (Global Variable)

Contents

[STATUS]

Date/Time = 3694422439.399489

#BSM Tube = 9.000000

#CM Tube = 0.000000

Read Global Test Version and display

view report

Launch IE to SigmaSure

exit

Close DaqMx tasks

Close window if executable

ATE

Launch test program (below)

Diagnostics

Launch Diagnostics (below)

Utilities

Launch Utilities (below)

ATE

Select model (FL FTN ARC) or cancel

Launch appropriate ATE vi

Set serial numbers of Stack minus UUT (A, B or C123456789 based on model) *Why different for stack?

Turn on TUV lamp and set column temperature (30)

Initialize

Set all controls to default

Stop Acquity Processes

Delete xml and html files from Report Dir

Clear report footer

Reset Start time

Start DHCP (schtasks?)

Home SSV

Set Location and Test Version

Test for Monday or no file – exit if so

Main

Wait for user to press test button or abort

Test

Scan

Launch Enter Test Info Dialog (PN, SN, FO, OP, TMDs)

Load Runtime Parameters

Depending on part number selected, load runtime parameters (specs).  Although this is in the FTN program, it covers all models, so this vi is used in all 3 test programs.

Look for firmware in SAP and match to current version on system.  If it can’t connect to SAP, it verifies the version you have on system is the version the operator wants.

Electrical Safety

Instruct OP to run and grade Hi-Pot and ground continuity.  Unit is off at this point.

Power on initialization

Show instructions to OP as to connections etc.

Home SSV

Step valve 3 steps (check functionality?)

Initialize UUT

Ask Op to power it up

Look for bootloader – if so need to load fw

If not bootloader, wait for initialization complete (Cap::SetState)

Get fw version, turn on enhanced output display

Look for alarms (Controller.extState)

Fans

Have Op check the 3 fans and check boxes on dialog

Firmware Download

Check serial port, IP address

Ping IP

If bootloader prompt(hard coded) Load firmware

Watch serial port to see when complete

Power-on Initialization

Re-Start UUT

Read serial in every 500ms

Look for bootloader – if so firmware is not loaded

Wait for Console:> prompt

Check version to be sure it’s a SM

Check Pc Bd Rev

Set serial output for enhanced

Wait for init complete (Cap::SetState)

Read firmware Version

Compare to ini file value

Process and return Errors

Door and Lights

Has operator open and close door and reads register

Check Jumpers

Reads jumper registers and verifies

Test Event Inputs

Activates Event Ins with Dig I/O card – reads back w/firmware peek

Test Event Outputs

Sets Event Outs with firmware poke and reads back with Dig I/O card

Test SPI

Reads ID chips on APH and Column heater

Test Leak Detectors

Reads sensor present registers

Simulates leak w/dig I/O

Reads leak detected

Test Ecord

Suspends  ecord to generate alarm – waits for alarm

Turns it back on – waits for present message

Initializes ecord