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SPECIFICATIONS UNIQUE TO H-CLASS FTN SAMPLE MANAGERS
Applies to Part Numbers: 186015017, 186015040, 186015085, 186015086, N2971031, 186018002

TEST STEP

SPECIFICATION

721001299 REFERENCE

Jumper Configuration

Jumpers = 129

PSC - This step verifies the build
Jprocess which is not validated

Needle Seal Calibration

Action — Completes without
errors

51528

Needle Seal Force

48 to 68 Ibs/in

5.15.1.36

Needle Seal Position

55 to 58 mm

5.15.1.41
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SPECIFICATIONS UNIQUE TO I-CLASS FTN SAMPLE MANAGERS
Applies to Part Numbers: 186015046, 186015088

TEST STEP

SPECIFICATION

721001299 REFERENCE

Jumper Configuration

Jumpers = 133

PSC - This step verifies the build
process which is not validated

Needle Seal Calibration

Action — Completes without errors

51528

Needle Seal Force

48 to 68 Ibs/in

5.15.1.36

Needle Seal Position

55to 58 mm

5.15.1.41
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SPECIFICATIONS UNIQUE TO VEYRON FTN (APC) SAMPLE MANAGER

Applies to Part Numbers: 186015051

NOTE: FSD 721003992 Section 1 “The Veyron SM-FTN will conserve all functionality except if stated in the
contents of this FSD performance and functionality of an H-Class FTN”.

TEST STEP SPECIFICATION 721003992 REFERENCE

Jumper Config 129 PSC - This step verifies the build
process which is not validated

Needle Seal Calibration Action — Completes without errors 5.16.2.8

Needle Seal Force 48 to 68 Ibs/in 5.16.1.17

Needle Seal Position 55t0 58 mm 5.16.1.11





