# INSTALLATION QUALIFICATION REPORT

For the manual firmware download performed at the PDA / eLambda Linearity Tester (TE82969-1)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Luo HuaZhen | Senior Test Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |
| Eric Turner | Principal Test Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Senior Operations Quality Engineer |  |  |

##### The “Prepared By” signature denotes that this document accurately describes the qualification activities.

##### The “Reviewed By” signatures denote agreement with the description of the qualification activities.

##### The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Emma Pierce | First document for approval |
| 01 | Luo HuaZhen | Firmware Update |

# Purpose

##### The PDA/ eLambda Linearity Tester (TE82969-1) is located in the Flextronics Manufacturer site in Singapore and is the computer used to download firmware to the following part numbers: 186015034, 186015053, 279002224, and 279002897.  The manual firmware download for these part numbers is unrelated to the Linearity and Wavelength Accuracy tests that are the primary purpose of the PDA/ eLambda Linearity Tester (TE82969-1).  After the instrument firmware download is complete, the UUT undergoes wavelength calibration and linearity testing using custom software at the TE82969 system, then it is moved to the 2998 Final Test System (TE82882) to test noise, drift, and optics

##### This Installation Qualification is being performed as required by the Validation Assessment Form TE82969-06JUL2018-VAF which can be found in the TE0082969QX directory of SAP.  This qualification is only of the manual firmware download steps as specified in the UPC2 PDA Detector and Veyron PDA Detector Final Test Procedure 735002531, Rev. J, Ver. 09.  The linearity test and custom software are not affected and have been qualified; the previous qualification, TE82969-1-IQR_Iteration 00, can be found in the TE0082969QX directory of SAP.

# Scope of Report

##### This Installation Qualification Report details the qualification activities undertaken for the manual firmware download performed at the PDA / eLambda Linearity Tester (TE82969-1) based on the requirements of the Installation Qualification Procedure 730004135, Ver. 00 and TE82969-06JUL2018-VAF.

##### This Installation Qualification Report defines the documents used to qualify the manual firmware download performed at the PDA / eLambda Linearity Tester (TE82969-1).

# Summary of Results - Statement of Fitness for Intended Use

##### The manual firmware download performed at the PDA / eLambda Linearity Tester was tested in accordance with the Installation Qualification Procedure 730004135, Ver. 00 and TE82969-06JUL2018-VAF.  All required Test Scripts were performed and all objective evidence requirements have been met.

##### Waters now has the necessary documentation in place to ensure that the manual firmware download performed at the PDA / eLambda Linearity Tester complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

##### Following the approval of this Qualification Report, the manual firmware download performed at the PDA / eLambda Linearity Tester (TE82969-1) can be considered fully qualified and fit for intended use.

# Glossary

##### The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| ECR | Engineering Change Request |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| N/A | Not Applicable |
| P/N | Part Number |
| PC | Personal Computer |
| PDA | Photo Diode Array |
| SAP | Enterprise Resource Planning Software |
| S/N | Serial Number |
| TSP | Test Specification and Procedure |
| UUT | Unit Under Test |
| VAF | Validation Assessment Form |

# Training

##### The operator of PDA / eLambda Linearity Tester (TE82969-1) was trained according to the UPC2 PDA and Veyron PDA Final Test Procedure 735002531, Ver. 09.  The executor of this qualification was trained according to the Installation Qualification Procedure 730004135, Ver. 00. Both training records are provided as Attachment 7.1 and 7.2, respectively.

# Qualification Results

## Test Script IQ-01: Associated Documentation Verification, Test Script IQ-02: Site Verification, and Test Script IQ-03: Instrumentation Calibration Verification were not performed according to  TE82969-06JUL2018-VAF.

## Test Script IQ-04: Software and Instrument Firmware Version was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.3.

#### The firmware version listed in ECR #9222852 for part numbers 186015034 and 186015053 is 1.70.6367 and was verified to match the version listed on the screen shots of device console.  The screen shots of device console are provided as Attachments 7.4 and 7.5.

## Test Script IQ-05: General Safety and Alarms was not performed according to TE82969-06JUL2018-VAF.

## Test Script IQ-06: Functional Verification was performed according to Installation Qualification Procedure 730004135, Ver. 00.  The completed Test Script Data Sheet is provided as Attachment 7.6.

#### One (1) UUT (P/N: 186015034, S/N: G18C2P393A) was tested using the PDA / eLambda Linearity Tester (TE82969-1) with all tests passing.  Refer to Attachment 7.7 for the passing Test Record.

#### One (1) UUT (P/N: 186015053, S/N: G18ADT145A) was tested using the PDA / eLambda Linearity Tester (TE82969-1) with all tests passing.  Refer to Attachment 7.8 for the passing Test Record.

## This table provides a summary of each test script as performed according to 730004135, Ver. 00 along with the attachment location of all objective evidence.  The table also indicates any deviation taken from the scripts during qualification.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-04 | Pass | Attachments 7.4, and 7.5 | None |
| IQ-06 | Pass | Attachments 7.7 and 7.8 | None |

# Attachments

## Completed Operator Training Record 735002531, Ver. 09

## Completed Executor Training Record 730004135, Ver. 00

## Completed Test Script IQ-04: Test Software and Instrument Firmware Version

## Objective Evidence IQ-04: The Screen Shot for UUT Part #186015034 (G18C2P393A)

## Objective Evidence IQ-04: The Screen Shot for UUT Part #186015053 (G18ADT145A)

## Completed Test Script IQ-06: Functional Verification

## Objective Evidence IQ-06: Passing Test Record for UUT Part #186015034 (G18C2P393A)

## Objective Evidence IQ-06: Passing Test Record for UUT Part #186015053 (G18ADT145A)

# Deviation from Test Equipment Qualification

##### There were no deviations generated during the execution of the Installation Qualification Procedure 730004135, Ver. 00.

# References

##### The current approved versions of the following documents are referenced within this report:

| Document Number | Document Title |
| --- | --- |
| 730004135, Ver. 00 | Installation Qualification Procedure |
| 735002531, Ver. 09 | UPC2 PDA and Veyron PDA Final Test Procedure |
| ###### TE0082969QX | Validation Assessment Form TE82969-06JUL2018-VAF |
| ###### TE0082969QX | Installation Qualification TE82969-1-IQR-Iteration 00 |
| ###### ECR #9222852 | Engineering Change Request for Firmware Release |