DOCUMENT TITLE: ALLIANCE HEATER COOLER TEST CAL/PM PROCEDURE

DOCUMENT TYPE: Calibration Preventive Maintenance

EQUIPMENT NAME: 269x alliance heater cooler test system

EQUIPMENT NUMBER: te0084018

| AUTHOR(S): tyler modelski | DATE: August 28, 2018 |
| --- | --- |
| REVISED BY: N/A | DATE: N/A |

# Equipment Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| N/A | N/A | N/A |

# Disposables Required

| Quantity | Part Number | Description |
| --- | --- | --- |
| N/A | N/A | N/A |

# reference documentation

| Document Number | Description |
| --- | --- |
| 730001799 | Operating Procedure, Calibration and PM |
| 730000295MP | Quality System Document Control |
| 735004582 | Alliance Heater Cooler Test Procedure |
| TE0084018EQ | Alliance Heater Cooler Test System Design Documents |

# DEFINITIONS AND ABBREVIATIONS

| Term | Definition/Abbreviation |
| --- | --- |
| Calibration | A process to ensure that equipment is capable of producing accurate results. This includes verifying measurements and/or making adjustment(s) to ensure measurement accuracy. |
| PM | Preventive Maintenance: The servicing, for the purpose of maintaining equipment in satisfactory working order, by performing systematic inspection, detection, and correction of potential failures. PM activities may include tests, measurements, and parts replacement. |
| NI | National Instruments |

# Procedure

NI-9217 Calibration

Remove the side panel of the tester. Send the NI-9217 Temperature Input Module out for calibration by an accredited calibration house. - NOTE: Unscrew the black wired connector and remove the entire module from the cDAQ-9174 chassis. The NI-9217 module is removable by squeezing the clips on the top and bottom (see picture below).

| Computer Maintenance |  |
| --- | --- |
| Maintain Test Computer per 735003678, Generic PC Maintenance Procedure. | PERFORMED SEE NOTE |

| Vacuum and Cleaning |  |
| --- | --- |
| General Cleaning, inside and out. | PERFORMED SEE NOTE |

| Electrical |  |
| --- | --- |
| Inspect cables for wear or physical damage. Replace as needed. | PERFORMED SEE NOTE |
| Inspect connectors for wear or physical damage. Replace as needed. | PERFORMED SEE NOTE |

| Mechanical |  |
| --- | --- |
| Inspect packaging physical damage/deformation. | PERFORMED SEE NOTE |
| Inspect packaging or sharp edges. | PERFORMED SEE NOTE |
| Check inside and around tester enclosure for loose hardware. | PERFORMED SEE NOTE |

| Functional Check |  |
| --- | --- |
| Run the test using a known good UUT. Verify a passing result is achieved. | PASS FAIL SEE NOTE |

- WATERS CORPORATION

269X Alliance Heater Cooler Test System
CALIBRATION and PREVENTIVE MAINTENANCE RECORD

| Date: | Performed By: |  |
| --- | --- | --- |
| Equipment I.D. | Location: |  |
| Due Date: | Interval (Months): | Next Due Date: |

| Test Equipment Used | Serial Number/Identification Number | Calibration/PM Date |
| --- | --- | --- |

| - SECTION 1: NI-9217 Calibration | - RESULT | - RESULT |
| --- | --- | --- |
| - Step 1.1: Calibration NI-2917 | - Calibration Status Valid | - PASS FAIL SEE NOTE |

| - SECTION 2: Computer Maintenance | - RESULT |
| --- | --- |
| - Step 2.1: Maintain Test Computer per 735003678, Generic PC Maintenance Procedure. | - PERFORMED SEE NOTE |

| - SECTION 3: Vacuum and Cleaning | - RESULT |
| --- | --- |
| - Step 3.1: General Cleaning, inside and out. | - PERFORMED SEE NOTE |

| - SECTION 4: Electrical | - RESULT |
| --- | --- |
| - Step 4.1: Inspect cables for wear or physical damage. Replace as needed. | - PERFORMED SEE NOTE |
| - Step 4.2 Inspect connectors for wear or physical damage. Replace as needed. | - PERFORMED SEE NOTE |

| - SECTION 5: Mechanical | - RESULT |
| --- | --- |
| - Step 5.1: Inspect packaging physical damage/deformation. | - PERFORMED SEE NOTE |
| - Step 5.2: Inspect packaging or sharp edges | - PERFORMED SEE NOTE |
| - Step 5.3: Check inside and around tester enclosure for loose hardware. | - PERFORMED SEE NOTE |

| - SECTION 6: Functional Check | - RESULT |
| --- | --- |
| - Step 6.1: Run the test using a known good UUT. Verify a passing result is achieved. | PASS FAIL SEE NOTE |

- NOTES