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Equipment Required
	Quantity
	Part Number
	Description

	N/A
	N/A
	N/A



Disposables Required
	Quantity
	Part Number
	Description

	N/A
	N/A
	N/A



reference documentation
	Document Number
	Description

	730001799
	Operating Procedure, Calibration and PM

	730000295MP
	Quality System Document Control

	735004582
	Alliance Heater Cooler Test Procedure

	TE0084018EQ
	Alliance Heater Cooler Test System Design Documents



DEFINITIONS AND ABBREVIATIONS
	Term
	Definition/Abbreviation

	Calibration
	A process to ensure that equipment is capable of producing accurate results. This includes verifying measurements and/or making adjustment(s) to ensure measurement accuracy.

	PM
	Preventive Maintenance: The servicing, for the purpose of maintaining equipment in satisfactory working order, by performing systematic inspection, detection, and correction of potential failures. PM activities may include tests, measurements, and parts replacement.

	NI
	National Instruments



Procedure
	NI-9217 Calibration

	Remove the side panel of the tester. Send the NI-9217 Temperature Input Module out for calibration by an accredited calibration house.
NOTE: Unscrew the black wired connector and remove the entire module from the cDAQ-9174 chassis. The NI-9217 module is removable by squeezing the clips on the top and bottom (see picture below). 
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	Computer Maintenance

	Maintain Test Computer per 735003678, Generic PC Maintenance Procedure.
	PERFORMED |_|    SEE NOTE |_|



	Vacuum and Cleaning

	General Cleaning, inside and out.
	PERFORMED |_|    SEE NOTE |_|



	Electrical

	Inspect cables for wear or physical damage. Replace as needed.
	PERFORMED |_|    SEE NOTE |_|

	Inspect connectors for wear or physical damage. Replace as needed.
	PERFORMED |_|    SEE NOTE |_|



	Mechanical

	Inspect packaging physical damage/deformation.
	PERFORMED |_|    SEE NOTE |_|

	Inspect packaging or sharp edges.
	PERFORMED |_|    SEE NOTE |_|

	Check inside and around tester enclosure for loose hardware. 		
	PERFORMED |_|    SEE NOTE |_|



	Functional Check

	Run the test using a known good UUT. Verify a passing result is achieved.
	PASS |_|  FAIL	|_|   SEE NOTE |_|
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	SECTION 1: NI-9217 Calibration
	RESULT
	RESULT

	Step 1.1: Calibration NI-2917
	Calibration Status Valid |_|
	PASS |_|  FAIL	|_|   SEE NOTE |_|



	SECTION 2: Computer Maintenance 
	RESULT

	Step 2.1: Maintain Test Computer per 735003678, Generic PC Maintenance Procedure.
	PERFORMED	|_|   SEE NOTE |_|



	SECTION 3: Vacuum and Cleaning  
	RESULT

	Step 3.1: General Cleaning, inside and out.
	PERFORMED	|_|   SEE NOTE |_|



	SECTION 4: Electrical 
	RESULT

	Step 4.1: Inspect cables for wear or physical damage. Replace as needed.
	PERFORMED	|_|   SEE NOTE |_|

	Step 4.2 Inspect connectors for wear or physical damage. Replace as needed.
	PERFORMED	|_|   SEE NOTE |_|



	SECTION 5: Mechanical 
	RESULT

	Step 5.1: Inspect packaging physical damage/deformation.
	PERFORMED	|_|   SEE NOTE |_|

	Step 5.2: Inspect packaging or sharp edges
	PERFORMED	|_|   SEE NOTE |_|

	Step 5.3: Check inside and around tester enclosure for loose hardware. 		
	PERFORMED	|_|   SEE NOTE |_|




	SECTION 6: Functional Check
	RESULT

	Step 6.1: Run the test using a known good UUT. Verify a passing result is achieved.
	PASS |_|  FAIL	|_|   SEE NOTE |_|





	NOTES
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