###### CHANGE HISTORY:

| Document Iteration | Description of Change |
| --- | --- |
| 00 | Initial Release |

Add rows as necessary.

###### PROJECT INFORMATION:

| Product / Project Name | Project Manager | TE Lead |
| --- | --- | --- |
| 2690/2790 Heater Cooler Replacement Tester | WPPL | Tyler Modelski |

# Introduction

  - This Test Equipment Design Plan under the authority of Waters’ management to detail the testing approach for the 2690/2790 Heater Cooler Tester based on the description of its operation and functions as detailed in the following documents:

735000183, TEST SPEC/PROC, 2690 HEATER/CHILLER ASSY Revision 5

WAT270645, PERF SPEC, 2690 SEPARATIONS MODULE Revision A

810000150, PERF.SPEC.,ALLIANCE SAMPLE HEATER/COOLER, Revision A

WAT270801, PERF SPEC., HEATER/COOLER,2690, Revision B

This document has been written in accordance with the Test Engineering Policy Management Procedure [730001833].

  - SCOPE

The scope of this Test Equipment Design Plan is to provide manufacturing test capability for the 2690/2790 Heater Cooler.  The tester will be located at Flextronics in Singapore  The type of test equipment required will fall under the category Subassembly test.   Data Collection will be implemented.  Defining the scope of this Test Equipment Design Plan will provide the necessary objectives, schedule, budget, requirements, qualification, and deliverables required for releasing the 2690/2790 Heater Cooler Tester, TE0084018.

###### PROPOSED SCHEDULE:

###### The schedule to complete this project can be broken down into several phases. Initial concept, lab experimentation, parts procurement, tester build, software development, debug, test qualification, documentation, (list as necessary), etc.  The following table lists these phases with an approximate time, considering the risks and contingencies of section 8.0.

| Tasks | Provider | Dependencies | Time(wks) |
| --- | --- | --- | --- |
| Initial Concept | TE | Equipment | 1 |
| Part procurement | Internal, external | Funding/Availability | 4 |
| Tester build | TE Technicians | Tech. schedule | 2 |
| Software development | TE | Equipment, TE schedule | 8 |
| Debug | TE | HW & SW complete, UUT | 1 |
| Qualification per 730001835 | TE | Available UUTs | 3 |
| Training Operators | TE | Available Ops/UUTs | 1 |
| Test Specification Procedure | TE | N/A | 1 |
| Calibration and Preventive Maintenance Procedure | TE | N/A | 1 |

| Planned Initial Product Shipment | Est. Start Date | Est. Total Time to Completion |
| --- | --- | --- |
| N/A | Q3-2018 | 15 Wks |

This is a replacement test system.  The tester that was in place failed and was not repairable due to an unavailable instrument.  It was decided to completely redesign the system instead of designing out the unavailable instrument as the system was controlled by a now obsolete embedded control with a hybrid touch screen interface that required an obsolete embedded environment.  While 2690/2790 is an older product it has been recently “refreshed” and is still a high volume product.  The Heater Cooler assemblies are currently being tested in instruments using the final test system, this is not viable long term as it takes extra time and ties up costly chromatographic systems.  The estimated time to completion is considered a low risk to supporting the planned initial shipment, because there is currently a test method in place, and the product is and has been shipping

###### REQUIRED INPUTS

  - This project will require certain key deliverables, provided in a timely manner, in order to accomplish this plan within the projected schedule.

RDE	-Performance Specification

FSD	-Functional Specification

CEA	-Funding Availability

Specifications are available in SAP as this is an established product.  The CEA was generated by WPPL, funding is available.

###### ESTIMATED BUDGET AND HARDWARE REQUIREMENTS:

| Qty | UM | Item Description | Project Phase | Cost Ea | Cost $ |
| --- | --- | --- | --- | --- | --- |
| 1 | EA | Waters Parts | N/A | 1500 | 1500. |
| 2 | EA | External Parts (data acquisition / temperature measurement | N/A | 2000. | 2000. |
| 2 | EA | Sample UUT | N/A | 450. | 400. |
| 2 | EA | PC and peripherals | N/A | 1500. | 1500. |
| 1 | EA | Packaging / fixtures | N/A | 7100. | 7100. |
|  |  | Total Project |  |  | 12500. |

Acceptable options for “Project Phase” are “EM”, “Alpha”, “Beta”, or “PP”.  This tester is a replacement for an existing tester for an existing / shipping product, so project phase does not apply.

###### ABBREVIATIONS AND DEFINITIONS

| Term | Explanation |
| --- | --- |
| ICT | In Circuit Test |
| FT | Functional Test |
| TE | Test Engineering or Test Equipment |
| PCB | Printed Circuit Board |
| RDE | Research and Development Engineering |
| NPI | New Products Introduction |
| CEA | Capital Expenditure Authorization |
| MeOH | Methanol |
| UUT | Unit Under Test |
| PP | Pre-Production |
| RG | Release General |
| PM | Preventative Maintenance |
| MRD | Marketing Requirements Document |
| HW | Hardware |
| SW | Software |
| PRF | Performance Specification |
| TSP | Test Procedure |
| FSD | Functional Specifications Document |
| CTQ | Critical to Quality |
| QSPI | Quad Serial Peripheral Interface |

###### RISKS AND CONTINGENCIES:

###### Scope – Significant changes in project scope will result in changes of this plan, which may delay product release. This may include new features or functionality or significant design changes.

###### Test Equipment – Should any of the required test equipment fail qualification; the product will be delayed until the test equipment passes.

###### Resources – Since resources are shared and not necessarily dedicated, conflicts are expected and will be addressed as they occur. Resources shall be fully familiarized with the associated hardware/software options prior to qualification commencing.

  - Budget – The cost of the test equipment is an estimate. The accuracy may vary during development of the test equipment.

Note: The schedule outlined here assumes no major issues arise during TE development.  Should a failure (or failures) occur, that require a design change to one or more components, the planned schedule will increase by some number of days jeopardizing the project schedule.

###### PROJECT DELIVERABLES:

###### The completion of this project will produce the following deliverables:

(List as necessary)

  - This Document (Test Equipment Design Plan)

  - Test Specification and Procedure revised to support new design

  - Calibration/Maintenance Procedure revised to support new design

  - Test Station and all related equipment Calibrated/PM’d

  - Test Station Qualified

  - Qualification Documents per procedure 730001835

  - TE Design History Record completed and backed up

- TRACEABILITY MATRIX

The section will captures all the requirements that the test equipment will verify on the UUT. This includes but not limited to requirements from RDE, NPI, Manufacturing, Test Engineering and Production.

| Item No. | Feature or Specification | Acceptance criteria | Reference |
| --- | --- | --- | --- |
| 1 | UUT ID | ID correct | PRF 810000150 Rev B, Section 4.2 |
| 2 | External Fans | Correct operation as reported via fan tachometer signal | PRF WAT270801 Rev B, Section 3.1.2, 3.3.2 |
| 3 | Thermal switch | Closed | PRF WAT270801 Rev B, Section 3.3.1 |
| 4 | Watchdog Timer | Functioning | PRF 810000150 Rev A, 2.3.5.1.6 |
| 5 | QSPI | Used to control UUT | PRF 810000150 Rev A, 2.3.5.1 |
| 6 | Temperature Accuracy | +/- 3 degrees C | PRF WAT270801 Rev B Section 3.3.1 |
|  | Precision** | +/- 1 degrees C | PRF WAT270801 Rev B Section 3.3.1 |
|  | Cooling* | Time to set point from 25°C 25°C to 4°C: <= 90 Minutes | PRF WAT270801 Rev B Section 3.3.1 |
|  | Heating* | Time to set point from 25°C to 40°C: <= 60 Minutes | PRF WAT270801 Rev B Section 3.3.1 |
| 7 | Cooling | Ambient - 10°C <= 8 minutes | TSP 735000183 Rev 5 section 3.6 |
| 8 | Heating | (Ambient - 10°C) to ambient <= 7 minutes | TSP 735000183 Rev 5 section 3.6 |

*The heating and cooling specifications referenced in the shaded sections of the traceability matrix above are the “customer facing” specifications as published in the performance specification and marketing documents.  The heating and cooling specifications that are not shaded are those employed by TE82620 and referenced in the test procedure (735000183).  This test system is being designed to replace TE82620, which has reached end of life.  The TE82620 test design employed a time to temperature slope test to verifying the heating / cooling functionality of the UUT.  By performing this test to a much tighter slope than specified by the design specifications, significant test time is saved while still assuring that the UUT performs adequately to meet design specifications, and assure quality. Refer to the Test Method Descriptions document within TE0084018EQ for further explanation of these tests.

The charts on the next page visually depict the slopes used for each test detailed below:

Cooling:

Performance specification: 25°C to 4°C in 90 minutes or less, slope = -0.23333

Test specification: ambient -10°C in 8 minutes or less, slope = -1.25

Heating:

Performance specification: 25°C to 40°C in 60 minutes or less, slope = 0.25

Test specification: (ambient - 10°C) to ambient in 7 minutes or less, slope = 1.43

**A precision test is not run while testing the heater cooler sub-assembly. This is a design evaluation specification, as it is dependent on the firmware control loop which tells the heater cooler to heat and cool in order to maintain a set point temperature. It is also dependent on the final instrument design, which ensures sealing of the entire sample compartment within the Alliance chassis. The manufacturing test of the heater cooler ensure that’s the UUT has been properly assembled. The accuracy test is acceptable to verify that the temperature probe within the unit is functioning properly.  A precision test is not necessary to indicate an issue with the heater cooler itself.

- APPENDICES AND REFERENCES:

  - Subassembly Performance Specification: WAT270801 PERF SPEC., HEATER/COOLER,2690, Revision B

  - PCB Performance Specification: 810000150 PERF.SPEC.,ALLIANCE SAMPLE HEATER/COOLER, Revision A

  - Test Procedure: 735000183 TEST SPEC/PROC,2690 HEATER/CHILLER ASSY, Revision 5

END OF DOCUMENT