# test method validation REPORT

For the 269X Alliance Heater Cooler Test System

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Tyler Modelski | Test Engineer II |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Emma Pierce | Test Validation Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |

The “Prepared By” signature denotes that this document accurately describes the validation activities.

The “Reviewed By” signature[s] denotes agreement with the description of the validation activities.

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Tyler Modelski | First document for approval |

# Purpose

The 269X Alliance Heater Cooler Test System (TE0084018) is located at Flextronics, Singapore.  The 269X Alliance Heater Cooler Test System (TE0084018) is used to perform sub-assembly testing of the 289000560 part number. The test system is comprised of a Alliance Driver / CPU board set and a National Instrument Compact-DAQ Chassis installed with the NI-9217 Temperature Input Module and NI-9482 Relay Output Module. These are connected to a test computer installed with custom software written in National Instruments LabVIEW 2017. The Alliance board set is used to send commands and receive information from the UUT. The NI-9482 Relay Output Module is used as a switch to power on and off the test system and UUT. The NI-9217 Temperature Input Module is used to measure the temperature inside the UUT while it is heating and cooling. Per 730001835, the NI-9217 Temperature Input Module is an external measurement device and is the subject of this TMV report.

There is no operator induced variation when using the temperature input module; therefore, only Appendix A Tolerance Budgeting was performed to validate the NI-2917 module for use in TE0084018.

# Scope of Report

This Test Method Validation Report details the Validation activities undertaken for the 269X Alliance Heater Cooler Test System based on the requirements of the executed Test Method Validation protocol, 730001936 Version 03.

This TMV Report defines the documents used to validate the NI-9217 Temperature Input Module for use within the 269X Alliance Heater Cooler Test System (TE0084018)

# Summary of Results - Statement of Fitness for Intended Use

The NI-9217 Temperature Input Module was documented in accordance with the Test Method Validation protocol, 730001936 Version 03.  All requirements were documented and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the NI-9217 Temperature Input Module used as part of the TE0084018 system complies with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Test Method Validation Report, the NI-9217 Temperature Input Module can be considered fully validated and fit for intended use within the 269X Alliance Heater Cooler Test System (TE0084018).

# Glossary

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| Cal/PM | Calibration/Preventive Maintenance |
| cGMP | Current Good Manufacturing Practice |
| CTQ | Critical To Quality |
| NI | National Instruments |
| TMV | Test Method Validation |
| TSP | Test Specification and Procedure |
| USL | Upper Specification Limit |
| LSL | Lower Specification Limit |
| UUT | Unit Under Test |

# Training

The executor of the Test Method Validation protocol, 730001936 Version 03 was trained to the protocol. Refer to Attachment #1 for a screenshot of the training database.

# Test Method Validation Results

The table below summarizes the results of the Tolerance Budgeting performed as part of Appendix A of the Test Method Validation Protocol, 730001936. Below the table is a summary of the result for each requirement.

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| Range | Pass | Attachment #3 | None |
| Resolution | Pass | Attachment #3 | None |
| Accuracy | Pass | Attachment #3 | None |

# Range:

The required range to assess the critical to quality specification is 4°C to 40°C. The test system must be capable of measuring temperature throughout the range of the operating temperature of the UUT (refer to performance specification WAT270801).

The NI-9217 Temperature Input Module has a temperature measurement range of -200°C to 850°C. This is module is more than capable of measuring the entire operating temperature of the UUT. A copy of the NI-9217 performance datasheet can be found as Attachment #3.

Resolution:

The required resolution to assess the critical to quality specification is 0.01°C. 0.01°C provides a detectable change that is acceptable to measure the two digit resolution of the heating and cooling specifications (-1.25°C/min for cooling and 1.43°C/min for heating) as documented in TSP 735004582.

The NI-9217 Temperature Input Module provides 0.003°C resolution in high resolution mode which meets the requirement. A copy of the NI-9217 performance datasheet can be found as Attachment #3 specifying the resolution of the module.

Accuracy:

The required accuracy to assess the critical to quality specification is +/- 10% of the accuracy range required for the Heater Cooler Probe within the UUT (+/- 3°C). This indicates that the temperature may be inaccurate by up to 0.3°C. The 10% accuracy allows for reasonable margin on the accuracy test.

The NI-9217 Temperature Input Module provides 0.20°C accuracy using a 3-wire RTD at ambient temperature, which is better than the required accuracy.

A copy of the NI-9217 performance datasheet can be found as Attachment #3 specifying the resolution of the module.

# Attachments

Attachment #1: Screenshot of Training Database for 730001936 Version 03

Attachment #2: Completed Appendix A: Tolerance Budgeting for the NI 2917 C Series Temperature Input Module

Attachment #3: NI 2917 C Series Temperature Input Module Datasheet

# Deviation from Test Method Validation

There were no deviations generated during the execution of the Test Method Validation Protocol 730001936 Version 03.

# References

The current approved versions of the following documents are referenced within this document:

| Document Number | Document title |
| --- | --- |
| 730001936 | Test Method Validation Protocol |
| WAT270801 | Performance Specification, Heater/Cooler, 2690 |
| 735004582 | Alliance Heater Cooler Test Procedure |