# Installation qualification REPORT

For the 2998 Final Test System (TE82882-15)

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| James J. Stors | Senior Test Validation Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Eric Turner | Principal Test Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Andre Kopoyan | Senior Operations Quality Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

The “Reviewed By” signature[s] denotes agreement with the description of the qualification activities.

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Emma Pierce | First document for approval. |
| 01 | James J. Stors | New software version to print serial number in the Report footer. |
| 02 | James J. Stors | New software version to add Optics Bench part numbers |

# Purpose

The 2998 Final Test System, TE82882-15, is located in Warranty and Repair in Milford, MA.  The 2998 Final Test System is used to perform functionality testing of Waters Corporation’s 2998 Photo Diode Arrays (PDA’s) as defined in TSP 735001613, to ensure compliance to the defined Manufacturing Test Specifications.

The system uses serial ports to communicate to up to four (4) UUT’s in order to test Noise, Drift, and Optics. Upon completion a Final Test Record is automatically generated and a hard copy is printed. In addition, the Final Test Record is sent to the SigmaProd Server.

# Scope of Report

This Installation Qualification Report details the qualification activities undertaken for the 2998 Final Test System 
(TE82882-15) based on the requirements of the performed Installation Qualification, 730004135, Version 02, as outlined by the Test Software Version 4.1 risk assessment, per TE82882-10JUN19-VAF.

All Test Scripts were executed and are summarized within this report.

# Summary of Results - Statement of Fitness for Intended Use

The 2998 Final Test System (TE82882-15) was tested in accordance with Installation Qualification Protocol, 730004135, Version 02. All required Test Scripts were performed and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the 2998 Final Test System (TE82882-15) remains compliant with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Qualification Report, the 2998 Final Test System (TE82882-15) remains fully qualified and fit for intended use.

# Glossary

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| ºC | Degrees Celsius |
| BOM | Bill of Material |
| cGMP | Current Good Manufacturing Practice |
| Hz | Hertz |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| N/A | Not Applicable |
| SAP | Enterprise Resource Planning (ERP) software |
| TSP | Test Specification and Procedure |
| UUT | Unit Under Test |
| VAC | Volts Alternating Current |

# Training

The executor of the 730004135, Installation Qualification Procedure, is the author and is therefore trained to the procedure.  The operator who performed the testing per IQ-06, Functional Verification, per the PDA Final Test - Test and Specification Procedure, 735001613 and the UPC2 PDA Detector and Veyron PDA Detector Final Test - Test and Specification Procedure, 735002531, is the author and is therefore trained to the procedures.

# Qualification Results

## Test Script IQ-01, Associated Documentation Verification, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02.

### IQ-01, Associated Documentation Verification Data Sheet, is provided as Attachment 7.1.

### The PDA Final Test - Test and Specification Procedure, 735001613, Version 22, was confirmed to be in SAP with a status of Ready. A screen capture is provided below:

### The UPC2 PDA Detector and Veyron PDA Detector Final Test - Test and Specification Procedure, 735002531, Version 11, was confirmed to be in SAP with a status of Ready. A screen capture is provided below:

### The PDA Linearity Test System Preventive Maintenance Procedure, 735003030, Version 00/Revision A, was confirmed to be in SAP with a status of Release General. A screen capture is provided below:

## Test Script IQ-02, Site Verification, was performed in accordance with Installation Qualification Procedure, 730004135, Version 02.

### IQ-02, Site Verification Test Script Data Sheet, is provided as Attachment 7.2.

### The site verification acceptance criteria for temperature and humidity are specified within the Velox 2998 PDA Detector Functional Specification Document, 721007611 within Section 5.5, Environmental.

### The ambient temperature and Relative Humidity were measured using a Fluke Model 971 Temperature/Humidity Meter, Control Number: 99157. The ambient temperature and relative humidity were in specification.

#### The calibration data sheet for the Temperature/Humidity Meter is provided as Attachment 7.3.

### The system space and the test bench used for the TE82882-15, 2998 Final Test System, were inspected and confirmed to have sufficient room for the system, testing and service, and be large and strong enough to support the station.

#### A photo of the TE82882-15, 2998 Final Test System is shown below:

### The site verification acceptance criterion for electrical supply is specified within the Velox 2998 PDA Detector Functional Specification Document, 721007611 within Section 5.2.2, Electrical.

### The Electrical Supply was measured using a Fluke 289 Digital Multimeter, Control Number: 99155. The voltage and frequency were in specification.

#### The calibration data sheet for the Multimeter is provided as Attachment 7.4.

### The station was confirmed to meet the requirement of having a minimum of six (6) electrical outlets dedicated.

### The station was confirmed to meet the requirement of having a minimum of one (1) network dedicated.

## Test Script IQ-03, Instrumentation Calibration Verification, was performed in accordance with Installation Qualification Procedure, 730004135, Version 02.

### IQ-03, Instrumentation Calibration Verification Test Script Data Sheet, is provided as Attachment 7.5.

### The TE82882-15, 2998 Final Test System, was confirmed to have had a preventive maintenance performed and to be in the Calibration Manager Database with a twelve (12) month maintenance cycle.

#### A copy of the Preventive Maintenance Certificate is provided as Attachment 7.6.

#### A screen capture of the Calibration Manager Database entry for the TE82882-15, 2998 Final Test System, with a twelve (12) month preventive maintenance cycle, is shown below:

## Test Script IQ-04, Test Software and Instrument Firmware Version, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02.

### IQ-04, Test Software and Instrument Firmware Version Data Sheet, is provided as Attachment 7.7.

### The Test Equipment Software version listed in 735001613, PDA Final Test - Test and Specification Procedure (Version 22), is 4.2 and was verified to match the version displayed on the TE82882-15, 2998 Final Test System program display, as shown in the screen capture below:

## Test Script IQ-05, General Safety and Alarms, was performed in accordance with Installation Qualification Procedure, 730004135, Version 02

### IQ-05, General Safety and Alarms Test Script Data Sheet, is provided as Attachment 7.8.

### It was confirmed that there were no mechanical safety, pinch points, or sharp edges.

## Test Script IQ-06, Functional Verification, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02.

### IQ-06, Functional Verification Test Script Data Sheet, is provided as Attachment 7.9.

### To cover all system functionality one of each type of Detector that the TE82882-15, 2998 Final Test System tests was run.  The 700012243 was selected for the 2998 Detector type, the 700012245 was selected for the eLamda type, and 186015032 was selected for the PDA type.

### The Firmware Version specified within the BOM in SAP for each part number that the TE82882-15, 2998 Final Test System tests was verified to be Version 1.70.6367, as shown in the representative screen capture for part number 186015032 provided below:

### One (1) UUT, Part Number 700012243, was tested using TE82882-15, 2998 Final Test System, and met the requirement of all required CTQ Specification measurements result in a Pass and Firmware Version 1.70.6367 being loaded on the unit.

#### The Instrument Test Record is provided as Attachment 7.10.

### One (1) UUT, Part Number 700012245, was tested using the TE82882-15, 2998 Final Test System, and met the requirement of all required CTQ Specification measurements result in a Pass and Firmware Version 1.70.6367 being loaded on the unit.

#### The Instrument Test Record is provided as Attachment 7.11.

### One (1) UUT, Part Number 186015032, was tested using the TE82882-15, 2998 Final Test System, and met the requirement of all required CTQ Specification measurements result in a Pass and Firmware Version 1.70.6367 being loaded on the unit.

#### The Instrument Test Record is provided as Attachment 7.12.

## A summary of the attached objective evidence that applies to each Test Script is provided in the Table below:

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachment 7.1 | None |
| IQ-02 | Pass | Attachments 7.2, 7.3, and 7.4 | None |
| IQ-03 | Pass | Attachment s7.5 and 7.6 | None |
| IQ-04 | Pass | Attachment 7.7 | None |
| IQ-05 | Pass | Attachments 7.8 | None |
| IQ-06 | Pass | Attachment 7.9 through 7.12 | None |

# Attachments

# IQ-01, Associated Documentation Verification Test Script Data Sheet

## IQ-02, Site Verification Test Script Data Sheet

## Fluke Model 971 Temperature/Humidity Meter, Control Number: 99157, Calibration Data Sheets

## Fluke 289 Multimeter, Control Number: 99155, Calibration Data Sheets

## IQ-03 – Instrumentation Calibration Verification Test Script Data Sheet

## TE82969-2 Preventive Maintenance Certificate

## IQ-04, Test Software and Instrument Firmware Version Test Script Data Sheet

## IQ-05 – General Safety and Alarms Test Script Data Sheet

# IQ-06, Functional Verification Test Script Data Sheet

## Instrument Final Test Record, Part Number: 700012243

## Instrument Final Test Record, Part Number: 700012245

## Instrument Final Test Record, Part Number: 186015033

# Deviation from Test Equipment Qualification

There were no deviations generated during the execution of the Installation Qualification, 730004135, Version 02, for the TE82882-15, 2998 Final Test System.

# References

The current approved versions of the following documents are referenced within this Report:

| Document Number | Document title |
| --- | --- |
| 721007611 | Velox 2998 PDA Detector Functional Specification Document |
| 730004135 | Installation Qualification Protocol |
| 735001613 | PDA Final Test System Test and Specification Procedure |
| 735002531 | UPC2 PDA Detector and Veyron PDA Detector Final Test Procedure |
| 735003030 | PDA Linearity Test System Preventive Maintenance Procedure |
| TE82882-10JUN19-VAF | Test Software Version 4.2 risk assessment |