# Installation qualification REPORT

For the PDA/eLambda Linearity Tester

PREPARED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| James J. Stors | Senior Test Validation Engineer |  |  |

REVIEWED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Eric Turner | Principal Test Engineer |  |  |

APPROVED BY

| Name | Role | Signature | Date |
| --- | --- | --- | --- |
| Megan Marshall | Senior Quality System Engineer |  |  |

The “Prepared By” signature denotes that this document accurately describes the qualification activities.

The “Reviewed By” signature[s] denotes agreement with the description of the qualification activities.

The “Approved By” signature denotes that this document has been reviewed and approved for overall quality and completeness.

REVISION HISTORY

| Version | Author | Summary of change |
| --- | --- | --- |
| 00 | Emma Pierce | First document for approval. |
| 01 | James J. Stors | New software version to add the 30mg concentration test. |
| 02 | Emma Pierce | Firmware update |
| 03 | James J. Stors | New software version to add Optics Bench part numbers |

# Purpose

The PDA/ eLambda Linearity Test System (TE82969-2) is located in the Warranty and Repair Area in Milford, MA.  The system shares a bench and a monitor with the 2998 Final Test System, TE82882.  This station is used to verify that linearity of PDA Detectors of the following part numbers is within specification: 186002998, 186015026, 186015030, 186015032, 186015033, 186015034, 186015053, 186017003, 186017013, 186299800, 279002655, N2971013, and N2971035; as well as the spares and refurbished equivalent parts as listed in the PDA Final Test TSP, 735001613, or the UPC2 PDA Detector and Veyron PDA Detector Final Test Procedure, 735002531.  Note that not all part numbers have a refurbished equivalent or spare.  After the UUT has completed the firmware download according to the corresponding procedure and has passed the linearity test at TE82969, it is moved to the 2998 Final Test System (TE82882) to test noise, drift, and optics.

# Scope of Report

This Installation Qualification Report details the qualification activities undertaken for the PDA/eLambda Linearity Test System (TE82969-2) based on the requirements of the performed Installation Qualification, 730004135, Version 02, as outlined by the Test Software Version 2.2 risk assessment, per TE82969-10JUN19-VAF.

All Test Scripts were executed and are summarized within this report.

# Summary of Results - Statement of Fitness for Intended Use

The PDA/eLambda Linearity Test System (TE82969-2) was tested in accordance with Installation Qualification Protocol, 730004135, Version 02. All required Test Scripts were performed and all objective evidence requirements have been met.

Waters now has the necessary documentation in place to ensure that the PDA/eLambda Linearity Test Ststem (TE82969-2) remains compliant with Waters’ requirements, the latest relevant quality system regulations, and cGMP practices.

Following the approval of this Qualification Report, the PDA/eLambda Linearity Test System (TE82969-2) remains fully qualified and fit for intended use.

# Glossary

The following definitions are used in this document:

| Term | Explanation |
| --- | --- |
| cGMP | Current Good Manufacturing Practice |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| mg | Milligram |
| PDA | Photo Diode Array |
| SAP | Enterprise Resource Planning (ERP) software |
| TSP | Test Specification and Procedure |
| UPC | UltraPerformance Convergence Chromatography |
| UUT | Unit Under Test |
| VAF | Validation Assessment Form |

# Training

The executor of the 730004135, Installation Qualification Procedure, is the author and is therefore trained to the procedure.  The operator who performed the testing per IQ-06, Functional Verification, per the PDA Final Test - Test and Specification Procedure, 735001613 and the UPC2 PDA Detector and Veyron PDA Detector Final Test - Test and Specification Procedure, 735002531, is the author and is therefore trained to the procedures.

# Qualification Results

## Test Script IQ-01, Associated Documentation Verification, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02.

### IQ-01, Associated Documentation Verification Data Sheet, is provided as Attachment 7.1.

### The PDA Final Test - Test and Specification Procedure, 735001613, Version 22, was confirmed to be in SAP with a status of Ready. A screen capture is provided below:

### The UPC2 PDA Detector and Veyron PDA Detector Final Test - Test and Specification Procedure, 735002531, Version 11, was confirmed to be in SAP with a status of Ready. A screen capture is provided below:

### The PDA Linearity Test System Preventive Maintenance Procedure, 735003437, Version 00, was confirmed to be in SAP with a status of Release General. A screen capture is provided below:

## Test Script IQ-02, Site Verification, was performed in accordance with Installation Qualification Procedure, 730004135, Version 02.

### IQ-02, Site Verification Test Script Data Sheet, is provided as Attachment 7.2.

### The site verification acceptance criteria for temperature and humidity are specified within the Velox 2998 PDA Detector Functional Specification Document, 721007611 within Section 5.5, Environmental.

### The ambient temperature and Relative Humidity were measured using a Fluke Model 971 Temperature/Humidity Meter, Control Number: 99157. The ambient temperature and relative humidity were in specification.

#### The calibration data sheet for the Temperature/Humidity Meter is provided as Attachment 7.3.

### The system space and the test bench used for the TE82969-2, PDA/eLambda Linearity Test System, were inspected and confirmed to have sufficient room for the system, testing and service, and be large and strong enough to support the station.

#### A photo of the TE82969-2, PDA/eLambda Linearity Test System is shown below:

### The site verification acceptance criterion for electrical supply is specified within the Velox 2998 PDA Detector Functional Specification Document, 721007611 within Section 5.2.2, Electrical.

### The Electrical Supply was measured using a Fluke 289 Digital Multimeter, Control Number: 99155. The voltage and frequency were in specification.

#### The calibration data sheet for the Multimeter is provided as Attachment 7.4.

### The station was confirmed to meet the requirement of having a minimum of three (3) electrical outlets dedicated.

### The station was confirmed to meet the requirement of having a minimum of one (1) network dedicated.

## Test Script IQ-03, Instrumentation Calibration Verification, was performed in accordance with Installation Qualification Procedure, 730004135, Version 02.

### IQ-03, Instrumentation Calibration Verification Test Script Data Sheet, is provided as Attachment 7.5.

### The TE82969-2, PDA/eLambda Linearity Test System, was confirmed to have had a preventive maintenance performed and to be in the Calibration Manager Database with a twelve (12) month maintenance cycle.

#### A copy of the Preventive Maintenance Certificate is provided as Attachment 7.6.

#### A screen capture of the Calibration Manager Database entry for the TE82969-2, PDA/eLambda Linearity Test System, with a twelve (12) month preventive maintenance cycle, is shown below:

## Test Script IQ-04, Test Software and Instrument Firmware Version, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02.

### IQ-04, Test Software and Instrument Firmware Version Data Sheet, is provided as Attachment 7.7.

### The Test Equipment Software version listed in 735001613, PDA Final Test - Test and Specification Procedure (Version 22), is 2.2 and was verified to match the version displayed on the PDA/eLambda Linearity Tester TE82969 program display, as shown in the screen capture below:

## Test Script IQ-05, General Safety and Alarms, was performed in accordance with Installation Qualification Procedure, 730004135, Version 02

### IQ-05, General Safety and Alarms Test Script Data Sheet, is provided as Attachment 7.8.

### It was confirmed that there were no mechanical safety, pinch points, or sharp edges.

## Test Script IQ-06, Functional Verification, was performed in accordance with Installation Qualification Protocol, 730004135, Version 02.

### IQ-06, Functional Verification Test Script Data Sheet, is provided as Attachment 7.9.

### One (1) UUT, Part Number 700012243, was tested using the TE82969-2, PDA/eLambda Linearity Test System, and met the requirement of all required CTQ Specification measurements result in a Pass.

#### The Instrument Test Record is provided as Attachment 7.10.

### One (1) UUT, Part Number 700012245, was tested using the TE82969-2, PDA/eLambda Linearity Test System, and met the requirement of all required CTQ Specification measurements result in a Pass.

#### The Instrument Test Record is provided as Attachment 7.11.

### One (1) UUT, Part Number 186015032, was tested using the TE82969-2, PDA/eLambda Linearity Test System, and met the requirement of all required CTQ Specification measurements result in a Pass.

#### The Instrument Test Record is provided as Attachment 7.12.

### The above part numbers were selected for testing to cover all three Detector types that the TE82969-2, PDA/eLambda Linearity Test System tests.  The 700012243 was selected for the 2998 Detector type, the 700012245 was selected for the eLamda type, and 186015032 was selected for the PDA type.

## A summary of the attached objective evidence that applies to each Test Script is provided in the Table below:

| Script or Section | Pass/Fail | Objective Evidence | Deviations |
| --- | --- | --- | --- |
| IQ-01 | Pass | Attachment 7.1 | None |
| IQ-02 | Pass | Attachments 7.2, 7.3, and 7.4 | None |
| IQ-03 | Pass | Attachment 7.5 and 7.6 | None |
| IQ-04 | Pass | Attachment 7.7 | None |
| IQ-05 | Pass | Attachments 7.8 | None |
| IQ-06 | Pass | Attachment 7.9 through 7.12 | None |

# Attachments

# IQ-01, Associated Documentation Verification Test Script Data Sheet

## IQ-02, Site Verification Test Script Data Sheet

## Fluke Model 971 Temperature/Humidity Meter, Control Number: 99157, Calibration Data Sheets

## Fluke 289 Multimeter, Control Number: 99155, Calibration Data Sheets

## IQ-03 – Instrumentation Calibration Verification Test Script Data Sheet

## TE82969-2 Preventive Maintenance Certificate

## IQ-04, Test Software and Instrument Firmware Version Test Script Data Sheet

## IQ-05 – General Safety and Alarms Test Script Data Sheet

# IQ-06, Functional Verification Test Script Data Sheet

## Instrument Final Test Record, Part Number: 700012243

## Instrument Final Test Record, Part Number: 700012245

## Instrument Final Test Record, Part Number: 186015033

# Deviation from Test Equipment Qualification

There were no deviations generated during the execution of the Installation Qualification Protocol, 730004135, Version 02, for the TE82969-2, PDA/eLambda Linearity Test System.

# References

The current approved versions of the following documents are referenced within this Report:

| Document Number | Document title |
| --- | --- |
| 721007611 | Velox 2998 PDA Detector Functional Specification Document |
| 730004135 | Installation Qualification Protocol |
| 735001613 | PDA Final Test System Test and Specification Procedure |
| 735002531 | UPC2 PDA Detector and Veyron PDA Detector Final Test Procedure |
| 735003437 | PDA Linearity Test System Preventive Maintenance Procedure |
| TE82969-10JUN19-VAF | Test Software Version 2.2 risk assessment |