# Purpose

The purpose of this Installation Qualification (IQ) protocol is to provide a series of Test Scripts to be used to qualify the installation of any Milford site process equipment and ancillary system to ensure compliance to manufacturer’s guidelines and/or Waters requirements.

Once completed, the applicable executed IQ protocol Test Script Data Sheets will be included within an Installation Qualification Report (IQR), which will serve as objective evidence that all key aspects of the equipment and ancillary system are correctly installed.

# scope

This Installation Qualification protocol shall be used for all equipment within a method or process that makes a decision which can affect product, or that the improper installation and/or configuration would impact the assessment or production of product, as defined in 730001835, Process Validation Procedure for Milford Site.

# ΔAbbreviations

| Term | Definition / Abbreviation |
| --- | --- |
| % | Percent |
| # | Number |
| ºC | Degrees Celsius |
| Cal | Calibration |
| CTQ | Critical to Quality |
| Doc. | Document |
| EH&S | Environmental Health and Safety |
| GDP | Good Documentation Practices |
| Hz | Hertz |
| ID | Identification |
| IQ | Installation Qualification |
| IQR | Installation Qualification Report |
| LED | Light Emitting Diode |
| N/A | Not Applicable |
| PM | Preventive Maintenance |
| SAP | Enterprise Resource Planning (ERP) software |
| TSP | Test and Specification Procedure |
| UUT | Unit Under Test |
| VAC | Voltage Alternating Current |
| VAF | Validation Assessment Form |
| WI | Work Instruction |

# roles and responsibilities

| Roles | Responsibilities |
| --- | --- |
| Engineering Managers | Ensure implementation of this protocol and affiliated operating procedures. Review and/or approve all necessary documentation deliverables, as required. Assign resources as required to support validation activities. |
| Originator | Coordinate the execution of the applicable IQ activities. Author the IQ Report summarizing results, upon completion of IQ activities. |
| Executor | Execute applicable IQ activities per instructions from Originator using the applicable Test Scripts. Responsible for providing the Originator with objective evidence needed to meet the requirements outlined in the applicable Test Scripts. Follow all GDP requirements per Waters’ Good Documentation Practices 730000523. |
| Reviewer | Review the finalized IQ Report. |
| Approver | Review and approve the finalized IQ Report. |

# related documents

| Doc # | Document Title |
| --- | --- |
| 730000837 | Waters Global Validation Policy |
| 730001833 | Test Engineering Management Policy |
| 730001835 | Process Validation Procedure for Milford Site |
| 730001799 | Operating Procedure, Calibration and Preventive Maintenance |
| 730000295MP | Quality System Document Control |
| WAT000050MP | Waters Record Retention Corporate Policy |
| 730000523 | Waters’ Good Documentation Practices |
| 730000483 | Global WI/Operating Procedure Template |
| 730002232 | Process Validation Master Plan - Milford |

# forms/Templates

| Doc # | Document Title |
| --- | --- |
| 730002845 | Validation Report Template |

## ΔRecords

The executed Test Script Data Sheets shall be used as objective evidence within the associated IQ Report. The IQ Report shall be stored in SAP in accordance with 730001835, Process Validation Procedure for Milford Site.

# ΔHealth & safety

### All personnel performing any part of this IQ shall be trained to this protocol, and/or the operating procedure(s) when applicable, prior to the start of any testing. A Training Record shall be filled out and attached to the associated IQR when training has been performed. A blank Training Record is provided as Appendix A and may be used. If used, the Originator and/or Executor shall digitally edit/enter the Document Number, Document Title, and the Version/Revision/Iteration in the light gray cells.

# ΔProcedure

The applicable Test Script Data Sheets provided after this Section shall be executed as follows:

## The Originator and/or Executor shall digitally edit/enter any descriptive information and acceptance criteria, as applicable, in the light gray cells of the Test Script prior to execution. Rows may be added as needed.

## Black text requirements are not to be edited or removed and shall be verified when applicable.

## N/A may be entered for any blank parameter or cell that is not used or not applicable.

## Blue text examples are provided and should be changed to font color black and/or edited/removed as applicable prior to execution.

## Blank and/or unused (blue text) rows may be removed.

## The Executor shall perform the applicable Test Scripts, hand writing in all results in the white cells following proper GDP, as defined in 730000523, Good Documentation Practices.

## Each page of the applicable Test Script Data Sheets shall have the Report Document Number, Report Iteration, Attachment Number, and be paginated, as defined in 730002845, Report Template.

### An example of an IQR Attachment stamp/marking is shown below:

### The stamp shown above is available for use from the Validation Team for unmarked attachments.

## Acceptance Criteria

The IQ will be considered successful when all applicable sections have been completed and an IQR has been written in accordance with 730002845, Validation Report Template, and has been approved. Only the Test Scripts that apply to the Equipment being qualified shall be performed.

Objective evidence will be included in, or attached to, the associated IQR. Examples of objective evidence include, but are not limited to:

All completed protocol Test Script Data Sheets

Screenshots

Calibration/PM Certificates and/or Records

Photos

Test reports

### Failure to Meet Acceptance Criteria

If the standard IQ execution is not able to be performed follow the following steps:

If a discrepancy is found which can be resolved, make the change and then continue with the execution. Discrepancies and their resolutions shall be summarized in the Deviation section of the IQR.

In the event that a discrepancy is not able to be resolved the Originator must decide to either submit the IQR to the appropriate signatories with a recommendation to accept, with justification, or reject the qualification. In the case of a rejected qualification the applicable engineering personnel shall assess the issue and determine if changes can be made in order to move forward with the IQ.

## Installation Qualification Report

### An Installation Qualification Report shall be written in accordance with 730002845, Validation Report Template, upon completion of all applicable IQ Test Script requirements. The IQR shall summarize the results of the testing performed and any deviations that occurred.

### ΔThe following Test Scripts shall be used, as applicable, for the execution of the Installation Qualification:

| Test Script # |  | IQ-01 |  |  |  |  |  |  | IQ Test Name: |  |  | Associated Documentation Verification |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment # |  | TE82882-11 |  |  |  |  |  |  | Equipment Name: |  |  | PDA Final Test System |  |  |  |  |  |  |  |  |  |
| Installation Site: |  |  |  | Flextronics Contract Manufacturing Site of WPPL in Singapore |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that supporting documentation is available prior to the completion of qualification. Test Execution: - List all applicable documents (Preventive Maintenance, Calibration, Operating Procedures, Test Specification Procedures, Manufacturing Procedures, etc.) under the Document type header, then complete the table below. Rows may be added/removed as needed. - All applicable documents must have a status of Ready, Pre-Rel-Prod, or Release General in SAP. If applicable documents are unavailable in SAP state why in the comments section. Objective Evidence: Completed IQ-01 Test Script Data Sheet SAP screenshots of below available document(s) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Document Type |  |  |  |  |  | Title, Doc. #, Version/Revision # |  |  |  |  |  |  | Status in SAP |  |  |  |  | Initial & Date |  |  |  |
| TSP |  |  |  |  |  | PDA Final Test - Test and Specification Procedure; 735001613; Version 22 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # |  | IQ -02 |  |  |  |  |  | IQ Test Name: |  |  |  | Site Verification |  |  |  |  |  |  |  |  |
| Equipment # |  | TE82882-11 |  |  |  |  |  | Equipment Name: |  |  |  | PDA Final Test System |  |  |  |  |  |  |  |  |
| Installation Site: |  |  |  | Flextronics Contract Manufacturing Site of WPPL in Singapore |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify the location in which the test equipment is to be installed is suitable for use. Test Execution: - List additional applicable environmental/input parameters in the Parameters column. Rows may be added as needed. - Using the applicable equipment documentation, define the specification(s) in the Acceptance Criteria column, as appropriate. - Verify and/or measure the applicable parameters and complete the table below. Do not delete any of the listed defined parameters. Enter N/A for any parameter or cell that is not used or applicable. The reason for why the parameter is not used or applicable shall be explained within the corresponding Installation Qualification Report. Objective Evidence: Completed IQ-02 Test Script Data Sheet Photo(s) of the work area, as applicable The source of each Acceptance Criterion shall be referenced in the IQR if stored in SAP. If not in SAP objective evidence specifying the Acceptance Criterion shall be provided in or as an attachment to the IQR. Copies of calibration certificates for measurement tools, if applicable |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Parameters |  |  |  |  | Acceptance Criteria |  |  |  |  | As Found |  |  |  |  | Pass/Fail |  |  |  | Initial & Date |  |
| Ambient Temperature |  |  |  |  | 18 ºC to 28 ºC |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Relative Humidity |  |  |  |  | 10 % to 80 %, Non-Condensing |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Workspace |  |  |  |  | Sufficient room for the test station, UUT and service |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Station Bench |  |  |  |  | Large and strong enough to support the station |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Electrical Supply |  |  |  |  | 100 VAC to 240 VAC, 50 Hz/60 Hz ±5% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Electrical Outlets |  |  |  |  | Minimum of 4 dedicated outlets |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Network Available |  |  |  |  | Minimum of 1 dedicated network |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: • Ambient temperature/humidity measured using an ExTech Model 445703 Thermohygrometer, Control Number 95648. • Electrical supply measured using a Fluke 189 Multimeter, Control Number 98209. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # | IQ-03 |  |  |  |  |  |  | IQ Test Name: |  |  | Instrumentation Calibration Verification |  |  |  |  |  |  |  |  |
| Equipment # | TE82882-11 |  |  |  |  |  |  | Equipment Name: |  |  | PDA Final Test System |  |  |  |  |  |  |  |  |
| Installation Site: |  |  | Flextronics Contract Manufacturing Site of WPPL in Singapore |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that all critical instruments are in the calibration database and that a calibration and/or a preventive maintenance has been performed, as appropriate. Test Execution: List all critical instruments/components and/or sub-systems that require calibration and/or preventive maintenance in the Description column. Rows may be added as needed. Verify that all listed critical instruments/components have had a preventive maintenance and/or calibration performed in accordance with the PM and/or Calibration procedure(s) or by an outside source, and are listed within the associated calibration database, then complete the table below. Objective Evidence: Completed IQ-03 Test Script Data Sheet Copies of calibration and/or maintenance certificates Screen captures from the associated Calibration database |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description |  |  |  |  |  |  | Asset # |  |  |  | Cal/PM Dates |  |  |  |  | Initial & Date |  |  |  |
| Enter the equipment or tooling nomenclature |  |  |  |  |  |  |  |  |  |  | Last Cal/PM: Cal/PM Due: |  |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Script # |  | IQ-04 |  |  |  |  |  | IQ Test Name: |  |  |  | Test Software and Instrument Firmware Version |  |  |  |  |  |  |  |  |
| Equipment # |  | TE82882-11 |  |  |  |  |  | Equipment Name: |  |  |  | PDA Final Test System |  |  |  |  |  |  |  |  |
| Installation Site: |  |  |  | Flextronics Contract Manufacturing Site of WPPL in Singapore |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purpose: To verify that the installed software/firmware is correct per the applicable design and/or operational document(s). Test Execution: List all relevant Equipment software and/or instrument firmware in the Description column. Specify the associated part number for the instrument firmware, if applicable. Rows may be added/removed as needed. List the associated required equipment software and/or critical instrument/internet connection sharing firmware version(s), as defined in the applicable design/operational document(s), in the Acceptance Criteria column. Verify that the software version displayed on the Operator Interface of the program and/or the firmware loaded on the instruments/PC match the acceptance criteria identically, if applicable, then complete the table below. Objective Evidence: Completed IQ-04 Test Script Data Sheet Identify and/or provide a screen capture/print out/copy of the applicable design/operational document that defines the required software/firmware version(s) within or as an attachment to the associated Installation Qualification Report. Screen capture and/or test record displaying/indicating the software and/or firmware version(s), as applicable |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description |  |  |  |  |  |  | Acceptance Criteria |  |  |  |  | As Found |  | Pass/Fail |  |  | Initial & Date |  |  |  |
| Equipment Software |  |  |  |  |  |  | Version: 4.2 |  |  |  |  | Version: |  |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  |  |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Test Script # | IQ-05 |  | IQ Test Name: | General Safety and Alarms |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Equipment # | TE82882-11 |  | Equipment Name: | PDA Final Test System |  |  |  |  |  |  |
| Installation Site: |  | Flextronics Contract Manufacturing Site of WPPL in Singapore |  |  |  |  |  |  |  |  |
| Purpose: To verify that the equipment can operate in a safe manner, within its designated location. Test Execution: List any additional safety related conditions or alarms in the Safety Inspection column, such as: guarding, interlocks, thermal or mechanical safety, chemical, secured gas tanks, restart or power loss, emergency stop, environmental exhaust, electrical safety, noise, safety labels, etc. Rows may be added as needed. EH&S shall be consulted as necessary. List the associated safety requirement(s)/specification(s) in the Acceptance Criteria column. Perform an inspection/test of the equipment to assess its safety related conditions, then complete the table below. Objective Evidence: Completed IQ-05 Test Script Data Sheet Photo(s) of the secured glass bottles, tanks, etc., if applicable The source of any applicable Acceptance Criterion shall be summarized in the IQR if stored in SAP. If the source is not stored in SAP, at a minimum the page specifying the applicable Acceptance Criterion shall be provided as an attachment. |  |  |  |  |  |  |  |  |  |  |
| Safety Inspection |  |  | Acceptance Criteria |  |  |  |  | Pass/Fail |  | Initial & Date |
| Mechanical Safety/ Pinch Points/Sharp Edges |  |  | None Found (or mitigated appropriately) |  |  |  |  |  |  |  |
| Glass Bottles |  |  | Secure |  |  |  |  |  |  |  |
| Waste Tanks |  |  | Secure and located below the system |  |  |  |  |  |  |  |
| Emergency Stop |  |  | System ceases operation upon depressing the Emergency Stop button |  |  |  |  |  |  |  |
| Guarding |  |  | Shielding in place to prevent hazard during operation |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |
| Test Script # | IQ-06 |  | IQ Test Name: | Functional Verification |  |  |  |  |  |  |
| Equipment # | TE82882-11 |  | Equipment Name: | PDA Final Test System |  |  |  |  |  |  |
| Installation Site: |  | Flextronics Contract Manufacturing Site of WPPL in Singapore |  |  |  |  |  |  |  |  |
| Purpose: To verify that the system operates as expected. Test Execution: Select the UUT(s) or tests to be performed as defined within the Installation Qualification Section of 730001835. List the Part Number and Serial Number, if applicable, in the UUT Part Number/Serial Number column. Define the appropriate Acceptance Criteria within the Acceptance Criteria column. Perform a test of the UUT(s), verify the Acceptance Criteria, then complete the table below. If an associated Test Record is being used as objective evidence the Header/Footer information should be verified for accuracy. Objective Evidence: Completed IQ-06 Test Script Data Sheet Copy of the Test Record(s) for the UUT(s), as applicable The UUT(s) selected for this Test Script and the reasoning for why it/they were selected shall be summarized within the associated Installation Qualification Report. If this Acceptance Criteria is other than “All required CTQ Specification measurements result in a Pass” then it shall be summarized within the associated VAF, Process Validation Plan, and/or Installation Qualification Report. For systems that download firmware to product the BOM showing the firmware requirement shall be provided within the IQR. |  |  |  |  |  |  |  |  |  |  |
| UUT Part Number/Serial Number |  |  | Acceptance Criteria |  |  |  | Pass/Fail |  | Initial & Date |  |
| Part Number: 700012243 Serial Number: |  |  | The UUT Serial Number is correctly listed at the bottom of the Test Record. |  |  |  |  |  |  |  |
| Part Number: 700012245 Serial Number: |  |  | The UUT Serial Number is correctly listed at the bottom of the Test Record. |  |  |  |  |  |  |  |
| Comments: |  |  |  |  |  |  |  |  |  |  |
| Overall Test Conclusion |  |  |  |  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| Pass Fail |  |  |  |  |  |  |  |  |  |  |

Appendix A:  TRAINING RECORD

| Document # | 735001613 | Document Title | PDA Final Test - Test and Specification Procedure | Version | 20 |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | Revision | N/A |

Read and Understood

By signing you acknowledge that you have reviewed, understand, and agree to comply with the policies and practices of the procedure defined above.

| Printed Name | Signature | Initials | Date |
| --- | --- | --- | --- |

|  | Document No.: ______________________. Iteration No.: ________. Attachment: _______ Page: ___ of ___. |
| --- | --- |

DOCUMENT HISTORY

| Version: | 02 | Originator: | James Stors | Change ID: | N/A |
| --- | --- | --- | --- | --- | --- |
| Summary of Change: Edited for overall grammar and terminology. Updated the Abbreviation section. Changed Performer to Executor. Updated the Records section to reference 730001835. Added Health & Safety Section. Section 8 updated to guide the Originator to type in descriptive/acceptance criteria in the Test Script Data Sheets and Training Form. Test Script Data Sheets updated to utilize gray cells for guidance and to provide a specified location for the Attachment Stamp/Marking. Removed blank rows from the Test Script Data Sheets. Removed the Electrical Safety Hi Pot/Ground Continuity test from IQ-05 Test Script Data Sheet. Added additional safety considerations to the Test Execution of IQ-05 Test Script Data Sheet. Added the Document Control Section marking to the Test Script Data Sheets and the Training Record appendix. |  |  |  |  |  |